1018233-2023-05365 | 2023-07-20 | 2023-06-23 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER WAS DIFFICULT TO REMOVE. THE CATHETER WAS PLACED ON 18TH JUNE AND WHEN THE CATHETERIZATION WAS NO LONGER NEEDED ON 23RD JUNE, UPON ATTEMPTING TO REMOVE THE CATHETER, THE STERILE WATER COULD NOT BE REMOVED. THEN THE CATHETER WAS CUT UNDER THE DIRECTION OF A UROLOGIST AND A SONDE WAS INSERTED INTO THE INFLATION LUMEN TO DRAIN WATER, BUT STILL THE WATER COULD NOT BE DRAINED. MEANWHILE, THE CATHETER FELL OUT OF THE PATIENT WITH THE BALLOON STILL INFLATED. THEN THE WATER WAS DRAINED AND ONE SIDE OF THE CATHETER WHICH WAS CUT WAS DISCARDED. ACCORDING TO THE INQUIRY SHEET ATTACHED, THE BALLOON WAS DEFLATED BY PUMPING AND INSERTION OF WIRE. USUALLY THE BALLOON COULD BE DEFLATED WITHIN SECONDS. | THE REPORTED EVENT WAS CONFIRMED - CAUSE UNKNOWN. SAMPLE WAS EVALUATED AND OBSERVED HEAVY SALT ACCUMULATION ALONG THE DRAINAGE LUMEN OF CATHETER THAT CAUSE DIFFICULT TO REMOVE ISSUE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE HEAVY SALT ACCUMULATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING. THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR. 1. FOLEY CATHETER 2. ACCESSORIES. CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER, WATER SOLUBLE LUBRICANT, ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION, TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS , GLOVES, STATLOCK® FOLEY. [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE. DIFFICULTY OR FAILURE TO REMOVE THE DEVICE. OCCLUSION OF CATHETER INNER LUMENS. ENCRUSTATION. ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE . DEVICE DAMAGE DUE TO INAPPROPRIATE USE. 2) ADVERSE EVENTS: URINARY-TRACT INFECTION, HEMORRHAGE, HEMATURIA, ALLERGY REACTION TO THE DEVICE, CALCULUS FORMATION , EDEMA, PAIN, DISCOMFORT, INJURY OF BLADDER OR URETHRAL, URETHRITIS, URINARY INCONTINENCE, RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE. STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05345 | 2023-07-20 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165818 |  | IT WAS REPORTED THAT THE USER HAD EXPERIENCED ISSUES WITH INDWELLING URINARY CATHETER "FELLING OUT" TWICE SINCE USER¿S LAST CHANGE. IT WAS STATED THAT THE ON MONDAY DURING THE PROCEDURE OF CHANGING FOLEY CATHETER THE BALLOON BURST WHILE IN PROCESS OF FILLING BALLOON WITH THE SPECIFIED AMOUNT OF WATER. THIS RESULTED IN ANOTHER INDWELLING URINARY CATHETER NEEDING TO BE PLACED WITH RESULTING TRAUMA REQUIRING THE CLIENT TO ATTEND ED FOR REVIEW. IT WAS MY UNDERSTANDING THAT THIS CATHETER ALSO "FELL OUT" ON TUESDAY RESULTING IN CLIENT AGAIN NEEDING TO GO TO ED FOR ATTENTION. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6)2023, IT WAS STATED THAT THE CUSTOMER HAD ALSO INFORMED THAT TWO OTHER CATHETERS FROM THAT SAME BOX HAD FAILED BY FALLING OUT IN THE DAYS PRIOR TO THE TWO CATHETERS. THERE WAS NO TRAUMA EXPERIENCED WITH THOSE FIRST TWO. THIS MEANS OUT OF A BOX OF TEN, SIX WERE USED OF WHICH FOUR ARE NOW REPORTED TO HAVE BEEN FAULTY. THE OTHER FOUR WERE STILL IN THE BOX READY FOR PICK UP, THERE WERE NO ACTUAL USED FAULTY SAMPLES AVAILABLE FOR PICK UP. IT WAS STATED THAT THE FIRST CATHETER WAS CHECKED AND THERE WERE NO MISSING PIECES. WHEN INSERTING THE SECOND CATHETER USER AND NURSE NOTICED BLOOD DRIPPING FROM THE EXTERNAL PART OF THE CATHETER AS WELL AS SEEPING AROUND THE CATHETER WERE INSERTED INTO PENIS. THE USER WAS TREATED IN ED, THEY DID NOT REMOVE THE CATHETER, THEY BANDAGED UP PENIS TO STOP BLEEDING AND KEPT USER OVERNIGHT FOR OBSERVATION BEFORE RELEASING THE NEXT DAY. THE CIRCUMSTANCES SURROUNDING THE SECOND CATHETER FALLING OUT WAS USER FELT LIKE THERE WAS ONLY A SMALL AMOUNT OF URINE IN BAG WHEN AT HOME SO CHECKED AND FOUND THAT THE CATHETER HAD FALLEN OUT. THE USER THEN RETURNED TO ED WHERE THE NURSE CHECKED THE CATHETER AND DETERMINED THE BALLOON WOULD NOT INFLATE. THERE WAS NO INDICATION ANY PIECES WERE LEFT INSIDE. THE NURSE THEN TESTED THE BALLOON OF A NEW DIFFERENT BRANDED CATHETER BEFORE INSERTION AND RETURNED HOME WITHOUT INCIDENT OF THAT CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03085 | 2023-07-20 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG TPS STYLET 4F MAXIMAL BARRIER TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED BY CUSTOMER THAT A 4FRENCH PICC LINE INTRODUCER HAD A PLASTIC BURR ON THE END AND NOT ALLOWING THEM TO THREAD IT INTO PATIENT. THEY NEEDED TO DROP A WHOLE NEW INTRODUCER FOR THE PATIENT. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05360 | 2023-07-20 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE CATHETER BALLOON WAS DAMAGED. AFTER PLACEMENT TO THE PATIENT, THE BALLOON WAS RUPTURED AND STERILE WATER WAS LEAKED. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS EVALUATED AND OBSERVED THERE WAS NO BALLOON BURST/RUPTURE. HOWEVER, THERE WERE MULTIPLE TEAR ON THE CATHETER SHAFT THAT ALLOWED WATER TO LEAK OUT. TEAR WAS EXAMINED UNDER MICROSCOPE AND NOTED TO BE ROUND BUT NOT SMOOTH. EXACT CAUSE OF SAMPLE CONDITION HAVE BEEN DETERMINED AND CONFIRMED AS A MANUFACTURING RELATED PROCESS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO BUBBLE DURING DIPPING PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE 1. MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05396 | 2023-07-21 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L22 |  | IT WAS REPORTED THAT WATER LEAKAGE OCCURRED FROM THE FOLEY CATHETER BALLOON SECTION (LOT# MYFU1698). IT WAS STATED THAT THE USED PRODUCTS HAD BEEN DISCARDED. THE CUSTOMER REQUESTED THAT ALL OTHER LOTS WITH THE SAME CATALOG NUMBER BE INVESTIGATED, SO THEY WERE COLLECTED. IT WAS STATED THAT DUE TO THE FREQUENT OCCURRENCE OF DEFECTS (OTHER EXTERNAL REFERENCE # (B)(4) AND OTHER EXTERNAL REFERENCE #(B)(4)) THIS YEAR, REPRESENTATIVES RECEIVED A REQUEST TO CHECK ALL OTHER LOTS OF THE SAME PART NUMBER(LOT# MYFY3438, LOT# MYEW1667, LOT# MYFP2501), SO THEY COLLECTED THEM. IT WAS STATED THAT THERE WERE MANY PROBLEMS WITH THIS PRODUCT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03108 | 2023-07-21 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) (W/MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7957505J |  | IT WAS REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, WHEN THE CATHETER WAS REPLACED AND THE CONNECTOR PART WAS PULLED OUT, THE CATHETER WAS BROKEN AT THE CONNECTION BETWEEN THE CONNECTOR PART AND THE CATHETER. TENSILE STRESS WAS NOT APPLIED STRONG. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER BREAK IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. TWO PHOTO SAMPLES OF A 5 FR DL GROSHONG NXT CATHETER WERE PROVIDED FOR EVALUATION. THE FIRST IMAGE SHOWS THE IMPLANTED CATHETER SECURED TO THE PATIENT. A COMPLETE BREAK AT THE PROXIMAL END OF THE CATHETER EXTENDING FROM THE BIFURCATION HUB WAS PRESENT. THE SECOND IMAGE SHOWED THE FRACTURE SURFACES OF THE BREAK LOCATIONS. THE BREAK FEATURES COULD NOT BE CLOSELY INSPECTED FROM THE SAMPLES PROVIDED. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE DAMAGE DURING HANDLING, MATERIAL FATIGUE CAUSED BY REPEATED KINKING, AND SHARP INSTRUMENT DAMAGE. SINCE A COMPLETE BREAK IN THE CATHETER WAS NOTED, THE COMPLAINT IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05397 | 2023-07-21 | 2023-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT ON (B)(6) 2023. THE CATHETER FELL OUT DUE TO A BALLOON BURST. ON (B)(6) 2023, THE NURSE FOLLOWED A DOCTOR'S ADVICE TO INDWELL A URINARY CATHETER FOR THE PATIENT. THE CATHETERIZATION PROCESS WENT SMOOTHLY, AND 30ML OF NORMAL SALINE WAS INJECTED INTO THE BALLOON, WHICH WAS DOUBLE-FIXED, DRAINED SMOOTHLY, AND HAD NO URINE LEAKAGE. THE MODEL WAS A DOUBLE-CAVITY, 16FR. ON (B)(6) 2023, WHEN THE NURSE CHECKED THE PATIENT IT WAS FOUND THAT THE CATHETER HAD FALLEN OUT. THE BALLOON WAS INSPECTED FOR A DAMAGE OF ABOUT 2CM IN SIZE AND THE DOUBLE FIXATION WAS STILL IN PLACE. THE NURSE IMMEDIATELY REPORTED TO THE DOCTOR AND HEAD DUTY NURSE TO CHANGE THE URINARY CATHETER THAT WAS DOUBLE FIXED AND INJECTED 30ML OF SALINE AND FOUND THE DRAINAGE WAS SMOOTH WITHOUT LEAKAGE OF URINE. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON 15JUL2023, CLARIFIED THAT THERE WERE NO MISSING PIECES OF BALLOON INSIDE THE PATIENT'S BODY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05399 | 2023-07-21 | 2023-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT WHEN THE DOCTOR PULLED OUT THE CATHETER, IT WAS FOUND THAT THERE WAS NO WATER IN THE AIR BAG. THEN THE AIR BAG WAS REINJECTED WITH WATER AND FOUND THAT THE WATER LEAKED. ACCORDING TO THE DEVICE FAILURE PERFORMANCE, THE CATHETER BALLOON LEAKED. THE DESCRIPTION OF EVENT CAUSE ANALYSIS STATED THAT THERE WAS A SMALL HOLE IN THE CATHETER DURING INDWELLING. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ MECHANICAL FAILURE/ OPERATOR ERROR/ THIN RUBBERIZE LAYER. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03111 | 2023-07-21 | 2023-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED HOLE IN PICC LINE. WHEN REPOSITIONING THE PICC LINE WHEN FLUSHING THE CATHETER, A HOLE IS DISCOVERED BETWEEN THE INSERTION SITE AND THE WINGS OF THE CATHETER. UNCLEAR AT WHAT LEVEL. THEY DECIDE TO PULL THE PICC LINE AND IT WAS THEN THROWN AWAY. THE PATIENT IS BOOKED FOR A NEW PICC LINE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION AS THE DEVICE WAS DISCARDED AFTER THE EVENT OCCURRED. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05400 | 2023-07-21 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT THE CUSTOMER EXPERIENCED A PRODUCT COMPLAINT WITH PRODUCT NUMBER (B)(6)AT LABOR & DELIVERY UNIT. THE FOLEY CATHETER FELL OUT OF THE PATIENT AND THE NURSE RECOGNIZED THAT THE TIP WAS MISSING (B)(6). THE CUSTOMER SAVED THE CATHETER AND IT WAS AVAILABLE TO BE RETURNED. NO PATIENT COMPLICATION WAS REPORTED. THE CUSTOMER WAS CONCERNED ABOUT THE FOLEY TRAYS AND OPEN LOT NUMBER, THEY NOTICED OTHER ISSUES WITH FOLEY CATHETERS, KINK TUBING (B)(6)AND CUT FOLEY CATHETER (B)(6). THESE FOLEY CATHETERS WERE NOT INSERTED IN PATIENTS. THEY HAD SAMPLES OF THESE TRAYS TOO. THEY QUARANTINED LOT NUMBER ALL THE THREE LOTS WHICH HAD BEEN REMOVED FROM CIRCULATION. AFTER SPEAKING WITH CUSTOMER, THEY PULLED ABOUT 3 CASES AND INDIVIDUAL FOLEY TRAYS FROM THE FLOOR. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 2-WAY CATHETER WITH ATTACHED SAMPLE PORT AND INLET TUBING. VISUAL INSPECTION NOTED NO BREAKAGE OR DEFECTS ON THE CATHETER. THE TIP OF THE CATHETER WAS INTACT. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED, RETURNING 10ML OF SOLUTION. THE BALLOON PASSIVELY DEFLATED WITH NO CUFFING OR LEAKS NOTED. THE DRAINAGE LUMEN WAS FLUSHED WITH D.I. WATER. AND FLOWED THROUGH WITH NO ISSUES OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. CORRECTION: D,H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05394 | 2023-07-21 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L22 |  | IT WAS REPORTED THAT WATER LEAKAGE OCCURRED FROM THE FOLEY CATHETER BALLOON SECTION (LOT# MYFU1698). IT WAS STATED THAT THE USED PRODUCTS HAD BEEN DISCARDED. THE CUSTOMER REQUESTED THAT ALL OTHER LOTS WITH THE SAME CATALOG NUMBER BE INVESTIGATED, SO THEY WERE COLLECTED. IT WAS STATED THAT DUE TO THE FREQUENT OCCURRENCE OF DEFECTS (OTHER EXTERNAL REFERENCE # 23050320 ) THIS YEAR, REPRESENTATIVES RECEIVED A REQUEST TO CHECK ALL OTHER LOTS OF THE SAME PART NUMBER(LOT# MYFY3438, LOT# MYEW1667, LOT# MYFP2501), SO THEY COLLECTED THEM. IT WAS STATED THAT THERE WERE MANY PROBLEMS WITH THIS PRODUCT. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 10JUL2023, IT WAS CONFIRMED THAT THE DEVICE DID NOT USED ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03112 | 2023-07-21 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG 4F SINGLE-LUMEN PICC PERIPHERALLY INSERTED CVC BASIC TRAY | CATHETER, PERCUTANEOUS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7715405 |  | IT WAS REPORTED BY THE CUSTOMER THAT DURING INJECTING SALINE PRIOR TO PLACEMENT, THE CATHETER WAS FOUND DAMAGED AND SEVERELY LEAKED LIQUIDS. IT WAS REPORTED THIS OCCURRED WITH 2 DEVICES THAT WERE RETURNED FOR INVESTIGATION. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF LEAKING CATHETERS WAS UNCONFIRMED BECAUSE THE PROBLEM COULD NOT BE REPRODUCED. THE PRODUCT RETURNED FOR EVALUATION WAS TWO 4FR S/L GROSHONG CATHETERS. BOTH SAMPLES WERE RECEIVED WITH INLAID STYLETS. THE FIRST SAMPLE (SAMPLE 1) EXHIBITED BENDS ALONG THE LENGTH OF THE STYLET NEAR THE 20CM, 26CM, AND 30CM DEPTH MARKINGS. THE SECOND SAMPLE (SAMPLE 2) EXHIBITED A BEND IN THE STYLET JUST DISTAL OF THE 5CM MARKING. ATTEMPTS TO INFUSE WATER THROUGH BOTH SAMPLES USING A 12ML SYRINGE REVEALED BOTH SAMPLES TO BE PATENT TO INFUSION AND ASPIRATION WITH NO OBSERVED LEAKS. NO LEAKS WERE OBSERVED DURING PRESSURIZATION OF THE SAMPLES. MICROSCOPIC INSPECTION OF THE SAMPLES DID NOT REVEAL ANY DAMAGE OR EVIDENCE OF LEAKAGE. NO DEFICIENCIES WERE DISCOVERED DURING EVALUATION OF THE RETURNED SAMPLES. CONSEQUENTLY, THIS COMPLAINT IS UNCONFIRMED AT THIS TIME. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03116 | 2023-07-21 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT APPROXIMATELY ONE MONTH POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER WAS ALLEGEDLY CRACKED. IT WAS FURTHER REPORTED THAT BLOOD EXUDATION WAS ALLEGEDLY OCCURRED WITH THE CATHETER BODY. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: (EXPIRATION DATE: 03/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05398 | 2023-07-21 | 2023-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT ON (B)(6), 2023, THE PATIENT WAS CATHETERIZED AND WAS INJECTED 18ML OF NORMAL SALINE INTO THE CATHETER BALLOON. AT 15:00 ON (B)(6), 2023, THE PATIENT INFORMED THE NURSE THAT THEY HEARD A "BANG" IN THE LOWER ABDOMEN. UPON CHECKING, THE URINE FLOWED OUT FROM THE URETHRAL OPENING AND A PART OF THE CATHETER CAME OUT. AFTER PULLING OUT THE CATHETER, UPON CHECKING THE BALLOON, IT WAS FOUND THAT THE BALLOON WAS DAMAGED AND THE STERILE WATER INSIDE HAD DRAINED. THE PATIENT WAS REPLACED WITH A NEW CATHETER AND THERE WAS NO IMPACT ON THE PATIENT. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6)2023, STATED THAT THE BALLOON WAS BURST INSIDE THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05392 | 2023-07-21 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE USER MADE A COMPLAINT REGARDING THE FOLEY CATHETER. IT WAS STATED THAT THE BALLOON OF THE CATHETER COULD NOT BE INFLATED. AS A RESULT, USER HAD TO INSERT ANOTHER CATHETER INTO THE PATIENT. IT WAS STATED THAT THE DATE WHEN THE COMPLAINT OCCURRED WAS (B)(6) 2023. THE COMPLAINT WAS NOTED DURING COMMISSIONING OF THE FOLEY CATHETER, AND THIS DID NOT AFFECT THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05386 | 2023-07-21 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THERE WAS A REGULAR PROBLEM WITH BARD FOLEY CATHETERS FOR A WHILE WITHIN THE INTENSIVE CARE DEPARTMENT. STATED THAT THEY WERE NOT DRAINING PROPERLY (BATCH# UNK) OR WERE CLOGGED (BATCH # UNK) AND THEREFORE THE PATIENTS GOT A FULL BLADDER. LAST WEEK, A PATIENT WHO HAD 1300 CC IN THE BLADDER AND HAD A CATHETER, WHEN REMOVED THE BALLOON LOOKED A BIT STRANGE (BATCH# UNK). IT WENT WELL FOR A WHILE, BUT LAST FALL THEY ALSO HAD THREE PATIENTS (BATCH# UNK) PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 13JUL2023, STATED THAT THE BALLOON LOOKED A BIT STRANGE MEANT THAT THE BALLOON LOOKED ADHERED AND WAS ALSO BLOODY BUT THAT COULD ALSO BE FROM THE PATIENT. FOR THIS CASE, THEY HAD NOW SEEN ONLY ONE PATIENT BUT THE LOT NUMBER COULD NOT BE TRACED. AS FOR THE CATHETERS THAT WERE NOT DRAINING PROPERLY, THEY DID NOT KNOW THE LOT NUMBER OR THE NUMBER OF PATIENTS THAT HAD THIS ISSUE. HOWEVER, THEY OFTEN HAD TO SIPHON TO EMPTY THE BLADDER PROPERLY, COULD BE DUE TO THE CATHETER BUT ALSO THE URIMETER. AS FOR THE CATHETERS THAT WERE CLOGGED, THEY DID NOT HAD ANY IDEA ABOUT THE LOT NUMBER OR THE NUMBER OF PATIENTS THAT HAD THIS ISSUE. AS FOR THE PREVIOUS PROBLEMS, NONE WERE REPORTED EXCEPT FOR ONE THAT WAS. IN THAT PATIENT, IT WAS TYPICAL THAT THEY DID URINATE 30-40 CC, BUT HAD ABDOMINAL PAIN. THEY TOOK OUT THE CATHETER AND AT THAT POINT THE URINE POURED OUT AND THE ABDOMINAL PAIN DISAPPEARED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: SALT ACCUMULATION)/BLOCK DRAINAGE LUMEN/NO DRAINAGE EYE). THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THEREFORE, NO ADDITIONAL ACTION REQUIRED AT THIS TIME. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05390 | 2023-07-21 | 2023-06-26 | I, n, j, u, r, y | Manufacturer report | BARD® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 366714 | 366714 |  | IT WAS REPORTED THAT THE CUSTOMER WAS CLIMBING THROUGH A FENCE ON SUNDAY EVENING AND HEARD A POP NOISE. SHORTLY AFTER, THE CUSTOMER'S FOLEY CATHETER HAD FALLEN OUT. THE CUSTOMER EXAMINED SAME AND FOUND THE BALLOON ON THE END OF THE CATHETER HAD RUPTURED. THE CUSTOMER THEN PHONED THE HOME CARE OFFICE THE NEXT MORNING AND REPORTED SAME TO THE HOME CARE NURSE. THE CUSTOMER CAME TO THE OFFICE AND HAD SAME REINSERTED. NO APPARENT HARM REACHED THE PATIENT. THE IMPACT OF THE INCIDENT WAS PATIENT WAS UNABLE TO VOID. THE PATIENT REQUIRED UNEXPECTED OR PROLONGED CARE. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05381 | 2023-07-21 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT A COUPLE OF MAGIC 3 GO MALE INTERMITTENT CATHETER SALINE PACKETS WERE EMPTY, AND PATIENT WAS NOT ABLE TO USE THE CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05379 | 2023-07-21 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT HAD 3 CATHETERS WITH BENT TIPS AND ONE THAT DID NOT HAVE ANY FLUID IN IT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INCORRECT PROCESS PARAMETERS". THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES, SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05382 | 2023-07-21 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT 2 HOURS AFTER PLACEMENT, THE BALLOON WAS NOT INFLATED, AND THERE WAS NO FIXATION WATER IN IT. WHEN AIR WAS TURNED ON, IT DEFLATED AND WAS FOUND TO HAVE A TEAR. PER PHOTOS RECEIVED FROM IBC ON 10JUL2023, STATED THAT THE USER RETURNED FROM THE OPERATING ROOM AT 22:10 PM, AND WAS INFORMED THAT THE URINARY CATHETER WAS INSERTED, AND 680 ML OF URINE HAD BEEN URINATED DURING THE OPERATION. THERE WAS NO URINATION AFTER THE OPERATION, BUT SINCE THE TUBE WAS OF A TYPE THAT WAS PRONE TO AIRLOCK, THE TUBE WAS MANIPULATED, AND THE SITUATION WAS OBSERVED. THE BALLOON PART WAS NOT INFLATED AND CONTAINED NO STATIONARY WATER. WHEN THE USER PUT AIR INTO IT, THEY FOUND THE BALLOON HAD SHRUNK AND WAS TORN. THE PATIENT REPORTED IT TO THE DOCTOR, AND INSTEAD OF REINSERTION, THE PATIENT WAS MEASURED WITH A DIAPER AND TREATED ACCORDINGLY. AS PER IBC NOTIFICATION RECEIVED VIA ON 10JUL2023, THE BALLOON RUPTURE WAS NOT OBSERVED TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05384 | 2023-07-21 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 63612 |  | IT WAS REPORTED THAT THE INSERTION WAS FINE BUT INTERMITTENT CATHETER DRAGGED WHEN PATIENT WAS PULLING IT OUT. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 13JUL2023, STATED THAT THE PATIENT WAS UNCOMFORTABLE. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿CATHETER SWELLS, SOFTENS, ETC., DUE TO INCOMPATIBILITY WITH COATING¿. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿FOR UROLOGICAL USE ONLY. INTENDED FOR USE BY PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICES ARE PASSED TO THE URINARY BLADDER VIA THE URETHRA. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. PRIOR TO OPENING THE SEALED CATHETER POUCH, APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. WET THE CATHETER BY HOLDING THE PACKAGE WITH THE PRINTED SIDE UP AND TIP THE PACKAGE END-TO-END THREE TO SIX TIMES TO WET THE CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. PEEL OPEN THE PACK AT THE FUNNEL/HANDLE END JUST ENOUGH TO EXPOSE 5 CM (2¿) OF THE CATHETER, OR FOR PRODUCTS WITH AN INSERTION SLEEVE, EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. USING THE CATHETER WITH INSERTION SLEEVE (GRIPPER): HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. WITHOUT INSERTION SLEEVE (GRIPPER): A. HOLD THE HANDLE/FUNNEL END AND REMOVE THE CATHETER FROM THE PACKAGING. GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA AND ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05376 | 2023-07-21 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THE 6 OUT OF THE 1ST BOX DID NOT HAVE ENOUGH LUBRICANT ON THE CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PROVIDED LOT NUMBER WAS INVALID THEREFORE DHR REVIEW CAN¿T BE COMPLETED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05378 | 2023-07-21 | 2023-06-30 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAS HAD 2 URINARY TRACT INFECTIONS AND WAS HOSPITALIZED FOR ONE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA AND IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-07
17370295 | 2023-07-21 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP FOLEY TRAY SYSTEM LUBRI-SIL IC COMPLETE CARE | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | A903416A | A903416A |  | FROM STAFF: IT WAS NOTICED THAT THE PATIENT HAD NOT BEEN PUTTING OUT ANY URINE. RN ATTEMPTED TO MANIPULATE THE FOLEY CATHETER. IT WAS NOTICED THAT THE CATHETER WAS NO LONGER IN THE PATIENT. THE BALLOON WAS NOT INFLATED AND THE TOP ON THE CATHETER WAS PARTIALLY DETACHED. ALL PARTS OF THE CATHETER DID COME OUT OF THE PATIENT. |  | N | Y |  |  | 2023-07
1018233-2023-05375 | 2023-07-21 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT WATER LEAKAGE OCCURRED FROM THE BALLOON OF THE FOLEY CATHETER. THE USED PRODUCTS HAD BEEN DISCARDED . THE SAME LOT OF PRODUCT HAD BEEN COLLECTED FOR INVESTIGATION. THERE WERE MANY PROBLEMS WHICH HAS BEEN ALREADY REPORTED (23050320 AND 23010589). PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 18JUL2023, STATED THAT THE SAMPLE IN WHICH THE EVENT OCCURRED WAS DISCARDED AT THE FACILITY, BUT FOUR UNUSED SAMPLES FROM THE SAME LOT WERE RETURNED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 20JUL2023, CLARIFIED THAT THE ORIGINAL SAMPLE WAS DISCARDED AT THE FACILITY. DISCARDED DEVICE WAS USED ON THE PATIENT. UNUSED WAS SELECTED BECAUSE THE UNOPENED PRODUCT WAS RECEIVED AT THE REQUEST OF THE FACILITY TO INVESTIGATE THE SAME LOT OF PRODUCTS. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTIONS: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03105 | 2023-07-21 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN PORT SINGLE LUMEN | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0603880C |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE IN THE RIGHT INTERNAL JUGULAR VEIN, THE VASCULAR SHEATH WAS ALLEGEDLY UNABLE TO BE INSERTED DURING THE OPERATION. IT WAS FURTHER REPORTED THAT THE SHEATH WAS ALLEGEDLY DIFFICULT TO REMOVE, AND THE SHEATH WAS FOUND TO BE DEFORMED UPON REMOVAL. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE BARDPORT MRI IMPLANTABLE PORT KIT WAS RECEIVED FOR EVALUATION AND TWO PHOTOS WERE PROVIDED FOR REVIEW. VISUAL AND MICROSCOPIC EVALUATIONS WERE PERFORMED. THE DISTAL TIP OF DILATOR WAS NOTED TO BE DAMAGED. THEREFORE, THE REPORTED MATERIAL DEFORMATION ISSUE IS UNCONFIRMED AS THE SPECIFIC DAMAGE SUCH AS DEFORMATION DUE TO COMPRESSIVE STRESS ISSUE WAS IDENTIFIED AND CONFIRMED DURING INVESTIGATION AND THE PHOTO REVIEW ALSO CONFIRMS THE SAME. HOWEVER, THE REPORTED DIFFICULT TO INSERT AND DIFFICULT TO REMOVE ISSUES IS INCONCLUSIVE AS THE EXACT CIRCUMSTANCE AT THE TIME OF THE EVENT REPORTED WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 05/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00529 | 2023-07-21 | 2023-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQ5054 | CQ5054 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE OUTER PACKAGE OF THE BALLOON WAS OPENED AND FLUSHED WITH NORMAL SALINE. IT WAS FURTHER REPORTED THAT THE TIP WAS ALLEGEDLY RUPTURED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRATION DATE: 01/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03099 | 2023-07-21 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT LONG-TERM HEMODIALYSIS CATHETER, POLYURETHANE, KIT, 14.5F | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5703690 | 5703690 |  | IT WAS REPORTED THAT APPROXIMATELY LESS THAN TWO MONTHS POST DIALYSIS CATHETER PLACEMENT, BLOOD EXUDATION ALLEGEDLY OCCURRED FROM THE CATHETER BODY DURING ROUTINE DIALYSIS. IT WAS FURTHER REPORTED THAT THE CATHETER BODY HAD ALLEGEDLY CRACKED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE SAMPLE WAS NOT RETURNED FOR EVALUATION. TWO ELECTRONIC PHOTOS AND ONE VIDEO WERE REVIEWED. BOTH PHOTOGRAPHS SHOWED A PATIENT WITH A CLINICIAN HOLDING A SWAB AT OR NEAR THE CATHETER, NEAR THE CATHETER EXIT SITE. THERE WAS A SMALL AMOUNT OF BLOOD IN THE AREA IMMEDIATELY SURROUNDING THE EXIT SITE. IN THE VIDEO THE CLINICIAN IS STILL HOLDING THE SWAB, AND REPEATEDLY TOUCHES IT TO THE CATHETER JUST BELOW THE EXIT SITE WHILE PUSHING THE CATHETER INWARD (INTO THE PATIENT) SOMEWHAT, SO THAT IT BENDS SLIGHTLY. NO SOURCE OF BLEEDING OR LEAK WAS APPARENT. NO DAMAGE TO THE CATHETER WAS APPARENT. THEREFORE, BASED ON THE PHOTO AND VIDEO REVIEW, THE REPORTED CATHETER CRACK AND LEAKAGE CANNOT BE CONFIRMED. A DEFINITIVE ROOT CAUSE FOR THE REPORTED CRACKING AND BLOOD EXUDATION COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: (EXPIRY DATE: 02/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05389 | 2023-07-21 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THERE WAS A REGULAR PROBLEM WITH BARD FOLEY CATHETERS FOR A WHILE WITHIN THE INTENSIVE CARE DEPARTMENT. STATED THAT THEY WERE NOT DRAINING PROPERLY OR WERE CLOGGED AND THEREFORE THE PATIENTS GOT A FULL BLADDER. LAST WEEK, A PATIENT WHO HAD 1300 CC IN THE BLADDER AND HAD A CATHETER, WHEN REMOVED THE BALLOON LOOKED A BIT STRANGE. IT WENT WELL FOR A WHILE, BUT LAST FALL THEY ALSO HAD THREE PATIENTS. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE BALLOON LOOKED A BIT STRANGE MEANT THAT THE BALLOON LOOKED ADHERED AND WAS ALSO BLOODY BUT THAT COULD ALSO BE FROM THE PATIENT. FOR THIS CASE, THEY HAD NOW SEEN ONLY ONE PATIENT, BUT THE LOT NUMBER COULD NOT BE TRACED. AS FOR THE CATHETERS THAT WERE NOT DRAINING PROPERLY, THEY DID NOT KNOW THE LOT NUMBER OR THE NUMBER OF PATIENTS THAT HAD THIS ISSUE. HOWEVER, THEY OFTEN HAD TO SIPHON TO EMPTY THE BLADDER PROPERLY, COULD BE DUE TO THE CATHETER BUT ALSO THE URIMETER. AS FOR THE CATHETERS THAT WERE CLOGGED, THEY DID NOT HAVE ANY IDEA ABOUT THE LOT NUMBER OR THE NUMBER OF PATIENTS THAT HAD THIS ISSUE. AS FOR THE PREVIOUS PROBLEMS, NONE WERE REPORTED EXCEPT FOR ONE THAT WAS. IN THAT PATIENT, IT WAS TYPICAL THAT THEY DID URINATE 30-40 CC, BUT HAD ABDOMINAL PAIN. THEY TOOK OUT THE CATHETER AND AT THAT POINT THE URINE POURED OUT AND THE ABDOMINAL PAIN DISAPPEARED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE LOW LATEX VISCOSITY CAUSING THIN SAC. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05383 | 2023-07-21 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION (UTI) FROM (B)(6) BY USING THE MAGIC3 GO INTERMITTENT CATHETER OF DIFFERENT LOT NUMBERS LOT# JUFV0450, LOT# JUFZ0455, LOT# JUFZ0782, LOT# JUFZ0926, LOT# JUGP9222, LOT# JUGQ3272, LOT# JUGR0160, LOT# JUGR1384, LOT# JUGU1966. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05388 | 2023-07-21 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE PATIENT ALWAYS USED FOLEY CATHETERS, AND A NEW ORDER HAD JUST BEEN RECEIVED. IT WAS STATED THAT THERE WAS A BIG LENGTH DIFFERENCE (BACK ORDER, ORIGINAL ORDER PLACED IN APRIL). THE USER ADVISED THAT THEY WERE NORMALLY APPROX. 11" BUT ALL THE NEW 12 CATHETERS USER HAD RECEIVED WERE MORE LIKE 8" AND COMPLETELY UNUSABLE. THE REPRESENTATIVE CHECKED OUR SHELF STOCK, AND ALL WERE THE SAME BATCH AND SHORTER LENGTH THAT THE PATIENT HAD RECEIVED AND COMPLAINED ABOUT. THE PHOTO UPLOADED SHOWING NORMAL CATHETER AND FAULTY SHORTER ONES. | THE REPORTED EVENT IS CONFIRMED PRODUCT TRANSFER ISSUE. AFFECTED SAMPLE RETURNED/PHOTO HAVE CONFIRMED THAT 11¿ FEMALE FOLEY CATHETERS ARE BEING PRODUCED AS 8¿ REGULAR LENGTH FEMALE FOLEY CATHETERS. A DHR REVIEW IS NOT REQUIRED FOR THIS COMPLAINT. THE REPORTED EVENT IS UNDERSTOOD, CHARACTERIZED, AND CONFIRMED AS UNRELATED TO THE LABELING ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03101 | 2023-07-21 | 2022-12-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT TWO DAYS POST PORT PLACEMENT VIA RIGHT INTERNAL JUGULAR VEIN, BLOOD ALLEGEDLY COULD NOT BE DRAWN. IT WAS FURTHER REPORTED THAT THE CATHETER WAS FOUND TO BE RUPTURED IN THE MIDDLE UPON REMOVAL. REPORTEDLY, THE DEVICE ALLEGEDLY HAD DIFFICULTY IN FLUSHING AND THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 11/2023), G3, H6 (DEVICE, METHOD). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03104 | 2023-07-21 | 2023-06-05 | I, n, j, u, r, y | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT ONE YEAR, FOUR MONTHS AND NINETEEN DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN FOR CHEMOTHERAPY, THE DEVICE ALLEGEDLY COULDN'T HAD BLOOD RETURN. IT WAS FURTHER REPORTED THAT THE INFUSION WAS DIFFICULT. FURTHERMORE, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN NEAR THE INTERNAL JUGULAR VEIN. REPORTEDLY, THE DISTAL CATHETER SEGMENT AND THE PORT BODY WERE REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE MRI POWERPORT ISP IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED TO THE CATHETER. THE DISTAL END OF THE CATHETER WAS NOT RETURNED. THE EDGES OF THE CIRCUMFERENTIAL BREAK WERE NOTED TO BE UNEVEN AND HAVE LONGITUDINAL SPLITTING AT THE BORDER OF TWO REGIONS. THE SURFACE WAS NOTED TO BE GRANULAR AND ROUGH IN ONE REGION AND ROUND AND SMOOTH IN THE OTHER. THE CATHETER WAS ELLIPTICAL IN SHAPE. THE FIRST IMAGE DEPICTS THE DEVICE HAS BEEN INSERTED VIA THE LEFT INTERNAL JUGULAR VEIN AND WAS PRIOR TO FRACTURE OF THE CATHETER AND THE SECOND X-RAY SHOWS THE PORT POSITIONED SIMILAR TO THE FIRST X-RAY BUT THE DISTAL SEGMENT OF CATHETER IN THE NECK IS NOT SEEN. THE FRACTURED PART OF THE CATHETER PRESUMABLY MIGRATED TO THE RIGHT HEART. THE PHOTO SHOWS THE DISTAL END OF THE CATHETER SEGMENT ATTACHED TO A TUBE WITH A SYRINGE NOTED NEAR THE CATHETER SEGMENT. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION AND THE IDENTIFIED MIGRATION, CATHETER WEAR AND DEFORMATION ISSUES. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO FLUSH AND SUCTION ISSUES AS THE EXACT CIRCUMSTANCE AT THE TIME OF EVENT REPORTED WAS UNKNOWN. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03106 | 2023-07-21 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT ONE YEAR FIVE MONTHS AND FIVE DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN FOR CHEMOTHERAPY, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN NEAR THE INTERNAL JUGULAR VEIN. REPORTEDLY, THE DISTAL CATHETER SEGMENT AND THE PORT BODY WERE REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THIS SUPPLEMENTAL MDR IS BEING SUBMITTED TO REPORT THAT MFR RPT# 3006260740-2023-03106 WAS A DUPLICATE RECORD AND WAS OPENED IN ERROR. THE EVENT DETAILS ARE BEING CAPTURED UNDER COMPLAINT FILE #(B)(4) AND WAS REPORTED TO THE FDA UNDER MFR RPT# 3006260740-2023-03104. H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03100 | 2023-07-21 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT TWO MONTHS POST PORT PLACEMENT, RESISTANCE ALLEGEDLY FELT WHEN FLUSHING THE CATHETER. IT WAS FURTHER REPORTED THAT THE PORT WAS ALLEGEDLY LEAKED SALINE. IT WAS ALSO REPORTED THAT THE CATHETER ALLEGEDLY NOT ABLE TO BE FLUSHED AND WAS ALLEGEDLY RUPTURED. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03103 | 2023-07-21 | 2022-06-17 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SEVEN MONTHS AND TWENTY-SIX DAYS POST PORT PLACEMENT VIA THE RIGHT SUBCLAVIAN VEIN, THE PORT ALLEGEDLY DID NOT GIVE A BLOOD RETURN WHEN ACCESSED AND STOPPED FUNCTIONING. IT WAS FURTHER REPORTED THAT A FLUOROSCOPIC IMAGING REVEALED THAT A PIECE OF THE PORT HAD ALLEGEDLY BROKEN OFF AND WAS FOUND TO BE DISLODGED IN THE PATIENT. FURTHERMORE, THE BROKEN PIECE WAS ALLEGEDLY MIGRATED AND EMBOLIZED IN THE PATIENT'S RIGHT MAIN PULMONARY ARTERY. REPORTEDLY, AN EMERGENCY SURGERY WAS DONE TO REMOVE THE FRACTURED AND DISLODGED PORTION OF THE PORT AS WELL AS THE REMAINING PORTION OF THE DEVICE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3 H11: B5, D6 (MEDICAL DEVICE IMPLANT DATE) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03115 | 2023-07-21 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP IMPLANTABLE PORT, GROSHONG, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | CP00003 | CP00003 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE PORT WAS ACCESSED AND BLOOD RETURN WAS ALLEGEDLY UNABLE TO BE OBTAINED. UPON FLUSHING THE CATHETER, SWELLING WAS NOTED AN INCH ABOVE THE PORT SITE AND A PATIENT NOTED PAIN WITH FLUSH. IT WAS FURTHER REPORTED THAT UPON X-RAY EXAMINATION, THE CATHETER WAS FOUND TO BE FRACTURED. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 08/2019). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03102 | 2023-07-21 | 2023-06-13 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP IMPLANTABLE PORT, GROSHONG, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | CP00003 | CP00003 |  | IT WAS REPORTED THAT APPROXIMATELY SIX YEARS POST-PORT PLACEMENT, THE PORT WAS ACCESSED AND FLUSHED WITHOUT DIFFICULTY; HOWEVER, BLOOD RETURN WAS ALLEGEDLY UNABLE TO BE OBTAINED, AND THE PATIENT REPORTED SLIGHT PAIN WITH FLUSHING. IT WAS FURTHER REPORTED THAT A CHEST X-RAY AND VENOGRAM WERE PERFORMED, WHICH REVEALED THE RIGHT CHEST PORT CONNECTED TO A PORTION OF THE TUBING THAT EXTENDS TOWARDS THE INTERNAL JUGULAR VEIN. FURTHERMORE, THE MAJORITY OF THE TUBING WAS DISPLACED DISTALLY, REJECTING THE EXPECTED LOCATION OF THE RIGHT ATRIUM AND INFERIOR VENA CAVA. REPORTEDLY, THE CATHETER FRAGMENT WAS RETRIEVED, AND THE PORT WAS REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05419 | 2023-07-24 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THERE WAS A CRACK IN THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03129 | 2023-07-24 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 5393230 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE VIA THE RIGHT INTERNAL JUGULAR VEIN, THE CATHETER ALLEGEDLY KINKED. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 06/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05420 | 2023-07-24 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT FOLEY CATHETER HAD PINHOLE OR DIFFICULT TO DEFLATE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 10JUL2023, THE SAMPLE WAS CONFIRMED TO BE BALLOON RUPTURE DURING PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03122 | 2023-07-24 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT LONG-TERM HEMODIALYSIS CATHETER, POLYURETHANE, KIT, 14.5F | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 5703690 |  | IT WAS REPORTED THAT FOUR MONTHS AND THIRTEEN DAYS POST A PORT PLACEMENT, THE CATHETER ALLEGEDLY LEAKED BLOOD. IT WAS FURTHER REPORTED THAT A CRACK WAS ALLEGEDLY FOUND IN THE CATHETER BODY AT THE SKIN SURFACE OF PUNCTURE THE CATHETER. REPORTEDLY, THE CATHETER WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 12/2023), G3, H6 (METHOD) H11: H6 (DEVICE) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05424 | 2023-07-24 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT DURING A PRETEST, THE FIXATION WATER LEAKED FROM THE CONNECTION OF THE INFLATION VALVE OF THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05421 | 2023-07-24 | 2023-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED, THAT 13 DAYS AFTER PLACEMENT. THE FOLEY CATHETER BALLOON GOT TORN AND THE BALLOON PORTION DROPPED OFF, DUE TO THE PATIENT SELF-EXTRACTION. IT WAS STATED, THAT REQUESTED AN INVESTIGATION TO CONFIRM, WHETHER THE BALLOON PORTION PEEL WAS EXPECTED AS A FAILURE MODE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05412 | 2023-07-24 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53612G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER DID NOT HAD AS MUCH LUBE ON THEM AS THEY USUALLY DID. THEREFORE, IT CAUSED THE PATIENT A LOT OF PAIN WHILE INSERTING. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05425 | 2023-07-24 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE FOLEY CATHETER. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. PHOTO: THERE ARE 2 PHOTOS THAT WERE SENT BY THE CUSTOMER. THOSE PHOTOS SHOW AN OVERVIEW OF THE 2-WAY CATHETER. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 2-WAY CATHETER. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE SOLUTION IMMEDIATELY WENT INTO THE DRAINAGE LUMEN AT BIFURCATION AND CREATED LUMEN TO LUMEN LEAK. THIS IS OUT OF SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE LUMEN COMPROMISED BY A PUNCTURE OR CUT. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03128 | 2023-07-24 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, AIR WAS ALLEGEDLY ASPIRATED FROM THE INTRODUCER NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 01/2025) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05422 | 2023-07-24 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT DURING THE PRETEST, THERE WAS RESISTANCE WHILE INJECTING STERILE WATER INTO THE CATHETER AND AFTER SEVERAL ATTEMPTS, THE WATER WAS INJECTED BUT THE CATHETER WAS INFLATED ASYMMETRICALLY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05402 | 2023-07-24 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT WATER LEAKAGE OCCURRED FROM THE BALLOON OF THE FOLEY CATHETER. THE USED PRODUCTS HAD BEEN DISCARDED. THE SAME LOT OF PRODUCTS HAD BEEN COLLECTED FOR INVESTIGATION. THERE WERE MANY PROBLEMS WHICH HAS BEEN ALREADY REPORTED (23050320 AND 23010589). PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 18-JUL-2023, STATED THAT THE SAMPLE IN WHICH THE EVENT OCCURRED WAS DISCARDED AT THE FACILITY, BUT FOUR UNUSED SAMPLES FROM THE SAME LOT WERE RETURNED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 20-JUL-2023, CLARIFIED THAT THE ORIGINAL SAMPLE WAS DISCARDED AT THE FACILITY. DISCARDED DEVICE WAS USED ON THE PATIENT. UNUSED WAS SELECTED BECAUSE THE UNOPENED PRODUCT WAS RECEIVED AT THE REQUEST OF THE FACILITY TO INVESTIGATE THE SAME LOT OF PRODUCTS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: '[WARNINGS] 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES: SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE: (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS: URINARY-TRACT INFECTION HEMORRHAGE, HEMATURIA ALLERGY REACTION TO THE DEVICE CALCULUS FORMATION EDEMA PAIN DISCOMFORT INJURY OF BLADDER OR URETHRAL URETHRITIS, URINARY INCONTINENCE RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05423 | 2023-07-24 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT DURING THE PRETEST, THERE WAS RESISTANCE WHILE INJECTING STERILE WATER INTO THE CATHETER AND AFTER SEVERAL ATTEMPTS, THE WATER WAS INJECTED BUT THE CATHETER WAS INFLATED ASYMMETRICALLY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05410 | 2023-07-24 | 2023-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT LEAKAGE WAS OBSERVED AT THE ROOT OF THE INFLATION VALVE DURING PRE-TEST. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 11JUL2023, IT WAS STATED THAT CONFIRMED TEAR ON INFLATION FUNNEL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05418 | 2023-07-24 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THERE WAS A REGULAR PROBLEM WITH BARD FOLEY CATHETERS FOR A WHILE WITHIN THE INTENSIVE CARE DEPARTMENT. STATED THAT THEY WERE NOT DRAINING PROPERLY (BATCH# UNK) OR WERE CLOGGED (BATCH # UNK) AND THEREFORE THE PATIENTS GOT A FULL BLADDER. LAST WEEK, A PATIENT WHO HAD 1300 CC IN THE BLADDER AND HAD A CATHETER, WHEN REMOVED THE BALLOON LOOKED A BIT STRANGE (BATCH# UNK). IT WENT WELL FOR A WHILE, BUT LAST FALL THEY ALSO HAD THREE PATIENTS (BATCH# UNK) PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6)2023, STATED THAT THE BALLOON LOOKED A BIT STRANGE MEANT THAT THE BALLOON LOOKED ADHERED AND WAS ALSO BLOODY BUT THAT COULD ALSO BE FROM THE PATIENT. FOR THIS CASE, THEY HAD NOW SEEN ONLY ONE PATIENT BUT THE LOT NUMBER COULD NOT BE TRACED. AS FOR THE CATHETERS THAT WERE NOT DRAINING PROPERLY, THEY DID NOT KNOW THE LOT NUMBER OR THE NUMBER OF PATIENTS THAT HAD THIS ISSUE. HOWEVER, THEY OFTEN HAD TO SIPHON TO EMPTY THE BLADDER PROPERLY, COULD BE DUE TO THE CATHETER BUT ALSO THE URIMETER. AS FOR THE CATHETERS THAT WERE CLOGGED, THEY DID NOT HAD ANY IDEA ABOUT THE LOT NUMBER OR THE NUMBER OF PATIENTS THAT HAD THIS ISSUE. AS FOR THE PREVIOUS PROBLEMS, NONE WERE REPORTED EXCEPT FOR ONE THAT WAS. IN THAT PATIENT, IT WAS TYPICAL THAT THEY DID URINATE 30-40 CC, BUT HAD ABDOMINAL PAIN. THEY TOOK OUT THE CATHETER AND AT THAT POINT THE URINE POURED OUT AND THE ABDOMINAL PAIN DISAPPEARED. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05405 | 2023-07-24 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE NURSE STATED THAT THEY HAD A PATIENT ON THE ARCTIC SUN DEVICE. THEY JUST SWAPPED OUT THE DEVICE, THE FIRST DEVICE WAS NEEDING TO BE SERVICED. NURSE STATED THAT THEY PLUGGED IN THE FOLEY TEMPERATURE PROBE TO THE DEVICE, AND IT WAS READING 31.2C. ON THE PREVIOUS DEVICE THE FOLEY PROBE WAS READING 36.8C. NURSE STATED THAT THEY ALSO HAVE AN ESOPHAGEAL PROBE CONNECTED TO THE BEDSIDE MONITOR AND IT WAS CURRENTLY READING 36C. HAD NURSE FLEX THE TEMPERATURE CABLE, NURSE CONFIRMED THE PATIENT TEMPERATURE WAS JUMPING AROUND. INFORMED NURSE THE CABLE LIKELY HAD A SHORT IN IT. RECOMMENDED TO CHECK WITH BIOMED FOR AN EXTRA TEMPERATURE CABLE TO SWAP OUT OR SWAP THE CABLE WITH ANOTHER DEVICE IF NECESSARY. INFORMED NURSE TO CALL BACK WITH ANY ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05415 | 2023-07-24 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS WITHDRAWN SPONTANEOUSLY. IT WAS STATED THAT A CATHETER WAS PLACED IN AN ORAL SURGERY, AND WHEN ATTEMPTING TO REMOVE IT LATER IN THE GHCU, STERILE WATER COULD NOT BE DRAINED. AFTER SOME CONFUSION, THE CATHETER WAS WITHDRAWN SPONTANEOUSLY. WHEN CHECKED, THE BALLOON WAS DAMAGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03130 | 2023-07-24 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | CUSTOMER REPORTED BLOOD WITHDRAWN IN THE CATHETER, LEAKAGE AT THE CONNECTION. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03132 | 2023-07-24 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED IN THE USE OF THE SHEATH CANNOT BE INSERTED NORMALLY ALONG THE GUIDE WIRE IN THE PUNCTURE PROCESS, THE CUSTOMER RESPONSE IS THE END OF THE GUIDE WIRE HAS A METAL EXPANSION, THE GENERAL SOLUTION IS TO USE SCISSORS TO CUT THE END, BUT BECAUSE THE END OF THE METAL SITE STRUCTURE, OFTEN NEED TO BE CUT SEVERAL TIMES TO HAVE A SMOOTHER CUT CAN BE FED INTO THE SHEATH. IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05428 | 2023-07-24 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0119L22 |  | IT WAS REPORTED THAT WATER LEAKAGE OCCURRED FROM THE FOLEY CATHETER BALLOON SECTION (LOT# MYFU1698). IT WAS STATED THAT THE USED PRODUCTS HAD BEEN DISCARDED. THE CUSTOMER REQUESTED THAT ALL OTHER LOTS WITH THE SAME CATALOG NUMBER BE INVESTIGATED, SO THEY WERE COLLECTED. IT WAS STATED THAT DUE TO THE FREQUENT OCCURRENCE OF DEFECTS (OTHER EXTERNAL REFERENCE # (B)(4) AND OTHER EXTERNAL REFERENCE #23010589) THIS YEAR, REPRESENTATIVES RECEIVED A REQUEST TO CHECK ALL OTHER LOTS OF THE SAME PART NUMBER(LOT# MYFY3438, LOT# MYEW1667, LOT# MYFP2501), SO THEY COLLECTED THEM. IT WAS STATED THAT THERE WERE MANY PROBLEMS WITH THIS PRODUCT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03140 | 2023-07-24 | 2023-07-06 | I, n, j, u, r, y | Manufacturer report | PPICC HF 5F TL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3385335QJ |  | IT WAS REPORTED THAT THE PATIENT HAD A POWERPICC IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, THE CATHETER WAS FOUND TO BE BROKEN. THE DISTAL CATHETER SEGMENT WAS MIGRATED TO THE VESSEL AND REMOVED VIA THE FEMORAL VEIN. THE DOCTOR GUESSED THAT THE PATIENT'S DELIRIUM-INDUCED SELF-REMOVAL OF THE CATHETER WAS MORE LIKELY THAN A PROBLEM WITH THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BROKEN CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 5FR T/L POWERPICC CATHETER. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. THE CATHETER WAS RECEIVED IN TWO SEGMENTS WHICH SHARED A COMPLETE BREAK BETWEEN THE 6CM AND 7CM DEPTH MARKINGS. MATERIAL WHITENING WAS OBSERVED IN THE VICINITY OF THE BREAK. TACTILE INSPECTION OF THE SAMPLE REVEALED TENSILE WEAKNESS AND NECKING IN THE VICINITY OF THE BREAK. MICROSCOPIC INSPECTION OF THE BREAK SITE REVEALED A GRANULAR FRACTURE SURFACE. THE OUTER CATHETER WALL AND THE INNER WEB WALL EXHIBITED DEFORMATION AT THE BREAK SITE. THE FRACTURE FEATURES, TENSILE WEAKNESS AND DEFORMATION WERE TYPICAL OF MATERIAL FAILURE DUE TO PULLING STRESS. THESE FINDINGS WERE CONSISTENT WITH THE EVENT DESCRIPTION WHICH INDICATED THAT ATTEMPTED CATHETER REMOVAL BY THE PATIENT LIKELY CONTRIBUTED TO CATHETER BREAKAGE. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03141 | 2023-07-24 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT A PICC LINE WAS "LEAKING" AT THE TOP OF THE HUB WHERE THE LUMENS ARE IN. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03138 | 2023-07-24 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 24CM (SHORT-TERM) (DUAL-LUMEN) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5593240 |  | CUSTOMER REPORTED THE PATIENT UNDERWENT BLOOD PURIFICATION TREATMENT WITH FEMORAL VEIN HEMOFILTRATION CATHETER PUNCTURE AND CATHETERIZATION UNDER COLOR ULTRASOUND POSITIONING. THE GUIDE WIRE WAS DISCOUNTED DURING THE CATHETERIZATION PROCESS. REPLACING THE GUIDE WIRE AND PLACING THE TUBE AGAIN DELAYS TREATMENT AND MAY CAUSE SERIOUS INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03139 | 2023-07-24 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED CATHETER RUPTURE AT 6 AND 10 CENTIMETERS. INCIDENT OCCURRED DURING USE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05445 | 2023-07-24 | 2023-07-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 50616 |  | IT WAS REPORTED THAT AFTER USING THIS CATHETER PATIENT ENDED UP IN THE HOSPITAL AND HAD TO HAVE PROSTATE SURGERY. IT WAS STATED THAT THE PATIENT USED ONE OF THE SAMPLES SENT AND IT GAVE AN INJURY WHILE IN THE SHOWER RESULTING IN PATIENT LOSING ALLOT OF BLOOD AND PASSING OUT. THE PATIENT THOUGHT THEY MIGHT HAVE USED THE CATHETER WRONG CAUSING INJURY BUT NOT SURE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
3006260740-2023-03133 | 2023-07-24 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED IN THE USE OF THE SHEATH CANNOT BE INSERTED NORMALLY ALONG THE GUIDE WIRE IN THE PUNCTURE PROCESS, THE CUSTOMER RESPONSE IS THE END OF THE GUIDE WIRE HAS A METAL EXPANSION, THE GENERAL SOLUTION IS TO USE SCISSORS TO CUT THE END, BUT BECAUSE THE END OF THE METAL SITE STRUCTURE, OFTEN NEED TO BE CUT SEVERAL TIMES TO HAVE A SMOOTHER CUT CAN BE FED INTO THE SHEATH. IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05443 | 2023-07-24 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER WAS DIFFICULT TO REMOVE FROM THE BALLOON OF THE FOLEY CATHETER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. PHOTO: THERE ARE 2 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE PHOTOS SHOW AN OVERVIEW OF THE PARTIALLY INFLATED 2-WAY CATHETER. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED THAT THE BALLOON WAS PARTIALLY INFLATED UPON RECEIVAL OF SAMPLE. DEFLATED THE BALLOON PASSIVELY WITH NO CUFFING OR LEAKS NOTED. USING THE RETURNED SYRINGE, THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 3 MINUTES WITH NO CUFFING OR LEAKS NOTED, RETURNING 10ML OF SOLUTION. THIS IS WITHIN SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05435 | 2023-07-24 | 2023-06-26 | I, n, j, u, r, y | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165818 |  | IT WAS REPORTED THAT THE USER HAD EXPERIENCED ISSUES WITH INDWELLING URINARY CATHETER "FELLING OUT" TWICE SINCE USER¿S LAST CHANGE. IT WAS STATED THAT THE ON MONDAY DURING THE PROCEDURE OF CHANGING FOLEY CATHETER THE BALLOON BURST WHILE IN PROCESS OF FILLING BALLOON WITH THE SPECIFIED AMOUNT OF WATER. THIS RESULTED IN ANOTHER INDWELLING URINARY CATHETER NEEDING TO BE PLACED WITH RESULTING TRAUMA REQUIRING THE CLIENT TO ATTEND ED FOR REVIEW. IT WAS MY UNDERSTANDING THAT THIS CATHETER ALSO "FELL OUT" ON TUESDAY RESULTING IN CLIENT AGAIN NEEDING TO GO TO ED FOR ATTENTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 12JUL2023, IT WAS STATED THAT THE CUSTOMER HAD ALSO INFORMED THAT TWO OTHER CATHETERS FROM THAT SAME BOX HAD FAILED BY FALLING OUT IN THE DAYS PRIOR TO THE TWO CATHETERS. THERE WAS NO TRAUMA EXPERIENCED WITH THOSE FIRST TWO. THIS MEANS OUT OF A BOX OF TEN, SIX WERE USED OF WHICH FOUR ARE NOW REPORTED TO HAVE BEEN FAULTY. THE OTHER FOUR WERE STILL IN THE BOX READY FOR PICK UP, THERE WERE NO ACTUAL USED FAULTY SAMPLES AVAILABLE FOR PICK UP. IT WAS STATED THAT THE FIRST CATHETER WAS CHECKED AND THERE WERE NO MISSING PIECES. WHEN INSERTING THE SECOND CATHETER USER AND NURSE NOTICED BLOOD DRIPPING FROM THE EXTERNAL PART OF THE CATHETER AS WELL AS SEEPING AROUND THE CATHETER WERE INSERTED INTO PENNIS. THE USER WAS TREATED IN ED, THEY DID NOT REMOVE THE CATHETER, THEY BANDAGED UP PENNIS TO STOP BLEEDING AND KEPT USER OVERNIGHT FOR OBSERVATION BEFORE RELEASING THE NEXT DAY. THE CIRCUMSTANCES SURROUNDING THE SECOND CATHETER FALLING OUT WAS USER FELT LIKE THERE WAS ONLY A SMALL AMOUNT OF URINE IN BAG WHEN AT HOME SO CHECKED AND FOUND THAT THE CATHETER HAD FALLEN OUT. THE USER THEN RETURNED TO ED WHERE THE NURSE CHECKED THE CATHETER AND DETERMINED THE BALLOON WOULD NOT INFLATE. THERE WAS NO INDICATION ANY PIECES WERE LEFT INSIDE. THE NURSE THEN TESTED THE BALLOON OF A NEW DIFFERENT BRANDED CATHETER BEFORE INSERTION AND RETURNED HOME WITHOUT INCIDENT OF THAT CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05409 | 2023-07-24 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 33428 |  | IT WAS REPORTED THAT THERE WAS A HAIR INSIDE THE CLOSED PACKAGING OF THE FOLEY CATHETER. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. THE DEVICE HAD NOT MET SPECIFICATIONS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE UNOPENED (WITH ORIGINAL PACKAGING), SILASTIC FOLEY CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED NO VISIBLE NONCONFORMITIES. A STRAND OF HAIR WAS OBSERVED INSIDE THE UNOPEN PACKAGE. THIS DOES NOT MEET SPECIFICATION WHICH STATES THAT " LOOSEN HAIR IN PACKAGE AND/OR TRAPPED IN SEALING IS NOT ALLOWED". A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05440 | 2023-07-24 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 |  | INTP14 |  | IT WAS REPORTED THAT INTERMITTENT CATHETER WAS PARTIALLY PACKAGED IN BETWEEN THE FOLDED LAYERS OF THE STERILE WRAPPER. UNABLE TO PROPERLY UNWRAP PACKAGE IN A STERILE MANNER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT LINE CLEARANCE / INADEQUATE WORK STATION FOR OPERATION". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW IS NOT REQUIRED BECAUSE A REVIEW OF THE LABEL COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03134 | 2023-07-24 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED IN THE USE OF THE SHEATH CANNOT BE INSERTED NORMALLY ALONG THE GUIDE WIRE IN THE PUNCTURE PROCESS, THE CUSTOMER RESPONSE IS THE END OF THE GUIDE WIRE HAS A METAL EXPANSION, THE GENERAL SOLUTION IS TO USE SCISSORS TO CUT THE END, BUT BECAUSE THE END OF THE METAL SITE STRUCTURE, OFTEN NEED TO BE CUT SEVERAL TIMES TO HAVE A SMOOTHER CUT CAN BE FED INTO THE SHEATH. IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05441 | 2023-07-24 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED, THAT THE WATER IN THE FOLEY CATHETER BALLOON DID NOT DRAIN DURING PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05430 | 2023-07-24 | 2023-06-29 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD DEVELOPED IRRITATION AND SORENESS IN THEIR VAGINAL AREA FROM PUREWICK FEMALE EXTERNAL CATHETER. THE REPRESENTATIVE ADVISED THEM TO STOP USING IT AND SEE A DOCTOR. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW UP VIA PHONE ON 05JUL2023, IT WAS REPORTED THAT THE CUSTOMER CONTACTED THEIR PHYSICIAN AND WAS ADVISED TO USE SALVE TO ASSIST WITH SORENESS AND IRRITATION. THE CUSTOMER WAS STILL USING THE WICKS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE OR MATERIAL SURFACE IS ROUGH, ABRASIVE OR UNCOMFORTABLE ". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: " AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER". "TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL". "DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS". "EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE". RECOMMENDATIONS: ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS". NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. PERI-CARE AND PLACEMENT: PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05437 | 2023-07-24 | 2023-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE NURSE STATED THE PATIENT¿S TEMPERATURE WAS LOWER THAN THE PADS. THEY REPLACED THE TEMPERATURE PROBE AND NOW IT WAS WORKING. IT WAS CONFIRMED WITH NURSE THEY DID NOT NEED ANY ASSISTANCE AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05438 | 2023-07-24 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT NURSE STATED THAT THEY HAVE A PATIENT COOLING ON THE DEVICE, THERE WAS ABOUT 40 MINUTES WAS REMAINING IN COOLING THERAPY. NURSE STATED THE PATIENTS TEMPERATURE HAD DROPPED BELOW TARGET AND THE DEVICE STARTED ALARMING AND TURNED OFF DUE TO BEING IN BELOW TARGET. THE TARGETED TEMPERATURE WAS 36C, PATIENT TEMPERATURE 34.4C WITH A FOLEY PROBE. THEY HAD CHECKED A TEMPORAL TEMPERATURE WHICH WAS 34.9C BUT THEY WERE GOING TO ALSO VERIFY A RECTAL TEMPERATURE. NURSE CONFIRMED THEY HAD FOUR PADS CONNECTED WITH ADEQUATE COVERAGE. NURSE CONFIRMED THEY HAD NOT ADMINISTERED ANY MEDICATIONS RECENTLY. THEY HAD NOT FLUSHED OR IRRIGATED THE FOLEY. THEY DID RECENTLY TURN THE PATIENT FOR AN X-RAY. NURSE STATED PATIENT WAS SHIVERING, CURRENTLY WAS NOT NOTICEABLY SHIVERING. HAD NURSE ACCESSED THE EVENT LOG, DEVICE ALARMED 50 AND 15 AROUND 2235. INFORMED NURSE WITHOUT THERAPY RUNNING, IT WAS HARD TO DETERMINE WHAT CAUSE THE PATIENTS TEMPERATURE TO DROP. IT WAS DISCUSSED IF THEY RESUMED THERAPY UNDER COOLING, THE DEVICE WOULD ATTEMPT TO REWARM PATIENT FROM 34.4C TO THE TARGETED TEMPERATURE OF 36 AS QUICKLY AS POSSIBLE. EXPLAINED THE NURSE MIGHT WANT TO CONSULT THE PROVIDER TO CONFIRMED IF THAT WAS OKAY OR IF THEY SHOULD DO A CONTROLLED REWARM. NURSE CONFIRMED WITH PROVIDER AT BEDSIDE TO REWARM THE PATIENT TO 37C AT A RATE OF 0.2C/HOUR. WALKED NURSE THROUGH ADJUSTING THE REWARM TO REWARM FROM THE CURRENT PATIENT TEMPERATURE AND STARTED THERAPY. AFTER A FEW MINUTES WATER FLOW RATE SETTLED TO 3.1 L/MIN, WATER TEMPERATURE 24.9C, SYSTEM DIAGNOSTICS: OUTLET MONITOR TEMPERATURE 22.6C, OUTLET CONTROL TEMPERATURE 22.6C, INLET TEMPERATURE 22.7C, CHILLER TEMPERATURE 6.7C, WATER FLOW RATE 3 L/MIN, IP -7PSI, CIRCULATION PUMP COMMAND 69PERCENTAGE, MIXING PUMP COMMAND 18PERCENTAGE, HEATER WAS 0PERCENTAGE, WATER RESERVOIR LEVEL 4, SYSTEM HOURS 82, AND PUMP HOURS 56. WALKED NURSE THROUGH PLACED DEVICE IN MANUAL CONTROL TO WARM WATER TO 40C, AFTER ABOUT 8 MINUTES, WATER TEMPERATURE REACHED 36.6C. HAD NURSE DISABLE MANUAL CONTROL AND RESUMED REWARMING, AT THIS POINT PATIENT TEMPERATURE WAS 35.4C. EXPLAINED A TEMPERATURE CHANGE OF ONE DEGREE WAS UNUSUAL. INFORMED NURSE TO FLEX THE TEMPERATURE CABLE, PATIENT TEMPERATURE DID NOT FLUCTUATE. INFORMED NURSE IT WAS POSSIBLE WHEN THEY TURNED THE PATIENT THE FOLEY PROBE WAS DAMAGED, SUGGESTED NURSE CONFIRMED A SECONDARY RECTAL TEMPERATURE TO CONFIRMED CORRELATION. RECOMMENDED THEY MIGHT CONSIDER CHANGING THE TEMPERATURE SENSING FOLEY TO ENSURE IT WAS WORKING PROPERLY. ALSO DISCUSSED WITH NURSE COUNTER WARMING MEASURES PER PROTOCOL SUCH AS A BAIR HUGGER SINCE THE PATIENT WAS GENERATING HEAT AS THIS MAY AFFECT REWARMING ABILITY. NURSE STATED THEY WOULD CHECK A RECTAL TEMPERATURE AND CHANGE THE FOLEY AND THEN MONITOR THE PATIENT FOR AWHILE TO CONFIRMED REWARMING AT 0.2C/HOUR, NURSE ALSO CHANGED REWARM FROM TO THEN MATCH CURRENT PATIENT TEMPERATURE OF 35.4C. NURSE WOULD THEN CALL BACK IF CONTINUING TO HAD ISSUES. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "SENSOR WIRE TOO WEAK." THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03135 | 2023-07-24 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED IN THE USE OF THE SHEATH CANNOT BE INSERTED NORMALLY ALONG THE GUIDE WIRE IN THE PUNCTURE PROCESS, THE CUSTOMER RESPONSE IS THE END OF THE GUIDE WIRE HAS A METAL EXPANSION, THE GENERAL SOLUTION IS TO USE SCISSORS TO CUT THE END, BUT BECAUSE THE END OF THE METAL SITE STRUCTURE, OFTEN NEED TO BE CUT SEVERAL TIMES TO HAVE A SMOOTHER CUT CAN BE FED INTO THE SHEATH. IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE FOURTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05436 | 2023-07-24 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT DUE TO A SMALL HOLE IN THE BALLOON ABOUT 2 HOURS AFTER THE SURGERY. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, THE BALLOON APPEARED TO BE BURST IN THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03136 | 2023-07-24 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED IN THE USE OF THE SHEATH CANNOT BE INSERTED NORMALLY ALONG THE GUIDE WIRE IN THE PUNCTURE PROCESS, THE CUSTOMER RESPONSE IS THE END OF THE GUIDE WIRE HAS A METAL EXPANSION, THE GENERAL SOLUTION IS TO USE SCISSORS TO CUT THE END, BUT BECAUSE THE END OF THE METAL SITE STRUCTURE, OFTEN NEED TO BE CUT SEVERAL TIMES TO HAVE A SMOOTHER CUT CAN BE FED INTO THE SHEATH. IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE FIFTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05442 | 2023-07-24 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123618C |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY RETENTION. AT 8:30 ON (B)(6) 2023, THE PATIENT WAS OPERATED NORMALLY AND CHECKED THAT THE FOLEY CATHETER WAS NORMAL. AFTER THE CATHETER WAS INSERTED, URINE LEAKAGE WAS FOUND, AND THE CATHETER WAS REMOVED IMMEDIATELY AND NO MISSING PIECES. THEN REPLACED THE CATHETER WITH A NEW ONE, AND THE LEAKAGE OF URINE DID NOT OCCUR AGAIN AFTER IT WAS INSERTED, WHICH CAUSED PAIN FOR THE PATIENT DUE TO THE SECOND CATHETERIZATION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE BAD FIT WITH SHAFT/NO DRAINAGE EYE/ POORLY SEATED BALLOON /BLADDER NECK INTERFACE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
17376024 | 2023-07-24 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC SOLO CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 1022786 | 1194108D |  | PATIENT ARRIVED IN ER AT ANOTHER HOSPITAL WITH A LEAKING PICC (PERIPHERALLY INSERTED CENTRAL CATHETER) LINE, NO CAP IN PLACE. PATIENT STATES THAT IT WAS REMOVED BY THE HOME CARE SERVICES BECAUSE THE NEW SOLO PICCS ARE MADE TO NO LONGER REQUIRED A PRESSURIZED CAP. THE VAT (VASCULAR ACCESS TEAM) AT THIS FACILITY WAS NOTIFIED, AND ADVISED THEY COULD NOT GET THE PICC LINE TO STOP LEAKING BLOOD WITHOUT A CAP. PATIENT TRANSFERRED HERE FOR VAT TO PLACE A NEW PICC LINE. AFTER REVIEW, IT APPEARS THE VALVE ON THE PICC FAILED. THIS IS THE 2ND INCIDENT WITH THIS PARTICULAR PICC AND LOT# TODAY. THIS LINE WAS ORIGINALLY PLACED WHILE THE PATIENT WAS INPATIENT LAST MONTH. |  | N | Y |  |  | 2023-07
3006260740-2023-03137 | 2023-07-24 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174355 |  | IT WAS REPORTED BY THE CUSTOMER INTRODUCER WITH MALFORMATION ON GRAY MAKING IT INOPERABLE. LONG TERM ANTIBIOTICS SO PICC LINE NEEDED. SKIN KNICK GIVEN BEFORE INTRODUCER WAS ATTEMPTED AND IT WOULD NOT PASS OVER THE WIRE AS THERE WAS A MALFORMATION OF THE PLASTIC. ADDITIONAL INFORMATION RECEIVED 13 JULY, 2023: IT WAS STATED THE INTRODUCER WOULDN¿T THREAD INTO PATIENT BECAUSE OF THE BURR ON THE END OF IT. THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. THE ATTEMPTS MADE TO THREAD INTRODUCER FAILED AND TRAUMA IN TRYING TO USE IT AS WE NORMALLY WOULD MAY HAVE CAUSED SOME DAMAGE AND I¿M SURE INCREASED THE RISK OF DVT. WE WERE ABLE TO DROP A INTRODUCER KIT ONTO STERILE FIELD AND USE A NEW INTRODUCER WHICH THREAD EASILY INTO PATIENT AND ALLOWED US TO FINISH HIS PICC LINE INSERTION AT BEDSIDE. THE EVENT DID NOT INTERRUPT THE ADMINISTRATION OF ANY MEDICATION, OR CAUSE A CLINICALLY SIGNIFICANT DELAY IN MEDICATION, THAT NEGATIVELY IMPACTED THE PATIENT. EVENT OCCURRED ONCE DURING PLACEMENT OF PICC LINE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03125 | 2023-07-24 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1806060J |  | IT WAS REPORTED THAT SOMETIMES POST A PORT PLACEMENT, IT WAS ALLEGEDLY DIFFICULT TO GET BLOOD RETURN THROUGH THE PORT. IT WAS FURTHER REPORTED THAT WHEN THE PORT SYSTEM WAS FLUSHED WITH SALINE SOLUTION, THE PATIENT COMPLAINED OF PAIN AROUND THE PORT. BECAUSE THERE WAS A POSSIBILITY OF LEAKAGE OF ANTICANCER AGENTS, THE PATIENT WAS TREATED WITH NECROTIC DRUGS. REPORTEDLY, THE PORT SYSTEM WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE MRI POWERPORT IMPLANTABLE PORT ATTACHED TO A CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A LONGITUDINAL SPLIT WAS NOTED APPROXIMATELY 5.0CM FROM THE DISTAL END OF THE CATH-LOCK. A LEAK FROM THE LONGITUDINAL SPLIT WAS NOTED UPON INFUSION. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED FRACTURE ISSUE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED SUCTION ISSUE AS THE EXACT CIRCUMSTANCE AT THE TIME OF THE EVENT REPORTED WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: (EXPIRATION DATE: 12/2022). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03118 | 2023-07-24 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 9808560 |  | IT WAS REPORTED THAT THREE YEARS, TEN MONTHS AND TWENTY DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN FOR CHEMOTHERAPY, A CONTRAST COMPUTED TOMOGRAPHY EXAMINATION WAS PERFORMED THROUGH THE PORT, THE PATIENT ALLEGEDLY COMPLAINED OF NECK PAIN AND LEAKAGE OF CONTRAST MEDIUM ALLEGEDLY OCCURRED IN A SUBCUTANEOUS TUNNEL. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. MULTIPLE BREAKS WERE NOTED TO THE CATHETER PROXIMAL TO THE CATH-LOCK. A PARTIAL COMPOUND CIRCUMFERENTIAL BREAK WAS NOTED FROM THE DISTAL END OF THE CATH-LOCK. THE EDGES OF THE PARTIAL COMPOUND BREAK WERE NOTED TO BE UNEVEN AND THE SURFACE WAS OBSERVED TO BE GRANULAR AND DEGRADED. UPON INFUSION, WATER WAS NOTED EXITING THE PARTIAL COMPOUND BREAK. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED LEAK AND THE IDENTIFIED FRACTURE AND DEGRADED ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 10/2021), G3, H6 (DEVICE). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03119 | 2023-07-24 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS TEN MONTHS AND ONE DAY POST PORT PLACEMENT FOR OVARIAN CANCER CHEMOTHERAPY, UPON X-RAY EXAMINATION THE CATHETER WAS ALLEGEDLY BROKEN AND SEPARATED. IT WAS FURTHER REPORTED THAT THE DISTAL CATHETER SEGMENT WAS STILL INSIDE THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
17375761 | 2023-07-24 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP FOLEY TRAY SYSTEM BARDEX LUBRICATH FOLEY CATHETER | CATHETER, COUDE | C. R. BARD, INC. | A899916 | A899916 |  | PATIENT SEEN IN OUTPATIENT INFUSION CLINIC FOR GEMCITABINE MAINTENANCE BLADDER INSTILLATION. MEDICATION ADMINISTERED VIA FOLEY TO BLADDER FOR 1-HOUR DWELL TIME. AFTER DWELL COMPLETED, FOLEY WAS UNCLAMPED AND DRAINED INTO BAG. FOLEY BAG WAS OBSERVED TO BE LEAKING AT CONNECTION BETWEEN DRAINAGE BAG AND CLAMPED GREEN TUBE FOR EMPTYING THE BAG. CHEMO SPILL KIT PROCESS INITIATED AND FOLLOWED WITH SPILL SITE CLEAN, AND ALL MATERIAL DISPOSED OF APPROPRIATELY. NO HARM TO PATIENT OR STAFF. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03168 | 2023-07-25 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THE PATIENT ATTENDED FOR CHEMOTHERAPY TREATMENT. PICC LINE ACCESSED AND VENOUS RETURN ACHIEVED. COMMENCED ON PRE-HYDRATION WITH FLUIDS OVER 2 HOURS. PATIENT COMPLAINED OF PAIN AT PICC LINE DURING INFUSION. INFUSION STOPPED AND PICC LINE RECHECKED FOR VENOUS RETURN. NO VENOUS RETURN ACHIEVED. WARD DOCTOR CONTACTED FOR REVIEW. CHEST X-RAY REQUESTED BY WARD DOCTOR. PATIENT ATTENDED FOR X-RAY. REVIEWED BY WARD DOCTOR. X-RAY OF PICC SHOWED PICC IN POSITION. INFUSION RE STARTED . PATIENT FURTHER COMPLAINED OF PAIN. WARD DOCTOR REVIEWED PICC LINE SITE, PICC LINE REMOVED. PICC LINE SENT TO INFUSIONAL SERVICES FOR INVESTIGATION AND NIPAC DOCUMENTATION REVIEWED. PATIENT COMPLAINED OF PAIN IN RIGHT ARM AND SWELLING NOTED DURING INFUSION. INFUSION STOPPED. LIKELY FRACTURE AND UNABLE TO BOOK LINOGRAM THAT DAY. PICC REMOVED AND FRACTURE AT 6.5CM NOTED. PATIENT ATTENDING FOR LINOGRAM NEXT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05462 | 2023-07-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V24S |  | IT WAS REPORTED THAT WHEN THE PATIENT TURNED OVER, THE BALLOON OF THE FOLEY CATHETER BURST. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿SHORT LATEX DWELL TIME" OR "LATEX DIP SPEED OUT TOO SLOW". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03165 | 2023-07-25 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS PRIMED PRIOR TO THE PLACEMENT PROCEDURE, IT WAS NOTICED THAT THE CATHETER WAS DAMAGED IN A PINHOLE SHAPE. THERE WAS NO REPORTED PATIENT INVOLVEMENT. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03161 | 2023-07-25 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED AFTER PLACED, THE CATHETER FAILED TO FLOW / REFLOW, IT LOOKED LIKE IT WAS OBSTRUCTED. THE CATHETER WAS REPLACED BY ANOTHER ONE. ONCE THE CATHETER WAS REPLACED, THE DEVICE WORKS AS EXPECTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03159 | 2023-07-25 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY CUSTOMER THAT A NURSE PLACING AN ACCUCATH IN THE ANTECUBITAL WAS UNABLE TO REMOVE THE DEVICE AFTER THE FAILED ATTEMPT, DUE TO THE GUIDEWIRE NOT RETRACTING, AND WAS MET WITH RESISTANCE. NO OTHER INFORMATION WAS PROVIDED. | BECTON DICKINSON AND COMPANY'S (BD) MEDICATION DELIVERY SOLUTIONS (MDS) BUSINESS UNIT WILL DISCONTINUE MALFUNCTION MDR REPORTING FOR CERTAIN DEVICE FAILURES THAT HAVE NOT CAUSED OR CONTRIBUTED TO DEATHS OR SERIOUS INJURIES IN THE PAST TWO YEARS, AND WHERE THE LIKELIHOOD OF A DEATH OR SERIOUS INJURY AS A RESULT OF THESE MALFUNCTIONS IS REMOTE. THIS DECISION FOLLOWS FDA GUIDELINES (MEDICAL DEVICE REPORTING FOR MANUFACTURERS GUIDANCE FOR INDUSTRY AND FOOD AND DRUG ADMINISTRATION STAFF ISSUED NOV 8, 2016, REFERENCE SECTION 2.15). BD NOTIFIED FDA OF THIS DECISION ON MAR 3, 2025. FDA HAS REVIEWED AND RESPONDED TO BD¿S NOTIFICATION (REFERENCE FDA DOCUMENT # (B)(4)) ON MARCH 7, 2025. THIS DOCUMENTATION IS AVAILABLE IN BD DOCUMENT MANAGEMENT SYSTEM (SAP). THIS SUPPLEMENTAL MDR IS BEING FILED TO DOCUMENT THAT THE BELOW DEVICE FAILURE WILL NO LONGER BE CONSIDERED A REPORTABLE MALFUNCTION MDR FOR THE PRODUCT FAMILY BELOW: PRODUCT FAMILY: ACCCATH ACE. DEVICE FAILURE: CATHETER DIFFICULT TO REMOVE. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05457 | 2023-07-25 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE PACKAGING OF THE FOLEY CATHETER APPEARED TO BE DAMAGED DURING INSPECTION AND WAS NOT USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03160 | 2023-07-25 | 2023-05-19 | I, n, j, u, r, y | Manufacturer report | GROSHONG 4F SL PICC PERIPHERALLY INSERTED CENTRAL VENOUS CATHETER | CATHETER, PERCUTANEOUS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7715405 |  | CUSTOMER REPORTED THE PATIENT COMPLAINED OF PAIN AT THE PUNCTURE DURING TREATMENT AND FOUND THAT THE LENGTH OF THE BROKEN CATHETER WAS ABOUT 8 CM DURING DRESSING CHANGE, SO THE INTERVENTIONAL DEPARTMENT WAS CONTACTED FOR ARREST. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05446 | 2023-07-25 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 175816 |  | IT WAS REPORTED THAT THE CUSTOMER EXPERIENCED A PRODUCT COMPLAINT WITH PRODUCT NUMBER A942216. THE FOLEY CATHETER FELL OUT OF THE PATIENT AND THE NURSE RECOGNIZED THAT THE TIP WAS MISSING (NGHQ2185). THE CUSTOMER SAVED THE CATHETER, AND IT WAS AVAILABLE TO BE RETURNED. NO PATIENT COMPLICATION WAS REPORTED. THE CUSTOMER WAS CONCERNED ABOUT THE FOLEY TRAYS AND OPEN LOT NUMBER, THEY NOTICED OTHER ISSUES WITH FOLEY CATHETERS, KINK TUBING (NGHG1002) AND CUT FOLEY CATHETER (NGGZ1611). THESE FOLEY CATHETERS WERE NOT INSERTED IN PATIENTS. THEY HAD SAMPLES OF THESE TRAYS TOO. THEY QUARANTINED LOT NUMBER ALL THE THREE LOTS WHICH HAD BEEN REMOVED FROM CIRCULATION. AFTER SPEAKING WITH CUSTOMER, THEY PULLED ABOUT 3 CASES AND INDIVIDUAL FOLEY TRAYS FROM THE FLOOR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05452 | 2023-07-25 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THERE WAS A REGULAR PROBLEM WITH BARD FOLEY CATHETERS FOR A WHILE WITHIN THE INTENSIVE CARE DEPARTMENT. STATED THAT THEY WERE NOT DRAINING PROPERLY (BATCH# UNK) OR WERE CLOGGED (BATCH # UNK) AND THEREFORE THE PATIENTS GOT A FULL BLADDER. LAST WEEK, A PATIENT WHO HAD 1300 CC IN THE BLADDER AND HAD A CATHETER, WHEN REMOVED THE BALLOON LOOKED A BIT STRANGE (BATCH# UNK). IT WENT WELL FOR A WHILE, BUT LAST FALL THEY ALSO HAD THREE PATIENTS (BATCH# UNK) PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE BALLOON LOOKED A BIT STRANGE MEANT THAT THE BALLOON LOOKED ADHERED AND WAS ALSO BLOODY BUT THAT COULD ALSO BE FROM THE PATIENT. FOR THIS CASE, THEY HAD NOW SEEN ONLY ONE PATIENT BUT THE LOT NUMBER COULD NOT BE TRACED. AS FOR THE CATHETERS THAT WERE NOT DRAINING PROPERLY, THEY DID NOT KNOW THE LOT NUMBER OR THE NUMBER OF PATIENTS THAT HAD THIS ISSUE. HOWEVER, THEY OFTEN HAD TO SIPHON TO EMPTY THE BLADDER PROPERLY, COULD BE DUE TO THE CATHETER BUT ALSO THE URIMETER. AS FOR THE CATHETERS THAT WERE CLOGGED, THEY DID NOT HAD ANY IDEA ABOUT THE LOT NUMBER OR THE NUMBER OF PATIENTS THAT HAD THIS ISSUE. AS FOR THE PREVIOUS PROBLEMS, NONE WERE REPORTED EXCEPT FOR ONE THAT WAS. IN THAT PATIENT, IT WAS TYPICAL THAT THEY DID URINATE 30-40 CC, BUT HAD ABDOMINAL PAIN. THEY TOOK OUT THE CATHETER AND AT THAT POINT THE URINE POURED OUT AND THE ABDOMINAL PAIN DISAPPEARED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: SALT ACCUMULATION)/BLOCK DRAINAGE LUMEN/NO DRAINAGE EYE.). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03164 | 2023-07-25 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS PRIMED PRIOR TO THE PLACEMENT PROCEDURE, IT WAS NOTICED THAT THE CATHETER WAS DAMAGED IN A PINHOLE SHAPE. THERE WAS NO REPORTED PATIENT INVOLVEMENT. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05458 | 2023-07-25 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT SPONTANEOUS WITHDRAWAL OF THE FOLEY CATHETER WAS CONFIRMED DURING DIAPER CHANGE. HOWEVER, WHEN THE USER TRIED TO INFLATE THE BALLOON OUTSIDE THE BODY, IT WAS INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05459 | 2023-07-25 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT WHEN STERILE WATER WAS INJECTED TO CHECK THE BALLOON, IT LEAKED FROM THE BALLOON. SO, THE FOLEY CATHETER WAS NOT USED. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. PHOTO: THERE ARE 3 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE FIRST PHOTO ATTACHED IS AN OVERVIEW OF THE 3-WAY CATHETER AND SYRINGE. THE SECOND PHOTO IS A CLOSER LOOK AT THE CATHETER BALLOON AND TIP. THE THIRD PHOTO IS A FOREIGN DOCUMENT. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND IT IMMEDIATELY LEAKED OUT OF A 0.012" PINHOLE FOUND ON THE BALLOON. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION CATHETER KINKING, DAMAGE, RUPTURE DIFFICULTY OR FAILURE TO REMOVE THE DEVICE OCCLUSION OF CATHETER INNER LUMENS ENCRUSTATION ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE DEVICE DAMAGE DUE TO INAPPROPRIATE USE FAILURE TO MEASURE TEMPERATURE IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS URINARY-TRACT INFECTION HEMORRHAGE, HEMATURIA ALLERGY REACTION TO THE DEVICE CALCULUS FORMATION EDEMA PAIN DISCOMFORT INJURY OF BLADDER OR URETHRAL URETHRITIS, URINARY INCONTINENCE RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ CORRECTIONS: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05455 | 2023-07-25 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER WAS DIFFICULT TO REMOVE FROM THE BALLOON OF THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05449 | 2023-07-25 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119108 |  | IT WAS REPORTED THAT NO OUTFLOW OF URINE AFTER CATHETER INSERTION. WHEN THE CATHETER WAS REMOVED AND WATER WAS INJECTED, IT COULD NOT BE DRAINED. POSSIBLE BLOCKAGE AT THE TIP OF THE CATHETER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. AS THE REPORTED EVENT WAS UNCONFIRMED, A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED. HOWEVER, ONE WAS COMPLETED BEFORE UNCONFIRMING THE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. AS THE REPORTED EVENT WAS UNCONFIRMED A LABELING REVIEW WAS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05454 | 2023-07-25 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE ON (B)(6)2023, A FOLEY CATHETER WAS INDWELLED FOR CONTINUOUS BLADDER IRRIGATION AFTER THE OPERATION. ON (B)(6)2023, THE URINARY CATHETER FELL OUT DUE TO THE RUPTURE OF THE URETERAL WATER SAC. AFTER REPORTING TO THE DOCTOR, A NEW FOLEY CATHETER WAS USED FOR THE PATIENT. IT WAS STATED THAT THE THIS OPERATION INCREASED THE PATIENT'S PAIN. NO MEDICAL INTERVENTION WAS REPORTED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6)2023, IT WAS STATED THAT THE NO INJURY CAUSED ON PATIENT AND THERE WAS NO PAIN FOUND IN THE PATIENT. IT WAS STATED THAT THE PATIENT DID NOT CONTRIBUTED PAIN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03169 | 2023-07-25 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, IT WAS FOUND THAT THE STYLET WAS PROTRUDING FROM THE CATHETER. IT IS PRESUMED THAT IT IS DUE TO THE STYLET CUT. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05450 | 2023-07-25 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 |  | FOL0102 |  | IT WAS REPORTED THAT THE WHITE CLIP ON STATLOCK WAS COMING APART AND FALLING OFF. STATED THAT THE STATLOCK WAS BEGINNING TO PEEL OFF QUITE EASILY. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 24JUL2023, STATED THAT THE DEFECT WAS FOUND AFTER USE ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
3006260740-2023-03162 | 2023-07-25 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY CUSTOMER THAT THE CATHETER THAT LEAKED FROM THE RUBBER-STOPPER WHEN PICC LINE WAS BEING INSERTED. ADDITIONAL INFORMATION FROM CUSTOMER ON 07-12-2023 THIS EVENT DID NOT INVOLVE A LIFE THREATENING SITUATION. THE LEAK WAS DISCOVERED DURING INSERTION. NO PATIENT HARM REPORTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03189 | 2023-07-25 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICC LEAKING. PATIENT ATTENDED FOR ASSESSMENT, HE STATED HE HAD FELT SOMETHING AT HIS SHOULDER WHEN THE LINE WAS FLUSHED. LINE REMOVED, HOLE PRESENT AT 23CM (SO INTERNAL). PICC USED FOR AVD TREATMENT. PICC REPLACED ON 07/07/2023. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03185 | 2023-07-25 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PPICC SOLO 5F DL FT 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | IT WAS REPORTED BY CUSTOMER THAT INTRODUCERS WILL NOT SLIDE INTO THE SKIN PROPERLY. CAUSING THE INTRODUCERS TO BEND AND IN SOME CASES THE SKIN TO BRUISE. CLINICIAN HAD TO OPEN ANOTHER KIT. ADDITIONAL INFORMATION RECEIVED 12 JULY 2023: IT WAS REPORTED BY THE CUSTOMER "EVENT DID NOT INVOLVE AN URGENT MEDICAL SITUATION. NO PATIENT HARM REPORTED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03181 | 2023-07-25 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 18GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1182250 |  | IT WAS REPORTED BY CUSTOMER THAT CATHETER HAD A HOLE IN IT FOLLOWING CT SCAN WITH CONTRAST. NO HARM TO PATIENT. NO OTHER INFORMATION WAS PROVIDED. | BECTON DICKINSON AND COMPANY'S (BD) MEDICATION DELIVERY SOLUTIONS (MDS) BUSINESS UNIT WILL DISCONTINUE MALFUNCTION MDR REPORTING FOR CERTAIN DEVICE FAILURES THAT HAVE NOT CAUSED OR CONTRIBUTED TO DEATHS OR SERIOUS INJURIES IN THE PAST TWO YEARS, AND WHERE THE LIKELIHOOD OF A DEATH OR SERIOUS INJURY AS A RESULT OF THESE MALFUNCTIONS IS REMOTE. THIS DECISION FOLLOWS FDA GUIDELINES (MEDICAL DEVICE REPORTING FOR MANUFACTURERS GUIDANCE FOR INDUSTRY AND FOOD AND DRUG ADMINISTRATION STAFF ISSUED NOV 8, 2016, REFERENCE SECTION 2.15). BD NOTIFIED FDA OF THIS DECISION ON MAR 3, 2025. FDA HAS REVIEWED AND RESPONDED TO BD¿S NOTIFICATION (REFERENCE FDA DOCUMENT # (B)(4)) ON MARCH 7, 2025. THIS DOCUMENTATION IS AVAILABLE IN BD DOCUMENT MANAGEMENT SYSTEM (SAP). THIS SUPPLEMENTAL MDR IS BEING FILED TO DOCUMENT THAT THE BELOW DEVICE FAILURE WILL NO LONGER BE CONSIDERED A REPORTABLE MALFUNCTION MDR FOR THE PRODUCT FAMILY BELOW: PRODUCT FAMILY: ACCCATH ACE. DEVICE FAILURE: EXTERNAL BREAK / CRACK / SPLIT. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03179 | 2023-07-25 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROVPPSOLO 4FDL MBP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1294108D4 |  | IT WAS REPORTED BY CUSTOMER THAT PATIENT PULLED OUT PICC. WHEN PATIENT PULLED PICC OUT HE PULLED ONE OF THE EXTENSION LEGS OFF. NO PATIENT HARM. ADDITIONAL INFORMATION FROM THE CUSTOMER ON (B)(6) 2023: IT WAS REPORTED "WE HAD A PATIENT IN ICU THAT DEVELOPED SOME ICU PSYCHOSIS ON NIGHT SHIFT AND PULLED THE RED LUMEN , COMPLETE SEVERING FROM THE FROM THE WING. THERE WERE NO COMPLICATIONS NOTED. THE PICC WAS REMOVED AND REPLACED." | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DETACHED EXTENSION TUBING WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS 4FR D/L POWERPICC SOLO PROVENA CATHETER. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. THE RED-LUERED EXTENSION TUBE WAS COMPLETELY DETACHED FROM THE MOLDED JOINT. MICROSCOPIC INSPECTION OF THE DISTAL END OF THE RED-LUERED EXTENSION TUBING REVEALED A GLOSSY SURFACE CONSISTENT WITH THE MANUFACTURED END, INDICATING THAT THE TUBING WAS DETACHED AND NOT BROKEN. THE DETACHED EXTENSION TUBING EXHIBITED PLASTIC DEFORMATION AND CURVED SHAPE MEMORY NEAR THE DISTAL END. TACTILE INSPECTION OF THE DETACHED EXTENSION TUBING REVEALED TENSILE WEAKNESS AND MATERIAL NECKING NEAR THE DISTAL END. THE PLASTIC DEFORMATION, MATERIAL WEAKNESS AND NECKING WERE CONSISTENT WITH EXTENSION TUBING DETACHMENT CAUSED BY APPLICATION OF PULLING STRESS. THESE FINDINGS WERE CONGRUENT WITH THE EVENT DESCRIPTION, WHICH INDICATED THAT THE PATIENT PULLED ON THE CATHETER. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05484 | 2023-07-25 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53614G |  | IT WAS REPORTED THAT THE THREE OF THE INTERMITTENT CATHETERS HAD EMPTY BURST PACKS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INCORRECT PROCESS PARAMETERS." THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05481 | 2023-07-25 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS FOUND TO BE BURST ON THE NEXT DAY AFTER PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05483 | 2023-07-25 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165818 |  | IT WAS REPORTED THAT THERE WAS A RIDGING ON THE FOLEY CATHETER BALLOON WHEN REMOVING. AS PER FOLLOW UP VIA EMAIL ON 06JUL2023, IT WAS STATED THAT THE CUSTOMER DID NOT HAVE THE ORIGINAL PACKAGING, SO NO LOT OR BATCH NUMBER. THE CATHETER WAS INSERTED IN A MALE PATIENT WITH NO CONCERNS IDENTIFIED AND THEN REMOVED 28 DAYS LATER FOR A ROUTINE CHANGE. THERE WAS NO RESIDUAL FLUID LEFT IN THE BALLOON, BUT A MARKED RIDGING PRESENT. THIS CAUSED THE PATIENT PAIN AND BLEEDING AT THE TIME, BUT FORTUNATELY THIS SETTLED DOWN WITHIN A FEW HOURS, AND OUTPUT REPORTED AS CLEAR OR AMBER URINE SUBSEQUENTLY. NURSES DID NOT PRETEST BALLOONS, AND IT WAS FILLED WITH 10ML AS PER PRODUCT SPECIFICATIONS. PATIENT WAS A YOUNG MALE POST TRAUMATIC INJURY, AND WAS UNABLE TO VERBALIZE THE EXPERIENCE, BUT THERE WAS RESISTANCE UPON REMOVAL AND NONVERBAL EXPRESSION OF PAIN. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER. VISUAL INSPECTION NOTED THAT THE BALLOON WAS MUSHROOMED . THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. DEFLATED BALLOON RETURNING 10ML OF THE SOLUTION WITH NO LEAKS NOTED. THIS IS OUT OF SPECIFICATION WHICH STATES, " BALLOON MUST NOT CUFF AFTER DEFLATION. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. AS THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05476 | 2023-07-25 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE NURSE FOLLOWED THE DOCTOR'S ADVICE TO USE A DISPOSABLE STERILE CATHETER TO PERFORM CATHETERIZATION FOR THE PATIENT. AFTER THAT, DURING WATER INJECTION, IT WAS FOUND THAT THE BALLOON WAS BURST AND WAS REPLACED BY A NEW ONE. ACCORDING TO THE INJURY PERFORMANCE THE PATIENT EXPERIENCED LEAKAGE AND DISCOMFORT. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON 14JUL2023, CLARIFIED THAT THE PATIENT EXPERIENCED DISCOMFORT BY THE PRIMARY CATHETER USED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05407 | 2023-07-25 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 33616 |  | IT WAS REPORTED THAT THERE WAS NO ORANGE PART IN THE BEGINNING OF THE FOLEY CATHETER. THAT IS, THE PROBE WAS MISSING A PIECE. THE PROBE WAS NOT USED IN PATIENT, THE PROBE WAS INSIDE THE ORIGINAL CLOSED PACKAGING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05474 | 2023-07-25 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123614A |  | IT WAS REPORTED THAT THE CATHETER BALLOON WAS TOO LARGE, MAKING IT DIFFICULT TO INSERT. WHEN THE URINARY CATHETER WAS REMOVED, THE BALLOON WAS TOO LARGE AND HARD, CAUSING URETHRAL INJURY AND URETHRAL BLEEDING. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON 15JUL2023, IT WAS UNDETERMINED IF THE DEVICE WAS USED DESPITE KNOWN THAT THE PRODUCT WAS OF BIGGER SIZE THAN THE PATIENT'S ANATOMY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ''UNDER CURED SAC'' (EXAMPLE: POOR SHAPE/ BAGGY SAC-NONCONCENTRIC BALLOON). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE: SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05473 | 2023-07-25 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53614G |  | IT WAS REPORTED THAT THE 1 INTERMITTENT CATHETER IN BOX HAD A DEFECT AT THE TIP OF IT, SO PATIENT DID NOT USE THE INTERMITTENT CATHETER. PER FOLLOW UP VIA PHONE (B)(6) 2023, THE END HAD A PROTRUDING TIP. THE LOT NUMBER WAS JUGZ20901. THE DEFECT WAS NOTICED BEFORE USE. THERE WAS ALSO TWO IN ANOTHER SHIPMENT WITH REFERENCE NUMBER (B)(4) AND LOT JUGW0050. CUSTOMERS ADVISED THEY WOULD FORWARD PICTURES OF THE MOST RECENT INCIDENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿MANDRELS TOUCHING EACH OTHER WITHIN THE PALLET ¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS NOT USED ON THE PATIENT. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05470 | 2023-07-25 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE TEMPERATURE DISPLAY DID NOT CHANGE FROM 30°C WHILE USING THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05466 | 2023-07-25 | 2023-07-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 50814G |  | IT WAS REPORTED THAT PATIENT HAD 4 URINARY TRACT INFECTIONS SINCE MARCH, 3 RESULTING IN EMERGENCY ROOM VISITS AND THE OTHER WITH PATIENT BEING HOSPITALIZED, PATIENT STATED THAT PATIENT EXPERIENCED EXCRUCIATING PAIN, FEVER AND VOMITING. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05467 | 2023-07-25 | 2023-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123620A |  | IT WAS REPORTED THAT A PATIENT UNDERWENT INTRACERVICAL WATER SAC PLACEMENT FOR OXYTOCIN AT 21:00 ON 2023-(B)(6) WHEN INJECTED WITH STERILE WATER, VAGINAL DISCHARGE OCCURRED AND THEN THE CATHETER WAS REMOVED. THE NORMAL SALINE WAS INJECTED INTO THE BALLOON AGAIN AND A LEAKAGE WAS FOUND. THEN IT WAS REPLACED BY A NEW CATHETER. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6)2023, THERE WAS A DAMAGE ON THE BALLOON AND IT WAS A BALLOON BURST EVENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05464 | 2023-07-25 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT ON (B)(6) 2023, AFTER AN OPERATION THE PATIENT WAS CATHETERIZED WITH A THREE LUMEN CATHETER. ON (B)(6) 2023, THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT AND THE BALLOON WAS FOUND TO BE BURST. IMMEDIATELY THE DOCTOR WAS INFORMED AND THE PATIENT WAS PLACED WITH A NEW THREE LUMEN URINARY CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05463 | 2023-07-25 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE SYRINGE WAS ATTACHED TO THE INFLATION VALVE AND ATTEMPTED TO INJECT WATER, BUT STERILE WATER COULD NOT BE INJECTED INTO THE FOLEY CATHETER AND LEAKED OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03173 | 2023-07-25 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SITE~RITE 8 ULTRASOUND SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770550 |  | IT WAS REPORTED THE SCREEN LOOKS VERY DARK AND THE BRIGHTNESS AND CONTRAST CANNOT BE ADJUSTED ANY FURTHER, MAKING IT VERY DIFFICULT TO LOCATE THE VESSELS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF THE SCREEN LOOKS VERY DARK, AND THE BRIGHTNESS AND CONTRAST CANNOT BE ADJUSTED ANY FURTHER, MAKING IT VERY DIFFICULT TO LOCATE THE VESSELS WAS CONFIRMED. THE ROOT CAUSE OF THE REPORTED FAILURE WAS IDENTIFIED AS THE INTERNAL GLUE HAS WORN AND RUBBER HAS COME AWAY FROM PROBE CABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03174 | 2023-07-25 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SITE~RITE 8 ULTRASOUND SYSTEM CONSOLE WITH CONNECTOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770600 |  | IMAGES ARE VERY DARK MAKING IT VERY DIFFICULT TO FIND THE PATIENTS VEIN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05475 | 2023-07-25 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123618A |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE PATIENT WAS UNABLE TO URINATE ON THEIR OWN DUE TO THE FULLNESS OF THE BLADDER AND THE PATIENT NEEDED TO BE CATHETERIZED. JUST AFTER THE CATHETERIZATION, WHEN THE BALLOON WAS INJECTED WITH SALINE, IT WAS FOUND THAT THE BALLOON RUPTURED AND THE SALINE COULD NOT BE INJECTED. THE CATHETER WAS IMMEDIATELY PULLED OUT AND WAS REPLACED BY A NEW CATHETER TO COMPLETE TREATMENT. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 15JUL2023, CLARIFIED THAT THE WATER WAS DIFFICULT TO BE INJECTED DUE TO BALLOON BURST. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: DO NOT USE THE PRODUCT IF PATIENT HAS LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00534 | 2023-07-25 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | VA80124 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED PRIOR TO BEING INFLATED OR REACHING NOMINAL PRESSURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 01/2026). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00533 | 2023-07-25 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | AT120124 |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE, IT WAS OBSERVED THAT THE BAG IN WHICH THE PTA BALLOON WAS KEPT LOOKED TORN WHEN IT WAS REMOVED FROM ITS FIRST PACKAGE (CARDBOARD). THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 10/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03190 | 2023-07-25 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED STYLET INSIDE THE CATHETER ADHERED TO THE CATHETER WITH DIFFICULTY IN REMOVAL. EVENT OCCURRED BEFORE USE. ADDITIONAL INFORMATION RECEIVED (B)(6), 2023: CUSTOMER REPORTS THEY PERFORMED "ALL OF THE ABOVE PRACTICES" AND THAT THE CATHETER WAS DEFORMED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00535 | 2023-07-25 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | VA80124 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED PRIOR TO BEING INFLATED OR REACHING NOMINAL PRESSURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 01/2026) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03158 | 2023-07-25 | 2021-07-27 | I, n, j, u, r, y | Manufacturer report | POWERPORT IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1759600 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT FIVE MONTHS SEVENTEEN DAYS POST PORT PLACEMENT FOR CANCER TREATMENT, THE PORT ALLEGEDLY FRACTURED. IT WAS FURTHER REPORTED THAT A PIECE HAD BROKEN OFF AND MIGRATED TO PATIENT HEART. THE PATIENT WAS DEVELOPED WITH HEART ARRHYTHMIA DURING THE REMOVAL PROCEDURE. FURTHER AN ADDITIONAL SURGERY WAS DONE TO REMOVE TO FRACTURED CATHETER PIECE. THE CURRENT STATUS OF THE PATENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. MEDICAL RECORDS WERE PROVIDED AND REVIEWED. APPROXIMATELY FIVE MONTHS AND NINETEEN DAYS POST PORT PLACEMENT, A CT OF CHEST WAS PERFORMED WHICH DEMONSTRATES THE CATHETER FRAGMENT APPEARING TO BE WITHIN THE RIGHT VENTRICLE. PATIENT UNDERWENT ULTRASOUND GUIDED RIGHT COMMON FEMORAL VEIN ACCESS, A RIGHT CHEST PORT WAS REMOVED AND ATTEMPTED FOREIGN BODY RETRIEVAL PROCEDURE. ON THE SAME DAY THE PATIENT WAS PLACED ON RIGHT NEW POWERPORT CLEARVUE SLIM PLACEMENT WITH NEW VENOUS ACCESS AND PORT POCKET. A REFERRAL WAS MADE TO CARDIOTHORACIC SURGERY FOR EVALUATION OF A RETAINED CATHETER FRAGMENT IN THE RIGHT VENTRICLE. PATIENT CONDITION WAS STABLE AND DISCHARGED TO HOME. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION AND MIGRATION. FURTHERMORE, THE CLINICAL CONDITION ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: A1, A3, D4 (EXPIRATION DATE: 05/2022), G2, G3, H6 (METHOD) H11: B5, B7, D1, D4, H6 (RESULT) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03191 | 2023-07-25 | 2023-06-20 | I, n, j, u, r, y | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED CUSTOMER ATTEMPTED AN ACCUCATH INSERTION TODAY AND HAD DIFFICULTY RETRACTING THE WIRE TO CHANGE NEEDLE INSERTION. WHEN RETRACTING THE WIRE IN THE ARM AN APPROXIMATE 0.75 CM OF THE WIRE WAS SHEARED OFF AND REMAINED IN THE TISSUE OF THE ARM. WHEN EXAMINING THE DEVICE ONCE REMOVED AND WHEN DEPLOYING THE WIRE, IT WAS NOTED TO COME OUT OF THE NEEDLE MID-NEEDLE. THE EVENT DID NOT INVOLVE AN URGENT MEDICAL SITUATION. ADDITIONAL INFORMATION RECEIVED VIA MEDWATCH ON (B)(6), 2023 "VASCULAR ACCESS RN NOTICED THAT THE TIP OF THE WIRE WAS MISSING AFTER INSERTING A 20G X 2.25 IN BARD ACCUCATH ACE IV WITH NITINOL GUIDEWIRE WITH COILED TIP. A STAT X-RAY IDENTIFIED THAT THE WIRE WAS STILL IN THE PATIENT'S ARM. THE GUIDEWIRE TIP WAS REMOVED THE FOLLOWING DAY UNDER FLUOROSCOPY GUIDANCE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03143 | 2023-07-25 | 2022-03-11 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1808561 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, THE PORT ALLEGEDLY DID NOT FLUSH. IT WAS FURTHER REPORTED THAT THE PORT HAD ALLEGEDLY FRACTURED, AND THE FRACTURED PIECE HAD MIGRATED TO THE PATIENT'S HEART. THERE WAS DIFFICULTY IN REMOVING THE FRACTURE PORT. THE CURRENT STATUS OF THE PATENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. MEDICAL RECORDS WERE PROVIDED AND REVIEWED. APPROXIMATELY SIX MONTHS AND TEN DAYS POST PORT PLACEMENT, THE PATIENT WAS DIAGNOSED WITH A FOREIGN BODY IN THE HEART AND FOREIGN BODY RETRIEVAL FROM HEART WAS PERFORMED USING A SNARE. USING ULTRASOUND, THE RIGHT COMMON FEMORAL VEIN WAS ACCESSED. THEN USING SELDINGER TECHNIQUE, 6-FRENCH SHEATH WAS USED AND THEN BENTSON WIRE WAS ENTERED INTO THE VEIN AND THEN EASILY PASSED INTO THE SUPERIOR VENA CAVA. THE PORTION OF THE PORT WAS NOTED TO BE COILED IN THE RIGHT HEART. AN EN SNARE DEVICE WAS ATTEMPTED TO USE MULTIPLE TIMES TO ATTEMPT TO SNARE THE FOREIGN BODY, BUT IT WAS UNSUCCESSFUL. THE BENTSON WIRE WAS THEN REINSERTED, AND THE SHEATH WAS UPSIZED TO A 12FR SHEATH. FROM THERE, USING THE SNARE DEVICE AND A BENTSON WIRE, IT WAS ATTEMPTED TO MANIPULATE THE FOREIGN BODY INTO THE INFERIOR VENA CAVA AND WAS ABLE TO SNARE, AND PULLED IT OUT THROUGH THE 12-FRENCH SHEATH. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FACTURE, MATERIAL SEPARATION, MIGRATION AND DIFFICULTY IN REMOVE. HOWEVER, INVESTIGATION IS KEPT AS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO FLUSH AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: B5, G2, D4 (EXPIRY DATE: 12/2023), H6 (DEVICE, METHOD). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03142 | 2023-07-25 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM LOW-PROFILE IMPLANTABLE PORT WITH ATTACHABLE 8F GROSHONG SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0606150 |  | IT WAS REPORTED THAT DURING THE PREPARATION OF PORT PLACEMENT PROCEDURE, THE NEEDLE ALLEGEDLY DID NOT ASPIRATE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE INTRODUCER NEEDLE WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. A CRACK WAS NOTED WITHIN THE INTRODUCER NEEDLE HUB. UPON INFUSION, SMALL LEAKS WERE OBSERVED FROM THE INTRODUCER NEEDLE HUB. ASPIRATION OF WATER INTO THE SYRINGE WAS ATTEMPTED AND WAS UNSUCCESSFUL AS WATER DID NOT FULLY RETURN WITHIN THE IN-HOUSE SYRINGE. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, TWO ELECTRONIC VIDEOS WERE PROVIDED FOR REVIEW. THE VIDEO SHOWS THE CLINICIAN WAS ASPIRATING WATER IN THE SYRINGE AND IT WAS UNSUCCESSFUL. MANUFACTURING SITE EVALUATION OF THE SAMPLE FOUND A CRACK IN THE PINK HUB OF THE INTRODUCER NEEDLE, IT WAS OBSERVED THAT THE CAVITY THAT PRESENTED THIS DAMAGE WAS NUMBER THREE, A SYRINGE CONNECTED TO THE NEEDLE WAS OBSERVED WHERE LIQUID WAS BEING INFUSED THROUGH THE NEEDLE, HOWEVER, A LEAK OF THAT LIQUID WAS OBSERVED THROUGH THE HUB OF THE INTRODUCER NEEDLE. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED SUCTION ISSUE AND IDENTIFIED FRACTURE ISSUES. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 06/2026) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05314 | 2023-07-26 | 2023-06-23 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175816 |  | IT WAS REPORTED THAT REGISTER NURSE WENT TO PLACE INDWELLING URETHRAL FOLEY CATHETER POST EPIDURAL. AFTER PLACEMENT, REGISTER NURSE NOTICED CATHETER FELL OUT AND TIP WAS NOT INTACT. PATIENT ASKED TO ATTEMPT TO VOID SPONTANEOUSLY TO SEE IF THE TIP WOULD DISLODGE. DOCTOR OF MEDICINE NOTIFIED. MANAGEMENT NOTIFIED. PATIENT VOIDED. REGISTER NURSE DID NOT FIND CATHETER TIP. PLANNED PROCEDURE AFTER DELIVERY, PER DOCTOR OF MEDICINE, TO VERIFY CATHETER TIP WAS NOT INSIDE THE BLADDER. NEW INDWELLING CATHETER PLACED, NO COMPLICATIONS. SURVEILLANCE CYSTOSCOPY PERFORMED. SALINE USED TO DISTEND BLADDER, AND THE RIGID CYSTOSCOPE WAS USED TO SURVEY THE ENTIRE CAVITY USING BOTH THE 30 DEGREE AND THE 70-DEGREE LENSES. NO FOREIGN BODY IDENTIFIED. NO BLADDER INJURY SEEN. BOTH URETERAL ORIFICES SEEN AND EFFLUXING URINE WITH STRONG JETS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03200 | 2023-07-26 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT THE POWERPICC HUB WAS LEAKING, NO PROBLEMS DURING PLACEMENT BUT LEAKING FROM HUB LATER, ISSUE PERSISTED DESPITE TROUBLESHOOTING; THIS OCCURRED WITH MULTIPLE CATHETERS AROUND THE SAME TIME. THE CUSTOMER REPORTED MULTIPLE EVENTS HOWEVER THE EXACT NUMBER OF EVENTS WAS NOT PROVIDED TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03198 | 2023-07-26 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER 4F BASIC TRAY (SINGLE-LUMEN) (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6194118 |  | IT WAS REPORTED BY THE CUSTOMER THAT BEFORE USE, HCP NOTICE THAT PICC'S LEAD-IN WIRE IS BROKEN. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03197 | 2023-07-26 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 HF CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000394 |  | IT WAS REPORTED BY THE CUSTOMER THAT "ISSUE WITH ECG WIRE, CLEAR STYLET CONNECTION WITH BEIGE PORT LEAKING SALINE WHEN FLUSHED WHEN PICC LINE IN PATIENT AND ISSUES PULLING BACK BLOOD RETURN WHEN ECG WIRE IN PLACE THROUGH BEIGE STYLET SEEMS TO LET AIR IN. HAVE TO THEN DISCONNECT FLUSH AND EXPEL AIR AND RECONNECT IT AND PULL BACK TO ASSESS FOR BLOOD RETURN." ADDITIONAL INFO RECEIVED ON 13-JULY-23 CUSTOMER REPORTED THAT "THE PATIENT WAS IN ICU IN CENTRAL ACCESS. THEY HAD ISSUES WITH THE ECG WIRE BEING ALREADY BENT INSIDE THE NEW PICC, ALONG PLUNGER ON THE CLEAR ECG WIRE HOUSING SPRAYING OUT ONCE FLUSHES WHILE NEW PICC IN INTRODUCER WELL AS PULLING AIR AS THERE DOESN¿T SEEM TO BE A GOOD SEAL WITH THE CHANGE IN THE BEIGE PLUNGE. THE EVENT DELAYED THEM FROM BEING ABLE TO CORR BLOOD FROM RED LUMEN ONCE PICC WAS PLACED. THE WIRE BUMPING AS WE ARE PULLING IT IN AND OUT OF TRIM IT ETC. SLOW ME DOWN AS PATIENT WAS CRITICALLY ILL. SHE WAS ALSO IN RAPID AFIB REQUIRING A CXR TIP LOCATION SO THAT WAS ALSO A DELAY THAT COULDN¿T BE AVOIDED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05512 | 2023-07-26 | 2023-06-22 | I, n, j, u, r, y | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT AFTER THE PATIENT GOT BLADDER LITHOTOMY, THE URINE WAS DRAINED TO PREVENT URETHRAL OBSTRUCTION, AND THE BLADDER URINE DRAINAGE WAS STRICTLY MAINTAINED. HOWEVER THE CATHETER FELL OUT CAUSED EXCESSIVE URINE IN THE BLADDER RESULTING IN URINARY EXTRAVASATION AT THE BLADDER SUTURE OPENING. AS PER THE DEVICE FAILURE PERFORMANCE, THE BALLOON WAS RUPTURED AND A CT SCAN WAS TAKEN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT IF YOU HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03195 | 2023-07-26 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT DURING PLACEMENT OF THE PICC IN THIS PATIENT, WHEN INSPECTING THE MICRO-INTRODUCER KIT FOR INTEGRITY AFTER UNPACKING PRIOR TO USE, A SPLIT WAS FOUND AT THE PROXIMAL END OF THE GREY PEEL-AWAY SHEATH OF THE MICRO-INTRODUCER. THIS MADE IT IMPOSSIBLE TO PENETRATE THE SKIN. THE CUSTOMER OPENED ANOTHER MST FOR PUNCTURE. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03194 | 2023-07-26 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED THAT VIA VAM ASSESSMENT THAT THE SECUREMENT DEVICE WAS NOT APPLIED PER MANUFACTURERS IFU. NO OTHER INFORMATION PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 3 DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03193 | 2023-07-26 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED THAT VIA VAM ASSESSMENT THAT THE SECUREMENT DEVICE WAS NOT APPLIED PER MANUFACTURERS IFU. NO OTHER INFORMATION PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 3 DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03199 | 2023-07-26 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT THE POWERPICC GREY PEEL AWAY SHEATH OVER INTRODUCER BUCKLED/CRUMPLED DURING INSERTION, OCCURRED MULTIPLE TIMES. NOT RECENT OCCURRENCES, DETAILS LIMITED DUE TO HOW LONG AGO THESE OCCURRED. THE CUSTOMER REPORTED MULTIPLE EVENTS HOWEVER THE EXACT NUMBER OF EVENTS WAS NOT PROVIDED TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03192 | 2023-07-26 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED THAT VIA VAM ASSESSMENT THAT THE SECUREMENT DEVICE WAS NOT APPLIED PER MANUFACTURERS IFU. NO OTHER INFORMATION PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 3 DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05511 | 2023-07-26 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT DURING SURGERY, THE FOLEY CATHETER BALLOON CAME OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05491 | 2023-07-26 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT DUE TO A BALLOON RUPTURE. AFTER INSERTION OF THE CATHETER, WATER WAS INJECTED INTO THE BALLOON, BUT THE CATHETER WAS NOT SECURED AND PULLED OUT. UPON CHECKING THE CATHETER, THE BALLOON WAS FOUND TO BE RUPTURED AND DID NOT INFLATE. THE EVENT OCCURRED JUST AFTER PLACEMENT. IT WAS STATED THAT THE COMPONENTS IN THE TRAY AND HAND GLOVES POSSIBLY CAME IN CONTACT TO THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05498 | 2023-07-26 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI20 |  | IT WAS REPORTED THAT WHEN ATTEMPTED TO USE THE FOLEY CATHETER, THERE WAS A BENT MARK FOUND ON THE CATHETER. THE EVENT HAS OCCURRED THREE TIMES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05499 | 2023-07-26 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165814 |  | IT WAS REPORTED THAT INDWELLING SILICONE FOLEY CATHETER 14F WAS NOT DRAINING AND WAS LEAKING, NEEDED TO BE DISCONTINUED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03196 | 2023-07-26 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PER-Q-CATH PLUS 2F SINGLE-LUMEN PICC WITH STIFFENING STYLET | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4132105 |  | IT WAS REPORTED BY THE CUSTOMER "INSIDE THE PACKAGING OF THE CATHETER IS A STRAND OF HAIR." ADDITIONAL INFORMATION WAS RECEIVED BY THE CUSTOMER 7/21/2023: "THE PRODUCT IS NOT DAMAGED, THE PACKAGING IS NOT TORN, THE HAIR IS INSIDE THE PACKAGING, WHICH IS SEALED. YOU CAN SEE IT THROUGH THE PLASTIC." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF APPLICABLE INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A HAIR FOUND IN THE PACKAGING IS CONFIRMED AND WAS DETERMINED TO BE RELATED TO MANUFACTURING. ONE SEALED PER-Q-CATH PLUS CATHETER KITS WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION REVEALED A SMALL, DARK HAIR INSIDE THE SEALED PER-Q-CATH KIT. NO OBVIOUS USE RESIDUES WERE OBSERVED INSIDE THE KIT. THE KIT WAS OPENED AFTER OBSERVING THE HAIR IN THE SEALED KIT, AND THE HAIR WAS CONFIRMED TO BE INSIDE THE KIT. BECAUSE THE DARK HAIR COULD BE SEEN IN THE KIT BEFORE IT WAS OPENED, THE COMPLAINT OF A HAIR FOUND INSIDE A PER-Q-CATH KIT IS CONFIRMED AND DETERMINED TO BE RELATED TO MANUFACTURING. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. THE INVESTIGATION WAS FORWARDED TO THE MANUFACTURING FACILITY FOR FURTHER EVALUATION, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05501 | 2023-07-26 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53614G |  | IT WAS REPORTED THAT THE PATIENT HAD WENT TO HOSPITAL DUE TO BLOCKAGE WITH THE MAGIC3 HYDROPHILIC MALE INTERMITTENT CATHETER. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, CUSTOMER WAS INSTRUCTED THAT THEIR DOCTOR STATED THIS CATHETER DID NOT WORK FOR THEIR ANATOMY AND THAT CUSTOMER NEEDED A COUDE CATHETER. CUSTOMER STATED THAT THEY WERE GIVEN ANTIBIOTICS DURING THEIR STAY AT THE HOSPITAL. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿INSUFFICIENT EYELET AREA BLOCKED EYE OR INNER DIAMETER OF DRAINAGE EYE SMALL ENOUGH TO ALLOW FOR OCCLUSION DURING USE¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT BUT IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE, INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05507 | 2023-07-26 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT URINE LEAKED FROM SLIGHTLY BELOW THE FOLEY CATHETER BIFURCATION. THE CATHETER WAS FIXED AT THE RIGHT THIGH. THE LEAK WAS DISCOVERED ON THE DAY OF INSERTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05505 | 2023-07-26 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT CUSTOMER EXPERIENCED BELLY PAIN AND WENT TO EMERGENCY ROOM. THE DOCTOR REMOVED THE CATHETER AND SHOWED THE PATIENT A WRINKLE IN THE CATHETER THAT THE DOCTOR STATED WAS THE CAUSE OF THE PAIN. DOCTOR PLACED A NEW CATHETER. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE COULD BE TO THIN SAC COULD CAUSE POOR SHAPE/ BAGGY SAC & NON- CONCENTRIC BALLOON. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON BURST. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE PATIENT USE ONLY. DO NOT REUSE AND RESTERILIZE. FOR UROLOGICAL USE ONLY. USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PROTOCOL, THE MAY BE CUT OFF. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05509 | 2023-07-26 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L18 |  | IT WAS REPORTED THAT URINE LEAKED FROM THE FOLEY CATHETER. IT WAS STATED THAT THE IN THE MORNING OF THE DAY FOLLOWING THE SURGERY, A URINE LEAK WAS DISCOVERED. IT WAS ALSO STATED THAT WHEN THE CATHETER WAS CHECKED, IT WAS NOT LEAKING FROM THE URETHRAL OPENING, BUT FROM A CRACK IN THE CATHETER SHAFT AREA. IT WAS NOTED THAT THE FOLEY CATHETER HAD BEEN INSERTED TO THE PATIENT IN THE OPERATING ROOM. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO (EXAMPLE: USER RELATED/OPERATOR ERROR/ROUGH HANDLING OR EXPOSE TO SHARP OBJECT/MECHANICAL FORCE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]. 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.]. (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.]. 2. APPLICABLE PATIENTS. (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.]. [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.]. (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.]. 2. APPLICABLE PATIENTS: (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX. (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05500 | 2023-07-26 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 50612 |  | IT WAS REPORTED THAT THE PATIENT HAD ABOUT 3 CATHETERS THAT HAD A CROOKED TIP AND PATIENT STATED IT WAS SIDEWAYS AND PATIENT DISCARDED. REPRESENTATIVE ADVISED PATIENT WHEN PATIENT HAS DEFECTIVE PRODUCTS TO LET REPRESENTATIVE KNOW SO REPRESENTATIVE CAN FIX IT. REPRESENTATIVE SENT 3 REPLACEMENT CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT DOES NOT APPEAR TO HAVE BEEN USED. HOWEVER, THE PRODUCT IS INTENDED TO BE USED FOR TREATMENT PURPOSE BUT IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05536 | 2023-07-27 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® RUBBER UTILITY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 277714 |  | IT WAS REPORTED THAT THE BARD 14FR WAS LARGER THAN THE 14FR THE PATIENT RECEIVED FROM THEM DOCTOR. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿OPERATOR ERROR¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05537 | 2023-07-27 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN FOLEY CATHETER | UNKNOWN FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 3 FOLEY KITS THAT THE BALLOON WAS MALFUNCTIONING. PRIOR TO PLACING THE FOLEY, THE BALLOON WAS TESTED. SALINE WAS ABLE TO BE INSTILLED BUT NOT REMOVED. PER FOLLOW UP VIA EMAIL ON (B)(6), 2023, STATED THAT THIS WAS MERELY A REPORTED EVENT REGARDING RECENT FOLEYS AND THEY WOULD WORK WITH THE HOSPITAL STAFF TO DETERMINE WHERE THE PROBLEMS OCCURRED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05535 | 2023-07-27 | 2023-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON RUPTURED ON THE FIRST POSTOPERATIVE DAY AND FELL OUT FROM THE BLADDER DURING IRRIGATION. THE BLADDER FLUSHING FLUID FLOWED TO THE BEDSIDE AND LOWER BODY OF THE PATIENT. IMMEDIATELY REPLACED THE FOLEY CATHETER AND CONTINUED BLADDER IRRIGATION. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6), 2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿SHORT LATEX DWELL TIME" OR "LATEX DIP SPEED OUT TOO SLOW". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05528 | 2023-07-27 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE NURSE TESTED THE PRODUCT BEFORE USING IT ON A PATIENT AND THE FIRST CATHETER WOULD NOT DEFLATE. SO THEY PULLED ANOTHER CATHETER WITH THE SAME LOT NUMBER AND GOT THE SAME RESULT, CATHETER WOULD NOT DEFLATE. THE DEVICE WAS NOT USED ON A PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX.. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05526 | 2023-07-27 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS USED FOR URINE DRAINAGE BUT THE BALLOON WAS BURST. THE CATHETER WAS REMOVED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 21JUL2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03214 | 2023-07-27 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0607540 |  | IT WAS REPORTED THAT PRIOR TO A PORT PLACEMENT PROCEDURE, THE HUMAN HAIR WAS ALLEGEDLY FOUND IN THE PACKAGING. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 08/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05527 | 2023-07-27 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTERMITTENT ROCHESTER CATHETERS | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT'S ADAPTATION TO THE MAGIC 3 (HYDROPHILIC) WAS VERY BAD AND COMPLAINTS WERE RECEIVED ON ACTIVATION AND LUBRICATION OF THE CATHETER. AS PER THE ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023 STATED THAT IT WAS NOT A DEFECT OR ANY DEVIATION FROM THE QUALITY OF THE PRODUCTS. THE PURPOSE OF THE REPORT WAS TO INFORM THAT DUE TO THE POOR MANUAL DEXTERITY OF MOST USERS (WHEELCHAIR USERS WITH SPINAL CORD INJURIES ¿ PARAPLEGICS AND QUADRIPLEGICS), THEY ENCOUNTERED DIFFICULTIES IN ACTIVATING THE HYDROPHILIC MAGIC3 CATHETER, FAILING TO PERFORM THE PROCEDURE, IN REGARDS TO A LIMITATION DUE TO THE CHARACTERISTIC OF THE PRODUCT AND NOT ANY MANUFACTURING DEFECT. AS PER THE ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023, THE CUSTOMER PROVIDED THE REASONS FOR DIFFICULTIES WITH ACTIVATION IN DETAIL AS FOLLOWS: AS PREVIOUSLY REPORTED, MANY OF THE USERS WERE QUADRIPLEGIC, THAT IS, FOR SOME REASON THEY HAD A HIGH SPINAL CORD INJURY COMPROMISING TOTALLY OR PARTIALLY THE MOVEMENTS OF THE UPPER LIMBS. IN THE CASE OF TOTAL LOSS, THERE WAS A NEED FOR ANOTHER PERSON TO PERFORM THE PROCEDURE. HOWEVER, OTHERS MANAGE, ALBEIT WITH A LOT OF EFFORT, TO PERFORM THE TWEEZING MOVEMENT ENOUGH TO HOLD THE CATHETER BUT WITHOUT ENOUGH DEXTERITY TO BREAK THE SACHET WITH WATER TO PROMOTE LUBRICATION. THERE WERE SOME POINTS TO BE CONSIDERED THAT INCLUDING THE POSITION OF THE SACHET WITH WATER WAS NOT THE MOST APPROPRIATE BASED ON THE EXPERIENCE REPORTED BY USERS. THEN THE SIZE AND VOLUME OF WATER WAS NOT ENOUGH TO COVER THE ENTIRE SURFACE OF THE CATHETER, REQUIRING THE DISTRIBUTION MANEUVER, WHICH WAS NOT SO SIMPLE FOR THOSE WITH LIMITATIONS. THEY ALSO SUGGESTED A FEW POINTS FOR IMPROVEMENT AS FOLLOWS: POSITIONING THE SACHET IN THE DISTAL PART OF THE CATHETER, A LARGER SIZE AND WITH A GREATER VOLUME OF WATER. THEY ALSO PROVIDED THE REASONS FOR THE DIFFICULTY WITH LUBRICATION AS FOLLOWS: MANY USERS WERE ALREADY USERS OF HYDROPHILIC CATHETERS FROM OTHER MANUFACTURERS (SPEEDICATH, LOFRIC ORIGO, LOFRIC SENSE). THE ABOVE PRODUCTS HAD AN EXTERNAL COATING OF PVP (POLYVINYLPYRROLIDONE) WHICH WAS A COMPONENT THAT HAS A HIGH CAPACITY FOR ABSORBING AND RETAINING WATER WITHIN ITS STRUCTURE, MAKING ITS SURFACE EXTREMELY SMOOTH, FACILITATING THE INTRODUCTION OF THE CATHETER INTO THE URETHRA WITHOUT CAUSING FRICTION AND/OR USER DISCOMFORT. THE REPORTED PERCEPTION WAS THAT THE MAGIC3 DOES NOT HAVE THE SAME LUBRICATION QUALITY AND SOME WERE AFRAID TO USE IT, AS THEY FEARED INJURIES AND URINARY TRACT INFECTIONS AS PREVIOUSLY HAPPENED WHEN USING CONVENTIONAL PVC CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND LOT NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | OTHER |  | 2023-07
1018233-2023-05519 | 2023-07-27 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT AFTER INJECTING FIXATION WATER INTO THE BALLOON DURING A PRETEST, THE WATER WAS DIFFICULT TO REMOVE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), AN ALL-SILICONE FOLEY CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED, RETURNING 10ML OF SOLUTION. THE BALLOON PASSIVELY DEFLATED IN UNDER 3 MINUTES WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. AS THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05517 | 2023-07-27 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED INTO THE PATIENT AT THE BEGINNING OF THE SURGERY. IT WAS ALSO STATED THAT WHILE TRANSFERRING THE PATIENT TO THE WARD AFTER SURGERY, THE URINE DRAINAGE BECAME POOR AND THE FOLEY CATHETER BALLOON WAS REINSERTED. IT WAS NOTED THAT THERE WAS NO URINE DRAINAGE PROBLEM DURING THE SURGERY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03212 | 2023-07-27 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT THIRTEEN YEARS AND FIVE MONTHS POST PORT PLACEMENT, WHEN THE PORT SYSTEM WAS ATTEMPTED TO BE REMOVED, IT WAS ALLEGEDLY DIFFICULT TO REMOVE THE CATHETER. IT WAS FURTHER REPORTED THAT WHEN THE REMOVED CATHETER WAS CHECKED, THE SURFACE OF THE CATHETER WAS UNEVEN AND ROUGH. REPORTEDLY, THERE WAS NO BREAKAGE OF THE CATHETER AND NO REMAINS IN THE PATIENT'S BODY. FURTHERMORE, THE PORT AND THE CATHETER WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE X-PORT IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL TESTING WERE PERFORMED. DURING MICROSCOPIC EVALUATION, TISSUE ADHESION WAS NOTED THROUGHOUT THE CATHETER AND THE CATHETER WAS NOTED TO BE ROUGH TO THE TOUCH. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED DEFORMATION AND PATIENT DEVICE INTERACTION ISSUE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR DIFFICULT TO REMOVE AS THE CONDITIONS OF USE (DEVICE IMPLANTED FOR OVER 13 YEARS) COULD NOT BE RECREATED. IT WAS REPORTED THAT THE DEVICE WAS IMPLANTED FOR MORE THAN THIRTEEN YEARS. THEREFORE, IT IS POSSIBLE THAT THE LENGTH OF THE IMPLANT CONTRIBUTED TO THE REMOVAL DIFFICULTIES AND THE TISSUE ADHESION ALONG THE LENGTH OF THE CATHETER. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 07/2012), G3. H11: H6 (DEVICE, RESULT). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03201 | 2023-07-27 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | SAMPLE RECEIVED BY THE MANUFACTURER. AIR ENTERS THROUGH THE INSERTION STYLET DURING THE PICC INSERTION WHEN ASPIRATING TO CHECK THE BLOOD RETURN, AND WHEN FLUSHING SALINE DROPS THROUGH THE INSERTION STYLET. THIS HAPPENS SINCE THE RUBBER CAP OF THE STYLET HAS CHANGED. ADDITIONAL INFORMATION RECEIVED 07/18/2023: WHEN ASPIRATING THROUGH THE EXTENSION BAR OF THE STYLET, ONCE THE CATHETER IS INSERTED TO CHECK THAT IT IS IN THE VEIN, AIR ENTERS AND WHEN WASHING, BLOOD OR SERUM LEAKS THROUGH THE STYLET. NO URGENT OR LIFE-THREATENING EVENT OCCURRED, AND NO PATIENT HARM WAS REPORTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05530 | 2023-07-27 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT THE BALLOON RUPTURED AND THE CATHETER FELL OUT DURING BLADDER IRRIGATION ON THE FIRST DAY AFTER SURGERY FOR BENIGN PROSTATIC HYPERPLASIA. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03213 | 2023-07-27 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0602830 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY FRACTURED. REPORTEDLY, THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO AND AN IMAGE WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05539 | 2023-07-27 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE BALLOON WAS DIFFICULT TO INFLATE. A FOLEY CATHETER WAS INSERTED BY A PHYSICIAN AND WHEN THE NURSE TRIED FILLING THE CATHETER WITH STERILE WATER, THE WATER LEAKED FROM THE CATHETER AND THE BALLOON COULD NOT BE INFLATED. THEN THE CATHETER WAS REMOVED AND THE STERILE WATER WAS TRIED TO BE FILLED OUTSIDE THE PATIENT'S BODY IN THE SAME WAY, BUT THE BALLOON COULD NOT BE INFLATED. THEN A NEW CATHETER WAS INSERTED TO THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05547 | 2023-07-27 | 2023-05-30 | I, n, j, u, r, y | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT ON (B)(6)2023, THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO PROGRESSIVE URINATION FATIGUE FOR 2 YEARS, WHICH AGGRAVATED FOR 1 DAY. A THREE-LUMEN URINARY CATHETER WAS PLACED AND THE BLADDER WAS CONTINUOUSLY IRRIGATED, AND THE LIGHT RED LIQUID WAS DRAWN OUT SMOOTHLY. ON (B)(6)2023, THE CATHETER DRAINAGE WAS NOT SMOOTH, AND THE BLADDER WAS FLUSHED MANUALLY ONCE. THE CATHETER BALLOON WAS FOUND TO BE RUPTURED. THE DOCTOR TRIED TO CHANGE A NEW THREE-LUMEN CATHETER BUT FAILED, SO THE PATIENT CONTINUED TO BE OBSERVED AND STOPPED THE BLADDER IRRIGATION. AT 2:00, THE PATIENT CAN URINE BY HIMSELF, AND A LARGE NUMBER OF BLOOD CLOTS WERE DISCHARGED. AS PER THE INJURY PERFORMANCE, THE PATIENT EXPERIENCED A URINARY TRACT BLEEDING. AS PER THE FOLLOW UP INFORMATION RECEIVED ON (B)(6)2023, IT WAS UNDETERMINED IF THE CATHETER CAUSED BLEEDING AND CONFIRMED THAT THE DATE OF EVENT OCCURRENCE WAS ON (B)(6)2023. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05542 | 2023-07-27 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE DISTRIBUTOR RECEIVED SOME COMPLAINTS ABOUT ONE OF THE FOLEY CATHETER. PER ADDITIONAL INFORMATION VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THIS DEVICE WAS NOT USED ON A PATIENT (HENCE NO REPORTED INJURY OR MEDICAL INTERVENTION REQUIRED). THE NURSE TESTED THE PRODUCT PRIOR TO PATIENT USE, AND THE CATHETER WOULD NOT DEFLATE. THE NURSE PULLED ANOTHER CATHETER WITH THE SAME LOT NUMBER AND GOT THE SAME RESULT, CATHETER WOULD NOT DEFLATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00541 | 2023-07-27 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE OF A CHRONIC TOTAL OCCLUSION IN THE SUPERFICIAL FEMORAL ARTERY, THE TIP OF THE CATHETER ALLEGEDLY SEPARATED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 07/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03221 | 2023-07-27 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0602680 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA THE RIGHT SUBCLAVIAN VEIN, THE GUIDEWIRE ALLEGEDLY PARTIALLY BROKE IN THE DISTAL REGION. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 02/2026). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05544 | 2023-07-27 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THERE WAS A PINHOLE IN THE TUBING OF THE URINE COLLECTION BAG NEAR THE FOLEY CATHETER, CAUSING URINE LEAKAGE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 11JUL2023, REPRESENTATIVE HAD IDENTIFIED TEAR IN THE INFLATION FUNNEL. NO DAMAGE WAS OBSERVED TO THE BAG'S INLET TUBE. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. PHOTO: THERE ARE 3 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE PHOTOS ATTACHED ARE AN OVERVIEW OF A URINARY DRAINAGE BAG ATTACHED TO A 2-WAY CATHETER. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), AN OPEN URINARY DRAINAGE BAG ATTACHED TO A 2-WAY CATHETER. VISUAL INSPECTION NOTED A TEAR MEASURING 0.1165" FOUND UNDER THE INFLATION CAP. THIS IS OUT OF SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INCORRECT MOLDING OF EITHER THE INFLATION VALVE OR RETENTION CAP. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05550 | 2023-07-27 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 50616 |  | IT WAS REPORTED THAT THE PATIENT HAD TO PUT LUBE ON THE MAGIC3 HYDROPHILIC COUDE MALE INTERMITTENT CATHETERS, BUT IT WAS PAINFUL TRYING TO GET THEM TO PASS. IT WAS STATED THAT THE PATIENT WAS WONDERING WHY THE PACKET WAS AT THE BOTTOM, WHEN THEY GOT THE 1ST 50616. THE PATIENT DID NOT GET INSTRUCTIONS. AND NO ONE TOLD THEM THAT. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05540 | 2023-07-27 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT FIXATION WATER WAS UNABLE TO REMOVE. AFTER INJECTING FIXATION WATER AND INFLATING THE BALLOON DURING A PRETEST, THE USER ATTEMPTED TO REMOVE THE WATER, BUT IT WAS UNABLE TO REMOVE. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, THE BALLOON INFLATED WELL WITHOUT NO ISSUES. AFTER THE BALLOON INFLATION, THE WATER IN THE BALLOON WAS UNABLE TO REMOVE WELL AND THE BALLOON BECAME IN ONE-SIDED BULGE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03215 | 2023-07-27 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROVPPSOLO 4FDL MBP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1294108D5 |  | IT WAS REPORTED BY CUSTOMER THAT SALINE LEAKS FROM T LOCK AREA ECG AND SEPTUM AREA. ADDITIONAL INFORMATION FROM CUSTOMER RESPONSE ON 7-19-2023: IT WAS REPORTED "WE NOTICED THAT WHEN WE WENT TO FLUSH THE PICC LINE, THE FLUSH WOULD COME OUT AROUND THE WIRE ON THE END CAP INSTEAD OF FLUSHING THE LINE. WE HAVE NOTICED THIS ON MANY PICC LINES. WE WOULD PUT OUR FINGER OVER THE END CAP TO FLUSH." CUSTOMER REPORTED THIS HAS HAPPENED ON MANY LINES HOWEVER THE EXACT NUMBER OF OCCURRENCES WAS NOT PROVIDED TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03228 | 2023-07-28 | 2023-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F DUAL-LUMEN CATHETER BASIC TRAY (NITINOL GW 70CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8275355 |  | IT WAS REPORTED PATIENT INDICATED FOR A POWER PICC PLACEMENT. IT WAS REPORTED THAT AT THE BEGINNING OF A CATHETER PLACEMENT PROCEDURE, WITH THE MINIMUM EFFORT THE RED DILATOR GOT DEFORMED (TWISTED), AND THEREFORE THE DEVICE COULD NOT BE USED, AND WAS EXCHANGED TO A NEW ONE. THERE WAS NO PATIENT IMPACT, PATIENT CURRENT STATE IS STABLE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05591 | 2023-07-28 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® FLIP-FLO¿ CATHETER VALVE | FLIP FLOW CATHETER VALVE | C.R. BARD, INC. (COVINGTON) -1018233 |  | BFF20 |  | IT WAS REPORTED THAT IN THE ONLINE SURVEY A PHYSICIAN STATED THAT NO, THEY WERE NOT ABLE TO SUCCESSFULLY USE THE DEVICE FOR URINE DRAINAGE BECAUSE IT HAD TROUBLE DRAINING THE URINE IN RELATION TO BFF20 ¿ FLIP-FLO CATHETER VALVE WITH 180° LEVER TAP. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05594 | 2023-07-28 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175812 |  | IT WAS REPORTED THAT THE REPORTER WAS THE STAFF NURSE, WHO HANDLED THE INCIDENT. THE PATIENT WAS DESCRIBED AS A FIDGETY LADY, THE CATHETER WAS NOT CORRECTLY SECURED AS THE STAFF HAD PRESUMED THAT THE CATHETER BAG LEG STRAPS WERE SUFFICIENT. WHEN THE NURSE WENT TO WASH THE PATIENT AND TURNED HER, THEY NOTICED THAT THE CATHETER WAS BETWEEN THE PATIENT'S LEGS AND THAT THE CATHETER BAG STRAPS HAD COME OFF. THE CATHETER BALLOON WAS ONLY PARTIALLY INFLATED AT THIS POINT. THE CATHETER HAD INITIALLY BEEN INSERTED ON (B)(6) 2023. THE PATIENT DID NOT REPORT ANY PAIN AND SEVERAL HOURS LATER THEY HAD STILL NOT BEEN RE-CATHETERISED AS THEY NO LONGER APPEARED TO BE EXPERIENCING URINARY RETENTION. THE ORIGINAL REPORTER DID NOT RETAIN THE PRODUCT FOR INVESTIGATION. IT WAS NOTED THAT THE CATHETER WAS MALE CATHETER BUT IT WAS USED ON FEMALE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05569 | 2023-07-28 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS BURST. AS PER THE ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON 12JUL2023, THERE WAS NO BALLOON RUPTURE OBSERVED ON THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03230 | 2023-07-28 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SLIM-CATH STRAIGHT SHORT-TERM DIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5853240J |  | IT WAS REPORTED THAT BECAUSE OTHER MANUFACTURER'S CATHETER ALREADY IMPLANTED IN THE RIGHT FEMORAL VEIN HAD TO BE REPLACED DUE TO POOR BLOOD ASPIRATION ETC., WHEN THE DILATOR WAS ATTEMPTED TO BE INSERTED OVER THE GUIDEWIRE, THE DILATOR COULD NOT BE ADVANCED FROM THE MIDDLE. THEREFORE, THE INSERTION PROCEDURE OF THE CATHETER WAS STOPPED. THERE WAS NO REPORTED PATIENT INJURY. THE CUSTOMER REPORTED TWO INCIDENTS HOWEVER ONLY ONE DEVICE WAS RETURNED FOR EVALUATION. THIS FILE ADDRESSES THE DEVICE THAT WAS RETURNED. (B)(6) 2023 - THE RETURNED SAMPLE EXHIBITED MULTIPLE VISIBLE KINKS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DIFFICULTY INSERTING A DILATOR OVER A GUIDEWIRE IS CONFIRMED. THE RETURNED PRODUCT WAS ONE 24 CM POWERTRIALYSIS DIALYSIS CATHETER, ONE 14 CM DUALATOR DILATOR, ONE 13 CM DUALATOR DILATOR, AND ONE 0.035 IN J-TIP GUIDEWIRE IN IT¿S PLASTIC HOOP. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUES THROUGHOUT THE GUIDEWIRE WITH MULTIPLE VISIBLE KINKS ALONG THE GUIDEWIRE. BOTH DILATORS EXHIBITED TIP DEFORMATION AND HAD USE RESIDUES ALONG THE DEVICES. A MICROSCOPIC OBSERVATION REVEALED THE TIP DEFORMATION ALONG THE DILATORS TO APPEAR PUSHED INWARD, BUT NOTHING ELSE REMARKABLE WAS OBSERVED REGARDING THE DILATORS. A MICROSCOPIC OBSERVATION OF THE RETURNED GUIDEWIRE REVEALED THAT THE DISTAL TIP OF THE GUIDEWIRE WAS INTACT WITH BOTH WELD TIPS, AND NO BREAKS WERE OBSERVED ALONG THE LENGTH OF THE DEVICE. ABUNDANT WIRE DEFORMATION WAS VISIBLE INCLUDING ONE VERY SHARP KINK ABOUT MIDWAY ALONG THE LENGTH OF THE GUIDEWIRE. MULTIPLE OTHER KINKS WERE VISIBLE ALONG THE GUIDEWIRE WITHOUT ANY COMPLETE BREAKS IN THE DEVICE. A FUNCTIONAL TEST OF ATTEMPTING TO ADVANCE BOTH DILATORS AND THE RETURNED CATHETER OVER A NON-COMPLAINANT 0.035 IN J-TIP GUIDEWIRE WERE SUCCESSFUL AND UNREMARKABLE. THE COMPLAINT OF DIFFICULTY INSERTING DILATORS OVER A GUIDEWIRE IS CONFIRMED AND WAS DETERMINED TO LIKELY BE CAUSED BY WIRE MANIPULATION DURING REPLACEMENT OF THE CATHETER. ATTEMPTED GUIDEWIRE INSERTION AGAINST RESISTANCE MAY POTENTIALLY CAUSE DAMAGE TO THE GUIDEWIRE, MAKING IT DIFFICULT TO SLIDE OTHER DEVICES OVER THE GUIDEWIRE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05597 | 2023-07-28 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L20 |  | IT WAS REPORTED, THAT THE FOLEY WAS LEAKING FROM THE PORT WHERE IT WAS INSERTED. REPRESENTATIVE RECOMMENDED, PATIENT CALL DOCTOR AND SEE IF AN INCREASE WAS SIZE WAS NECESSARY. AND WE WILL FOLLOW DOCTOR'S INSTRUCTIONS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE, A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05595 | 2023-07-28 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175812 |  | IT WAS REPORTED, THAT THE CATHETER FELL OUT OF THE FEMALE PATIENT. THE REPORTER WAS NOT THE NURSE WHO DEALT WITH THE INCIDENT, BUT THIS WAS RELAYED TO THEM, DURING MORNING HANDOVER. THE INCIDENT OCCURRED AT AN UNSPECIFIED TIME, DURING NIGHT SHIFT. AND THE PATIENT DESCRIBED AS HAVING CAPACITY. AS THE REPORTER WAS NOT DIRECTLY INVOLVED IN THE EVENT, THEY WERE NOT ABLE TO PROVIDE FURTHER INFORMATION, INCLUDING THE BATCH NUMBER OF THE AFFECTED PRODUCT. THE REPORTER WAS UNSURE WHETHER. THE CATHETER HAD BEEN PART OF A FOLEY TRAY SET OR WHICH SIZE IT HAD BEEN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE, A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05588 | 2023-07-28 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT DURING A PRETEST PERFORMED PRIOR TO THE DEVICE PLACEMENT ON THE PATIENT, THE STERILE WATER LEAKED FROM THE CONNECTION BETWEEN THE SYRINGE AND THE INFLATION VALVE OF FOLEY CATHETER. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 12JUL2023, STATED THAT THE EVENT MIGHT HAVE BEEN A SYRINGE ISSUE BUT THERE WAS NO OBVIOUS BASIS FOR DETERMINING THAT THE SYRINGE WAS DEFECTIVE, SO BOTH THE CATHETER AND SYRINGE SHOULD BE INVESTIGATED AND THERE WAS NO FITTING ISSUE REPORTED. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. DHR IS NOT REQUIRED FOR THIS COMPLAINT. A RISK REVIEW AND LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED, | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05575 | 2023-07-28 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE BALLOON, WHEN THE BALLOON WAS CHECKING DURING INFLATION. THE EVENT OCCURRED DURING PRETEST. THERE WAS NO BALLOON RUPTURE OR TEAR. NONE OF THE MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05572 | 2023-07-28 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165816 |  | IT WAS REPORTED THAT THE SURVEY MEMO WAS ENCLOSED WITH THE ACTUAL SAMPLE. IT WAS STATED THAT WATER LEAKED FROM THE FUNNEL BIFURCATION. IT WAS ALSO STATED THAT WHEN THE FOLEY CATHETER WAS REPLACED WITH A SILICONE TYPE, THERE WERE SYMPTOMS LIKE AIR LEAKAGE OBSERVED, BUT URINE DRAINAGE WAS CONFIRMED. WHEN THE ANTICANCER DRUG WAS INJECTED, WATER LEAKAGE WAS CONFIRMED. THERE WAS A PINHOLE AT THE FUNNEL BIFURCATION. IT WAS NOTED THAT AN ILLUSTRATION WAS PROVIDED IN THE MEMO ENCLOSED WITH THE SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05571 | 2023-07-28 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | WIDEBAND¿ SELF-ADHERING MALE EXTERNAL CATHETER | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 36103 |  | IT WAS REPORTED THAT THE GLUE SUBSTANCE WAS STUCK TO APPLICATION AREA, CAUSING RAW SKIN AND WAS UNCOMFORTING. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT PATIENT'S PROSTATE WAS REMOVED TWO MONTHS AGO. THE BATCH OF MALE EXTERNAL CATHETERS THEY RECEIVED HAD FAR TOO MUCH ADHESIVE TO THE POINT WHERE THEY HAD TO USE SCISSORS TO GENTLY CUT AWAY ANY RESIDUAL RESIDUE. THEY USED RUBBING ALCOHOL AS WELL, BUT IT DID LITTLE TO RESOLVE THE ISSUE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05592 | 2023-07-28 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119316 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05568 | 2023-07-28 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT THE FOLEY TRAY WAS NOT SEALED PROPERLY. THEREFORE, IT WAS UNUSABLE WHEN PATIENT WENT TO HAVE IT CHANGED. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PRODUCT PACKAGING OBVIOUSLY NOT SEALED OR TORN, I.E. OUT OF BOX FAILURE'. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05579 | 2023-07-28 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE LUER LOCK PORT BROKE OFF OF THE URINARY CATHETER WHILE APPLYING AN EQUASHIELD, HAD TO REMOVE URINARY CATHETER AND PLACE A NEW ONE TO ADMINISTER IVES CHEMOTHERAPY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05553 | 2023-07-28 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53614G |  | IT WAS REPORTED THAT THE 1 INTERMITTENT CATHETER IN BOX HAD A DEFECT AT THE TIP OF IT, SO PATIENT DID NOT USE THE INTERMITTENT CATHETER. PER FOLLOW UP VIA PHONE (B)(6) 2023, THE END HAD A PROTRUDING TIP. THE LOT NUMBER WAS JUGZ20901. THE DEFECT WAS NOTICED BEFORE USE. THERE WAS ALSO TWO IN ANOTHER SHIPMENT WITH REFERENCE NUMBER 53614G AND LOT JUGW0050. CUSTOMERS ADVISED THEY WOULD FORWARD PICTURES OF THE MOST RECENT INCIDENT. | THE REPORTED EVENT IS CONFIRMED - MANUFACTURING RELATED. VISUAL INSPECTION NOTED TWO (2) PHOTO SAMPLE. THE FIRST PHOTO SAMPLE SHOWCASES A NICK ON THE EDGE OF THE ISC. THE SECOND PHOTO SAMPLE SHOWCASES ISC IN CLOSED PACKAGING. BASED ON THE PHOTO SAMPLES RECEIVED, THE PRODUCT DOES NOT MEET SPECIFICATIONS. THE POTENTIAL ROOT CAUSE COULD BE MANDRELS TOUCHING EACH OTHER WITHIN THE PALLET. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05560 | 2023-07-28 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT STERILE WATER DID NOT DRAIN OUT FROM THE FOLEY CATHETER DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05564 | 2023-07-28 | 2023-06-05 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT WATER IN THE FOLEY CATHETER BALLOON COULD NOT BE DRAINED. ATTEMPTS WERE MADE TO CUT THE CATHETER AND PERFORATE THE BALLOON WITH A GUIDEWIRE, BUT THE CATHETER COULD NOT BE REMOVED. THEN, THE CATHETER WAS WITHDRAWN SPONTANEOUSLY. PER ADDITIONAL INFORMATION RECEIVED FROM PHOTO SAMPLE PROVIDED BY IBC ON 19JUL2023, STATED THAT THEY REPLACED DISTILLED WATER TO CHECK FIXED WATER ON (B)(6) 2023. 10 ML WAS COLLECTED, BUT ONLY 5 ML WAS FILLED. THERE WAS RESISTANCE. THE SAME APPLIED WHEN PERFORMED BY THE PHYSICIAN ON DUTY. ON (B)(6) 2023,THE UROLOGIST ATTEMPTED TO REPLACE THE CATHETER, BUT AS ON THE PREVIOUS DAY, THE FIXED WATER CANNOT BE RETRIEVED AND ATTEMPTS WERE MADE TO CUT THE CATHETER AND PUNCTURE THE BALLOON WITH A GUIDEWIRE, BUT WITHOUT SUCCESS. EARLY MORNING ON (B)(6) 2023, THE CATHETER SPONTANEOUSLY DROPPED OUT AND THE BALLOON DEFLATED. ULTRASOUND CONFIRMED THAT THE BALLOON WAS INFLATED WITHIN THE PROSTATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05561 | 2023-07-28 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ALL SILICONE CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CATHETER WAS LEAKING IN POST-ANESTHESIA CARE UNIT (PACU). WHEN THE PATIENT GOT TO THE FLOOR, THERE WAS A HOLE IN THE Y SILICONE SECTION OF THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05556 | 2023-07-28 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT DURING A PRETEST, THE STERILE WATER LEAKED FROM THE CATHETER TIP. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 1(B)(6) 2023, THERE WAS NO VISIBLE DAMAGE ON THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05557 | 2023-07-28 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT IMMEDIATELY AFTER PLACEMENT, THE FOLEY CATHETER FELL OFF DUE TO STERILE WATER LEAK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05559 | 2023-07-28 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE FOLEY CATHETER DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03226 | 2023-07-28 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE GUIDEWIRE WAS ALLEGEDLY STUCK INTO THE INTRODUCER NEEDLE. IT WAS FURTHER REPORTED THAT WHEN THE GUIDEWIRE WAS ATTEMPTED TO BE WITHDRAWN, THE GUIDEWIRE WAS ALLEGEDLY DAMAGED AND STRETCHED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 01/2025). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03225 | 2023-07-28 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT THREE MONTHS POST A DIALYSIS CATHETER PLACEMENT, SILICON EXTENSION TUBING OF THE CATHETER ALLEGEDLY BROKE ON THE ONE SIDE. IT WAS FURTHER THAT THE BLOOD ALLEGEDLY LEAKED DURING DIALYSIS. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE GLIDEPATH D/L 23CM CATHETER WAS RECEIVED FOR EVALUATION AND ONE ELECTRONIC PHOTO AND VIDEO WAS PROVIDED FOR REVIEW. VISUAL, TACTILE, MICROSCOPIC AND FUNCTIONAL EVALUATIONS WAS PERFORMED. A PARTIAL C-SHAPED BREAK WAS NOTED TO THE DISTAL END OF THE BLUE LUER EXTENSION LEG. UPON INFUSION, A LEAK FROM THE PARTIAL C-SHAPE BREAK ON THE EXTENSION LEG WAS OBSERVED WHILE DYE WATER EXITED THE DISTAL END OF THE CATHETER. ASPIRATION WAS ATTEMPTED AND WAS UNSUCCESSFUL. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND FLUID LEAK ISSUES AND THE PHOTO AND VIDEO REVIEW ALSO CONFIRMS THE SAME. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 07/2024), G3. H11: H6 (METHOD, RESULT, CONCLUSION) . H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03224 | 2023-07-28 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE GUIDEWIRE WAS ALLEGEDLY STUCK ON THE PACKAGE. IT WAS FURTHER REPORTED THAT, THE DEVICE WAS ALLEGEDLY UNABLE TO ADVANCE. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 19CM GLIDEPATH D/L CATHETER KIT. FOLLOWING COMPONENTS WERE RECEIVED: ONE 19CM GLIDEPATH D/L CATHETER, ONE 8FR DILATOR, ONE 12FR DILATOR, TWO STYLET SHAFTS, ONE J-TIP GUIDEWIRE LOADED TO AN INTRODUCER NEEDLE IN A GUIDEWIRE HOOP, ONE EMPTY GUIDEWIRE HOOP, ONE TUNNELER AND ONE INTRODUCER NEEDLE WERE RETURNED FOR EVALUATION. THE J-TIP GUIDEWIRE WAS NOTED TO HAVE RESIDUE ON THE EXPOSED DISTAL END AND WAS NOTED TO BE STUCK WITHIN THE LOADED INTRODUCER NEEDLE. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED PHYSICAL RESISTANCE AND IDENTIFIED STRETCHED ISSUE. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE TO ADVANCE ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF REPORTED EVENT WAS UNKNOWN. BASED ON THE MEASUREMENTS RECORDED DURING SAMPLE EVALUATION AND APPLICABLE DRAWINGS, NO MEASUREMENTS RECORDED COULD BE CONFIRMED TO BE OUT OF TOLERANCE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 06/2024). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05581 | 2023-07-28 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETERS CAUSED RASH TO THE PATIENT. PER FOLLOW UP VIA PHONE (B)(6) 2023, CUSTOMER WAS ON ANTIBIOTICS FOR THREE WEEKS WITH USE OF THE PUREWICK BECAUSE THEY HAD A SEVERE UTI (URINARY TRACT INFECTION). CUSTOMER WAS NO LONGER SING THE PUREWICK AS THEIR DOCTOR TOLD THEM THAT IT CAUSED THEIR UTI. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
3006260740-2023-03231 | 2023-07-28 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SLIM-CATH STRAIGHT SHORT-TERM DIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5853240J |  | IT WAS REPORTED, THAT BECAUSE OTHER MANUFACTURER'S CATHETER ALREADY IMPLANTED IN THE RIGHT FEMORAL VEIN HAD TO BE REPLACED, DUE TO POOR BLOOD ASPIRATION ETC.. WHEN THE DILATOR WAS ATTEMPTED TO BE INSERTED OVER THE GUIDEWIRE, THE DILATOR COULD NOT BE ADVANCED FROM THE MIDDLE. THEREFORE, THE INSERTION PROCEDURE OF THE CATHETER WAS STOPPED. THERE WAS NO REPORTED PATIENT INJURY. THE CUSTOMER REPORTED TWO INCIDENTS. HOWEVER, ONLY ONE DEVICE WAS RETURNED FOR EVALUATION. THIS FILE ADDRESSES THE DEVICE THAT WAS NOT RETURNED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
MW5120349 | 2023-07-28 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | POWERGLIDE PRO MIDLINE CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. | F218108PT |  |  | PER THE USER, "INSIDE THE POWERGLIDE HOUSING, THERE IS SUPPOSED TO BE A GUIDEWIRE. THE GUIDEWIRE IS MISSING." |  | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03227 | 2023-07-28 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716000J |  | IT WAS REPORTED THAT AFTER THE PORT SYSTEM WAS IMPLANTED, BLOOD RETURN COULD NOT BE CONFIRMED. IT WAS FURTHER REPORTED THAT WHEN THE PORT BODY AND THE CATH LOCK WERE CHANGED, BLOOD RETURN COULD BE CONFIRMED. REPORTEDLY, THE PORT SYSTEM WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE IMPLANTABLE SLIM POWERPORT WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPLETE DIAGONAL BREAK WAS NOTED ON THE DISTAL END OF THE ATTACHED CATHETER AND THE DISTAL CATHETER SEGMENT WAS NOT RETURNED. THE EDGES OF THE BREAK WERE NOTED TO BE UNEVEN AND THE SURFACE APPEARED TO BE GRANULAR. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED FRACTURE AND MATERIAL SEPARATION ISSUES. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED SUCTION ISSUE AS THE EXACT CIRCUMSTANCE AT THE TIME OF THE EVENT REPORTED WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 08/2024) . H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
MW5120354 | 2023-07-28 | 2023-07-22 | I, n, j, u, r, y | Voluntary report | POWERGLIDE PRO MIDLINE CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. | F322087T |  |  | WAS PLACING A 22G X 8CM POWERGLIDE PRO MIDLINE CATHETER IN THE LEFT FOREARM USING ULTRASOUND GUIDANCE. ONCE IN THE VEIN WIRE THREADED IN WITH NO ISSUES, ATTEMPTED TO THREAD IN CATHETER AND MET RESISTANCE ABOUT 1/3 OF THE WAY IN. ASSESSED THAT THE NEEDLE WAS STILL IN THE VEIN BUT CATHETER WOULD NOT ADVANCE THE CATHETER WAS ATTEMPTED TO BE RETRACTED BACK ON THE NEEDLE. RESISTANCE FELT AND HAD TO RETRACT CATHETER, NEEDLE AND GUIDEWIRE ALL TOGETHER WITH MODERATE PRESSURE TO REMOVE DEVICE. ONCE REMOVED NOTICED THAT 2/3 OF THE CATHETER WAS RETAINED IN THE PATIENT. PATIENT WAS TAKEN TO SURGERY AND A VASCULAR SURGEON REMOVED THE RETAINED CATHETER. |  | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03255 | 2023-07-31 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE SIXTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03253 | 2023-07-31 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE FOURTH DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03256 | 2023-07-31 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE SEVENTH DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03250 | 2023-07-31 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03251 | 2023-07-31 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION . | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03259 | 2023-07-31 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE TENTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03257 | 2023-07-31 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE EIGHTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03264 | 2023-07-31 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED, THAT THE DILATOR WAS RUPTURED. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED, THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03261 | 2023-07-31 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE TWELFTH DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03260 | 2023-07-31 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE ELEVENTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03263 | 2023-07-31 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE DILATOR WAS RUPTURED. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03262 | 2023-07-31 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE DILATOR WAS RUPTURED. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05624 | 2023-07-31 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT FIXATION WATER LEAKED FROM THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05617 | 2023-07-31 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE CATHETER WAS BENT DURING THE TRAY PACKAGING. ACCORDING TO THE PACKAGING CONSIDERATIONS, THE CATHETER WAS BENT AND HAD A CREASE (BENT MARK) ON IT. THE BENT MADE IT DIFFICULT FOR THE URINE TO FLOW AND SO THE URINE LEAKED BETWEEN THE CATHETER AND URETHRAL OPENING. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 20JUL2023, STATED THAT WHEN THE PRODUCT REACHED THE CUSTOMER, THE CATHETER WAS ALREADY BENT IN ITS PACKAGING. IT WAS UNCLEAR AT WHAT STAGE OF THE PACKAGING PROCESS THE CATHETER WAS BENT. THE JAPANESE DESCRIPTION STATES ", AT THE TIME OF PACKAGING," MEANING "IN THE PACKAGED STATE¿. CLARIFIED THAT THE CATHETER WAS USED EVEN AFTER THE BENT WAS NOTICED AND THE CUSTOMER BELIEVES THAT THE CATHETER WAS LEAKING DUE TO BENT. AS THE ORIGINAL SAMPLE WAS DISCARDED AT THE FACILITY, IT COULD NOT BE CONFIRMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03248 | 2023-07-31 | 2023-07-07 | I, n, j, u, r, y | Manufacturer report | POWERGROSHONG CATHETER WITH SHERLOCK 3CG TPS STYLET FULL KIT (5F) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000517 |  | IT WAS REPORTED CATHETER BREAK AT 5 CM. PATIENT HAD TO UNDERGO THE REMOVAL OF THE INTRAVASCULAR PART OF THE DEVICE (37 CM) WITH AN ANGIORADIOLOGICAL PROCEDURE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03249 | 2023-07-31 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT AFTER INSERTION OF THE CATHETER, IT WAS NOT POSSIBLE TO WASH, THE CATHETER WAS NOT WORKING. IT WAS NECESSARY TO REPUNCTURE THE PATIENT. NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL DETAILS PROVIDED 11.SEP.2023: I NO LONGER HAVE THE PATIENT'S PERSONAL DATA. THIS WAS AN ONCOLOGY PATIENT, WHO WAS EASILY PUNCTURED USING THE MODIFIED SELDINGER TECHNIQUE, HOWEVER, WHEN PATENCY WAS TESTED, THE CATHETER SHOWED GOOD FLOW AND HAD NO REFLUX. WE PULLED THE CATHETER AND REINSERTED IT FOUR TIMES, WITHOUT SUCCESS. WE REMOVED THE CATHETER AND TESTED REFLUX WITH SALINE SOLUTION AND OBSERVED THAT IT DID NOT PRESENT REFLUX. WE OPENED ANOTHER CATHETER FOR IMPLANTATION IN THE CHILD. | H11: THE INITIAL COMPLAINT APPEARED TO BE A REPORTABLE OCCURRENCE. ONCE THE SAMPLE WAS RETURNED AND EVALUATED IT WAS DETERMINED THAT A REPORTABLE EVENT HAD NOT OCCURRED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03246 | 2023-07-31 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TPS STYLET | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED, "WHEN PREPARING THE CATHETER, THE OPERATOR FEELS RESISTANCE IN PULLING THE GUIDE UP INTO THE DEVICE TO REACH THE TIP, AND THE MOMENT THE GUIDE IS PULLED BACK, THE CATHETER 'CURLS'. THEREFORE THIS DEVICE WAS NOT USED AS A PRECAUTION AND WAS REPLACED". 1 UNIT, THE DEFECT OCCURRED ONCE, DURING USE. NO SERIOUS INJURY, ERRONEOUS RESULTS, COURSE OF TREATMENT CHANGED, EXPOSURE TO BLOOD/BODILY FLUID, MEDICAL INTERVENTION OTHER THAN FIRST AID, NEEDLE/PROBE STICK, SAFETY ISSUE, OTHER ACTIONS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03245 | 2023-07-31 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT A CLINICIAN ON SOCIAL MEDIA IS PROMOTING DAMAGED/LEAKING PICC CATHETERS. A SPECIFIC NUMBER OF AFFECTED DEVICES WAS NOT PROVIDED BY THE COMPLAINANT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05623 | 2023-07-31 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OFF. IT WAS STATED THAT THE CATHETER WAS PLACED TO THE PATIENT IN THE UROLOGY WARD. AFTER INSERTED, IT WAS REMOVED SPONTANEOUSLY FOR A WHILE. WHEN THE CATHETER WAS CHECKED, THE BALLOON PART WAS DEFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03234 | 2023-07-31 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BD TRIALYSIS/HD CATHETER | CATHETER, HEMODIALYSIS, DUAL LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED DURING CLINICAL ASSESSMENT: 'FLUID LEAKING FROM CENTRAL LINE INSERTION SITE." DOUBLE LUMEN BD TRIALYSIS/HD CATHETER IN THE GROIN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03233 | 2023-07-31 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC FT CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F MAXIMAL BARRIER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275108FD |  | IT WAS REPORTED BY CUSTOMER, THAT DL PICC KITS, WHERE THE WIRE COMES OUT OF. THE TOP PART (RUBBER PIECE), ON THE RED PORT AT THE TOP, WHERE THE WIRE COMES OUT. WHEN WE FLUSH SALINE IS LEAKING OUT OF THE TOP. WHEN WE DRAW BACK OUR BLOOD RETURN, AIR IS COMING INTO OUR SYRINGES. THIS WAS REPORTED TO ME, THIS MORNING. AND I HAD THEM OPEN A KIT AND SHOW ME. WHEN YOU FLUSH, SALINE JUST LEAKS OUT OF THE TOP OF THAT RUBBER PIECE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05628 | 2023-07-31 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 50614 |  | IT WAS REPORTED THAT THE HYDRO PACKETS WERE EMPTY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05625 | 2023-07-31 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAD BROKEN INSIDE OF PATIENT ON TWO DIFFERENT OCCASIONS, BALLOON LOOKED TO HAVE A LARGE SLIT DOWN THE SIDE. PATIENT DID NOT PULL ON CATHETER, ONE TIME THE CATHETER FELL OUT WHILE PATIENT WAS SHOWERING. ON A SEPARATE OCCASION THE NURSE WAS GOING TO INFLATE THE BALLOON AND THE PORT WHERE THE SALINE WAS INSTILLED FELL OFF AND SALINE SPRAYED EVERYWHERE. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINHOLE IN BALLOON OR CUFF. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER ATTACHED TO A DRAINAGE BAG. VISUAL INSPECTION NOTED BALLOON BURST MEASURING 0.2990" WITH NO MISSING PIECES NOTED. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE WHICH STATES, "BALLOON MUST NOT BE TORN". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿BARD® SILVER LUBRI-SIL® FOLEY TRAY: SINGLE USE ONLY: [WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES]: BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. MATERIAL: BALLOON CATHETER: SILICONE; SILVER COATING; THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO 22 EVERY 2FR. BALLOON CATHETER; FOLEY CATHETER; TEMPERATURE-SENSING CATHETER. ACCESSORIES: CLOSED DRAINAGE BAG; (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.); SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT; ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION; TWEEZERS; GAUZE PADS; WATERPROOF SHEET; COTTON BALLS; GLOVES; STATLOCK FOLEY. [INTENDED USE & EFFECT- EFFICACY]: THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE]: METHOD OF USE: CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. PRECAUTIONS FOR USE: TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT: OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING: CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY.] DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING: REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING: DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS]: PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) BALLOON-RUPTURE METHOD: WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. NON-RUPTURE METHOD: ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. BALLOON-RUPTURE METHOD: INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. MRI SAFETY: BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. MALFUNCTION AND ADVERSE EVENTS: MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE; DIFFICULTY OR FAILURE TO REMOVE THE DEVICE; OCCLUSION OF CATHETER INNER LUMENS; ENCRUSTATION; ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE; DEVICE DAMAGE DUE TO INAPPROPRIATE USE; FAILURE TO MEASURE TEMPERATURE; IMPROPER TEMPERATURE INDICATION. ADVERSE EVENTS: URINARY-TRACT INFECTION; HEMORRHAGE, HEMATURIA; ALLERGY REACTION TO THE DEVICE; CALCULUS FORMATION; EDEMA; PAIN; DISCOMFORT; INJURY OF BLADDER OR URETHRAL; URETHRITIS, URINARY INCONTINENCE; RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]: STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05621 | 2023-07-31 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE BASE OF THE FUNNEL DURING PRE-TEST IN THE OPERATING ROOM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05609 | 2023-07-31 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE BALLOON OF THE FOLEY CATHETER WAS CRACKED AND THE STERILE WATER COULD NOT BE FILLED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03235 | 2023-07-31 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC FT CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F MAXIMAL BARRIER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275108FD |  | IT WAS REPORTED BY CUSTOMER THAT DL PICC KITS WHERE THE WIRE COMES OUT OF THE TOP PART (RUBBER PIECE) - ON THE RED PORT AT THE TOP WHERE THE WIRE COMES OUT, WHEN FLUSH SALINE IS LEAKING OUT OF THE TOP. WHEN DRAW BACK OUR BLOOD RETURN, AIR IS COMING INTO OUR SYRINGES. THIS WAS REPORTED TO ME THIS MORNING AND I HAD THEM OPEN A KIT AND SHOW ME. WHEN FLUSH, SALINE JUST LEAKS OUT OF THE TOP OF THAT RUBBER PIECE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03238 | 2023-07-31 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5613150J |  | IT WAS REPORTED THAT AFTER CONFIRMING THE PATENCY OF THE CATHETER, BLOOD LEAKED NEAR THE BASE OF THE EXTENSION TUBE OF THE RED LUER CONNECTOR SIDE APPROXIMATELY 5 MINUTES AFTER THE REVERSE CONNECTION. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03237 | 2023-07-31 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT CATHETER LEAKED FROM THE RUBBER-STOPPER WHEN PICC LINE WAS BEING INSERTED. ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023: LEAK DISCOVERED DURING INSERTION. CATHETER SECUREMENT WAS STAT LOCK. IT WAS STATED THE EVENT DID NOT INVOLVE AN URGENT OR LIFE THREATENING MEDICAL SITUATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03242 | 2023-07-31 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC HF CATHETER W/ SHERLOCK 3CG (TPS) STYLET (5F) (TL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000795 |  | IT WAS REPORTED BY CUSTOMER THAT ITEM WITHIN GAUZE IN STERILIZED PACKAGING. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03244 | 2023-07-31 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED THAT 4FR PROVENA PICC CLAMP ON PURPLE LUMEN CRACKED CAUSING CLAMP NOT TO FASTEN. ADD INFO RECEIVED 17-JUL-2023 EVENT DID NOT INVOLVE URGENT/LIFE THREATENING EVENT THE CRACK WAS DISCOVERED WHEN THE NURSE TRIED TO CLAMP THE CLAMP MULTIPLE TIMES AND IT KEPT COMING OPEN. WHEN SHE LOOKED AT IT, SHE NOTICED THE CRACK. THE EVENT CAUSED THE PATIENT TO HAVE AN OVER THE WIRE EXCHANGE OF THE PICC. NO DELAY IN CARE. NO SIGNS, SYMPTOMS OR COMPLICATIONS TO PATIENT. STAT LOCK SECUREMENT WAS USED. NO LEAKAGE OF MEDICATION/FLUID BECAUSE THE CRACK IN THE CLAMP WAS NOTICED IMMEDIATELY AFTER LINE WAS PLACED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED CLAMP WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR DUAL LUMEN POWERPICC PROVENA CATHETER. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. A FRACTURE WAS OBSERVED IN THE PURPLE-LUERED LUMEN CLAMP. MICROSCOPIC INSPECTION OF THE FRACTURE SITE REVEALED A GRANULAR FRACTURE SURFACE. NO DEFORMATION WAS OBSERVED IN THE VICINITY OF THE FRACTURE. SLIGHT DISCOLORATION WAS OBSERVED. ATTEMPTS TO REPLICATE THE OBSERVED DAMAGE BY MANIPULATING THE CLAMP OF A NON-COMPLAINANT DEVICE RESULTED IN ABUNDANT DEFORMATION AND DISCOLORATION. THE FRACTURE FEATURES WERE CONSISTENT WITH BRITTLE MATERIAL FAILURE. ATTEMPTS TO REPLICATE THE FRACTURE FEATURES WERE UNSUCCESSFUL, SUGGESTING THAT AN ADDITIONAL UNIDENTIFIED ENVIRONMENTAL FACTOR(S) LIKELY AFFECTED THE MATERIAL PROPERTIES OF THE CLAMP. CONSEQUENTLY, THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05603 | 2023-07-31 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT IMMEDIATELY AFTER CATHETER PLACEMENT TO THE PATIENT IN THE OPERATING ROOM, THE CATHETER FELL OUT DUE TO STERILE WATER LEAK WHICH COULD HAVE BEEN DUE TO A HOLE IN THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05639 | 2023-07-31 | 2023-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DUE TO THE BALLOON RUPTURE THREE DAYS AFTER SURGERY. | THE REPORTED EVENT WAS CONFIRMED - CAUSE UNKNOWN. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. VISUAL INSPECTION NOTED IRREGULAR BURST WITHOUT MISSING PIECES. INTRODUCED WATER INTO THE INFLATION LUMEN AND DID NOT EXPERIENCE ANY OBSTRUCTIONS TO IMPEDE FLOW. MICROSCOPIC EXAMINATION FOUND THERE WAS PRESENT SALT ACCUMULATION ON THE BALLOON AREA. THE EXACT CAUSE ON HOW AND WHEN PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR) AND ALSO MIGHT BE CONTRIBUTED WITH "HIGH MECHANICAL STABILITY TIME (MST) OF INCOMING LATEX" OR "COAGULANT DIP 2- DIPPING SPEED TOO SLOW". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "LUBRI-SIL® IC INFECTION CONTROL. ALL-SILICONE FOLEY CATHETER. SINGLE USE ONLY. [WARNINGS] 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS: (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS: (1) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] MATERIAL: BALLOON CATHETER: SILICONE; SILVER COATING. THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO 24 EVERY 2FR. 1. BALLOON CATHETER: [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. THE SURFACE OF CATHETER IS COATED FIRST WITH A TRACE AMOUNT OF METALLIC SILVER, AND THEN WITH POLYURETHANE ONTO THAT, TO INHIBIT PROLIFERATION OF BACTERIA THAT MAY ADHERE TO THE CATHETER. [DIRECTIONS FOR USE] 1. METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLE-LESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE: (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD: WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) B. BALLOON-RUPTURE METHOD: 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. 1) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) 2) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) 3) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) 3. MALFUNCTION AND ADVERSE EVENTS: 1) MALFUNCTION: - CATHETER KINKING, DAMAGE, RUPTURE. - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE. - OCCLUSION OF CATHETER INNER LUMENS. - ENCRUSTATION. - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE. - DEVICE DAMAGE DUE TO INAPPROPRIATE USE. 2) ADVERSE EVENTS: - URINARY-TRACT INFECTION. - HEMORRHAGE, HEMATURIA. - ALLERGY REACTION TO THE DEVICE. - CALCULUS FORMATION. - EDEMA. - PAIN. - DISCOMFORT. - INJURY OF BLADDER OR URETHRAL. - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03236 | 2023-07-31 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERGLIDE PRO 18G | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED BY THE CUSTOMER THAT THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT: FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. POWERGLIDE PRO 18G, UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION AND SPECIFIC PATIENT DETAILS WERE NOT DISCLOSED AS PART OF THE SURVEY. AS SUCH, ADDITIONAL INFORMATION AND THE SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05608 | 2023-07-31 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE BALLOON AREA OF THE FOLEY CATHETER WHEN THE WATER WAS INJECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05558 | 2023-07-31 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® URETHRAL CATHETER WHISTLE TIP | URETHRAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 055016 |  | IT WAS REPORTED THAT THEY WERE LONG TIMER USERS OF THE WHISTLE TIP AND UPON STARTING TO USE THE CATHETER IT DID NOT FIT AS A 16FR SHOULD. AS PER THE FOLLOW UP RECEIVED ON 06JUL2023, THE SALES REPRESENTATIVE RESPONDED THAT THE ONE THAT WAS USED AND CLEARLY NOTICED IN USE WAS NOT A 16FR. IT WAS THROWN AWAY. AND THE CUSTOMER WAS GIVEN FOUR UNOPENED CATHETERS THAT THEY DESCRIBED AS 16FR, BUT WHEN COMPARED WITH A DIFFERENT LOT OF 16FR, THEY WERE CLEARLY SMALLER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05605 | 2023-07-31 | 2023-07-07 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WAS GETTING URINARY TRACT INFECTIONS WHILE USING THE PUREWICK FEMALE EXTERNAL CATHETER. THE PATIENT HAS HAD TWO URINARY TRACT INFECTIONS AND NOT SURE WHAT THEY WERE FROM AND THE KIT HAS FILM AND CLOUDY. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. PER FOLLOW UP VIA PHONE ON 12JUL2023, IT WAS REPORTED THAT PATIENT'S MOTHER DID HAVE TO RECEIVE AN ANTIBIOTIC FOR BACK-TO-BACK URINARY TRACT INFECTIONS CAUSED BY USING THE PUREWICK FEMALE EXTERNAL CATHETER. THE MEDICINE WAS PRESCRIBED BY A DOCTOR IN THE EMERGENCY ROOM. CUSTOMER STATED THAT PATIENT CLEANED EVERYTHING DAILY AS INSTRUCTED. CUSTOMER WAS ADVISED TO CLEAN TUBING AND CANNISTER WITH ALCOHOL TO DISINFECT ACCESSORIES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05601 | 2023-07-31 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE BASE OF THE FUNNEL DURING PRE-TEST IN THE OPERATING ROOM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05604 | 2023-07-31 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165SI12 |  | IT WAS REPORTED THAT 32 DAYS AFTER PLACEMENT, THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DURING USE. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 12JUL2023, THE BALLOON RUPTURE WAS OBSERVED ON THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03269 | 2023-08-01 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER, THAT THE 4F DL PROVENA PICC KINKED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05680 | 2023-08-01 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT STERILE WATER COULD BE INJECTED INTO THE FOLEY CATHETER BALLOON DURING THE LEAK TEST, BUT COULD NOT BE DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05678 | 2023-08-01 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS ASYMMETRICAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05676 | 2023-08-01 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THERE WAS A CRACK AT THE BASE OF THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05675 | 2023-08-01 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS INFLATED DURING PRETEST, BUT IT WAS DIFFICULT TO DEFLATE WHEN DRAINING STERILE WATER. WHEN THE USER CHANGED TO A NEW ONE, IT CAME OUT SMOOTHLY. THE USER TOLD THAT IT WAS DIFFICULT TO DRAIN THE WATER AFTER CHANGING TO SILICONE CATHETER, BUT IT SEEMS THAT THERE ARE INDIVIDUAL DIFFERENCES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05674 | 2023-08-01 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS INFLATED DURING PRETEST, BUT IT WAS DIFFICULT TO DEFLATE WHEN DRAINING STERILE WATER. WHEN THE USER CHANGED TO A NEW ONE, IT CAME OUT SMOOTHLY. THE USER TOLD THAT IT WAS DIFFICULT TO DRAIN THE WATER AFTER CHANGING TO SILICONE CATHETER, BUT IT SEEMS THAT THERE ARE INDIVIDUAL DIFFERENCES. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. PHOTO: THERE ARE 2 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE FIRST PHOTO ATTACHED IS AN OVERVIEW OF THE CATHETER WIT CUT PORTION OF INLET TUBING AND SYRINGE. THE SECOND PHOTO SHOW A CLOSER LOOK AT THE CATHETER BALLOON AND THE TIP. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND SYRINGE. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. USING THE RETURNED SYRINGE, THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 2 MINUTES WITH NO CUFFING OR LEAKS NOTED, RETURNING 10ML OF SOLUTION. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. AS THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05668 | 2023-08-01 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05671 | 2023-08-01 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE STERILE WATER WAS SPONTANEOUSLY REMOVED FROM THE BALLOON AND UPON CHECKING THE CATHETER, THE BALLOON WAS TORN. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(4) 2023, CLARIFIED THAT UPON CHECKING THE RETURNED SAMPLE, A BALLOON BURST WAS CONFIRMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05669 | 2023-08-01 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS WITHDRAWN SPONTANEOUSLY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05664 | 2023-08-01 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETERS WERE BLOCKED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-08
1018233-2023-05663 | 2023-08-01 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THERE WAS NO URINE FLOW AND SO WAS STOPPED FROM USE. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, THE DETAILS WEREN'T PROVIDED FROM CUSTOMER IF THERE WAS ANY BLOCKAGE THAT AFFECTED URINE FLOW FROM THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05665 | 2023-08-01 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05661 | 2023-08-01 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT DURING PRETEST, THE WATER LEAKED FROM THE SIDE OF THE CATHETER AND THE EXACT LOCATION OF LEAKAGE WAS UNKNOWN. THE USER REQUESTED THE INVESTIGATOR TO INVESTIGATE THE ENTIRE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05662 | 2023-08-01 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY BALLOON DID NOT INFLATE DURING THE PRE-TEST. THE INFLATION FUNNEL DID NOT SWELL. THERE WAS NO DENT MARK OR ABNORMALITY OBSERVED ON THE CATHETER. THE WATER WAS INJECTED VIGOROUSLY. USUALLY THE BALLOON CAN BE INFLATED WITHIN AROUND 5 SECONDS. THE SYRINGE PLUNGER WAS UNABLE TO PUSH AND SO THE BALLOON WAS DIFFICULT TO INFLATE. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 19JUL2023, THERE WAS NO DETAILS PROVIDED FROM THE CUSTOMER REGARDING IF THE INOPERABLE SYRINGE LED TO DIFFICULTY IN INFLATING THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05660 | 2023-08-01 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON RUPTURED INSIDE THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05656 | 2023-08-01 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAD BROKEN INSIDE OF PATIENT ON TWO DIFFERENT OCCASIONS, BALLOON LOOKED TO HAVE A LARGE SLIT DOWN THE SIDE. PATIENT DID NOT PULL ON CATHETER, ONE TIME THE CATHETER FELL OUT WHILE PATIENT WAS SHOWERING. ON A SEPARATE OCCASION THE NURSE WAS GOING TO INFLATE THE BALLOON AND THE PORT WHERE THE SALINE WAS INSTILLED FELL OFF AND SALINE SPRAYED EVERYWHERE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05653 | 2023-08-01 | 2023-07-11 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT USING THE PUREWICK FEMALE EXTERNAL CATHETERS CAUSED THE PATIENT BEDSORE. CAREGIVER SAID THAT PATIENT WAS CURRENTLY ON MEDICATIONS. PER FOLLOW UP VIA PHONE ON 18JUL2023, IT WAS REPORTED THAT THE PATIENT GOT A BEDSORE FROM USING THE DEVICE. CUSTOMER STATED THAT THE PATIENT WAS COMPLETELY BEDRIDDEN AND USED PRESCRIBED MEDICATION THAT THEY ALREADY HAD FROM PREVIOUS BEDSORES. CUSTOMER STATED THAT THEY WERE NOT USING THE DEVICE AT THE MOMENT, BUT PLAN TO USE IT AGAIN IN THE FUTURE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-08
1018233-2023-05670 | 2023-08-01 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, THE BALLOON WAS FOUND RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03273 | 2023-08-01 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT AFTER INSERTION OF THE CATHETER, THE EXPOSED PORTION OF THE CATHETER COULD NOT BE ATTACHED TO THE OVERSLEEVE PORTION SMOOTHLY. AFTER FORCEFUL ATTACHMENT, SO STRONG RESISTANCE WAS FELT DURING ASPIRATION THAT MADE IT IMPOSSIBLE TO COMPLETE ASPIRATION. ANOTHER OVERSLEEVE PORTION WAS USED AND INSTALLED SMOOTHLY. ASPIRATION WAS CONDUCTED PROPERLY. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03271 | 2023-08-01 | 2023-07-13 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED "PATIENT WAS GIVEN GROSHONG PICC CATHETER FOR INFUSION NEEDS, (B)(6) 2023 WAS PLACED, (B)(6) 2023 WAS FOUND BROKEN AT HOME, IT WAS CONFIRMED ON RADIOGRAPHS THAT THE CATHETER BREAK OCCURRED IN THE PATIENT'S BODY, CATHETER WAS FREE INTO THE LUNGS, THE LENGTH OF THE BREAK IN THE BODY WAS PREDICTED TO BE 40+CM, IT WAS TRANSFERRED TO THE HOSPITAL, THE WHOLE CATHETER WAS REMOVED INTACT UNDER THE DSA, THE PATIENT WAS IN GOOD CONDITION." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A BROKEN CATHETER IS CONFIRMED, BUT THE EXACT CAUSE COULD NOT BE DETERMINED FROM THE RETURNED X-RAY PHOTO. ONE X-RAY IMAGE OF A CHEST CAVITY WITH A PICC WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED THERE IS THE EXTERNAL PORTION OF A PERIPHERALLY INSERTED CATHETER LOCATED OVER THE RIGHT ANTECUBITAL FOSSA. THERE IS DISPLACED CATHETER TUBING THAT EXTENDS FROM UNDERNEATH THE CLAVICLE THROUGH THE HEART INTO THE RIGHT LOWER LOBE PULMONARY ARTERY. THERE IS DISPLACED CATHETER TUBING THAT EXTENDS FROM UNDERNEATH THE CLAVICLE THROUGH THE HEART INTO THE RIGHT LOWER LOBE PULMONARY ARTERY. THE EXTERNAL PORTION OF THE PICC MAY BE LOCATED OVER THE ANTECUBITAL FOSSA. THE CATHETER LENGTH IS SUGGESTIVE THAT IT WAS BROKEN CLOSE TO THE INSERTION SITE, BUT THE PROXIMAL FRACTURE SITE LOCATION IS OVER THE SCAPULA WHICH MAKE SUBTLE EVALUATION FOR THE EXACT LOCATION DIFFICULT. BECAUSE THE CHARACTERISTICS OF THE CATHETER BREAK COULD NOT BE OBSERVED, THE COMPLAINT OF A BREAK IN THE CATHETER IS CONFIRMED, BUT THE ROOT CAUSE COULD NOT BE DETERMINED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03270 | 2023-08-01 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD DL POWERPICCSOLO | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED DURING CLINICAL ASSESSMENT FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. THIS EVENT INVOLVED A DOUBLE LUMEN IN THE UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
MW5120479 | 2023-08-01 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | VENUS CATHETER POWERPICC PROVENA | CATHETER, INTRAVASCULAR, THERAPEUTIC, LONG-TERM GREATER THAN 30 DAYS | BECKTON DICKSON / BARD ACCESS SYSTEMS, INC. |  |  |  | PICC LINE NEEDLE WAS BENT IN PACKAGING BEFORE USE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03295 | 2023-08-01 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR KIT CATHETER (22G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1221252 |  | IT WAS REPORTED, THAT ON THE MORNING OF (B)(6) 23. THE SUPERVISOR WAS NOTIFIED BY A TEAM MEMBER TRAINED IN ACCUCATH, THAT THERE WAS A MALFUNCTION WHEN USING THE DEVICE. THE TEAM MEMBER REPORTS HE WAS IN THE VEIN, THREADED THE GUIDE WIRE, AND THE CATHETER WENT IN SMOOTH "LIKE BUTTER" OVER THE GUIDE WIRE. AND WHEN HE PRESSED THE WHITE SAFETY BUTTON TO RETRACT NEEDLE, THE NEEDLE ONLY PARTIALLY RETRACTED, AND HE COULD NOT REMOVE GUIDEWIRE/NEEDLE. THE ENTIRE CATHETER HAD TO BE PULLED. AND THE PATIENT HAD TO BE RESTUCK, DUE TO ERROR. TEAM MEMBER FEELS IF NOT FOR THIS ISSUE, THE IV WOULD HAVE BEEN SUCCESSFUL ON INITIAL ATTEMPT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03294 | 2023-08-01 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR KIT CATHETER (22G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1221252 |  | IT WAS REPORTED THAT THE NEEDLE DOES NOT RETRACT COMPLETELY INTO THE SAFETY CHAMBER. ON 7/12 A TEAM MEMBER HAD IT OCCUR IN TRAINING AND WAS TOLD TO JUST REMOVE THE ENTIRE LINE. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03287 | 2023-08-01 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | THE CUSTOMER REPORTED THERE ARE ISSUES WITH THEIR SHERLOCK 3CG TCS CHEST SENSOR NOT GIVING THEM PROPER WIRE/PICC GUIDANCE/PLACEMENT. UPON REVIEW, IT LOOKS LIKE THE LAST 4 PICCS PLACED BY ALL 3 OF THE CLINICIANS ON SEPARATE OCCASIONS, NEEDED CHEST XRAYS. THIS REPORT ADDRESSES THE FIRST OCCURRENCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF ISSUES WITH THEIR SHERLOCK 3CG TCS CHEST SENSOR NOT GIVING THEM PROPER WIRE/PICC GUIDANCE/PLACEMENT WAS CONFIRMED. THE SHERLOCK SENSOR STOPS FUNCTIONING WHEN THE USB CABLE END IS MOVED. THE USB PORT IS BENT AT AN UPWARD ANGLE. THE ROOT CAUSE OF THE FAILURE WAS IDENTIFIED AS USE-RELATED DAMAGE. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03292 | 2023-08-01 | 2023-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/ SHERLOCK 3CG TPS STYLET W/ MICROINTRODUCER (4F) (SL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174118 |  | IT WAS REPORTED THE FLOOR NURSE EXPERIENCED ISSUES FLUSHING AND CALLED PICC TEAM. AN UPPER ARM X RAY SHOWED THE PICC WAS KINKED. NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023: DEVICE WAS A SL 4 FRENCH, PLACED IN A 3.5CM DEEP BASILIC VEIN. IT WAS CUT 48CM AND WAS LEFT OUT 2CM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03289 | 2023-08-01 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | THE CUSTOMER REPORTED THERE ARE ISSUES WITH THEIR SHERLOCK 3CG TCS CHEST SENSOR NOT GIVING THEM PROPER WIRE/PICC GUIDANCE/PLACEMENT. UPON REVIEW, IT LOOKS LIKE THE LAST 4 PICCS PLACED BY ALL 3 OF THE CLINICIANS ON SEPARATE OCCASIONS, NEEDED CHEST XRAYS. THIS REPORT ADDRESSES THE THIRD OCCURRENCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03291 | 2023-08-01 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F W/ MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT THEY HAVE DL PICC LINE WITH A LEAKING T LOCK ASSEMBLY. IT WAS USED ON ONE PATIENT. THERE WAS NO HARM REPORTED. THE EVENT DID NOT INVOLVE AN URGENT OR LIFE THREATENING SITUATION. THE LEAK WAS DISCOVERED WHEN THE CLINICIAN PRIMED THE LINE. THE EVENT DID NOT INTERRUPT TREATMENT. NO SIGNS, SYMPTOMS OR COMPLICATIONS EXPERIENCE BY PATIENT. CLINICIAN DO NOT HAVE THE NAME OF PLANNED MEDICATIONS TO BE INFUSED. CATHETER WAS SECURED THE DEVICE WITH A STATLOCK. THERE WAS NO PATIENT HARM NOTED DUE TO CLINICIAN INTERVENTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03283 | 2023-08-01 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN PUNCTURING FOR THE GROSHONG CATHETER, THE NURSE SPECIALIST FOUND THAT THE SELDINGER COULD NOT BE INSERTED SMOOTHLY AFTER INSERTION OF THE GUIDE WIRE. IT WAS OBSERVED THAT THE PROXIMAL END OF THE GUIDE WIRE BULGED. THE GUIDE WIRE KIT WAS REPLACED FOR PUNCTURE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05651 | 2023-08-01 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI18 |  | IT WAS REPORTED THAT THESE FOLEY CATHETER CAUSED BLEEDING AND WERE DIFFICULT TO REMOVE. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INCORRECT BALLOON DESIGN (BALLOON WALL THICKNESS EXCESSIVE)". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "LUBRI-SIL® I.C. ANTI-MICROBIAL ALL-SILICONE FOLEY CATHETER WITH BARD® HYDROGEL AND BACTI-GUARD® SILVER ALLOY COATING INFECTION CONTROL ALL-SILICONE FOLEY CATHETER 100% LATEX FREE. WARNING: ON CATHETERS, DO NOT USE OINTMENT OR LUBRICANT. THEY CAN DAMAGE THE SILICONE AND MAY CAUSING THE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY CONTAIN POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. CAUTION: FEDERAL (U.S.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. CAUTION: DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. STORAGE: STORE THE CATHETER AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. NOTE: AGGRESSIVE TRACTION, PARTICULARLY IN THE PRESENCE OF SUTURING, IS NOT RECOMMENDED FOR 100% SILICONE FOLEY CATHETERS. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. SINGLE PATIENT USE ONLY. DO NOT REUSE. DO NOT RE-STERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIDE SYRINGES. DO NOT USE NEEDLES. CATHETERS SHOULD BE REPLACED IN ACCORDANCE WITH THE CDC GUIDELINES ¿GUIDELINE FOR PREVENTION OF CATHETER-ASSOCIATED URINARY TRACT INFECTION¿. AT THE ONSET OR FIRST SIGNS OF A URINARY TRACT INFECTION, CATHETER ENCRUSTATION, OR ANY OTHER CATHETER-RELATED ADVERSE EFFECT, THE CATHETER SHOULD BE REPLACED. PREVENTING CATHETER-RELATED TO DEFLATE THE CATHETER BALLOON: GENTLY INSERT THE SYRINGE INTO THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. ALLOW THE BALLOON TO DEFLATE SLOWLY ON ITS OWN. DO NOT ASPIRATE OR MANUALLY ACCELERATE THE DEFLATION OF THE BALLOON. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. RECOMMENDED INFLATABLE CAPACITIES 5CC BALLOONS: USE 10CC STERILE WATER 30CC BALLOON: USE 35CC STERILE WATER DO NOT EXCEED RECOMMENDED CAPACITY." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05700 | 2023-08-01 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE TEMPERATURE DISPLAY HAD FAILED ON TEMPERATURE SENSING FOLEY CATHETER. IT WAS STATED THAT DURING A LAPAROSCOPIC SURGERY, THE TEMPERATURE DISPLAY SHOWED A LOW OF 34°C OR DISAPPEARED. FOR THIS REASON, THEY WERE ASKED TO INVESTIGATE IF THERE WAS A PROBLEM ON THE CATHETER SIDE. THE RELAY CABLE COULD NOT BE RETRIEVED, BUT WE HAD IT CHECKED BY THE CALIBRATOR AND FOUND NO RELAY CABLE PROBLEM. THE MONITOR WAS ALSO CHECKED AND THEY FOUND NO PROBLEM. THE CLIENT REQUESTED THAT WE CONFIRM THE PROBLEM BECAUSE THE SAME PROBLEM HAD BEEN OCCURRING SINCE LAST YEAR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03280 | 2023-08-01 | 2023-07-14 | I, n, j, u, r, y | Manufacturer report | PPICC 4F S/L 3CG BASIC, | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 22184118 |  | IT WAS REPORTED RUPTURE OF THE CATHETER. DURING THE ADMINISTRATION OF CHEMOTHERAPY THE PATIENT REPORTS A BURNING SENSATION IN THE ARM. THERE IS OEDEMA OF THE ARM ABOVE THE INSERTION POINT. THE DEVICE IS NON-FUNCTIONAL. ADDITIONAL INFORMATION (B)(6) 2023: THE SPECIFIC INTERVENTION THE COMPANY PROCEDURE IN THE EVENT OF CHEMOTHERAPY DRUG OVERFLOW WAS ADOPTED, ADOPTING TARGETED INTERVENTIONS INDICATED THEREIN. IN PARTICULAR, ICE WAS APPLIED TO THE PATIENT IMMEDIATELY AFTER THE EVENT AND FOR 48 HOURS AFTERWARDS, THREE TIMES A DAY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05686 | 2023-08-01 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAD BROKEN INSIDE OF PATIENT ON TWO DIFFERENT OCCASIONS, BALLOON LOOKED TO HAVE A LARGE SLIT DOWN THE SIDE. PATIENT DID NOT PULL ON CATHETER, ONE TIME THE CATHETER FELL OUT WHILE PATIENT WAS SHOWERING. ON A SEPARATE OCCASION THE NURSE WAS GOING TO INFLATE THE BALLOON AND THE PORT WHERE THE SALINE WAS INSTILLED FELL OFF AND SALINE SPRAYED EVERYWHERE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED OR CRACK/CRUSHED VALVE CAUSED BY INCORRECT AIR PRESSURE SETTING FOR VALVING. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05687 | 2023-08-01 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT FLUID LEAKED FROM THE FOLEY CATHETER. IT WAS STATED THAT THE STERILE DISTILLED WATER LEAKED FROM THE CATHETER IN THE OPERATION ROOM. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. AS THE REPORTED EVENT WAS UNCONFIRMED A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED. HOWEVER, A DEVICE HISTORY RECORD REVIEW WAS COMPLETED BEFORE UNCONFIRMING THE EVENT. AS THE REPORTED EVENT WAS UNCONFIRMED A LABELING REVIEW WAS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05692 | 2023-08-01 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165818 |  | IT WAS REPORTED THAT THE USER HAD EXPERIENCED ISSUES WITH INDWELLING URINARY CATHETER "FELLING OUT" TWICE SINCE USER¿S LAST CHANGE. IT WAS STATED THAT THE ON MONDAY DURING THE PROCEDURE OF CHANGING FOLEY CATHETER THE BALLOON BURST WHILE IN PROCESS OF FILLING BALLOON WITH THE SPECIFIED AMOUNT OF WATER. THIS RESULTED IN ANOTHER INDWELLING URINARY CATHETER NEEDING TO BE PLACED WITH RESULTING TRAUMA REQUIRING THE CLIENT TO ATTEND ED FOR REVIEW. IT WAS MY UNDERSTANDING THAT THIS CATHETER ALSO "FELL OUT" ON TUESDAY RESULTING IN CLIENT AGAIN NEEDING TO GO TO ED FOR ATTENTION. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 12JUL2023, IT WAS STATED THAT THE CUSTOMER HAD ALSO INFORMED THAT TWO OTHER CATHETERS FROM THAT SAME BOX HAD FAILED BY FALLING OUT IN THE DAYS PRIOR TO THE TWO CATHETERS. THERE WAS NO TRAUMA EXPERIENCED WITH THOSE FIRST TWO. THIS MEANS OUT OF A BOX OF TEN, SIX WERE USED OF WHICH FOUR ARE NOW REPORTED TO HAVE BEEN FAULTY. THE OTHER FOUR WERE STILL IN THE BOX READY FOR PICK UP, THERE WERE NO ACTUAL USED FAULTY SAMPLES AVAILABLE FOR PICK UP. IT WAS STATED THAT THE FIRST CATHETER WAS CHECKED AND THERE WERE NO MISSING PIECES. WHEN INSERTING THE SECOND CATHETER USER AND NURSE NOTICED BLOOD DRIPPING FROM THE EXTERNAL PART OF THE CATHETER AS WELL AS SEEPING AROUND THE CATHETER WERE INSERTED INTO PENNIS. THE USER WAS TREATED IN ED, THEY DID NOT REMOVE THE CATHETER, THEY BANDAGED UP PENNIS TO STOP BLEEDING AND KEPT USER OVERNIGHT FOR OBSERVATION BEFORE RELEASING THE NEXT DAY. THE CIRCUMSTANCES SURROUNDING THE SECOND CATHETER FALLING OUT WAS USER FELT LIKE THERE WAS ONLY A SMALL AMOUNT OF URINE IN BAG WHEN AT HOME SO CHECKED AND FOUND THAT THE CATHETER HAD FALLEN OUT. THE USER THEN RETURNED TO ED WHERE THE NURSE CHECKED THE CATHETER AND DETERMINED THE BALLOON WOULD NOT INFLATE. THERE WAS NO INDICATION ANY PIECES WERE LEFT INSIDE. THE NURSE THEN TESTED THE BALLOON OF A NEW DIFFERENT BRANDED CATHETER BEFORE INSERTION AND RETURNED HOME WITHOUT INCIDENT OF THAT CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05688 | 2023-08-01 | 2023-07-13 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 51812 |  | IT WAS REPORTED THAT MIDDLE OF CENTER PIECE WHEN USE OF INTERMITTENT CATHETER SEPARATED. THIS HAS CAUSED MORE URINARY TRACT INFECTION (UTIS) FOR MONTHS, DIFFICULT TO CURE DUE TO NEWFOUND BACTERIA. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03277 | 2023-08-01 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT BARB ON THE SHEATH. DURING CATHETER PLACEMENT, THREE UNITS USED SUCCESSIVELY WERE FOUND RUPTURED. AFFECTED CATHETER PLACEMENT OPERATIONS BY HEAD NURSE. NO OTHER INFORMATION WAS PROVIDED. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03278 | 2023-08-01 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT BARB ON THE SHEATH. DURING CATHETER PLACEMENT, THREE UNITS USED SUCCESSIVELY WERE FOUND RUPTURED. AFFECTED CATHETER PLACEMENT OPERATIONS BY HEAD NURSE. NO OTHER INFORMATION WAS PROVIDED. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03274 | 2023-08-01 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT AFTER INSERTION OF THE CATHETER, THE EXPOSED PORTION OF THE CATHETER COULD NOT BE ATTACHED TO THE OVERSLEEVE PORTION SMOOTHLY. AFTER FORCEFUL ATTACHMENT, SO STRONG RESISTANCE WAS FELT DURING ASPIRATION THAT MADE IT IMPOSSIBLE TO COMPLETE ASPIRATION. ANOTHER OVERSLEEVE PORTION WAS USED AND INSTALLED SMOOTHLY. ASPIRATION WAS CONDUCTED PROPERLY. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03281 | 2023-08-01 | 2023-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | IT WAS REPORTED THAT THE CATHETER WAS INSERTED IN THE PATIENT ON MAY 18. THE EXTENSION TUBING LEAKED LIQUIDS WHEN INFUSING CHEMOTHERAPEUTIC DRUGS. THE CUSTOMER REMOVED THE CATHETER AND REPLACED IT WITH A NEW ONE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05652 | 2023-08-01 | 2023-07-11 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT INFECTION AND THEN GOT SEPTIC SHOCK. DOCTOR SAID IT WAS FROM THE PUREWICK AND THE PATIENT HAD STOPPED USING IT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION AND SEPTIC SHOCK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03286 | 2023-08-01 | 2023-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE SENSOR IS INTERMITTENTLY GIVING INCORRECT READINGS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF SENSOR IS INTERMITTENTLY GIVING INCORRECT READINGS WAS CONFIRMED. THE SHERLOCK SENSOR STOPS FUNCTIONING WHEN THE USB PORT END OF THE CABLE IS MOVED. THE ROOT CAUSE OF THE REPORTED FAILURE WAS IDENTIFIED AS A MATERIAL FAILURE OF THE USB CABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05643 | 2023-08-01 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE CATHETER HAD BEEN DEFORMED AND WAS USED AFTER THE CONFIRMATION OF DEFORMATION. WHILE CHANGING AN INDWELLING BLADDER CATHETER FOR THE PATIENT, IT WAS FOUND THAT THE CATHETER WAS BULGING A FEW MILLIMETERS IN THE MIDDLE. THERE WAS NO HARDNESS AT THE BULGING POINT, IT COULD BE BENT, AND THE URINE FLOW WAS GOOD. SO, THE CATHETER WAS INSERTED INTO THE PATIENT AND THE PATIENT'S CONDITION WAS OBSERVED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO 1ST LATEX OUT-DEPOSIT LATEX TOO FAST. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL BALLOON CATHETER: NATURAL RUBBER LATEX SIZES OF CATHETERS AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE 1. MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION CATHETER KINKING, DAMAGE, RUPTURE DIFFICULTY OR FAILURE TO REMOVE THE DEVICE OCCLUSION OF CATHETER INNER LUMENS ENCRUSTATION ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS URINARY-TRACT INFECTION HEMORRHAGE, HEMATURIA ALLERGY REACTION TO THE DEVICE CALCULUS FORMATION EDEMA PAIN DISCOMFORT INJURY OF BLADDER OR URETHRAL URETHRITIS, URINARY INCONTINENCE RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05647 | 2023-08-01 | 2023-07-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53614G |  | IT WAS REPORTED THAT THE PATIENT HAD WENT TO HOSPITAL DUE TO BLOCKAGE WITH THE MAGIC3 HYDROPHILIC MALE INTERMITTENT CATHETER. PER FOLLOW UP VIA PHONE ON 12JULY2023, CUSTOMER WAS INSTRUCTED THAT THEIR DOCTOR STATED THIS CATHETER DID NOT WORK FOR THEIR ANATOMY AND THAT CUSTOMER NEEDED A COUDE CATHETER. CUSTOMER STATED THAT THEY WERE GIVEN ANTIBIOTICS DURING THEIR STAY AT THE HOSPITAL. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "DUE TO ANATOMIC CONDITION". THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-08
1018233-2023-05642 | 2023-08-01 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V24S |  | IT WAS REPORTED THAT THERE WAS ANOTHER INSTANCE OF THE FOLEY CATHETER BALLOON NOT DEFLATING ON THIS PATIENT WITH A VISIT TO THE EMERGENCY ROOM AND HAD TO BREAK THE BALLOON WITH A WIRE. THE CUSTOMER WAS PART OF THE HOME CARE PROGRAM (SELF MANAGED CARE). THE CATHETER HAD BLOCKED AT CUSTOMER'S HOME AND FAMILY WAS ASSISTING IN FLUSHING. THE CATHETER HAD FAILED FROM FLUSHING AND FAMILY ATTEMPTED TO CHANGE CATHETER. THE CATHETER WAS UNABLE TO DEFLATE. THEY HAD TO GO TO EMERGENCY AT HOSPITAL AND PHYSICIAN HAD VALVE CUT BUT NO FURTHER DEFLATION OF CATHETER. THE CATHETER HAD TO BE PIERCED WITH A WIRE AND PHYSICIAN PULLED CATHETER BALLOON VIA STOMA. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INCORRECT PARAMETER SETTING - BLOCKED DRAINAGE EYE/LUMEN OR NO DRAINAGE EYE OR USER RELATED (SALT ACCUMULATION). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
2020394-2023-00543 | 2023-08-01 | 2023-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | AT75144 |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE IN THE RIGHT BRACHIOCEPHALIC VEIN, THE PTA BALLOON ALLEGEDLY LEAKED AT THE TIP OF THE BONDING AREA. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ATLAS PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, THE PROXIMAL END OF THE BALLOON WAS NOTED HAVING FRAYED FIBERS AT THE BONDING SITE. NO KINKS WERE NOTED TO THE CATHETER, NO ANOMALIES TO THE LUERS/BIFURCATE. DURING FUNCTIONAL EVALUATION, AN IN-HOUSE PRESTO INFLATION DEVICE WAS USED TO INFLATE THE BALLOON AND WATER WAS NOTED LEAKING FROM THE BALLOON NEAR ITS DISTAL TIP. THE BALLOON FIBERS WERE STRIPPED AND UNDER MICROSCOPIC EXAMINATION, A PINHOLE RUPTURE WAS NOTED ON THE BALLOON. ALSO, FRAYED FIBERS WERE NOTED AT THE DISTAL TIP. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED LEAK AND IDENTIFIED PINHOLE BALLOON RUPTURE AS A PINHOLE RUPTURE WAS NOTED ON THE BALLOON WHICH COULD HAVE BEEN THE SOURCE OF THE LEAK. THE INVESTIGATION IS ALSO CONFIRMED FOR THE IDENTIFIED FRAYED FIBERS AS FRAYED FIBERS WERE NOTED AT BOTH THE PROXIMAL AND DISTAL ENDS OF THE BALLOON. THE PINHOLE BALLOON RUPTURE IDENTIFIED DURING THE MICROSCOPIC EXAMINATION MIGHT HAVE BEEN THE SOURCE OF THE REPORTED LEAK. HOWEVER, THE DEFINITIVE ROOT CAUSE FOR THE ALLEGED LEAK, IDENTIFIED PINHOLE BALLOON RUPTURE AND FRAYED FIBERS COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 06/2025). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03258 | 2023-08-01 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE NINTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03254 | 2023-08-01 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE FIFTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03252 | 2023-08-01 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT IN FACILITY, WHEN THE OPERATOR WAS PLACING 12 UNITS OF THE 7655405, NO ¿CLICK¿ REPRESENTING ENGAGEMENT WAS HEARD DURING ATTACHMENT OF THE CATHETERS TO THE EXTENSION TUBING. ANOTHER PULL WAS APPLIED AND IT WAS FOUND THAT THE CATHETERS COULD NOT BE ENGAGED WITH THE EXTENSION TUBING. OTHER EXTENSION TUBING HAD TO BE USED FOR RE-SECUREMENT." THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03266 | 2023-08-01 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | UNK GLIDEPATH |  | IT WAS REPORTED THAT SOMETIME POST A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CUFF OF THE DEVICE WAS ALLEGEDLY EXPOSED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05650 | 2023-08-01 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119316M |  | IT WAS REPORTED THAT FOLEY CATHETER BALLOON FAILED. FOLEY CAME OUT OF PATIENT. ON INSERTION THERE WAS URINE RETURN AND THE PREPARED 10ML SALINE INSTILLED. THE PRODUCT WAS NOT IN CUSTOMER OFFICE AS THERE WAS URINE IN BAG AND CONTAMINATION WAS NOT SOMETHING THAT THEY WANTED TO HANG OUT IN THEIR OFFICE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03303 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR KIT CATHETER (22G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1221252 |  | IT WAS REPORTED BY THE CUSTOMER "UPON NOTIFYING SUPERVISOR, HE REPORTED HE HAD ALSO HEARD FROM SEVERAL OTHERS TRAINED THAT THEY HAVE EXPERIENCED THIS OCCASIONALLY WHERE THE NEEDLE DOES NOT RETRACT COMPLETELY INTO THE SAFETY CHAMBER. STAFF MEMBER REPORTS HE HAS DONE ABOUT 16 SINCE TRAINING AND IT HAS HAPPENED AT LEAST 2 TIMES. OTHER TEAM MEMBERS TRAINED ALSO WERE REPORTING THIS BUT DID NOT LET LEADERSHIP KNOW UNTIL THIS WEEK. THESE TEAM MEMBERS RANGE FROM EMERGENCY, TRANSPORT, AND FLEX. EACH TEAM MEMBER HAVING EXPERIENCED THIS OCCUR ON A PATIENT AT LEAST ONCE. UNSURE IF USER ERROR OR MANUFACTURING CONCERN. CONCERN IS IF THERE IS AN ISSUE WITH THE ACTUAL SAFETY MECHANISM OR ARE WE HAVING SMALL USER ERRORS WHEN IT COMES TO PRESSING THE SAFETY BUTTON AND REMOVING GUIDE-WIRE AT THE SAME TIME. SUPERVISOR XXX ALSO CALLED MAIN ED AND RN AT MAIN CAMPUS WHO ALSO REPORT THIS OCCURRING BUT ARE UNSURE IF THIS IS USER ERROR OR MANUFACTURING ISSUE AS IT HAPPENS OCCASIONALLY. NURSES HAVE ALSO REPORTED IT HAS HAPPENED WITH ALL SIZE CATHETERS WE USE SUCH AS THE 22G 1.25IN. ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER "I DO NOT HAVE ANY MORE DETAILS TO SHARE. I HAVE FOLLOWED UP WITH A FEW NURSES WHO STATED THEY HAVE NOT HAD ANOTHER ISSUE WITH THE RETRACTOR BUTTON." THIS REPORT ADDRESSES THE EMERGENCY ROOM EVENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03302 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR KIT CATHETER (22G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1221252 |  | IT WAS REPORTED BY THE CUSTOMER "UPON NOTIFYING SUPERVISOR, HE REPORTED HE HAD ALSO HEARD FROM SEVERAL OTHERS TRAINED THAT THEY HAVE EXPERIENCED THIS OCCASIONALLY WHERE THE NEEDLE DOES NOT RETRACT COMPLETELY INTO THE SAFETY CHAMBER. STAFF MEMBER REPORTS HE HAS DONE ABOUT 16 SINCE TRAINING AND IT HAS HAPPENED AT LEAST 2 TIMES. OTHER TEAM MEMBERS TRAINED ALSO WERE REPORTING THIS BUT DID NOT LET LEADERSHIP KNOW UNTIL THIS WEEK. THESE TEAM MEMBERS RANGE FROM EMERGENCY, TRANSPORT, AND FLEX. EACH TEAM MEMBER HAVING EXPERIENCED THIS OCCUR ON A PATIENT AT LEAST ONCE. UNSURE IF USER ERROR OR MANUFACTURING CONCERN. CONCERN IS IF THERE IS AN ISSUE WITH THE ACTUAL SAFETY MECHANISM OR ARE WE HAVING SMALL USER ERRORS WHEN IT COMES TO PRESSING THE SAFETY BUTTON AND REMOVING GUIDE-WIRE AT THE SAME TIME. SUPERVISOR ALSO CALLED MAIN ED AND RN AT MAIN CAMPUS WHO ALSO REPORT THIS OCCURRING BUT ARE UNSURE IF THIS IS USER ERROR OR MANUFACTURING ISSUE AS IT HAPPENS OCCASIONALLY. NURSES HAVE ALSO REPORTED IT HAS HAPPENED WITH ALL SIZE CATHETERS WE USE SUCH AS THE 22G 1.25IN. ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER "I DO NOT HAVE ANY MORE DETAILS TO SHARE. I HAVE FOLLOWED UP WITH A FEW NURSES WHO STATED THEY HAVE NOT HAD ANOTHER ISSUE WITH THE RETRACTOR BUTTON." THIS REPORT ADDRESSES THE SECOND EVENT AFTER TRAINING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03300 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD ACCUCATH | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE FACILITY MAIN ED SUPERVISOR AND RN THAT THE NEEDLE DOES NOT RETRACT COMPLETELY INTO THE SAFETY CHAMBER, BUT THEY ARE UNSURE IF THIS IS USE ERROR OR A MANUFACTURING ISSUE AS IT HAPPENS OCCASIONALLY. NURSES HAVE REPORTED IT HAPPENED WITH ALL SIZE CATHETERS SUCH AS THE 20G X 2.25. ADDITIONAL INFORMATION FROM THE CUSTOMER "I HAVE FOLLOWED UP WITH A FEW OF THE NURSES WHO STATED THEY HAVE NOT HAD ANOTHER ISSUE WITH THE RETRACTOR BUTTON BUT I UNFORTUNATELY DO NOT HAVE ANY MORE INFORMATION TO SHARE." THE CUSTOMER REPORTED THIS OCCURRED SEVERAL TIMES WITH ALL CATHETERS HOWEVER THE EXACT QUANTITY OR THE SPECIFIC PRODUCT INFORMATION WAS NOT PROVIDED TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05723 | 2023-08-02 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 50614 |  | IT WAS REPORTED THAT THERE APPEARED TO BE A CLOG IN THE LINE, (LOT# JUHN2149 - DOE (B)(6) 2023) (LOT# UNK - DOE UNK) A DISCOLORATION AT THE TIP OF THE CATHETER (LOT# JUHN2149 - DOE (B)(6)2023) (LOT# UNK - DOE UNK) NOT UNLIKE A HARDENED SUBSTANCE WAS BLOCKING THE FLOW OF FLUID. THIS HAD BEEN OCCURRING ON AND OFF THE LAST TWO MONTHS, POSSIBLY WITH OTHER UNKNOWN ADDITIONAL LOT NUMBERS. MULTIPLE UNKNOWN DATE OF EVENTS BUT ALSO ON (B)(6)2023. PER FOLLOW UP INFORMATION RECEIVED VIA PHONE ON 21JUL2023, CUSTOMER STATED THAT THEY ONLY HAD ONE CATHETER SAMPLE AVAILABLE AND THEY WERE NOT SURE IF IT WAS FROM ONE OF THE REPORTED LOT NUMBERS AS THEY TOOK IT OUT OF THE TRASH AFTER SUBMITTING THE COMPLAINT. PER CUSTOMER, AFTER INSERTING THE CATHETER, THE TIP TURNED GREEN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-08
1018233-2023-05722 | 2023-08-02 | 2023-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT CAME TO THE HOSPITAL ON (B)(6) 2023, AND THE DOCTOR GAVE THE PATIENT A DISPOSABLE STERILE FOLEY CATHETER FOR URINE CATHETERIZATION. AFTER CATHETERIZATION, FLUID LEAKAGE WAS FOUND AND THE CATHETER WAS REPLACED. THE ISSUE DID NOT OCCUR AGAIN AFTER REPLACEMENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE POOR BALLOON SHAPE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05719 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) AND WAS GIVEN INDWELLING CATHETERIZATION AND CONTINUED BLADDER FLUSHING, WHICH FLUSHED OUT THE RED BLOOD CLOT. ON (B)(6), THE CATHETER BALLOON BURST, AND THE ATTENDING PHYSICIAN'S RE-INDWELLING CATHETERIZATION AND BLADDER FLUSHING WERE SMOOTH AND LIGHT RED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03299 | 2023-08-02 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER THAT A PICC FRACTURED. NO OTHER INFORMATION PROVIDED. ADDITIONAL INFORMATION RECEIVED 08 AUGUST 2023: THE PATIENT CAME TO NO HARM, NO EXTRAVASATION. PICC FRACTURE WAS INTERNAL, PATIENT WAS ON CHEMOTHERAPY AND THEY NEEDED NEW PICC LINE WHICH HAS NOW BEEN PLACED. EXTERNAL LENGTH: 20.5CM | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
17441702 | 2023-08-02 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER, INTRAVASCULAR, THERAPEUTIC, LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1274108D5 | S1274108D5 |  | SENSOR BENT AND BROKE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-08
17442136 | 2023-08-02 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC SOLO CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 1022786 | 1194108D |  | PATIENT SOLO PERIPHERALLY INSERTED CENTRAL CATHETER (PICC) NOT FLOWING CORRECTLY, VASCULAR ACCESS TEAM (VAT) TEAM CALLED BY MANAGER TO REVIEW. VAT REGISTERED NURSE (RN) REVIEWED LINE AND NOTED BLOOD BACKING UP IN THE PICC LINE. A PRESSURIZED CAP WAS IN PLACE, AND IT APPEARS WITHOUT THE CAP BLOOD WOULD HAVE DRIPPED OUT, BUT IT HAD CLOTTED, AND THE PICC WAS NO LONGER USABLE. A NEW PICC WAS PLACED. AFTER REVIEW, IT APPEARS THE VALVE IN THE PICC HAD FAILED. |  | N | Y |  |  | 2023-08
1018233-2023-05710 | 2023-08-02 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165816 |  | IT WAS REPORTED THAT THERE WAS AN ISSUE OVER THE WEEKEND WITH 2 CERTAIN LOTS OF FOLEY BALLOONS THAT WERE NOT WORKING (LOT# NGHP0791) AND (LOT# NGHQ1143). PER FOLLOW UP VIA PHONE ON 21JUL2023, IT WAS REPORTED THAT FOLEY CATHETER BALLOON WAS LEAKING AFTER IT WAS INSERTED INTO THE PATIENT (LOT# NGHP0791). IT WAS ALSO REPORTED THAT ANOTHER LOT WAS NOT DEFLATING PROPERLY (LOT# NGHQ1143). THE CUSTOMER WAS NOT ABLE TO CROSS-REFERENCE WHICH LOT HAD WHICH ISSUE, BUT LOT# NGHP079 WAS DISCARDED, AND 6-8 TRAYS FROM LOT# NGHQ1143 ARE AVAILABLE FOR A SAMPLE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "BALLOON NOT COMPLETELY SEALED TO SHAFT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® EZ-LOK® SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK (FIG. 1A) OR SLIP TIP SYRINGES (FIG. 1B). 1. OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. 2. SWAB SURFACE OF BARD® EZ-LOK® SAMPLING PORT WITH ANTISEPTIC WIPE. 3. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. 4. SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND REMOVE SYRINGE FROM SAMPLE PORT. 5. UNKINK TUBING IF NECESSARY AND TRANSFER URINE SPECIMEN INTO SPECIMEN CUP OR FOLLOW HOSPITAL PROTOCOL. DISCARD SYRINGE ACCORDING TO HOSPITAL PROTOCOL. 6. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. DIRECTIONS FOR USE: PROPER TECHNIQUES FOR URINARY CATHETER INSERTION. PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION. INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE. SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED. MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT. ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE. LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03304 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR KIT CATHETER (22G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1221252 |  | IT WAS REPORTED BY THE CUSTOMER "UPON NOTIFYING SUPERVISOR, HE REPORTED HE HAD ALSO HEARD FROM SEVERAL OTHERS TRAINED THAT THEY HAVE EXPERIENCED THIS OCCASIONALLY WHERE THE NEEDLE DOES NOT RETRACT COMPLETELY INTO THE SAFETY CHAMBER. STAFF MEMBER REPORTS HE HAS DONE ABOUT 16 SINCE TRAINING AND IT HAS HAPPENED AT LEAST 2 TIMES. OTHER TEAM MEMBERS TRAINED ALSO WERE REPORTING THIS BUT DID NOT LET LEADERSHIP KNOW UNTIL THIS WEEK. THESE TEAM MEMBERS RANGE FROM EMERGENCY, TRANSPORT, AND FLEX. EACH TEAM MEMBER HAVING EXPERIENCED THIS OCCUR ON A PATIENT AT LEAST ONCE. UNSURE IF USER ERROR OR MANUFACTURING CONCERN. CONCERN IS IF THERE IS AN ISSUE WITH THE ACTUAL SAFETY MECHANISM OR ARE WE HAVING SMALL USER ERRORS WHEN IT COMES TO PRESSING THE SAFETY BUTTON AND REMOVING GUIDE-WIRE AT THE SAME TIME. SUPERVISOR XXX ALSO CALLED MAIN ED AND RN AT MAIN CAMPUS WHO ALSO REPORT THIS OCCURRING BUT ARE UNSURE IF THIS IS USER ERROR OR MANUFACTURING ISSUE AS IT HAPPENS OCCASIONALLY. NURSES HAVE ALSO REPORTED IT HAS HAPPENED WITH ALL SIZE CATHETERS WE USE SUCH AS THE 22G 1.25IN. ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER "I DO NOT HAVE ANY MORE DETAILS TO SHARE. I HAVE FOLLOWED UP WITH A FEW NURSES WHO STATED THEY HAVE NOT HAD ANOTHER ISSUE WITH THE RETRACTOR BUTTON." THIS REPORT ADDRESSES THE TRANSPORT EVENT. | H11: SECTION A THROUGH F: THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05702 | 2023-08-02 | 2023-07-10 | I, n, j, u, r, y | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V24S |  | IT WAS REPORTED THAT THERE WAS ANOTHER INSTANCE OF THE FOLEY CATHETER BALLOON NOT DEFLATING ON THIS PATIENT WITH A VISIT TO THE EMERGENCY ROOM AND HAD TO BREAK THE BALLOON WITH A WIRE. THE CUSTOMER WAS PART OF THE HOME CARE PROGRAM (SELF MANAGED CARE). THE CATHETER HAD BLOCKED AT CUSTOMER'S HOME AND FAMILY WAS ASSISTING IN FLUSHING. THE CATHETER HAD FAILED FROM FLUSHING AND FAMILY ATTEMPTED TO CHANGE CATHETER. THE CATHETER WAS UNABLE TO DEFLATE. THEY HAD TO GO TO EMERGENCY AT HOSPITAL AND PHYSICIAN HAD VALVE CUT BUT NO FURTHER DEFLATION OF CATHETER. THE CATHETER HAD TO BE PIERCED WITH A WIRE AND PHYSICIAN PULLED CATHETER BALLOON VIA STOMA. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
2020394-2023-00548 | 2023-08-02 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CQF75124 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE DEVICE ALLEGEDLY MALFUNCTIONED. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE CONQUEST 40 PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, A KINK WAS NOTED TO THE CATHETER AND SLIGHT BENDS WERE NOTED THROUGHOUT THE CATHETER SHAFT. A BEND WAS ALSO NOTED TO THE BALLOON. DURING FUNCTIONAL EVALUATION, AN IN-HOUSE PRESTO INFLATION DEVICE WAS USED TO INFLATE THE BALLOON AND WATER WAS NOTED LEAKING FROM THE BALLOON. THE BALLOON FIBERS WERE STRIPPED AND UNDER MICROSCOPIC EXAMINATION A LONGITUDINAL RUPTURE WAS NOTED TO THE BALLOON. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED MATERIAL DEFORMATION AS A KINK AND SLIGHT BENDS WERE NOTED TO THE SHAFT, AND A BEND WAS ALSO NOTED TO THE BALLOON. THE INVESTIGATION IS ALSO CONFIRMED FOR THE IDENTIFIED LONGITUDINAL BALLOON RUPTURE AS A LONGITUDINAL RUPTURE WAS NOTED TO THE BALLOON UNDER MICROSCOPIC EXAMINATION. A DEFINITIVE ROOT CAUSE FOR THE IDENTIFIED MATERIAL DEFORMATION AND LONGITUDINAL BALLOON RUPTURE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. D4 (EXPIRATION DATE: 11/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
2020394-2023-00547 | 2023-08-02 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 014 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | US14150430 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE SUPERFICIAL FEMORAL ARTERY VIA A CONTRALATERAL APPROACH, THE PTA BALLOON ALLEGEDLY GOT STUCK WITH THE GUIDEWIRE AND WAS UNABLE TO BE REMOVED. IT WAS FURTHER REPORTED THAT THE DEVICE WAS ALLEGEDLY DIFFICULT TO REMOVE FROM THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED GUIDE WIRE REMOVAL DIFFICULTY COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 02/2024), G3, H6 (METHOD) . H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03290 | 2023-08-02 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | THE CUSTOMER REPORTED THERE ARE ISSUES WITH THEIR SHERLOCK 3CG TCS CHEST SENSOR NOT GIVING THEM PROPER WIRE/PICC GUIDANCE/PLACEMENT. UPON REVIEW, IT LOOKS LIKE THE LAST 4 PICCS PLACED BY ALL 3 OF THE CLINICIANS ON SEPARATE OCCASIONS, NEEDED CHEST XRAYS. THIS REPORT ADDRESSES THE FOURTH OCCURRENCE. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05716 | 2023-08-02 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0102L14 |  | IT WAS REPORTED THAT NURSE WENT TO INSERT A FOLEY CATHETER AND THE BALLOON KEPT DEFLATING. THEY FOUND THE GREEN PORT WAS LEAKING. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. 1 SAMPLE WERE CONFIRMED TO EXHIBIT THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), USED FOLEY CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED ATTEMPT TO INFLATE THE BALLOON WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND LEAK OBSERVED FROM PIN HOLE LOCATED NEAR THE CAP ON THE INFLATION VALVE- CONFIRMED. NO LEAK WAS OBSERVED FROM THE VALVE CAP. THIS DOES NOT MEET SPECIFICATION WHICH STATES THAT "CATHETER IN FINISHED PRODUCT 0128XXX AND 0109XXX, DAMAGES/PERFORATIONS ON CATHETER BONDING AND SUCTION CONTROL ARE NOT ALLOWED". A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE "INSUFFICIENT OR MISSING CYCLOHEXANONE". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE THE BALLOON TO BURST. CAUTION: DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. SINGLE PATIENT USE ONLY. DO NOT REUSE. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP TIP SYRINGE. DO NOT USE NEEDLE. RECOMMENDED INFLATION CAPACITIES: 3CC BALLOON: USE 5ML STERILE WATER; 5CC BALLOON: USE 10ML STERILE WATER; 15CC BALLOON: USE 20ML STERILE WATER; 20CC BALLOON: USE 25ML STERILE WATER; 30CC BALLOON: USE 35ML STERILE WATER; 40CC BALLOON: USE 45ML STERILE WATER; 75CC BALLOON: USE 80ML STERILE WATER; DO NOT EXCEED RECOMMENDED CAPACITIES. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE, AND FEDERAL LAWS AND REGULATIONS. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03305 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR KIT CATHETER (22G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1221252 |  | IT WAS REPORTED BY THE CUSTOMER "UPON NOTIFYING SUPERVISOR, HE REPORTED HE HAD ALSO HEARD FROM SEVERAL OTHERS TRAINED THAT THEY HAVE EXPERIENCED THIS OCCASIONALLY WHERE THE NEEDLE DOES NOT RETRACT COMPLETELY INTO THE SAFETY CHAMBER. STAFF MEMBER REPORTS HE HAS DONE ABOUT 16 SINCE TRAINING AND IT HAS HAPPENED AT LEAST 2 TIMES. OTHER TEAM MEMBERS TRAINED ALSO WERE REPORTING THIS BUT DID NOT LET LEADERSHIP KNOW UNTIL THIS WEEK. THESE TEAM MEMBERS RANGE FROM EMERGENCY, TRANSPORT, AND FLEX. EACH TEAM MEMBER HAVING EXPERIENCED THIS OCCUR ON A PATIENT AT LEAST ONCE. UNSURE IF USER ERROR OR MANUFACTURING CONCERN. CONCERN IS IF THERE IS AN ISSUE WITH THE ACTUAL SAFETY MECHANISM OR ARE WE HAVING SMALL USER ERRORS WHEN IT COMES TO PRESSING THE SAFETY BUTTON AND REMOVING GUIDE-WIRE AT THE SAME TIME. SUPERVISOR XXX ALSO CALLED MAIN ED AND RN AT MAIN CAMPUS WHO ALSO REPORT THIS OCCURRING BUT ARE UNSURE IF THIS IS USER ERROR OR MANUFACTURING ISSUE AS IT HAPPENS OCCASIONALLY. NURSES HAVE ALSO REPORTED IT HAS HAPPENED WITH ALL SIZE CATHETERS WE USE SUCH AS THE 22G 1.25IN. ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER "I DO NOT HAVE ANY MORE DETAILS TO SHARE. I HAVE FOLLOWED UP WITH A FEW NURSES WHO STATED THEY HAVE NOT HAD ANOTHER ISSUE WITH THE RETRACTOR BUTTON." THIS REPORT ADDRESSES THE FLEX EVENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03310 | 2023-08-02 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY W/ MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY CUSTOMER THAT THE CATHETER LEAKED FROM THE RUBBER-STOPPER WHEN PICC LINE WAS BEING INSERTED. NO PATIENT HARM. NO OTHER INFORMATION PROVIDED. ADDITIONAL INFORMATION RECEIVED 26 JULY 2023 : NO URGENT/LIFE THREATENING MEDICAL SITUATION. DID NOT INTERRUPT TREATMENT, OR CAUSE A CLINICALLY SIGNIFICANT DELAY IN TREATMENT, THAT NEGATIVELY IMPACTED THE PATIENT. STAT LOCK USED FOR CATHETER SECUREMENT. NO PATIENT HARM, INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03308 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC FT CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275108FD |  | IT WAS REPORTED BY CUSTOMER THAT LEAKING AT T-CONNECTOR OF PICC. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03301 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR KIT CATHETER (22G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1221252 |  | IT WAS REPORTED BY THE CUSTOMER "UPON NOTIFYING SUPERVISOR, HE REPORTED HE HAD ALSO HEARD FROM SEVERAL OTHERS TRAINED THAT THEY HAVE EXPERIENCED THIS OCCASIONALLY WHERE THE NEEDLE DOES NOT RETRACT COMPLETELY INTO THE SAFETY CHAMBER. STAFF MEMBER REPORTS HE HAS DONE ABOUT 16 SINCE TRAINING AND IT HAS HAPPENED AT LEAST 2 TIMES. OTHER TEAM MEMBERS TRAINED ALSO WERE REPORTING THIS BUT DID NOT LET LEADERSHIP KNOW UNTIL THIS WEEK. THESE TEAM MEMBERS RANGE FROM EMERGENCY, TRANSPORT, AND FLEX. EACH TEAM MEMBER HAVING EXPERIENCED THIS OCCUR ON A PATIENT AT LEAST ONCE. UNSURE IF USER ERROR OR MANUFACTURING CONCERN. CONCERN IS IF THERE IS AN ISSUE WITH THE ACTUAL SAFETY MECHANISM OR ARE WE HAVING SMALL USER ERRORS WHEN IT COMES TO PRESSING THE SAFETY BUTTON AND REMOVING GUIDE-WIRE AT THE SAME TIME. SUPERVISOR XXX ALSO CALLED MAIN ED AND RN AT MAIN CAMPUS WHO ALSO REPORT THIS OCCURRING BUT ARE UNSURE IF THIS IS USER ERROR OR MANUFACTURING ISSUE AS IT HAPPENS OCCASIONALLY. NURSES HAVE ALSO REPORTED IT HAS HAPPENED WITH ALL SIZE CATHETERS WE USE SUCH AS THE 22G 1.25IN. ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER "I DO NOT HAVE ANY MORE DETAILS TO SHARE. I HAVE FOLLOWED UP WITH A FEW NURSES WHO STATED THEY HAVE NOT HAD ANOTHER ISSUE WITH THE RETRACTOR BUTTON." THIS REPORT ADDRESSES THE FIRST EVENT AFTER TRAINING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03311 | 2023-08-02 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK APHERESIS PORT ACCESS SYS |  | IT WAS REPORTED THAT ONE MONTH AND TWENTY SEVEN DAYS POST PORT PLACEMENT, THE PORT ALLEGEDLY LEAKED BLOOD. IT WAS FURTHER REPORTED THAT BLEEDING WAS STOPPED BY APPLIED HEMOSTAT DRESSING. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3, H6 (METHOD). H11: H6 (DEVICE). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03321 | 2023-08-02 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERGLIDE | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT 3 POWERGLIDE ST 8CMS HAVE HAD ISSUES WITH INFUSING AND DRAWING AFTER PLACEMENT. SHE WAS ONLY ABLE TO TROUBLESHOOT ONE OF THEM COULD NOT GET IT TO DRAW OR INFUSE." THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03316 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | IT WAS REPORTED THAT DURING PLACEMENT OF THE PICC, THE PACKAGE WAS OPENED TO INSPECT THE CATHETER FOR FUNCTIONAL INTEGRITY. DURING THE INSPECTION, IT WAS FOUND THAT THE DISTAL END CONNECTOR OF THE CATHETER WAS DISLODGED FROM THE EXTENSION TUBING CONNECTION AND THE STYLET COULD NOT BE REMOVED. THE CATHETER WAS UNABLE TO BE USED PROPERLY. THE CATHETER PLACEMENT WAS COMPLETED WITH A NEW CATHETER REPLACEMENT. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED T-LOCK ASSEMBLY WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR SINGLE LUMEN POWERPICC CATHETER. THE SAMPLE WAS RECEIVED WITH AN INLAID STYLET AND ATTACHED T-LOCK ASSEMBLY. LIGHT RESIDUE WAS OBSERVED ON THE STYLET WIRE. A COMPLETE BREAK WAS OBSERVED IN THE T-LOCK ASSEMBLY PROXIMAL OF THE MALE LUER ADAPTER. MICROSCOPIC INSPECTION OF THE FRACTURE SITE REVEALED A SHARPLY DEFINED, FINELY GRANULAR FRACTURE SURFACE. A PORTION OF THE FRACTURE SURFACE EXHIBITED SLIGHT DEFORMATION AND WHITENING. AN ATTEMPT TO REPLICATE THE SAMPLE DAMAGE USING A SIMILAR NON-COMPLAINANT DEVICE RESULTED IN EXTENSIVE DEFORMATION AND DISCOLORATION. THE SLIGHT DEFORMATION AND DISCOLORATION SUGGESTED THAT MECHANICAL STRESSES LIKELY CONTRIBUTED TO THE BREAK; HOWEVER, ATTEMPTS TO REPLICATE THE DAMAGE RESULTED IN SIGNIFICANTLY MORE DEFORMATION, SUGGESTING THAT AN ADDITIONAL UNIDENTIFIED ENVIRONMENTAL FACTOR(S) LIKELY CONTRIBUTED TO MATERIAL FAILURE. CONSEQUENTLY, THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03320 | 2023-08-02 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERGLIDE | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT 3 POWERGLIDE ST 8CMS HAVE HAD ISSUES WITH INFUSING AND DRAWING AFTER PLACEMENT. SHE WAS ONLY ABLE TO TROUBLESHOOT ONE OF THEM COULD NOT GET IT TO DRAW OR INFUSE." THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03322 | 2023-08-02 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000375 |  | IT WAS REPORTED BY THE CUSTOMER "PICC LINE KINKED. LINE WAS REPORTED AS BLOCKED. WHEN INVESTIGATED BY ATTENDING CLINICIAN LINE WAS KINKED. RESOLVED BY CLINICIAN, LINE NOW PATENT AND WORKING." THE LINE IS STILL IN THE PATIENT AS WE PULLED THE LINE BACK AND IT IS NOW WORKING. THE LINE WAS INSERTED AND LEFT AT 0CM EXPOSED, FOLLOWING AN X-RAY IT NEEDED PULLING BACK TO 4CM EXPOSED. AT INSERTION AND AT THIS POINT IT WAS WORKING WELL HOWEVER IT BECAME PROBLEMATIC FOLLOWING THE PULL-BACK. MY TEAM ATTENDED THE NEXT DAY BECAUSE THE WARD STAFF HAD SAID IT WAS BLOCKED, ON INVESTIGATION IT APPEARED SOMEONE HOLDING THE SKIN NEAR TO THE INSERTION SITE HELPED THE PATENCY AND SO THE LINE WAS RETRACTED A FURTHER 1CM. WHEN IT WAS PULLED OUT THE SECOND TIME A KINK WAS NOTICED THAT HAD BEEN ABOUT 0.5CM INTO THE SKIN AND THIS WAS NOW STRAIGHTENED OUT. THE LINE NOW WORKS WELL. WE WILL TRY TO KEEP THE LINE ONCE IT IS READY FOR REMOVAL HOWEVER THIS MAY NOT BE POSSIBLE IF SHE GOES HOME TO FINISH HER ANTIBIOTICS." | NO SAMPLE WITH DOE: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03319 | 2023-08-02 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERGLIDE | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT 3 POWERGLIDE ST 8CMS HAVE HAD ISSUES WITH INFUSING AND DRAWING AFTER PLACEMENT. SHE WAS ONLY ABLE TO TROUBLESHOOT ONE OF THEM COULD NOT GET IT TO DRAW OR INFUSE." THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03307 | 2023-08-02 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR KIT CATHETER (22G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1221252 |  | IT WAS REPORTED THAT EVENT HAPPENED ON THE NIGHT OF (B)(6) 2023. WHEN PRESSING THE WHITE SAFETY BUTTON TO RETRACT NEEDLE, THE NEEDLE ONLY PARTIALLY RETRACTED AND HE COULD NOT REMOVE GUIDEWIRE/NEEDLE. ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER "I DO NOT HAVE ANY MORE DETAILS TO SHARE. I HAVE FOLLOWED UP WITH A FEW NURSES WHO STATED THEY HAVE NOT HAD ANOTHER ISSUE WITH THE RETRACTOR BUTTON." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03317 | 2023-08-02 | 2023-06-29 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THE PATIENT COMPLETED TREATMENT SESSIONS ON JUNE 29. DURING REMOVAL ON THE ORDER OF DOCTOR, THE CATHETER WAS REMOVED SMOOTHLY WITHOUT TRACTION OR OCCLUSIONS. AFTER REMOVAL, IT WAS FOUND THAT THE CATHETER WAS RUPTURED AT THE SCALE OF 11 CM WITH FLAT RUPTURE EDGES, WITHOUT ELASTIC DEFORMATION. THE PATIENT WAS ORDERED NOT FOR OFF-BED ACTIVITY. THE CARDIAC INTERVENTION DEPARTMENT WAS CONTACTED TO REMOVE THE RUPTURED PORTION OF THE PICC CATHETER UNDER DSA. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03315 | 2023-08-02 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8295118 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN PRE-FLUSHING THE CATHETER FOLLOWING UNPACKING, THE HCP FOUND THAT THE CATHETER LEAKED LIQUIDS AT THE SCALE OF 49 CM. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED CATHETER IS CONFIRMED. ONE 5 FR DUAL LUMEN POWERPICC SOLO CATHETER WITH A T-LOCK AND STIFFENING STYLET INSERTED INTO THE RED LUMEN, ONE 12 ML SYRINGE, ONE 10 ML SYRINGE, ONE 5.0 FR MICROINTRODUCER, ONE STATLOCK DEVICE, ONE 20 G IV INTRODUCER NEEDLE, ONE PAPER MEASURING TAPE, ONE NITINOL GUIDEWIRE IN A PLASTIC HOOP, AND TWO END CAPS WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED SOME USE RESIDUE ON THE STIFFENING STYLET JUST PROXIMAL TO THE T-LOCK AND ONE THE PAPER TAPE MEASURE. DAMAGE WAS OBSERVED IN THE CATHETER TUBING JUST DISTAL TO THE 49 CM DEPTH MARKER. A MICROSCOPIC OBSERVATION REVEALED TWO SMALL, CURVED SPLITS IN THE CATHETER JUST DISTAL TO THE 49 CM DEPTH MARKER. THE EDGES OF THESE SPLITS WERE OBSERVED TO BE DISTINCT. ONE SPLIT WAS OBSERVED TO BE GENERALLY PERPENDICULAR TO THE CATHETER TUBING AND THE OTHER SPLIT WAS OBSERVED TO BE SHARPLY ANGLED WITH A NARROW STRIP OF MATERIAL CONNECTING TO ONE EDGE OF THE OTHER SPLIT. STRIATIONS WERE OBSERVED ON THE FRACTURE SURFACES OF THE SPLITS, WHICH WERE MOSTLY FLAT AND STRAIGHT. THE CHARACTERISTICS OF THE SPLITS OBSERVED IN THE RETURNED SAMPLE SUGGEST THE CATHETER WAS DAMAGED BY A SHARP INSTRUMENT SUCH AS SCISSORS OR A SCALPEL. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03313 | 2023-08-02 | 2023-06-29 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | CUSTOMER REPORTED THAT AT THE END OF THE TREATMENT, THE PICC TUBE WAS REMOVED ACCORDING TO THE DOCTOR'S ADVICE. AFTER EXTUBATION, IT WAS FOUND THAT THE CATHETER WAS SHORT AND THERE WAS NO BLACK MARK AT THE FRONT END OF THE CATHETER. IMMEDIATELY PRESS THE CATHETER PUNCTURE POINT AND PRESERVE THE CATHETER. IT WAS FOUND THAT THE CATHETER BROKE AT 11CM, THE FRACTURE EDGE WAS FLAT, AND THERE WAS NO ELASTIC DEFORMATION. AFTER INQUIRING THE EXTUBATION NURSE, IT WAS FOUND THAT THE EXTUBATION PROCESS WAS SMOOTH, AND THERE WAS NO PULLING, BLOCK OR ELASTIC RETRACTION. IT IS PRELIMINARILY CONCLUDED THAT INTERNAL TUBE RUPTURE OCCURRED BEFORE EXTUBATION OF PICC CATHETER. THEN WENT TO THE WARD TO SEE THE PATIENT, ASKED THE PATIENT NO CHEST TIGHTNESS, SHORTNESS OF BREATH AND OTHER NO DISCOMFORT, FACE, EXTREMITY SKIN COLOR NORMAL. ASK THE PATIENT TO LIE FLAT AND STAY UP FROM BED. REPORT TO THE DOCTOR AND ORDER A CHEST X-RAY. THE SPECIALIST NURSE ACCOMPANIED THE FAMILY TO THE RADIOLOGY DEPARTMENT IN A WHEELCHAIR, AND QUICKLY CONTACTED THE CARDIAC INTERVENTIONAL DEPARTMENT OF (B)(6) TO REMOVE THE BROKEN PICC TUBE UNDER DSA. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03312 | 2023-08-02 | 2023-07-10 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | CUSTOMER REPORTED THAT THE PATIENT WAS DIAGNOSED WITH STAGE PT2N3BM0 IVB OF RIGHT SUBMAXILLARY LYMPHOEPITHELIOMATOID CARCINOMA AFTER SURGERY. FOR HYPERTENSION, IN ORDER TO FACILITATE INTRAVENOUS FLUID REPLENISHMENT, PICC TUBE WAS PLACED THROUGH THE LEFT MEDIAN VEIN AT 17:30 ON (B)(6) 2023 AND THE PROCESS WAS SMOOTH. THE PATIENT WAS RED, SWOLLEN, PAINFUL, SLIGHTLY HIGHER SKIN TEMPERATURE AND INDURATION ABOVE THE PICC TUBE OF THE LEFT UPPER LIMB. AFTER THE PICC TUBE OF THE LEFT UPPER LIMB WAS PLACED IN THE LEFT UPPER LIMB BY COLOR ULTRASOUND, THROMBOSIS OCCURRED AROUND THE PICC TUBE IN THE DISTAL CEPHALIC VEIN OF THE UPPER ARM. SWELLING AND THICKENING OF THE SURROUNDING SUBCUTANEOUS SOFT TISSUE. THERE WERE NO OBVIOUS ABNORMAL SONOGRAMS IN DEEP MOVEMENT AND VEIN OF LEFT UPPER LIMB. AFTER CONSULTATION, DEPUTY CHIEF PHYSICIAN OF VASCULAR SURGERY DEPARTMENT, IT WAS SUGGESTED THAT THE PATIENT CONSIDER THROMBOTIC SUPERFICIAL PHLEBITIS, RIVAROXABAN TABLET 10MGQD, CALCIUM OXYBENZOATE 0.5 ORAL TID, EXTERNAL USE OF POLYSULFONIC MUCOPOLYSCAN OINTMENT, IF PICC IS SMOOTH, THE PICC CAN CONTINUE BEING USED. | H11: SECTION A THROUGH F THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05731 | 2023-08-02 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THERE WAS A REGULAR PROBLEM WITH BARD FOLEY CATHETERS FOR A WHILE WITHIN THE INTENSIVE CARE DEPARTMENT. STATED THAT THEY WERE NOT DRAINING PROPERLY (BATCH# UNK) OR WERE CLOGGED (BATCH # UNK) AND THEREFORE THE PATIENTS GOT A FULL BLADDER. LAST WEEK, A PATIENT WHO HAD 1300 CC IN THE BLADDER AND HAD A CATHETER, WHEN REMOVED THE BALLOON LOOKED A BIT STRANGE (BATCH# UNK). IT WENT WELL FOR A WHILE, BUT LAST FALL THEY ALSO HAD THREE PATIENTS (BATCH# UNK). PER FOLLOW-UP INFORMATION RECEIVED FROM (B)(6) ON 13JUL2023, STATED THAT THE BALLOON LOOKED A BIT STRANGE MEANT THAT THE BALLOON LOOKED ADHERED AND WAS ALSO BLOODY BUT THAT COULD ALSO BE FROM THE PATIENT. FOR THIS CASE, THEY HAD NOW SEEN ONLY ONE PATIENT BUT THE LOT NUMBER COULD NOT BE TRACED. AS FOR THE CATHETERS THAT WERE NOT DRAINING PROPERLY, THEY DID NOT KNOW THE LOT NUMBER OR THE NUMBER OF PATIENTS THAT HAD THIS ISSUE. HOWEVER, THEY OFTEN HAD TO SIPHON TO EMPTY THE BLADDER PROPERLY, COULD BE DUE TO THE CATHETER BUT ALSO THE URIMETER. AS FOR THE CATHETERS THAT WERE CLOGGED, THEY DID NOT HAD ANY IDEA ABOUT THE LOT NUMBER OR THE NUMBER OF PATIENTS THAT HAD THIS ISSUE. AS FOR THE PREVIOUS PROBLEMS, NONE WERE REPORTED EXCEPT FOR ONE THAT WAS. IN THAT PATIENT, IT WAS TYPICAL THAT THEY DID URINATE 30-40 CC, BUT HAD ABDOMINAL PAIN. THEY TOOK OUT THE CATHETER AND AT THAT POINT THE URINE POURED OUT AND THE ABDOMINAL PAIN DISAPPEARED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE OPERATOR¿S ERROR- MISSED PLUG THE CATHETERS INTO THE SPIGOT ON THE MANIFOLD AND CAUSE OCCLUDED DRAINAGE EYE (REDUCED FLOW RATE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05730 | 2023-08-02 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 50614 |  | IT WAS REPORTED THAT THERE APPEARED TO BE A CLOG IN THE LINE, (LOT# JUHN2149 - DOE (B)(6) 2023) (LOT# UNK - DOE UNK) A DISCOLORATION AT THE TIP OF THE CATHETER (LOT# JUHN2149 - DOE (B)(6) 2023) (LOT# UNK - DOE UNK) NOT UNLIKE A HARDENED SUBSTANCE WAS BLOCKING THE FLOW OF FLUID. THIS HAD BEEN OCCURRING ON AND OFF THE LAST TWO MONTHS, POSSIBLY WITH OTHER UNKNOWN ADDITIONAL LOT NUMBERS. MULTIPLE UNKNOWN DATE OF EVENTS BUT ALSO ON (B)(6) 2023. PER FOLLOW UP INFORMATION RECEIVED VIA PHONE ON (B)(6) 2023, CUSTOMER STATED THAT THEY ONLY HAD ONE CATHETER SAMPLE AVAILABLE AND THEY WERE NOT SURE IF IT WAS FROM ONE OF THE REPORTED LOT NUMBERS AS THEY TOOK IT OUT OF THE TRASH AFTER SUBMITTING THE COMPLAINT. PER CUSTOMER, AFTER INSERTING THE CATHETER, THE TIP TURNED GREEN. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿INSUFFICIENT EYELET AREA BLOCKED EYE OR INNER DIAMETER OF DRAINAGE EYE SMALL ENOUGH TO ALLOW FOR OCCLUSION DURING USE¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT BUT IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03276 | 2023-08-02 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT BARB ON THE SHEATH. DURING CATHETER PLACEMENT, THREE UNITS USED SUCCESSIVELY WERE FOUND RUPTURED. AFFECTED CATHETER PLACEMENT OPERATIONS BY HEAD NURSE. NO OTHER INFORMATION WAS PROVIDED. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED MICROINTRODUCER SHEATH IS CONFIRMED; HOWEVER, THE ROOT CAUSE COULD NOT BE DETERMINED. ONE PHOTOGRAPH OF A MICROINTRODUCER SHEATH WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED A MICROINTRODUCER SHEATH IN A PLASTIC BAG. THE DISTAL END OF THE SHEATH APPEARS TO BE DAMAGED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05729 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53616G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE MAKING THE PATIENT SICK. LOT JUGT1738 HAD ALL BUT 7 WERE GOOD. LOT JUHQ9015 2 BOXES - WITH THOSE 75% TESTED WITH BACTERIA. ONLY ABOUT 25% WERE OK. THEY TESTED THEM AND WERE POSITIVE FOR BACTERIA, BUT THEY DID NOT DO FURTHER TESTING TO DETERMINE WHICH TYPE OF BACTERIA. THEY HAD ORDERED THEM FROM EXPRESS MEDICAL SUPPLY MOST RECENTLY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03314 | 2023-08-02 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | CUSTOMER REPORTED THE PATIENT HAD WATER LEAKAGE DURING THE PICC INFUSION PROCESS, AND THE NORMAL SALINE WAS FLUSHED INTO THE TUBE, AND THERE WAS WATER JET AT THE UPPER END OF THE CONNECTION PORT, AND INTRAVENOUS INFUSION COULD NOT BE PERFORMED. GIVE PERIPHERAL NEEDLE INFUSION AND INFORM THE PICC SPECIALIST NURSE TO PROTECT THE PICC PIPELINE UNDER ASEPTIC OPERATION. ON JULY 8, ASK THE PICC SPECIALIST NURSE TO CUT THE LEAKING SECTION UNDER ASEPTIC OPERATION AND REPLACE THE JOINT BEFORE SMOOTH USE. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03288 | 2023-08-02 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | THE CUSTOMER REPORTED THERE ARE ISSUES WITH THEIR SHERLOCK 3CG TCS CHEST SENSOR NOT GIVING THEM PROPER WIRE/PICC GUIDANCE/PLACEMENT. UPON REVIEW, IT LOOKS LIKE THE LAST 4 PICCS PLACED BY ALL 3 OF THE CLINICIANS ON SEPARATE OCCASIONS, NEEDED CHEST XRAYS. THIS REPORT ADDRESSES THE SECOND OCCURRENCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05717 | 2023-08-02 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT TRANSURETHRAL BLADDER TUMOR LASER RESECTION UNDER SPINAL ANESTHESIA ON (B)(6) 2023. AFTER THE OPERATION, A 20-GAUGE THREE-LUMEN URINARY CATHETER WAS INDWELLED FOR URINE DRAINAGE. ON (B)(6) 2023, THE URINARY CATHETER FELL OUT OF THE PATIENT DUE TO A BALLOON RUPTURE. AFTER REPORTING, THE DOCTOR ORDERED OBSERVATION AND SUGGESTED NOT TO INDWELL A NEW URINARY CATHETER. THEY ADVISED THE PATIENT TO DRINK MORE WATER AND NOT SUPPRESS URINE. AS PER THE INJURY PERFORMANCE, THE PATIENT WAS RE-CANNULATED. AS PER THE FOLLOW UP INFORMATION ON 21JUL2023, CLARIFIED THAT THE BALLOON BURST EVENT DID NOT HAVE ANY IMPACT (NO INJURY) TO THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03293 | 2023-08-02 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC FT CATHETER WITH SHERLOCK 3CG TPS STYLET 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275108FD |  | IT WAS REPORTED BY CUSTOMER THAT LEAKING AT T-CONNECTOR OF PICC. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05646 | 2023-08-02 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER FELL OUT OF THE PATIENT AFTER 3 DAYS OF PLACEMENT. WHEN THE USER ATTEMPTED TO INJECT STERILE WATER INTO THE BALLOON OUTSIDE THE PATIENT, THE WATER CAME OUT IMMEDIATELY. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6)2023, STATED THAT A PINHOLE WAS FOUND IN THE CATHETER TUBE UPON CHECKING THE RETURNED SAMPLE WHEN ASKED REGARDING THE LOCATION OF STERILE WATER LEAK. | THE REPORTED EVENT WAS CONFIRMED ¿ UNKNOWN CAUSE. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE CATHETER WAS EVALUATED AND OBSERVED MULTIPLE TEARS ON SHAFT. THE TEARS LOOK LIKE SCRATCHES CAUSED ANY SHARP OBJECT. INFLATED THE CATHETER WITH WATER. OBSERVED WATER LEAKAGE FROM THE TEARS ON THE SHAFT OF THE CATHETER. THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. HENCE, THE REPORTED EVENT IS CONFIRMED AS UNKNOWN CAUSE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE (EXAMPLE: USER RELATED/OPERATOR ERROR/ROUGH HANDLING OR EXPOSE TO SHARP OBJECT/MECHANICAL FORCE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05679 | 2023-08-02 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI12 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE FOLEY CATHETER. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. IT IS UNKNOWN IF THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND SOLUTION LEAKED FROM A TEAR UNDERNEATH THE INFLATION CAP MEASURING 0.1290". THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INCORRECT MOLDING OF EITHER THE INFLATION VALVE OR RETENTION CAP. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03279 | 2023-08-02 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT,4FR POWERMIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P4154108D |  | IT WAS REPORTED BY CUSTOMER THAT ATTEMPTED TO PLACE SINGLE LUMEN 4 FR MIDLINE CATHETER IN LEFT UPPER ARM IN LEFT BASILIC VEIN USING ULTRASOUND GUIDANCE AND 1.5CM DEPTH NEEDLE GUIDE. ESTABLISHED BLOOD RETURN BUT COULD NOT WIRE. THE WIRE AND NEEDLE WERE REMOVED. ATTEMPTED TO ACCESS SAME VEIN USING 1 CM NEEDLE GUIDE. APPEARED TO BE IN VESSEL WITH THE NEEDLE AND HAD BRISK BLOOD RETURN. WIRE INSERTED PART WAY AND RESISTANCE WAS MET. REMOVED NEEDLE OVER WIRE AS WIRE WAS NOT ABLE TO BE EASILY REMOVED WITH NEEDLE SIMULTANEOUSLY. WIRE NEEDED SOME FORCE IN ORDER TO BE REMOVED FROM PATIENT. WIRE IMMEDIATELY INSPECTED AND FOUND TO HAVE A KNOT AT ITS TIP. PROCEDURE ABORTED. PATIENT NOTIFIED. COVERING PROVIDER NOTIFIED AND WAS NOTIFIED TO OBTAIN IMAGING OF LUE TO LOOK FOR ANY WIRE FRAGMENTS POSSIBLY LEFT BEHIND. PROVIDER ASKED IF WE SHOULD TRY TO OBTAIN MIDLINE ON CONTRALATERAL SIDE AS PATIENT HAS NO IV ACCESS AT THIS TIME. PROVIDER INSTRUCTED PICC TEAM TO TRY FOR MIDLINE ON RIGHT ARM. PATIENT AGREEABLE TO THIS. RIGHT MIDLINE SUCCESSFULLY PLACED IN RUE. LIMB ALERT BRACELET LEFT IN PLACE ON LUE UNTIL IMAGING COMPLETED AND ANY FRAGMENTS HAVE BEEN RULED OUT. KNOTTED WIRE LINED UP WITH NEW, UNUSED WIRE FROM NEW KIT TO COMPARE LENGTH. TINY DISCREPANCY IN LENGTH BUT UNABLE TO DETERMINE IF KNOT UNDONE IF THAT WOULD ACCOUNT FOR THE MISSING LENGTH. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A KNOTTED GUIDEWIRE IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. THE PRODUCT RETURNED FOR EVALUATION WAS A 0.018¿ NITINOL GUIDEWIRE. USE RESIDUE WAS OBSERVED ON THE GUIDEWIRE. CURVED SHAPE MEMORY WAS OBSERVED NEAR THE DISTAL PORTION OF THE GUIDEWIRE. A KNOT WAS OBSERVED AT THE DISTAL END OF THE GUIDEWIRE. MICROSCOPIC EVALUATION REVEALED THE GUIDEWIRE WAS INTACT AND IT WAS CONFIRMED THERE WAS A KNOT AT THE DISTAL TIP OF THE GUIDEWIRE. THE FEATURES WERE CONSISTENT WITH THE GUIDEWIRE BECOMING KNOTTED DURING INSERTION. GUIDEWIRE KNOTS HAVE BEEN REPLICATED BY INSERTION AGAINST RESISTANCE WITH A TWISTING MOTION PRIOR TO WITHDRAWAL. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03282 | 2023-08-02 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F BASIC W/ MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT "PICC TEAM DISCOVERED KINKED 5F DUAL LUMEN PICC WHEN X RAY OF PATIENTS ARM AFTER FLOOR NURSE COULD NOT FLUSH LINE". NO OTHER INFORMATION WAS PROVIDED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A KINKED CATHETER WAS CONFIRMED AND APPEARS TO BE USE RELATED. THREE RADIOGRAPHIC IMAGES WERE RETURNED FOR EVALUATION. THE IMAGES SHOWED A PICC WITH A KINK AT THE ENTRANCE INTO THE VEIN. CONSISTENT WITH THE EVENT DESCRIPTION, THE KINK AT THE VEIN ENTRANCE SITE APPEARED TO BE TIGHT ENOUGH TO AFFECT FLOW THROUGH THE CATHETER. THERE WAS ALSO A POSSIBLE KINK AT THE SKIN ENTRANCE. THE CATHETER HAD TO ¿BACK TRACK¿ FROM THE SKIN ENTRANCE SITE TO THE MORE DISTAL VEIN ENTRANCE. THE ENTRANCE INTO THE SKIN AND VEIN CAN BE POINTS OF ENTRAPMENT AND KINKING. THE KINK AT THE VEIN ENTRANCE APPEARED TO BE RELATED TO THE ABNORMAL ANGLE OF APPROACH FOR THE VENOUS ACCESS. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03275 | 2023-08-02 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED ON (B)(6) 2023, THE POWER PICC SOLO WAS INSERTED FROM THE RIGHT BASILIC VEIN. THE INSERTION SITE WAS LOCATED AT 10 CM ABOVE THE CUBITAL CREASE, AND 36 CM OF THE CATHETER WAS INSERTED INTERNALLY, THE TIP AT T6. ON (B)(6), SEEPAGE THAT WAS CLEAR LIQUID OCCURRED AT THE INSERTION SITE. IT WAS CONSIDERED AS LIQUIDS FROM THE LYMPHANGION. THE GAUZE AND DRESSING WERE CHANGED EVERY DAY. ON (B)(6), THE SEEPAGE AT THE INSERTION SITE TURNED LIGHT GREEN. THE DRESSING WAS REMOVED AND AFTER PRE-FLUSHING WITH 14 ML OF SOLUTION, EVIDENT SEEPAGE OCCURRED. SUBSEQUENTLY, 4 CM OF THE CATHETER WAS RETRACTED, AND DAMAGE AND LIQUID LEAKS WERE FOUND. THE PICC CATHETER WAS REMOVED. ADDITIONAL INFORMATION RECEIVED 7/23/2023: NO BLOOD CULTURES WERE DRAWN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER LEAK IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE PHOTO SAMPLE OF A PRODUCT LABEL INDICATING LOT: REGT3140 AND ONE VIDEO SAMPLE WERE PROVIDED FOR EVALUATION. THE VIDEO SHOWS A 4 FR POWERPICC SOLO CATHETER OUTSIDE OF PATIENT USE. THE CLINICIAN IS INFUSING A CLEAR FLUID THROUGH THE HUB AND A BEAD OF FLUID IS OBSERVED TO LEAK NEAR THE 3 CM DEPTH MARKER. THE CHARACTERISTICS OF THE DAMAGE COULD NOT BE INSPECTED FROM THE VIDEO PROVIDED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03297 | 2023-08-02 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 5608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY HAD HIVES AROUND THE PORT. IT WAS FURTHER REPORTED THAT THE HEALTHCARE PROFESSIONAL ORDERED TO DISCONTINUE THE FLUSHING. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING ADDITIONAL MEDICAL INTERVENTION OR MEDICATION TO TREAT HIVES. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENT REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05681 | 2023-08-02 | 2023-07-11 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT USED PUREWICK FEMALE EXTERNAL CATHETERS FOR 3 DAYS AND THE PATIENT GOT A BAD URINARY TRACT INFECTION. PER THE FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE PATIENT WAS PRESCRIBED AN ANTIBIOTIC FOR PATIENT¿S URINARY TRACT INFECTION. THE CUSTOMER ALSO REPORTED THAT AFTER THE URINARY TRACT INFECTION CLEARED UP, THE PATIENT CONTINUED TO USE THE PUREWICK FEMALE EXTERNAL CATHETER. NO OTHER SERIOUS IMPACTS WERE REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PUREWICK¿ FEMALE EXTERNAL CATHETER INSTRUCTIONS FOR USE SETUP: 1. CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE PUREWICK¿ URINE COLLECTION SYSTEM, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE PUREWICK¿ URINE COLLECTION SYSTEM USER GUIDE FOR FURTHER INFORMATION. 2. USING STANDARD SUCTION TUBING, CONNECT THE PUREWICK¿ FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER. PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICK¿ FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICK¿ FEMALE EXTERNAL CATHETER IS POSITIONED. REMOVAL: 5. TO REMOVE THE PUREWICK¿ FEMALE EXTERNAL CATHETER, FULLY SEPARATE THE LEGS, GLUTEUS, AND LABIA. TO AVOID POTENTIAL SKIN INJURY UPON REMOVAL, GENTLY PULL THE PUREWICK¿ FEMALE EXTERNAL CATHETER DIRECTLY OUTWARD. ENSURE SUCTION IS MAINTAINED WHILE REMOVING THE PUREWICK¿ FEMALE EXTERNAL CATHETER. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. MAINTENANCE: 6. REPLACE THE PUREWICK¿ FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICK¿ FEMALE EXTERNAL CATHETER. EXTERNAL USE ONLY INDICATION FOR USE: THE PUREWICK¿ FEMALE EXTERNAL CATHETER IS INTENDED FOR NON-INVASIVE URINE OUTPUT MANAGEMENT IN FEMALE PATIENTS. CONTRAINDICATIONS: ¿ PATIENTS WITH URINARY RETENTION WARNINGS: ¿ DO NOT USE THE PUREWICK¿ FEMALE EXTERNAL CATHETER WITH BEDPAN OR ANY MATERIAL THAT DOES NOT ALLOW FOR SUFFICIENT AIRFLOW. ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICK¿ FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICK¿ FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. ¿ AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: - AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICK¿ FEMALE EXTERNAL CATHETER - HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE - EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE - EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON ¿ DO NOT USE BARRIER CREAM ON THE PERINEUM WHEN USING THE PUREWICK¿ FEMALE EXTERNAL CATHETER. BARRIER CREAM MAY IMPEDE SUCTION. ¿ PROCEED WITH CAUTION IN PATIENTS WHO HAVE UNDERGONE RECENT SURGERY OF THE EXTERNAL UROGENITAL TRACT. ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICK¿ FEMALE EXTERNAL CATHETER. ¿ MAINTAIN SUCTION UNTIL THE PUREWICK¿ FEMALE EXTERNAL CATHETER IS FULLY REMOVED FROM THE PATIENT TO AVOID URINE BACKFLOW. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ PRIOR TO CONNECTING THE PUREWICK¿ FEMALE EXTERNAL CATHETER TO HOSPITAL WALL SUCTION TUBING, VERIFY SUCTION FUNCTION BY COVERING THE OPEN END OF THE SUCTION TUBING WITH ONE HAND AND OBSERVING THE PRESSURE DIAL. IF THE PRESSURE DOES NOT INCREASE WHEN THE LINE IS COVERED, VERIFY THAT THE TUBING IS SECURED, CONNECTED, AND NOT KINKED. ¿ ENSURE THE PUREWICK¿ FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICK¿ FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. ¿ PROPERLY PLACING THE PUREWICK¿ FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICK¿ FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICK¿ FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICK¿ FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03328 | 2023-08-03 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH & PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D4 |  | IT WAS REPORTED BY CUSTOMER THAT HE CLEARED PLASTIC HUB WHICH THE INSERTION WIRE GOES THROUGH IS BOTH LEAKING AIR AND FLUID. WHEN HE DREW BACK TO CHECK FOR BLOOD FLOW, AIR GETS PULLED INTO SYRINGE. WHEN HE FLUSHED THE SALINE SQUIRTS OUT AROUND WIRE. IT IS MORE THAN JUST A BUBBLE OR TWO OF AIR. I OBSERVED BETWEEN 2-3 CC OF AIR DRAWN UP INTO FLUSH SYRINGE. IT WAS NOT DUE TO A LOOSE HUB. UNABLE TO QUANTIFY AMOUNT OF SALINE THAT LEAKS OUT. ADD INFO RECEIVED ON 21-JUL-2023. CUSTOMER REPORTED "DELAYS WITH INSERTION OF THE PICC LINE. NO URGENT/LIFE THREATENING MEDICAL SITUATION". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05744 | 2023-08-03 | 2023-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI12 |  | IT WAS REPORTED THAT A FEW HOURS AFTER INSERTION, NO URINE WAS FLOWING AND UPON CHECKING THE CATHETER, THE BALLOON HAD RUPTURED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] . METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX. PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 8 TO 24 EVERY 2FR: FOLEY CATHETER, ACCESSORIES, CLOSED DRAINAGE BAG. (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER, WATER SOLUBLE LUBRICANT, ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION. TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS, GLOVES, STATLOCK® FOLEY. [INTENDED USE & EFFECT- EFFICACY]. THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] . METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT. OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING: CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) PRECAUTIONS FOR USE: WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS). EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD: INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) MALFUNCTION AND ADVERSE EVENTS: MALFUNCTION. CATHETER KINKING, DAMAGE, RUPTURE, DIFFICULTY OR FAILURE TO REMOVE THE DEVICE, OCCLUSION OF CATHETER INNER LUMENS, ENCRUSTATION, ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE DEVICE DAMAGE DUE TO INAPPROPRIATE USE. ADVERSE EVENTS: URINARY-TRACT INFECTION, HEMORRHAGE, HEMATURIA, ALLERGY REACTION TO THE DEVICE, CALCULUS FORMATION, EDEMA, PAIN, DISCOMFORT, INJURY OF BLADDER OR URETHRAL, URETHRITIS, URINARY INCONTINENCE, RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05743 | 2023-08-03 | 2023-07-14 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WAS HAVING SKIN REACTION TO PRODUCT WHEN BEEN USING PUREWICK FEMALE EXTERNAL CATHETER SINCE NOVEMBER. THE CUSTOMER ASKED THE PATIENT THERE HAS BEEN ANY CHANGES IN THE WICK AND IF REPORTED OF OTHER PATIENTS WERE HAVING ISSUES. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW UP VIA PHONE ON 20JULY2023, PATIENT WAS BEING TREATED WITH BARRIER CREAMS FOR THE REACTION. CUSTOMER HAD BEEN USING THE WICKS SINCE (B)(6) 2022 WITHOUT ISSUE AND JUST STARTED EXPERIENCING THE REACTION IN THE LAST THREE WEEKS WHEN THEY RECEIVED THE MOST RECENT ORDER OF WICKS. CUSTOMER STATED PATIENT HAD GIVEN LOT INFORMATION AT ANOTHER TIME AND DOES NOT HAVE IT CURRENTLY. PER ADDITIONAL INFORMATION RECEIVED ON 20JUL2023, THE PATIENT HAD REPLACED 77 CATHETERS THAT WERE CAUSING THE PATIENT SKIN ISSUES THAT ALL HAD THE SAME LOT. PER NOTIFICATION FROM INVESTIGATOR ON 05OCT2023, DURING SAMPLE EVALUATION IT WAS FOUND THE FAILED DIMENSIONAL EVALUATION. | THE REPORTED EVENT IS INCONCLUSIVE, AS THE PATIENT'S ANATOMY IS UNKNOWN. NO VISUAL DEFECTS WERE NOTED ON ALL RETURNED WICKS. ALL WICKS WERE INSPECTED AND FOUND THAT SOME OF THEM WERE SHAPED IRREGULARLY WITH THE GAUZE NOT BEING UNIFORM ALL THE WAY DOWN THE WICK. THE GAUZE LOOKED LUMPY ON A FEW OF THE WICKS. NO FOREIGN MATTER WAS OBSERVED. A POTENTIAL ROOT CAUSES INCLUDE INADEQUATE MATERIAL SELECTION/MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE/MATERIAL SURFACE IS ROUGH, ABRASIVE OR UNCOMFORTABLE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿SETUP: CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE DRYDOCTM VACUUM STATION, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE DRYDOCTM VACUUM STATION USER GUIDE FOR FURTHER INFORMATION. USING STANDARD SUCTION TUBING, CONNECT THE PUREWICKTM FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER. PERI-CARE AND PLACEMENT: PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. REMOVAL: TO REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER, FULLY SEPARATE THE LEGS, GLUTEUS, AND LABIA. TO AVOID POTENTIAL SKIN INJURY UPON REMOVAL, GENTLY PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER DIRECTLY OUTWARD. ENSURE SUCTION IS MAINTAINED WHILE REMOVING THE PUREWICKTM FEMALE EXTERNAL CATHETER. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. MAINTENANCE: REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | N | LAY USER/PATIENT |  | 2023-08
1018233-2023-05742 | 2023-08-03 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT STERILE WATER COULD NOT BE DRAINED DURING PRETEST. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 20JUL2023, THE DETAILS ABOUT RESISTANCE WHILE INFLATING THE BALLOON WAS NOT PROVIDED FROM CUSTOMER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. PHOTO: THERE ARE 2 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE FIRST PHOTO ATTACHED IS AN OVERVIEW OF THE CATHETER WIT CUT PORTION OF INLET TUBING AND SYRINGE. THE SECOND PHOTO SHOW A CLOSER LOOK AT THE CATHETER BALLOON AND THE TIP. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING AND SYRINGE. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. USING THE RETURNED SYRINGE, THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 3 MINUTES WITH NO CUFFING OR LEAKS NOTED, RETURNING 10ML OF SOLUTION. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. AS THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05740 | 2023-08-03 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 53616G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE MAKING THE PATIENT SICK. LOT JUGT1738 HAD ALL BUT 7 WERE GOOD. LOT JUHQ9015 2 BOXES - WITH THOSE 75% TESTED WITH BACTERIA. ONLY ABOUT 25% WERE OK. THEY TESTED THEM AND WERE POSITIVE FOR BACTERIA, BUT THEY DID NOT DO FURTHER TESTING TO DETERMINE WHICH TYPE OF BACTERIA. THEY HAD ORDERED THEM FROM EXPRESS MEDICAL SUPPLY MOST RECENTLY. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "NO FOLLOW UP TO PRODUCTION AREAS CLEANING PROCEDURE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03327 | 2023-08-03 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT APPROXIMATELY ONE MONTH AND TWENTY SEVEN DAYS POST PORT PLACEMENT PROCEDURE, THE PORT ALLEGEDLY LEAKED BLOOD. IT WAS FURTHER REPORTED THAT BLEEDING WAS STOPPED BY APPLIED HEMOSTAT DRESSING. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05746 | 2023-08-03 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER COUDE TIP | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0168L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON BURST ONCE INSIDE PATIENT. PROPER AMOUNT OF WATER WAS INSTILLED, SO THE USER WAS UNSURE WHY THE BALLOON BURST. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023,IT WAS STATED THAT PATIENT STATUS WAS FINE. THE FOLEY CATHETER WAS INSERTED AND A FEW HOURS LATER THE PATIENT CALLED NURSING TO REPORT SINCE THEY FELT A POP AND THAT THEIR CATHETER WAS FALLING OUT. NURSE ASSESSED AND NOTED A SPLIT IN THE BALLOON PART OF THE CATHETER. 10CC WAS INSERTED INTO THE BALLOON AS PER INSTRUCTIONS ON THE CATHETER END. IT WAS STATED THAT THAT THERE WAS SOME SLIGHT TRAUMA TO THE URETHRA, BUT ALSO HAD TO HAVE ANOTHER CATHETER INSERTION SO NOT SURE WHICH CAME FIRST. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON (B)(6) 2023, THE CUSTOMER WAS UNSURE IF THE URETHRA TRAUMA WAS THERE BEFORE THE FIRST CATHETER WAS INSERTED. THERE WAS NO ISSUE WITH THE SECOND CATHETER, IT WAS STATED TO ALERT THAT ANOTHER CATHETER WAS ABLE TO BE PLACED IN THE CUSTOMER DUE TO THE ORIGINAL CATHETER BURSTING. IT WAS REPORTED THAT THE PATIENT STATUS WAS FINE, AND NO ADDITIONAL MEDICAL INTERVENTION WAS REPORTED | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05737 | 2023-08-03 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 |  | INTS14 |  | IT WAS REPORTED THAT THE BAG COULD NOT DROP DOWN FAR ENOUGH FOR THE URINE TO DRAIN. ALSO STATED THAT IT STARTED, THEN STOPPED AND DID NOT DRAIN FURTHER. SO IN THE MEANTIME, THE NURSES ON THAT FLOOR WERE OPENING A BRAND NEW FULL FOLEY KIT AND TAKING THE CATHETER OUT, THEN DISCARDING THE REST. (PCN INTS14, PCN INTS16) PER FOLLOW UP VIA EMAIL ON 21JUL2023, IT WAS STATED THAT THEY HAD WORKED WITH THE REPRESENTATIVE ON EDUCATION FOR THE UNIT. CURRENTLY THAT HAD A POSITIVE EFFECT ON STAFF¿S USAGE OF THE PRODUCT. PER FOLLOW UP VIA EMAIL ON 21JUL2023, STATED THAT THE CUSTOMER MADE IT CLEAR THAT THE ISSUE HAS BEEN RESOLVED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "HUMAN ERROR / LACK OF TRAINING". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT ¿ USE THE SMALLEST CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, AND INDIVIDUAL WHO INSERTED CATHETER INDICATIONS FOR USE: PVC & RED RUBBER: FOR UROLOGICAL USE ONLY SILICONE: INTENDED FOR USE FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION AND MEASUREMENT. THE DEVICES ARE GENERALLY PASSED TO THE URINARY BLADDER VIA THE URETHRA. NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA MINORA MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD BARD® EZ-LOK¿ SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK (FIG. 1A) OR SLIP TIP COLLECTION DEVICES (FIG. 1B). BD VACUTAINER® URINE PRODUCTS CAN BE USED FOR COLLECTION, STORAGE, AND TRANSPORT OF URINE SPECIMENS FROM A URINARY CATHETER. 1. SWAB SURFACE OF BARD® EZ-LOK¿ SAMPLING PORT WITH ANTISEPTIC (FIG. 2) E.G., SITE-SCRUB¿ IPA DEVICE WHICH IS AN EFFECTIVE DISINFECTING AGENT FOR LUER HUBS. 2. USING ASEPTIC TECHNIQUE, POSITION THE COLLECTION DEVICE IN THE CENTER OF THE SAMPLING PORT. PRESS THE DEVICE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. (FIG. 3) 3. IF USING BD VACUTAINER® LUER-LOK¿ ACCESS DEVICE (FIG. 4), COLLECT URINE INTO THE BD VACUTAINER® TUBES PER HOSPITAL PROTOCOL. IF USING SYRINGE, SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND TRANSFER TO COLLECTION CUP OR FOLLOW HOSPITAL PROTOCOL. 4. AFTER SAMPLE IS OBTAINED, DISCARD COLLECTION DEVICE ACCORDING TO HOSPITAL PROTOCOL. 5. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB. DIRECTIONS FOR USE 1. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS 2. WASH HANDS AND DON CLEAN GLOVES 3. DON STERILE GLOVES 4. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN 5. EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT 6. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERI-URETHRAL AREA 7. POSITION FENESTRATED DRAPE ON PATIENT 8. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND A. WHEN CATHETER TIP HAS ENTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE 9. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL 10. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL 11. LUBRICATE CATHETER 12. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03330 | 2023-08-03 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET (4F) (SL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000677 |  | IT WAS REPORTED BY CUSTOMER THAT LEAKAGE FOUND AT THE INSERTION SITE. PATIENT COMPLAINED OF BURNING AT THE SITE. EVALUATED AND FOUND LEAK AT THE 2CM MARK JUST BELOW THE SKIN LEVEL AT THE INSERTION SITE. REMOVED. RESPONSE RECEIVED 27 JULY 2023: THE EVENT DID NOT INVOLVE AN URGENT OR LIFE-THREATENING MEDICAL SITUATION, CAUSE A CLINICALLY SIGNIFICANT DELAY IN TREATMENT, INTERRUPTION IN TREATMENT, AND CAUSED NO HARM TO THE PATIENT. PATIENT WAS RECEIVING OR PLANNED TO RECEIVE CEFAZOLIN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK IN THE CATHETER IS CONFIRMED BUT THE CAUSE IS UNKNOWN. ONE 4FR SL POWERPICC CATHETER WAS RETURNED FOR EVALUATION. DURING THE INVESTIGATION, A LEAK WAS OBSERVED IN THE CATHETER BODY FROM A CRESCENT CIRCUMFERENTIAL SPLIT. THE FINDINGS WERE CONSISTENT WITH CATHETER DAMAGE CAUSED BY CONTACT WITH A SHARP INSTRUMENT. THE FRACTURE EDGES WERE SHARPLY FORMED AND HAD PLANAR (FLAT) SURFACES. THE EXACT CAUSE OF THE DAMAGE IS UNKNOWN AND THE SHORT TIME BETWEEN THE INSERTION DATE AND THE DATE OF DISCOVERY SUGGESTS THIS DAMAGE MAY HAVE OCCURRED DURING INSERTION. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO THE MANUFACTURE OF THE PRODUCT. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05777 | 2023-08-03 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT AFTER THE FOLEY CATHETER WAS INSERTED, IT WAS WITHDRAWN IN THE WARD AND THE BALLOON WAS FOUND TO BE RUPTURED. THE EVENT OCCURRED 12 HOURS AFTER PLACEMENT. IT WAS STATED THAT THE BALLOON RUPTURE OR TEAR WAS OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BALLOON. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05745 | 2023-08-03 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 50614 |  | IT WAS REPORTED THAT THE MAGIC 3 GO MALE INTERMITTENT CATHETER WERE NOT LUBRICATED AND THE PATIENT WAS HAVING A DIFFICULT TIME INSERTING THEM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-08
1018233-2023-05755 | 2023-08-03 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAD RUPTURED. THE PATIENT WAS REINTUBATED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05762 | 2023-08-03 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS INJECTED WITH WATER, STERILE WATER LEAKED FROM THE BALLOON PORTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
17450451 | 2023-08-03 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | DURING PLACEMENT, WHEN CHECKING FOR BLOOD RETURN, NOTED AIR IN T PIECE CONNECTOR AND NS SYRINGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-08
17450488 | 2023-08-03 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | PATIENT NOTED TO BE BLEEDING FROM PICC SITE. THE PORT FROM THE PICC HUB BECAME DISCONNECTED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05772 | 2023-08-03 | 2023-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT A DAY AFTER CATHETER PLACEMENT TO THE PATIENT, THE CATHETER FELL AND IT WAS FOUND THAT THE BALLOON HAD BURST. A PRETEST WAS DONE PRIOR TO THE PLACEMENT AND IT WAS CONFIRMED THEN THAT THERE WAS NO TEAR IN THE BALLOON. THE PATIENT DID NOT HAVE BLADDER STONES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05761 | 2023-08-03 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. ON (B)(6)2023, THE CATHETER WAS INSERTED INTO THE PATIENT. DURING THE PRETEST, THE BALLOON HAD NO DAMAGE AND INFLATED WITHOUT ANY ISSUES. AT 10: 30 PM, THERE WAS A URINE LEAK AND UPON CHECKING THE STERILE WATER IN THE BALLOON, THE DOCTOR FOUND ONLY ABOUT 6CC OF STERILE WATER, SO THE CATHETER WAS REINSERTED AND FILLED WITH 10CC OF STERILE WATER. ON (B)(6)2023, UPON CHECKING FOR SHEET CONTAMINATION, THE CATHETER WAS FOUND TO HAVE FALLEN OUT WITH NO STERILE WATER INSIDE. NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER AND THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 26JUL2023, THERE WAS NO ASYMMETRICAL BALLOON ISSUE AND THE CUSTOMER DID NOT REPORT ON THAT. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 01AUG2023, THERE WAS NO APPARENT ABNORMALITY FOR THE WATER TO LEAK FROM THE BALLOON AND DAMAGE ON THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05758 | 2023-08-03 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAD RUPTURED. THE PATIENT WAS REINTUBATED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03329 | 2023-08-03 | 2023-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH & PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D4 |  | IT WAS REPORTED BY CUSTOMER THAT HE CLEARED PLASTIC HUB WHICH THE INSERTION WIRE GOES THROUGH IS BOTH LEAKING AIR AND FLUID. WHEN HE DREW BACK TO CHECK FOR BLOOD FLOW, AIR GETS PULLED INTO SYRINGE. WHEN HE FLUSHED THE SALINE SQUIRTS OUT AROUND WIRE. IT IS MORE THAN JUST A BUBBLE OR TWO OF AIR. I OBSERVED BETWEEN 2-3 CC OF AIR DRAWN UP INTO FLUSH SYRINGE. IT WAS NOT DUE TO A LOOSE HUB. UNABLE TO QUANTIFY AMOUNT OF SALINE THAT LEAKS OUT. ADD INFO RECEIVED ON 24-JUL-2023 IT CAUSED DELAYS WITH INSERTION OF THE PICC LINE PROCEDURE. NO URGENT/LIFE THREATENING MEDICAL SITUATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05756 | 2023-08-03 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAD RUPTURED. THE PATIENT WAS INSERTED WITH ANOTHER CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05753 | 2023-08-03 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO HOSPITAL AND THE DOCTOR PERFORMED CATHETERIZATION. DURING THE OPERATION, THE DISPOSABLE STERILE FOLEY CATHETER WAS FOUND TO BE LEAKING FROM THE CATHETER LUMEN. IMMEDIATELY THE USER REPLACED IT WITH A NEW DISPOSABLE STERILE CATHETER, AND THEN USED IT NORMALLY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05760 | 2023-08-03 | 2023-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE BALLOON BURST IMMEDIATELY AFTER THE INSERTION OF FOLEY CATHETER. NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER. THERE WAS NO PATIENT'S URINE STONE AND THE CATHETER WAS NOT PULLED. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. EVALUATION CONFIRMED THE BALLOON WAS BURST. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05764 | 2023-08-03 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119310M |  | IT WAS REPORTED THAT CUSTOMER HAVING A FEW ISSUES WITH BARD TEMPERATURE SENSING CATHETERS DISCONNECTING FROM THE URI METERS RECENTLY. CUSTOMER ASKED THEY WERE JUST WONDERING IF THERE HAVE BEEN ANY RECENT CHANGES TO THE CATHETERS AS FAR AS THE PLASTICS BEING USED, ANY SLIGHT CHANGE IN SIZE OF CONNECTION END, OR ANYTHING ELSE CAN THINK OF THAT MIGHT BE CAUSING THE ISSUE. PER SAMPLE RECEIVED ON (B)(6) 2023, CUSTOMER RETURNED 4 USED ALL SILICONE FOLEY CATHETERS WITH MATERIAL NUMBER (B)(4) . | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), ALL-SILICONE TEMPERATURE SENSING FOLEY CATHETER. VISUAL INSPECTION NOTED THE FOLEY CATHETER WAS CONNECTED TO THE (IN-HOUSE) DRAINAGE BAG AND NO DISCONNECT OBSERVED. THIS IS WITHIN SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. CORRECTION- D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05763 | 2023-08-03 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON BURST INSIDE THE PATIENT'S BLADDER DURING INFLATION WITH 10ML NORMAL SALINE. STATED THAT NO IMPACT NOTED TO PATIENT AND FOLEY BALLOON IMMEDIATELY REMOVED WHEN BURST WAS HEARD. THE BALLOON WAS INSPECTED WITH REGISTERED NURSE AND A TEAR WAS NOTED ON THE OUTER LAYER OF THE FOLEY TIP, BELOW THE DRAINAGE HOLE AND A NEW FOLEY WAS PLACED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: INSERT FOLEY CATHETERS ONLY FOR APPROPRIATE INDICATIONS AND LEAVE IN PLACE ONLY AS LONG AS NEEDED. CDC GUIDELINES FOR APPROPRIATE INDICATIONS FOR INDWELLING URETHRAL CATHETER USE: PATIENT HAS ACUTE URINARY RETENTION OR BLADDER OUTLET OBSTRUCTION. NEED FOR ACCURATE URINE OUTPUT MEASUREMENTS. USE FOR SELECTED SURGICAL PROCEDURES. TO ASSIST IN HEALING OF OPEN SACRAL OR PERINEAL WOUNDS. PATIENT REQUIRES PROLONGED IMMOBILIZATION. TO IMPROVE COMFORT FOR END OF LIFE CARE. PROPER TECHNIQUES FOR URINARY CATHETER INSERTION: PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION. INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN THE PATIENT RECORD. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: SECURE THE FOLEY CATHETER. USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED. MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. -MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. EMPTY THE COLLECTION BAG REGULARLY USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT. GENERALLY, DRAINAGE IS ACCOMPLISHED BY INSERTING THE CATHETER THROUGH THE URETHRA AND INTO THE BLADDER. HOWEVER, DRAINAGE IS SOMETIMES ACCOMPLISHED BY SUPRAPUBIC OR OTHER PLACEMENT OF THE CATHETER, SUCH AS NEPHROSTOMY TRACT". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03331 | 2023-08-03 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC SOLO DL CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. FLUID LEAKING FROM A DOUBLE LUMEN POWERPICC SOLO CENTRAL LINE INSERTION SITE. INSERTED IN THE UPPER ARM. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05774 | 2023-08-03 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 |  | INTS16 |  | IT WAS REPORTED THAT THE BAG COULD NOT DROP DOWN FAR ENOUGH FOR THE URINE TO DRAIN. ALSO STATED THAT IT STARTED, THEN STOPPED AND DID NOT DRAIN FURTHER. SO IN THE MEANTIME, THE NURSES ON THAT FLOOR WERE OPENING A BRAND NEW FULL FOLEY KIT AND TAKING THE CATHETER OUT, THEN DISCARDING THE REST. (PCN INTS14), (PCN INTS16). PER FOLLOW UP VIA EMAIL ON 21JUL2023, IT WAS STATED THAT THEY HAD WORKED WITH THE REPRESENTATIVE ON EDUCATION FOR THE UNIT. CURRENTLY THAT HAD A POSITIVE EFFECT ON STAFF¿S USAGE OF THE PRODUCT. PER FOLLOW UP VIA EMAIL ON 21JUL2023, STATED THAT THE CUSTOMER MADE IT CLEAR THAT THE ISSUE HAS BEEN RESOLVED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "HUMAN ERROR / LACK OF TRAINING". THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿DIRECTIONS FOR USE PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT ¿ USE THE SMALLEST CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, AND INDIVIDUAL WHO INSERTED CATHETER INDICATIONS FOR USE: PVC & RED RUBBER: FOR UROLOGICAL USE ONLY SILICONE: INTENDED FOR USE FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION AND MEASUREMENT. THE DEVICES ARE GENERALLY PASSED TO THE URINARY BLADDER VIA THE URETHRA. 1. WASH HANDS AND DON CLEAN GLOVES 2. EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERI-URETHRAL AREA 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL 5. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP 6. DON STERILE GLOVES 7. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE 8. POSITION FENESTRATED DRAPE ON PATIENT 9. SATURATE 3 FOAM SWAB STICKS IN POVIDONE IODINE 10. LUBRICATE CATHETER 11. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA MINORA MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD 12. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND A. WHEN CATHETER TIP HAS ENTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE 13. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL EZ-LOK TM SAMPLING PORT BARD® EZ-LOK¿ SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK OR SLIP TIP COLLECTION DEVICES. BD VACUTAINER® URINE PRODUCTS CAN BE USED FOR COLLECTION, STORAGE, AND TRANSPORT OF URINE SPECIMENS FROM A URINARY CATHETER. 1. SWAB SURFACE OF BARD® EZ-LOK¿ SAMPLING PORT WITH ANTISEPTIC E.G., SITE-SCRUB¿ IPA DEVICE WHICH IS AN EFFECTIVE DISINFECTING AGENT FOR LUER HUBS. 2. USING ASEPTIC TECHNIQUE, POSITION THE COLLECTION DEVICE IN THE CENTER OF THE SAMPLING PORT. PRESS THE DEVICE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. 3. IF USING BD VACUTAINER® LUER-LOK¿ ACCESS DEVICE, COLLECT URINE INTO THE BD VACUTAINER® TUBES PER HOSPITAL PROTOCOL. IF USING SYRINGE, SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND TRANSFER TO COLLECTION CUP OR FOLLOW HOSPITAL PROTOCOL. 4. AFTER SAMPLE IS OBTAINED, DISCARD COLLECTION DEVICE ACCORDING TO HOSPITAL PROTOCOL. 5. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE.¿ SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05754 | 2023-08-03 | 2023-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE CATHETER WAS INSERTED INTO THE PATIENT ON (B)(6) 2023. THE CATHETER WAS FOUND TO DISASSEMBLE AUTOMATICALLY ON (B)(6) 2023, AND THE NURSE'S EXAMINATION FOUND THAT THE CATHETER WAS CAUSED BY THE RUPTURE OF THE AIR SAC. IT WAS STATED THAT THE DOCTORS RE-CATHETERIZED THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05776 | 2023-08-03 | 2023-07-06 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT'S DOCTOR BELIEVES THAT THIS PUREWICK FEMALE EXTERNAL CATHETER WAS CAUSING AN INCREASE OF URINARY TRACT INFECTION IN PATIENT. PATIENT ASKED TO STOP USING THE PUREWICK WHICH THEY HAD. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. PER FOLLOW UP VIA PHONE ON (B)(6)2023, IT WAS REPORTED THAT THE CUSTOMER SOUGHT MEDICAL INTERVENTION AND WAS ADVISED TO DISCONTINUE USE OF THE CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03337 | 2023-08-03 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD TRIALYSIS TL CATHETER | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. FLUID LEAKING FROM TRIPLE LUMEN CENTRAL LINE INSERTION SITE IN THE NECK. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03324 | 2023-08-03 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT, THE PATIENT ALLEGEDLY GOT HURT WHEN THE PORT WAS ACCESSED FOR THE FIRST TIME. IT WAS FURTHER REPORTED THAT THE PATIENT ALLEGEDLY EXPERIENCED PULLING AND DISCOMFORT IN THE CHEST. REPORTEDLY, THERE WAS NO ADDITIONAL INTERVENTION PERFORMED OR MEDICATION PRESCRIBED TO TREAT THE DISCOMFORT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03325 | 2023-08-03 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 5676300 |  | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED REDNESS AND ITCHINESS IN THE CHEST. REPORTEDLY, THERE IS NO ADDITIONAL INTERVENTION PERFORMED OR MEDICATION PRESCRIBED TO TREAT REDNESS AND ITCHINESS. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 11/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03326 | 2023-08-03 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP IMPLANTABLE PORT, GROSHONG, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | CP00003 |  | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY STRETCHED. IT WAS FURTHER REPORTED THAT THE CATHETER HAD OTHER DAMAGES TO THE PORTION OF THE CATHETER BETWEEN THE PORT STEM CONNECTION AND THE VENOMITY SITE. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05770 | 2023-08-03 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT WITH CATHETERIZATION, THE URINE DID NOT DRAIN OUT. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 21JUL2023, CLARIFIED THAT THERE WAS A BLOCKAGE THAT AFFECTED THE URINE FLOW FROM THE CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: SALT ACCUMULATION)/BLOCK DRAINAGE LUMEN/NO DRAINAGE EYE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05782 | 2023-08-03 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED DURING A PRETEST, THE CATHETER BALLOON WAS DIFFICULT TO INFLATE AND TO DEFLATE. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON 26JUL2023, THE INFORMATION OF THE BALLOON CONDITION IF INFLATED COMPLETELY OR PARTIALLY WAS NOT PROVIDED BY THE CUSTOMER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. PHOTO: THERE ARE 6 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE FIRST PHOTO ATTACHED IS AN OVERVIEW OF THE CATHETER WIT CUT PORTION OF INLET TUBING AND SYRINGE. THE SECOND PHOTO SHOW A CLOSER LOOK AT THE SAMPLE PORT CONNECTOR AND THE INFLATION CAP. THE THIRD PHOTO SHOW THE TOP OF THE INFLATION CAP. THE FOURTH PHOTO SHOW THE TIP OF THE SYRINGE. THE FIFTH PHOTO SHOW THE BALLOON PORTION AND THE TIP OF THE CATHETER AT A RESTING STATE. THE SIXTH PHOTO IS A FOREIGN DOCUMENT. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. WITH SOME RESISTANCE, THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) USING THE RETURNED SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 4 MINUTES WITH NO CUFFING OR LEAKS NOTED, RETURNING 10ML OF SOLUTION. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05771 | 2023-08-03 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT DURING A PRETEST IN AN OUTPATIENT DEPARTMENT, WHEN THE WATER WAS INJECTED INTO THE FOLEY CATHETER, THE BALLOON DID NOT INFLATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03336 | 2023-08-03 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 3F SL SOLO MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1193108D4 |  | IT WAS REPORTED BY CUSTOMER THAT "SALINE LEAKS FROM T LOCK AREA ECG AND SEPTUM AREA. WE NOTICED THAT WHEN WE WENT TO FLUSH THE PICC LINE, THE FLUSH WOULD COME OUT AROUND THE WIRE ON THE END CAP INSTEAD OF FLUSHING THE LINE. WE HAVE NOTICED THIS ON MANY PICC LINES. WE WOULD PUT OUR FINGER OVER THE END CAP TO FLUSH." ADDITIONAL INFORMATION FROM CUSTOMER RESPONSE ON (B)(6) 2023. "THIS HAPPENED ON ALMOST EVERY PICC WE USED. IT WAS NOTED BEFORE AND DURING USE. NO PT INVOLVEMENT. NO SERIOUS INJURY TO ANYONE INVOLVED. THIS WAS MORE OF A PAIN FOR THE INSERTER OF THE PICC. NO MEDICATION USED BESIDES NORMAL SALINE FLUSH WITH THE PROVIDED SYRINGES THAT COMES WITH THE PICC KIT. ALL THE PICCS WERE INSERTED AND WIRE WITH CAP DISCARDED." THE CUSTOMER REPORTED THIS OCCURRED WITH MANY PICC LINES HOWEVER THE EXACT QUANTITY WAS NOT PROVIDED TO BD FIELD ASSURANCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05791 | 2023-08-04 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129416M |  | IT WAS REPORTED THAT THE OPERATING ROOM EXPERIENCED TEMPERATURE FOLEY CATHETER HAS BEEN MALFUNCTIONING AT LEAST 2 TIMES IN THE LAST WEEK. THE TEMPERATURE SENSORS EITHER WERE NOT WORKING OR NOT PROPERLY READING THE PATIENT BLADDER TEMPERATURE. A KINK WAS NOTICED ON THE TEMPERATURE PROBE FROM PACKAGING FOR THE TWO CATHETERS THAT MALFUNCTIONED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS AVAILABLE FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THE FAILURE COULD BE "INADEQUATE STRAIN RELIEF". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
2020394-2023-00556 | 2023-08-04 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DORADO | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | DR8076 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN ILIAC, THE PTA BALLOON WAS FULLY INFLATED AND PRESSURE DROPPED TO 4 ATM. IT WAS FURTHER REPORTED THAT THE PTA BALLOON ALLEGEDLY WOULD NOT INFLATE HIGHER. IT WAS ALSO REPORTED THAT THE PTA BALLOON ALLEGEDLY HAD A PINHOLE RUPTURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 05/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03348 | 2023-08-04 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE PROXIMAL END OF THE DILATOR WAS NOT SMOOTH, MAKING IT IMPOSSIBLE TO INSERT THE SKIN. NO PATIENT INJURY REPORTED. NO OTHER INFORMATION PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 3 DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05822 | 2023-08-04 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119416M |  | IT WAS REPORTED THAT THE BALLOON IN TEMPERATURE SENSING FOLEY CATHETER LEAKED AND FELL OUT OF PATIENT. THEY TRIED TO INFLATE BALLOON AGAIN, BUT NO RESULT. THEY WERE REQUESTED TO RETURN IMMEDIATELY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05819 | 2023-08-04 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® URETHRAL CATHETER TRAY, CLEAR PLASTIC CATHETER | URETHRAL CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 |  | 802214 |  | IT WAS REPORTED THAT THE CUSTOMER OPENED THE FOLEY TRAY TO INSERT A CATHETER FOR THE PATIENT. UPON OPENING, IT WAS DISCOVERED THAT THE JELLY WAS MISSING. THE CUSTOMER REPORTED THAT THIS HAPPENED ALL THE TIME, AND THEY HAD TO KEEP SEVERAL PACKAGES WITH THEM IN CASE SOMETHING WAS MISSING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03342 | 2023-08-04 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT SOMETIMES POST DIALYSIS CATHETER PLACEMENT, THE CATHETER TUBE ALLEGEDLY BROKE. IT WAS FURTHER REPORTED THAT BLOOD WAS COMING THROUGH IT. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE SAMPLE WAS NOT RETURNED FOR EVALUATION, ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS ONE CATHETER IMPLANTED IN THE PATIENT BODY AND A FRACTURE WAS NOTED ON THE RED LUER EXTENSION LEG NEAR THE HUB. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND FLUID LEAK ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05808 | 2023-08-04 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON DID NOT INFLATE WHEN THE STERILE WATER WAS ATTEMPTED TO INJECT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05807 | 2023-08-04 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER TORN OFF WHILE TRANSFERRING PATIENT FROM WHEELCHAIR TO BED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05813 | 2023-08-04 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER CREEVY MODEL | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1267520 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
17457985 | 2023-08-04 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC SV CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 1274118 | 1274118 |  | ULTRASOUND GUIDANCE WAS USED TO IDENTIFY VEIN. DYNAMIC, REAL-TIME IMAGING OF VESSEL CANNULATION PERFORMED. STERILE GEL AND PROBE COVER USED IN ULTRASOUND-GUIDED CENTRAL VENOUS CATHETER INSERTION. CHLORAPREP ANTISEPTIC USED DURING CENTRAL VENOUS CATHETER INSERTION. SKIN PREP AGENT COMPLETELY DRIED PRIOR TO PROCEDURE. HAND HYGIENE PERFORMED PRIOR TO CENTRAL VENOUS. CAP, MASK, STERILE GOWN, STERILE GLOVES AND STERILE DRAPE USED. LEFT BASILIC VEIN WAS LOCATED USING ULTRASOUND, NEEDLE WAS PLACED WITH BLOOD RETURN, GUIDEWIRE WAS ADVANCED 5-10 CM UNTIL THERE WAS RESISTANCE AND SO ADVANCEMENT WAS STOPPED AND GUIDEWIRE WAS PULLED BACK ALONG WITH THE NEEDLE 2-3 CM UNTIL RESISTANCE WAS MET AND STOPPED. ATTEMPTED TO REMOVE NEEDLE AND GUIDEWIRE TOGETHER BUT WAS UNABLE TO DO SO. NEEDLE WAS EASILY REMOVED OVER THE GUIDEWIRE. PROCEDURE WAS ABORTED. WITH THE GUIDEWIRE REMAINING IN PLACE, PRESSURE WAS APPLIED TO THE SITE WITH MINIMAL BLEEDING. LIDOCAINE WAS ADMINISTERED AT SITE AND PATIENT WAS GIVEN IV MORPHINE FOR PAIN CONTROL. SURGEONS WERE NOTIFIED, DOCTORS ALL RESPONDED TO BEDSIDE TO EVALUATE PATIENT. X-RAY WAS COMPLETED. DECISION WAS MADE TO TAKE PATIENT TO OR FOR FOREIGN BODY REMOVAL. MOM AND PATIENT WERE UPDATED AND ALL QUESTIONS WERE ANSWERED. |  | N | Y |  |  | 2023-08
1018233-2023-05799 | 2023-08-04 | 2023-07-18 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER COULD NOT BE REMOVED. WHEN THE USER TRIED TO REMOVE THE CATHETER BY THE NORMAL METHOD BY ATTACHING A SYRINGE AND ALLOWING THE STERILE WATER TO DRAIN NATURALLY BUT WAS NOT ABLE TO REMOVE IT. THE PATIENT HAD BIT OFF THE DRAINAGE FUNNEL WHILE TRYING TO FIGURE OUT THE NEXT METHOD FOR NOT BEING ABLE TO REMOVE THE CATHETER. THE USER AGAIN ATTEMPTED TO REMOVE THE CATHETER USING THE NON-RUPTURED BALLOON METHOD, WHICH ALSO FAILED. THE CATHETER MIGHT HAVE BEEN PULLED INTO THE URETHRA WHEN THE PATIENT BIT OFF THE DRAINAGE FUNNEL, WHICH MADE IT EVEN MORE DIFFICULT TO REMOVE. THE CATHETER DID NOT BUDGE WHEN IT MOVED BACK AND FORTH. THE CATHETER WAS IN PLACE WITH THE DRAINAGE FUNNEL TORN OFF AND WAS SCHEDULED FOR CATHETER REMOVAL, BUT IF THE NEXT REMOVAL WAS NOT SUCCESSFUL, IT MIGHT BE A SURGICAL PROCEDURE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03346 | 2023-08-04 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE PROXIMAL END OF THE DILATOR WAS NOT SMOOTH, MAKING IT IMPOSSIBLE TO INSERT THE SKIN. NO PATIENT INJURY REPORTED. NO OTHER INFORMATION PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 3 DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03344 | 2023-08-04 | 2023-07-03 | I, n, j, u, r, y | Manufacturer report | HICKMAN PORT SINGLE LUMEN | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0603880C |  | IT WAS REPORTED THAT DURING THE PORT PLACEMENT PROCEDURE, THERE WAS ALLEGEDLY LIQUID OOZING FROM THE PUNCTURE POSITION. IT WAS FURTHER REPORTED THAT NO BLOOD RETURNED WHEN THE PREFILLED CATHETER FLUSHER WAS WITHDRAWN. IT WAS ALSO REPORTED THAT A CHEST X-RAY EXAMINATION SHOWED THE INFUSION PORT TUBE ON THE RIGHT SIDE OF THE PATIENT WAS REFLEXED, AND THE END WAS AT THE LEVEL OF THE LOWER EDGE OF THE CLAVICLE. FURTHERMORE, THE PATIENT ALLEGEDLY EXPERIENCED A THROMBUS THAT FORMED ON THE WALL. REPORTEDLY, THE PORT WAS REMOVED. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 12/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05820 | 2023-08-04 | 2023-07-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 50614 |  | IT WAS REPORTED THAT CUSTOMER WAS USING PUREWICK FEMALE EXTERNAL CATHETER FOR 2 YEARS. ON SATURDAY, THEY WERE USING THE CATHETER BUT THIS TIME THE CATHETER CUT THE INSIDE OF THE PATIENT URETHRA. PATIENT WENT TO URGENT CARE AND WAS PUT ON ANTIBIOTIC. A NEW CATHETER WAS PUT IN FOR THE PATIENT. CUSTOMER SAID IT LOOKED LIKE THE END OF THE CATHETER WAS NOT FORMED PROPERLY AND THE TIP CUT THE URETHRA. PER FOLLOW UP VIA PHONE ON 25JUL2023, IT WAS REPORTED THAT THE CUSTOMER WAS USING THE ANTIBIOTIC, BACTRIM, TO HELP ALLEVIATE THE URINARY TRACT INFECTION. THEY MADE IT CLEAR THAT THIS WAS AN INTERMITTENT CATHETER AND NOT A FEMALE EXTERNAL CATHETER. | THE REPORTED EVENT IS CONFIRMED - MANUFACTURING RELATED. RECEIVED ONE USED (1) MAGIC3 ISC IN OPENED PACKAGING. THE ISC WAS ANALYZED WITH GUIDELINES. VISUAL INSPECTION NOTED THE TIP OF ISC HAS EXCESS SILICONE MEASURING 0.0595". THEREFORE, PRODUCT DOES NOT MEET SPECIFICATIONS. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED, A POTENTIAL ROOT CAUSE COULD BE MANDRELS TOUCHING EACH OTHER WITHIN THE PALLET. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03347 | 2023-08-04 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE PROXIMAL END OF THE DILATOR WAS NOT SMOOTH, MAKING IT IMPOSSIBLE TO INSERT THE SKIN. NO PATIENT INJURY REPORTED. NO OTHER INFORMATION PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 3 DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05801 | 2023-08-04 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129416M |  | IT WAS REPORTED THAT THE DEFECTIVE TEMPERATURE SENSING FOLEY CATHETERS. CUSTOMER STATED THAT THE ITEM WAS NOT READING ACCURATE TEMPERATURES AND THEY HAVE ABOUT 50 EACH TO RETURN ON FROM MULTIPLE PO'S. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "SENSOR WIRE TOO WEAK." IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® EZ-LOK® SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK OR SLIP TIP SYRINGES: OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. SWAB SURFACE OF BARD® EZ-LOK® SAMPLING PORT WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND REMOVE SYRINGE FROM SAMPLE PORT. UNKINK TUBING IF NECESSARY AND TRANSFER URINE SPECIMEN INTO SPECIMEN CUP OR FOLLOW HOSPITAL PROTOCOL. DISCARD SYRINGE ACCORDING TO HOSPITAL PROTOCOL. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB. STERILE UNLESS PACKAGE IS OPENED OR DAMAGED, EXCEPT FOR ANY INDIVIDUALLY PACKAGED COMPONENTS WITHIN THE TRAY WHICH ARE NOT LABELED AS STERILE. THESE COMPONENTS ARE NOT TERMINALLY STERILIZED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05806 | 2023-08-04 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE BALLOON OF THE FOLEY CATHETER POSSIBLY DUE TO A PINHOLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03345 | 2023-08-04 | 2023-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN PORT SINGLE LUMEN | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0603880C |  | IT WAS REPORTED THAT TWO MONTHS AND NINE DAYS POST A PORT PLACEMENT, THE FLOW WAS ALLEGEDLY NOT SMOOTH UPON FLUSHING. IT WAS FURTHER REPORTED THAT THE BASE OF THE PORT AND THE CATHETER WERE ALLEGEDLY FOUND TO BE BROKEN. FURTHERMORE, THE PATIENT ALLEGEDLY EXPERIENCED SKIN SWELLING. HOWEVER, THERE IS NO ADDITIONAL INTERVENTION PERFORMED OR ADDITIONAL MEDICATION PRESCRIBED TO TREAT SKIN SWELLING. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE SAMPLE WAS NOT RETURNED FOR EVALUATION. TWO ELECTRONIC PHOTOS WERE PROVIDED FOR REVIEW. THE PHOTO SHOWS A HICKMAN PORT AND A BROKEN CATHETER SEGMENT. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND MATERIAL SEPARATION ISSUE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO FLUSH ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 07/2025), G3, H6 (DEVICE, METHOD) H11: H6 (RESULT, CONCLUSION) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05790 | 2023-08-04 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165816 |  | IT WAS REPORTED THAT THE PACKAGING OF THE FOLEY CATHETER INDICATED THAT THE BALLOON SIZE WAS A 5CC, HOWEVER THE INFLATION ARM OF THE CATHETER INDICATED A BALLOON SIZE OF 10CC (PCN# 165816). IT WAS CAUSING GREAT CONFUSION TO THE END USERS AND COULD POTENTIALLY CAUSE UNDER INFLATION OF THE BALLOON. PER MAIL RECEIVED FROM IBC ON 19JUL2023, CUSTOMER PROVIDED ADDITIONAL CATHETER CATALOG NUMBERS THAT WERE AFFECTED. THE PACKING ON THE CATHETERS 165812, 165814, 165816 WAS INDICATING A 5CC BALLOON, WHILE THE INFLATION ARMSTATED A 10CC BALLOON (PCN# 165812 AND PCN# 165816). THIS WAS CAUSING COMPLETE CONFUSION AND A LOT OF PROBLEMS WITHIN HOSPITALS. TO EXPLAIN TO ALL STAFF TO NOT GO BY THE PACKAGING BUT RATHER THAN THE INFLATION ARM, WAS NOT EASILY ACHIEVED AND IT WAS NOT SOMETHING END USERS WERE IN THE HABIT OF DOING. THEY WOULD CHECK THE PACKAGING FOR THE BALLOON SIZE AND WOULD NOT OPEN THE PACKAGING TO CHECK THE INFLATION ARM AS THIS WAS A WASTE OF STERILE PRODUCT AND NOT ALWAYS EASILY VISIBLE THROUGH THE PACKAGING. IT WAS CAUSING MASSIVE CONFUSION AND POTENTIAL PATIENT COMPLICATIONS RELATING TO INFLATION ERRORS. STATED THAT WHAT WAS ON THE PACKAGING, SHOULD BE ON THE CATHETER. PER MAIL RECEIVED FROM IBC ON 21JUL2023, STATED THAT THE CUSTOMER WILL NOT BE RETURNING THE STOCK AS THERE WAS NO ALTERNATIVE THAT THEY CAN SOURCE IN THE SHORT TERM. ALTHOUGH THERE WAS NO DEFECT WITH THE PRODUCT ITSELF, THEY WANTED TO MAKE A FORMAL COMPLAINT AS THEY WERE CONCERNED ABOUT THE POSSIBLE CONFUSION THE LABELLING WOULD CAUSE TO THE END USER. THE LABELLING ON THE PACKAGING AFFECTS A NUMBER OF SKU¿S WHICH THEY SUPPLY WITH US LABELLING. THIS WAS PICKED UP AS THEY HAD RECENTLY TRANSITIONED FROM A CE CODE TO A US CODE AS A PART OF RECODE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE MET RELEVANT SPECIFICATIONS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED 12 UNOPENED (WITHOUT ORIGINAL PACKAGING), ALL-SILICONE FOLEY CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED ALL 12 FOLEY CATHETER INDICATED THAT THE BALLOON SIZE WAS A 5CC, WITH THE INFLATION ARM OF THE CATHETER INDICATED A BALLOON SIZE OF 10ML. NO DISCREPANCIES NOTED. DHR IS NOT REQUIRED SINCE THE REPORTED FAILURE WAS UNCONFIRMED. THE INVESTIGATION IS CONCLUDED, AND NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05792 | 2023-08-04 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED DURING A PRETEST, THE CATHETER BALLOON WAS DIFFICULT TO INFLATE AND TO DEFLATE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THERE ARE 6 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE FIRST PHOTO ATTACHED IS AN OVERVIEW OF THE CATHETER WIT CUT PORTION OF INLET TUBING AND SYRINGE. THE SECOND PHOTO SHOW A CLOSER LOOK AT THE SAMPLE PORT CONNECTOR AND THE INFLATION CAP. THE THIRD PHOTO SHOW THE TOP OF THE INFLATION CAP. THE FOURTH PHOTO SHOW THE TIP OF THE SYRINGE. THE FIFTH PHOTO SHOW THE BALLOON PORTION AND THE TIP OF THE CATHETER AT A RESTING STATE. THE SIXTH PHOTO IS A FOREIGN DOCUMENT. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. WITH SOME RESISTANCE, THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) USING THE RETURNED SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 4 MINUTES WITH NO CUFFING OR LEAKS NOTED, RETURNING 10ML OF SOLUTION. THIS IS WITHIN SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05821 | 2023-08-04 | 2023-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT A FEW HOURS AFTER INSERTION, NO URINE WAS FLOWING AND UPON CHECKING THE CATHETER, THE BALLOON HAD RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05802 | 2023-08-04 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED DURING A PRETEST, THE CATHETER BALLOON WAS DIFFICULT TO INFLATE. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RICK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE PUNCH DEPTH TOO LARGE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A SYRINGE. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) USING THE RETURNED SYRINGE AND THE SOLUTION IMMEDIATELY LEAKED INTO THE DRAINAGE LUMEN INSIDE OF THE BALLOON. THE BALLOON WAS DISSECTED AND FOUND THAT THE NOTCH IS DOUBLE PERFORATED. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER.¿ CORRECTIONS: D,F,H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05784 | 2023-08-04 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT URINE LEAKED FROM THE CUFF AREA. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 24JUL2023, CLARIFIED THAT THERE WAS A CRACK OBSERVED NEAR THE FUNNEL. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT CAUSED THE REPORTED FAILURE. IT IS UNKNOWN IF THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING AND SYRINGE. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND SOLUTION IMMEDIATELY LEADED FROM A .014'' PINHOLE LOCATED 0.4525" FROM BIFURCATION. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE IP7603444, REVISION 0, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿PROPER TECHNIQUES FOR URINARY CATHETER INSERTION. PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION. INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE. SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT. ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE. LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED. FOLEY CATHETER REMOVAL. 1. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE 2. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY 3. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION 4. IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL 5. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03341 | 2023-08-04 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 9808560 |  | IT WAS REPORTED THAT ONE WEEK AND SIX DAYS POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03340 | 2023-08-04 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT TITANIUM ISP 6CF INT WSP ATT SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8706060 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CONNECTING NEEDLE TUBE ALLEGEDLY LEAKED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE SAMPLE WAS NOT RETURNED FOR EVALUATION. ONE ELECTRONIC VIDEO WAS PROVIDED FOR REVIEW. IN THE PROVIDED VIDEO, THE FLUID WAS NOTED TO BE LEAKED FROM THE CATHETER WHEN THE PHYSICIAN TRYING TO INFUSE THE PORT WITH AN INFUSION SET. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FLUID LEAK. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 05/2024), G3, H6 (METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03349 | 2023-08-06 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1618070 |  | IT WAS REPORTED THAT APPROXIMATELY ONE MONTH POST PORT PLACEMENT, THE STERI-STRIPS HAD NOT YET FALLEN OFF, AND THE PATIENT ALLEGEDLY HAD REDNESS AROUND THE PORT. IT WAS FURTHER REPORTED THAT THE DEVICE HAD BEEN ACCESSED TWICE WITHOUT DIFFICULTY IN FLUSHING OR GETTING A BLOOD RETURN. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 08/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05829 | 2023-08-07 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FLOOR WAS FOUND BEING WET 2 DAYS AFTER THE DEVICE PLACEMENT SINCE THE FOLEY CATHETER FELL OUT OF THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05846 | 2023-08-07 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT AFTER INSERTING THE CATHETER INTO THE PATIENT, CATHETER PLACEMENT INTO THE BLADDER WAS CONFIRMED, BUT NO OUTFLOW OF URINE WAS OBSERVED. THEN THE PATIENT WAS REPLACED WITH A CATHETER FROM A NEW TRAY AND THE URINE FLOWED. PER FOLLOW UP INFORMATION RECEIVED VIA (B)(6) ON 24JUL2023, THE DETAILS WEREN'T PROVIDED FROM CUSTOMER IF THERE WAS A BLOCKAGE IN THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03355 | 2023-08-07 | 2023-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT TI ISP, 8 FR CHRONOFLEX, SUTURE PLUG, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1708060 |  | IT WAS REPORTED THAT FIVE MONTHS AND SIXTEEN DAYS POST PORT PLACEMENT IN UPPER RIGHT CHEST, THE DEVICE ALLEGEDLY FAILED TO GET BLOOD FROM THE PORT. THE PROCEDURE WAS COMPLETED USING HAND AGAIN TO DRAW BLOOD AND HAD TO STICK PATIENT FIVE TIMES IN THE HAND AND FIVE TIMES IN ARM. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 07/2024), G3. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03352 | 2023-08-07 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA THE RIGHT INTERNAL JUGULAR VEIN, THE GUIDEWIRE ALLEGEDLY COULD NOT BE INSERTED INTO THE BODY, AND WHILE RETRACTING, THE GUIDEWIRE BECAME ENTANGLED WITH THE INTRODUCER NEEDLE. IT WAS FURTHER REPORTED THAT THE GUIDEWIRE WAS ALLEGEDLY DIFFICULT TO REMOVE AND WAS STRETCHED UPON WITHDRAWAL. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE SAMPLE WAS NOT RETURNED FOR EVALUATION. ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED STRETCHED AND IDENTIFIED FRACTURE AND MATERIAL SEPARATION ISSUES AS THE GUIDEWIRE WAS NOTED TO BE UNCOILED AND UNRAVELLED AND BOTH THE ROUND AND FLAT CORE WIRES WERE NOTED TO BE PROTRUDING. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO INSERT, PHYSICAL RESISTANCE AND DIFFICULT TO REMOVE ISSUE AS THERE WAS NO OBJECTIVE EVIDENCE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05845 | 2023-08-07 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THERE WAS A CRACK FOUND IN THE FOLEY CATHETER. AFTER CATHETER PLACEMENT, URINE WAS OBSERVED LEAKING FROM THE SHAFT. IT WAS STATED THAT UPON INSPECTION OF THE PRODUCT, THERE WAS A VISIBLE CRACK AT THE CONNECTION BETWEEN THE SHAFT AND FUNNEL. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION". THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿BARD® SILVER LUBRI-SIL® FOLEY TRAY SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED . <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿ PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) NON-CLINICAL STUDIES HAVE SHOWN THAT TEMPERATURE SENSING CATHETER IS MR CONDITIONAL. MR EXAMINATIONS CAN BE SAFELY PERFORMED ON PATIENTS WEARING THIS PRODUCT UNDER THE FOLLOWING CONDITIONS (SELF CERTIFICATION). - STATIC MAGNETIC FIELD STRENGTH: 1.5T,3.0T - GRADIENT OF STATIC MAGNETIC FIELD: 4,000GAUSS/CM OR LESS -WHOLE-BODY MAXIMUM SAR (SPECIFIC ABSORPTION RATE) (MODE) INDICATED BY THE MR SYSTEM:2-W/KG (NORMAL OPERATING MODE) THE MAXIMUM TEMPERATURE RISE THAT CAN OCCUR IN THIS PRODUCT DURING A SCAN TIME OF 15 MINUTES OR LESS PER PULSE SEQUENCE UNDER THE ABOVE CONDITIONS IS 1.7 °C OR LESS. THE ARTIFACT THAT CAN OCCUR WHEN THE PRODUCT IS IMAGED USING THE GRADIENT MAGNETIC FIELD ECHO METHOD IN A 3.0T MR SYSTEM IS 10MM FROM THE ACTUAL IMAGE OF THE PRODUCT. DO NOT PERFORM SCAN LONGER THAN 60 MINUTES UNDER THE ABOVE CONDITIONS. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL (I.E. THE REMOVABLE EXTENSION CABLE, ETC.) FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE PRIOR TO THE MR EXAMINATION. 4) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING INSULATING MATERIAL SUCH AS FOAM RUBBER PADDING BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
MW5120710 | 2023-08-07 | 2023-07-16 | I, n, j, u, r, y | Voluntary report | BARD POWERPORT SLIM IMPLANTABLE PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C. R. BARD, INC. /BARD ACCESS SYSTEMS, INC. |  | 1716000 |  | BARD POWERPORT ADVERSE EVENT: (B)(6) 2023 HAD BARD POWERPORT SURGICALLY PLACED FOR CHEMOTHERAPY INFUSIONS. AREA SEEMED AVERAGE AND EXPECTED FOR HEALING. PORT WAS ACCESSED (B)(6) 2023 FOR CHEMOTHERAPY INFUSION #1: NOTHING UNUSUAL NOTED. PORT WAS ACCESSED (B)(6) 2023 FOR LABS. PORT WAS ACCESSED (B)(6) 2023 FOR CHEMOTHERAPY INFUSION #2. (B)(6) 2023 - THE INSERTION/INCISION SITE WHERE PORT WAS PLACED UNDER THE SKIN WAS TENDER, BREAKING OPEN, AND DISCHARGING PUS. I CONTACTED MY CARE PROVIDER AND A BLOOD DRAW WAS TAKEN ON (B)(6) 2023 ACCESSING THE PORT. (B)(6) 2023 PORT WAS ACCESSED FOR LABS. (B)(6) 2023 PORT WAS ACCESSED FOR CHEMOTHERAPY INFUSION - THERE WAS INFLAMMATION AND REDNESS NOTED SEVERAL HOURS AFTER THE CHEMOTHERAPY INFUSION WAS COMPLETE, BUT INITIALLY I THOUGHT THIS TO JUST BE RELATED TO CHEMOTHERAPY SIDE EFFECTS. OVER THE NEXT 48 HOURS THERE WAS RAPID CHANGE AND WORSENING SYMPTOMS, ALL INDICATIVE TO A SEVERE INFECTION : FEVER, RED STREAKS, EXTREME TENDERNESS AND REDNESS, AND HEAD TO TOE PAIN, I COULD BARELY MOVE THE ARM ON THE AFFECTED SIDE, SEVERE WEAKNESS. (B)(6) 2023 I REPORTED TO THE CLINIC AND WITH ONLY A VISUAL OBSERVATION OF MY PHYSICAL STATE, IT WAS DETERMINED TO REMOVE THE PORT ASAP, AND I WAS STARTED ON A HIGH POWERED BROAD SPECTRUM ANTIBIOTIC TO BEGIN FIGHTING INFECTION. (B)(6) 2023 EMERGENCY HOSPITALIZATION SURGICAL PROCEDURE FOR PORT REMOVAL. (B)(6) 2023 RECOVERY CONTINUES AT HOME WITH CONTINUED HIGH LEVEL OF DISCOMFORT, SKIN REMAINS VERY RED AT THE LEFT CHEST AREA (RADIATING DOWN TO 1/3-1/2 THE BREAST) WHERE PORT WAS REMOVED AND THERE IS RESTRICTED LEFT SIDE ARM MOVEMENT. I AM CURRENTLY PERFORMING WOUND CHECKS AND DRESSING CHANGES AS PRESCRIBED IN MY POST-SURGICAL PAPERWORK AND TAKING A BROAD SPECTRUM ORAL ANTIBIOTIC FOR A 10DAY COURSE OF TREATMENT. I HAVE CONTACTED MY ONCOLOGY CLINIC, THE HOSPITAL, AND THE SURGEON TO MAKE THEM AWARE THAT I BELIEVE THIS NEEDS TO BE REPORTED AS AN ADVERSE EVENT TO THE PORT IMPLANT I HAD. I AM ALSO SELF-REPORTING THE ADVERSE EVENT AS A CONSUMER. |  | Y | Y |  |  | 2023-08
3006260740-2023-03351 | 2023-08-07 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY BROKEN INTO TWO PARTS WHILE PASSING THROUGH THE TUNNEL. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE 23CM GLIDEPATH D/L CATHETER WAS RECEIVED FOR EVALUATION. GROSS VISUAL, TACTILE, MICROSCOPIC AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED APPROXIMATELY 4.9CM FROM THE DISTAL END OF THE Y-BODY. THE EDGES OF THE DISTAL END OF THE CIRCUMFERENTIAL BREAK WERE NOTED TO BE ROUGH AND UNEVEN AND THE SURFACE WAS OBSERVED TO BE GRANULAR IN NATURE. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS ONE GLIDEPATH CATHETER AND A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED NEAR THE CUFF AREA. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND MATERIAL SEPARATION ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
17469000 | 2023-08-07 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDEX LUBRI-SIL ALL-SILICONE TEMP SENSING FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | 119216M | 119216M |  | BALLOON WILL NOT STAY INFLATED. CAUSING IT TO COME OUT OF PATIENT. |  | N | Y |  |  | 2023-08
1018233-2023-05843 | 2023-08-07 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OFF AND THE BALLOON WAS DAMAGED. AFTER SPINAL SURGERY, WHEN THE PATIENT WAS PLACED IN THE SUPINE POSITION AND TRIED TO FIX THE INDWELLING BLADDER CATHETER WITH KINOSOFT, THE CATHETER FELL OUT OF THE PATIENT WITHOUT RESISTANCE. IT WAS CONFIRMED THAT THE BALLOON PART OF THE CATHETER WAS DAMAGED. THIS EVENT WAS REPORTED TO THE DOCTOR. SINCE THERE WAS NO HEMATURIA OR CHIPPING (DAMAGE) OF THE BALLOON, THE CATHETER WAS REPLACED. THE TIME FROM THE START OF ANESTHESIA TO THE END OF SURGERY WAS ABOUT 2 HOURS AND 30 MINUTES, AND THE AMOUNT OF URINE WAS ABOUT 100 ML. AFTER REPLACING WITH A NEW INDWELLING BLADDER CATHETER, THE URINE VOLUME WAS ABOUT 300 ML. THIS WAS THE FIRST REPORT BECAUSE THERE WAS A CHANGE IN THE NURSE IN CHARGE OF THE OPERATION. A PREOPERATIVE REPORT WAS MADE TO THE PREOPERATIVE NURSE. THE EVENT OCCURRED THREE HOURS AFTER THE PLACEMENT. THE WATER LEAKED FROM THE BALLOON. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. EVALUATION CONFIRMED THE BALLOON WAS BURST. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE LOW LATEX VISCOSITY, SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05831 | 2023-08-07 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER TIP WAS DIRTY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03354 | 2023-08-07 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 5733690 |  | IT WAS REPORTED THAT TWO MONTHS AND FOUR DAYS POST DIALYSIS CATHETER PLACEMENT, THE CUFFS APPEARED PAPERY WITH NO FIBROSIS. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY CAME OUT FROM THE PATIENT SPONTANEOUSLY. REPORTEDLY, THE CATHETER WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 19CM HEMOSPLIT D/L CATHETER WAS RETURNED FOR EVALUATION. MICROSCOPIC VISUAL AND FUNCTIONAL TESTING WERE PERFORMED ON THE RETURNED DEVICE. ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS A HEMOSPLIT DIALYSIS CATHETER. NO OTHER ANOMALIES WERE OBSERVED. DURING VISUAL EVALUATION, THE TISSUE CUFF WAS NOTED TO BE PAPERY AND MATTED. DURING FUNCTIONAL TESTING, LURE CAPS WERE REMOVED, AND CLAMP DISENGAGED. BOTH THE RED AND BLUE LUERS WERE PATENT TO INFUSION AND ASPIRATION WITHOUT ISSUE. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED CATHETER EXPULSION AND LOOSENING ISSUES AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 05/2024), G3 H11: H6 (METHOD, RESULT, CONCLUSION) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05836 | 2023-08-07 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT IMMEDIATELY AFTER PLACEMENT, THE CATHETER FELL OUT OF THE PATIENT. AS PER THE NOTIFICATION RECEIVED ON 26JUL2023, THERE WAS A BALLOON RUPTURE OBSERVED ON THE RETURNED SAMPLE. | THE REPORTED EVENT WAS CONFIRMED AND CAUSE WAS UNKNOWN. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. VISUAL INSPECTION NOTED IRREGULAR BURST WITHOUT MISSING PIECES. INTRODUCED WATER INTO THE INFLATION LUMEN AND DID NOT EXPERIENCE ANY OBSTRUCTIONS TO IMPEDE FLOW. MICROSCOPIC EXAMINATION FOUND NO CONDITIONS THAT COULD BE ASSOCIATED WITH THE REPORTED EVENT. THE EXACT CAUSE ON HOW AND WHEN PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES]: BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 8 TO 24 EVERY 2FR. FOLEY CATHETER: ACCESSORIES: CLOSED DRAINAGE BAG; (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.); SYRINGE PREFILLED WITH STERILE WATER; WATER SOLUBLE LUBRICANT; ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION; TWEEZERS; GAUZE PADS; WATERPROOF SHEET; COTTON BALLS; GLOVES; STATLOCK® FOLEY. [INTENDED USE & EFFECT- EFFICACY]: THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE]: METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT: OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING: CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. PRECAUTIONS FOR USE: WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING: REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING: DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS]: PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS): EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.); BALLOON-RUPTURE METHOD: WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. NON-RUPTURE METHOD: ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. BALLOON-RUPTURE METHOD: INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. MALFUNCTION AND ADVERSE EVENTS: MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE; DIFFICULTY OR FAILURE TO REMOVE THE DEVICE; OCCLUSION OF CATHETER INNER LUMENS; ENCRUSTATION; ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE; DEVICE DAMAGE DUE TO INAPPROPRIATE USE. ADVERSE EVENTS: URINARY-TRACT INFECTION; HEMORRHAGE, HEMATURIA; ALLERGY REACTION TO THE DEVICE; CALCULUS FORMATION; EDEMA; PAIN; DISCOMFORT; INJURY OF BLADDER OR URETHRAL; URETHRITIS, URINARY INCONTINENCE; RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]: STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05837 | 2023-08-07 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT DURING PRETEST, THE BALLOON DID NOT INFLATE EVENLY AFTER THE INJECTION OF STERILE WATER DURING THE LEAK CHECK. UPON DRAINING THE STERILE WATER, A LARGE CUFF ROLL WAS FORMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05838 | 2023-08-07 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP¿ FOLEY TRAY SYSTEM WITH BARD® LUBRI-SIL® TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | A119216M |  | IT WAS REPORTED THAT THE WATER SYRINGE IN FOLEY TRAY HAD MISSING CAP AND THERE WAS NO WATER IN THE SYRINGE. PER FOLLOW UP VIA EMAIL RECEIVED ON 27JUL2023, CUSTOMER STATED THERE WAS NO HARM OR INJURY. THE EMPTY SYRINGE WAS DOING UPON OPENING A PREPPING THE KIT. ONCE DISCOVERED, THE NURSE DISCARDED THE KIT AND OPENED A NEW ONE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03353 | 2023-08-07 | 2023-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT, THE PORT ALLEGEDLY OCCLUDED BECAUSE THERE WAS RESIDUE IN THE PORT BODY. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
17469673 | 2023-08-07 | 2023-05-07 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | NOTED LEAKING AT THE RUBBER-LIKE STOPPER WITH FLUSH. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05851 | 2023-08-07 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE FOLEY CATHETER DURING A PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05849 | 2023-08-07 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE STERILE WATER IN THE BALLOON HAD BEEN DRAINED OVER TIME, ALTHOUGH NOT IMMEDIATELY AFTER THE WATER WAS INJECTED. IT WAS NOT KNOWN WHERE THE WATER WAS BEING DRAINED FROM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05856 | 2023-08-07 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS BLOCKED. THERE WAS NO URINE DISCHARGE IMMEDIATELY AFTER PLACEMENT IN THE PATIENT. THE USER FELT RESISTANCE WHEN INJECTING STERILE WATER INTO THE CATHETER. ALSO STATED THAT A LITTLE WHITE SUBSTANCE CAME OUT FROM THE CATHETER TIP WHEN STERILIZED WATER WAS INJECTED. PER MAIL RECEIVED FROM INVESTIGATOR ON 21AUG2023, STATED THAT THE WHITE FOREIGN MATERIAL WAS SALT ACCUMULATION WHICH HAD CAUSED THE BLOCKAGE IN THE CATHETER. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND TO BE A CASCADING FAILURE OF AN EVENT PREVIOUSLY REPORTED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03357 | 2023-08-07 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT APPROXIMATELY SIX MONTHS POST DIALYSIS CATHETER PLACEMENT PROCEDURE, CATHETER ALLEGEDLY BROKE AT CLAMPING SITE. REPORTEDLY, THE CATHETER WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION, ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTOS SHOWS A GLIDEPATH CATHETER IN WHICH THE BLUE LUER EXTENSION LEG WAS NOTED TO BE BROKEN IN THE CLAMPING SITE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 01/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05870 | 2023-08-07 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT WHEN WATER WAS INJECTED WITH THE FOLEY SYRINGE PROVIDED, IT FLOWED BACK AND LEAKED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
2020394-2023-00568 | 2023-08-07 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER IQ ULTRASONIC | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | XCTO146 |  | IT WAS REPORTED THAT DURING THE PREPARATION OF A RECANALIZATION PROCEDURE, THE TIP OF THE CATHETER WAS ALLEGEDLY FRAYED. IT WAS FURTHER REPORTED THAT CATHETER ALLEGEDLY BECAME HOT AND THE SALINE SPRAY OUT EVERYWHERE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 08/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05841 | 2023-08-07 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THERE WAS HOLE IN FOLEY CATHETER BALLOON AND WAS UNABLE TO HOLD THE STERILE FLUID. THE PRODUCT WAS NOT USABLE. AS PER FOLLOW-UP INFORMATION RECEIVED FROM CUSTOMER VIA EMAIL ON 27JUL2023, STATED THAT THE ONLY IMPACT TO THE PATIENT WAS HAVING ANOTHER FOLEY CATHETER PLACED IN, AS IT IS THEIR HOSPITAL POLICY NOT TO CHECK BALLOON INFLATION PRIOR TO INSERTION. THEY DO NOT HAVE ANY OTHER INFORMATION REGARDING THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. SWAB SURFACE OF BARD® EZ-LOK® SAMPLING PORT WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT (AT APPROXIMATELY 80 - 100 DEGREE ANGLE). PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND REMOVE SYRINGE FROM SAMPLE PORT. UNKINK TUBING AND TRANSFER URINE SPECIMEN INTO SPECIMEN CUP OR FOLLOW HOSPITAL PROTOCOL. DISCARD SYRINGE ACCORDING TO HOSPITAL PROTOCOL. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB. STERILE: CONTENTS OF INNER WRAP ARE STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RE-STERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05867 | 2023-08-07 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER WAS UNABLE TO REMOVE. DURING A PRETEST BEFORE CATHETER INSERTION, AFTER STERILE WATER INJECTION INTO THE BALLOON, WHEN ATTEMPTING TO REMOVE THE WATER, THE WATER DID NOT DRAIN. BALLOON WATER WAS OUT WHEN THEY RECEIVED IT. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. PHOTO: THERE ARE 2 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE FIRST PHOTO ATTACHED IS AN OVERVIEW OF THE CATHETER WIT CUT PORTION OF INLET TUBING AND SYRINGE. THE SECOND PHOTO SHOW A CLOSER LOOK AT THE CATHETER BALLOON AND THE TIP. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING AND SYRINGE. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) USING THE RETURNED SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 2 MINUTES WITH NO CUFFING OR LEAKS NOTED, RETURNING 10ML OF SOLUTION. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. AS THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-08
2020394-2023-00567 | 2023-08-07 | 2023-07-12 | I, n, j, u, r, y | Manufacturer report | CROSSER IQ | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | XCTO146 |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE IN THE TIBIOPERONEAL TRUNK VIA PERONEAL ARTERY ACCESS, THE CATHETER WAS ALLEGEDLY FOUND TO BE LODGED. IT WAS FURTHER REPORTED THAT THE TIP OF THE CATHETER ALLEGEDLY GOT FRAYED AND A PART GOT STUCK IN THE PATIENT'S VESSEL WALL. REPORTEDLY, AFTER TRYING TO RETRIEVE THE CATHETER NUMEROUS TIMES OUT OF THE PERONEAL, THE PHYSICIAN WAS ABLE TO GET THE SUPPORT CATHETER TO ADVANCE OVER THE TIP OF THE CATHETER, AND IT WAS REMOVED FROM THE BODY WITH EXTREME DIFFICULTY. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 11/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05858 | 2023-08-07 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE PATIENT HAD BEEN BACK IN TOUCH TO STATE THEIR NEW REPLACEMENT ORDER WAS EXACTLY THE SAME AS D236914S-ALL 8" (LOT# NGHN4359) INSTEAD OF THE USUAL 11-AND ALL HAD THE SAME LOT NUMBER AS STOCK WE HAD ISSUES WITH LAST MONTH (LOT# NGHN4359). IT WAS STATED THAT THEY HAD CHECKED THE STOCK ON SHELVES AND THAT IT WAS ALL THE SAME ISSUE AND LOT NUMBER. IT WAS ALSO STATED THAT THEY HAD CHECKED THE NEW ORDER OF 12CH D236912S THEY HAVE IN STOCK, AND THEY WERE ALSO 8" INSTEAD OF THE USUAL 11 (LOT# UNK)". PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 02AUG2023, IT WAS STATED THAT THE ALL OF THE SAME LOT. NUMBER SO 12 CATHETERS IN TOTAL. THE PATIENT HAD THE SAME ISSUE IN JUNE WITH ALL 12 CATHETERS BEING AFFECTED ALSO. IN TOTAL ALL 24 CATHETERS HAD BEEN THE SAME LOT NUMBER. IT WAS ALSO STATED THAT THE OUR STOCK OF 12CH HAVE THE SAME ISSUE ALSO AND NOT SURE IF THE LOT NUMBER WAS THE SAME ON THOSE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05869 | 2023-08-07 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L22 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE FOLEY CATHETER. IT WAS STATED THAT DURING THE CATHETER REPLACEMENT AT A HOME CARE VISIT, DISTILLED WATER WAS INJECTED TO CHECK THE NEW CATHETER, AND THE INFLATED BALLOON BECAME SMALLER. IT WAS NOTED THAT UPON CHECKING IT, THE CATHETER TIP WAS DAMAGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05910 | 2023-08-08 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT THERE WAS NO LEAKAGE DURING THE CUFF CHECK, BUT WHEN THE BALLOON WAS PULLED AFTER BEING INSERTED INTO THE PATIENT AND INJECTED WITH STERILIZED WATER, IT FELL OUT. AND A CRACK WAS CONFIRMED WHEN THE CUFF WAS CHECKED. THE EVENT OCCURRED JUST AFTER PLACEMENT AND THE BALLOON WAS RUPTURED. HAND (GLOVE) POSSIBLY CONTACTED TO THE BALLOON AND THE PATIENT'S URINE STONE WAS NONE. THE CATHETER WAS PULLED RARELY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03373 | 2023-08-08 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV SOLO 4F DL MAX | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S9294108D |  | IT WAS REPORTED THAT THE STIFFENING WIRE IN THE BARD PICC BENDS AND CANNOT BE ADVANCED ONCE THE STIFFENING WIRE IS ALTERED. HAD TO USE 3 PICC LINE KITS ON THIS PATIENT. STIFFENING WIRE INSIDE PICC LINE BENT AT A 90 DEGREE APPROXIMATELY 15 CM FROM HUB. UNABLE TO FULLY THREAD PICC LINE AFTER WIRE BENT. HAD TO TRIM THE PICC LINE TO WHERE THE STIFFENING WIRE BENT TO USE PICC LINE. IT WAS REPORTED THAT THREE DIFFERENT PICC KITS WITH THREE DIFFERENT MATERIAL NUMBERS WERE USED. THIS REPORT ADDRESSES THE THIRD KIT AND MATERIAL NUMBER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05900 | 2023-08-08 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | BARDEX® LUBRICATH® FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS BLOCKED. THERE WAS NO URINE DISCHARGE IMMEDIATELY AFTER PLACEMENT IN THE PATIENT. THE USER FELT RESISTANCE WHEN INJECTING STERILE WATER INTO THE CATHETER. ALSO STATED THAT A LITTLE WHITE SUBSTANCE CAME OUT FROM THE CATHETER TIP WHEN STERILIZED WATER WAS INJECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03371 | 2023-08-08 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV SOLO 4F DL MAX | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S3294108D |  | IT WAS REPORTED THAT THE STIFFENING WIRE IN THE BARD PICC BENDS AND CANNOT BE ADVANCED ONCE THE STIFFENING WIRE IS ALTERED. HAD TO USE 3 PICC LINE KITS ON THIS PATIENT. STIFFENING WIRE INSIDE PICC LINE BENT AT A 90 DEGREE APPROXIMATELY 15 CM FROM HUB. UNABLE TO FULLY THREAD PICC LINE AFTER WIRE BENT. HAD TO TRIM THE PICC LINE TO WHERE THE STIFFENING WIRE BENT TO USE PICC LINE. IT WAS REPORTED THAT THREE DIFFERENT PICC KITS WITH THREE DIFFERENT MATERIAL NUMBERS WERE USED. THIS REPORT ADDRESSES THE FIRST KIT AND MATERIAL NUMBER USED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05899 | 2023-08-08 | 2023-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT CAME TO THE CLINIC FOR PROSTATE RESECTION DUE TO URINARY RETENTION. WHEN THE NURSE INSPECTED THE WARD, THEY FOUND THAT THE BALLOON END OF THE FOLEY CATHETER HAD BURST AND THE CATHETER WAS REPLACED IMMEDIATELY WITHOUT CAUSING ADVERSE CONSEQUENCES. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03367 | 2023-08-08 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED, "CALLER STATES SHE HAD A PICC PLACED FOR CANCER TREATMENTS IN 2022, STATES THAT HER PICC WAS PULLED OUT ABOUT AN INCH AND THE CLINICIANS WERE UNSURE HOW TO PROCEED. CALLER STATES THEY LEFT IT LIKE THAT BUT SHOULD THEY PUSH IT BACK IN? SHOULD THEY PERFORM AN X-RAY? RESPONSE: EXPLAINED THAT BD DOES NOT HAVE A RECOMMENDATION FOR PUSHING A CATHETER BACK INTO THE ARM OR VESSEL IF IT COMES OUT. RECOMMENDED SHE TALK TO THE MD ABOUT AN X-RAY TO DETERMINE PICC TIP PLACEMENT EVEN THOUGH IT SEEMS TO BE WORKING FINE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05897 | 2023-08-08 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 998404 |  | IT WAS REPORTED THAT THE CONNECTION TO THE EAGLE INFLATION DEVICE WAS CRACKED, AND THE BALLOON COULD NOT BE EXPANDED. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED THE SAMPLE WAS RECEIVED IN A CLOSED ZIPLOC BAG IDENTIFIED WITH A LABEL. INSIDE OF THE BAG THERE WAS THE TYVEK POUCH WITH UNIT LABEL IDENTIFYING PN 998404 AND LOT NUMBER BMFPFM08. INSIDE OF THE TYVEK POUCH THERE WAS THE BALLOON CATHETER. THE BALLOON HUB OF THE CATHETER READS "4MM X 4CM" AND "30 ATM"; THE WIRE HUB READS "0.038(0.97MM) WIRE". VISUAL INSPECTION WAS PERFORMED TO THE CATHETER, NO DAMAGED WERE OBSERVER IN THE BALLOON CATHETER, HOWEVER IT WAS IDENTIFIED A CRACK IN THE BALLOON HUB CONNECTOR. THE BALLOON HUB WAS INSPECTED WITH MAGNIFICATION, AND IT WAS IDENTIFIED THAT THERE ARE STRESS MARKS IN THE TOP OF THE HUB CONNECTOR. FUNCTIONAL EVALUATION NOTED A 0.038" GUIDEWIRE WAS INTRODUCED THROUGH THE WIRE HUB AND IT WAS ABLE TO PASS FREELY THROUGH ALL CATHETER INNER LUMEN. THEN THE CATHETER WAS TRIED TO BE INFLATED WITH DISTILLED WATER USING AN INFLATION DEVICE, HOWEVER IT WAS NOT POSSIBLE TO INFLATE DUE TO A LEAK IN THE BALLOON HUB. THE BALLOON HUB IS COMPLETE ONLY WITH A CRACK STARTING ON THE TOP OF THE CONNECTOR AND ENDING CLOSE TO THE WINGS AND THERE IS NO EVIDENCE THAT IT HAS BEEN DEFORMED BY A CRUSH, THEREFORE IS LIKELY THAT THE CONNECTOR HAD BEEN OVER TORQUED. MANUFACTURING PROCESS WAS REVIEWED , AND IT WAS CONFIRMED THAT ONLY PLASTIC CONNECTORS ARE USED TO CONNECTOR BALLOON HUB FOR LEAK TEST AND FOLDING. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE OVER TORQUE. THE DHR REVIEW DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: CONTRAINDICATION: NONE. WARNINGS: ¿ USE ONLY LIQUID INFLATION MEDIA. DO NOT INFLATE WITH AIR. ¿ ALWAYS FOLLOW THE MANUFACTURER¿S DIRECTIONS ACCOMPANYING THE BALLOON DILATION CATHETER FOR INSTRUCTIONS FOR USE, MAXIMUM BALLOON INFLATION PRESSURE, PRECAUTIONS, AND WARNINGS FOR THAT DEVICE. ¿ THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE REGULATIONS. PRECAUTION: ¿ BEFORE USE, INSPECT THE DEVICE TO VERIFY THAT NO DAMAGE HAS OCCURRED DURING SHIPPING AND HANDLING. CAUTION: FEDERAL (USA) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. INSTRUCTIONS FOR USE: PREPARATION: MAKE ALL ASPIRATION AND INJECTION MANEUVERS WITH THE LOCK LEVER PUSHED LEFT, I.E., UNLOCKED. UNLOCK THE PISTON BY PUSHING THE LOCK LEVER LEFT. IN THIS POSITION, YOU CAN FREELY PULL THE PISTON BACK FOR ASPIRATION, OR PUSH IT FORWARD FOR INJECTION. TO LOCK THE PISTON IN POSITION, SLIDE THE LEVER RIGHT TO THE STRAIGHT UP POSITION. 1. PREPARE A SOLUTION OF CONTRAST MEDIUM AND NORMAL SALINE IN A SMALL STERILE BOWL OR IN THE WELL PROVIDED WITH THE PACKAGE. CHECK BALLOON CATHETER AND CONTRAST MEDIUM INSTRUCTIONS FOR SPECIFIC CONTRAST MIXTURE RECOMMENDATIONS. 2. ORIENT THE TUBING DOWNWARD INTO THE CONTRAST MEDIUM. 3. PUSH THE RELEASE LEVER LEFT AND ASPIRATE ENOUGH SOLUTION TO FILL THE SYRINGE. 4. HOLD THE DEVICE UPRIGHT TO PURGE THE AIR FROM THE SYRINGE AND CONNECTING TUBE. TAP THE SYRINGE LIGHTLY, IF NECESSARY, TO REMOVE ALL THE AIR BUBBLES AND FILL THE CONNECTING TUBE COMPLETELY. 5. INSPECT THE SYRINGE AND TUBING TO ENSURE THAT THE DEVICE HAS BEEN COMPLETELY PURGED. 6. ADJUST THE SYRINGE VOLUME TO 3 TO 4CC. IF MORE CONTRAST SOLUTION IS NEEDED, SUBMERGE THE SYRINGE TIP INTO THE BASIN OF SOLUTION AND ASPIRATE. ATTACHING THE INFLATION DEVICE TO THE BALLOON DILATION CATHETER: 1. PREPARE AND TEST THE BALLOON DILATION CATHETER ACCORDING TO THE MANUFACTURER¿S DIRECTIONS FOR USE. 2. IF A SEPARATE SYRINGE WAS USED TO PREPARE THE BALLOON CATHETER, REMOVE IT. CREATE A FLUID-FLUID CONNECTION BETWEEN THE BALLOON AND THE CONNECTING TUBE (MALE ROTATING ADAPTER) OF THE INFLATION DEVICE BY PLACING A DROP OF CONTRAST SOLUTION FROM THE SYRINGE INTO EACH HUB. 3. HAND TIGHTEN THE HUBS SECURELY. BALLOON INFLATION AND DEFLATION: 1. RELEASE THE LOCK LEVER AND ALLOW THE PISTON TO MOVE FORWARD INTO NEUTRAL POSITION (0 ATM). 2. TO INFLATE THE BALLOON, ENGAGE THE LOCK LEVER, TURN THE PALM GRIP ON THE PISTON CLOCKWISE SLOWLY UNTIL THE DESIRED INFLATION PRESSURE IS REACHED. THE LOCK LEVER MAINTAINS THE INCREASING PRESSURE. 3. TO DEFLATE THE BALLOON, PUSH THE LOCK LEVER LEFT, RELEASING THE PISTON, AND PULL BACK. SLIDE THE LOCK LEVER BACK TO LOCK, IF DESIRED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05894 | 2023-08-08 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE STERILE WATER COULD NOT BE DRAINED FROM THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05902 | 2023-08-08 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE PATIENT HAD BEEN BACK IN TOUCH TO STATE THEIR NEW REPLACEMENT ORDER WAS EXACTLY THE SAME AS D236914S-ALL 8" (LOT# NGHN4359) INSTEAD OF THE USUAL 11-AND ALL HAD THE SAME LOT NUMBER AS STOCK WE HAD ISSUES WITH LAST MONTH (LOT# NGHN4359). IT WAS STATED THAT THEY HAD CHECKED THE STOCK ON SHELVES AND THAT IT WAS ALL THE SAME ISSUE AND LOT NUMBER. IT WAS ALSO STATED THAT THEY HAD CHECKED THE NEW ORDER OF 12CH D236912S THEY HAVE IN STOCK, AND THEY WERE ALSO 8" INSTEAD OF THE USUAL 11 (LOT# UNK)". PER ADDITIONAL INFORMATION VIA EMAIL FROM (B)(6) ON 02AUG2023, IT WAS STATED THAT THE ALL OF THE SAME LOT. NUMBER SO 12 CATHETERS IN TOTAL. THE PATIENT HAD THE SAME ISSUE IN JUNE WITH ALL 12 CATHETERS BEING AFFECTED ALSO. IN TOTAL ALL 24 CATHETERS HAD BEEN THE SAME LOT NUMBER. IT WAS ALSO STATED THAT THE OUR STOCK OF 12CH HAVE THE SAME ISSUE ALSO AND NOT SURE IF THE LOT NUMBER WAS THE SAME ON THOSE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05918 | 2023-08-08 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE URINE WAS DRAINED AND THE BALLOON OF THE FOLEY CATHETER HAD RUPTURED, REQUIRING RE-INDWELLING CATHETERIZATION. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05909 | 2023-08-08 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT DURING THE PRETEST, STERILE WATER LEAKED FROM THE CONNECTION BETWEEN THE SYRINGE AND THE INFLATION VALVE OF THE FOLEY CATHETER. THERE WAS NO DAMAGE FOUND VISUALLY, AND THE WATER LEAKED FROM THE VALVE MENTIONED AS BACKFLOW. NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER. AS PER THE FOLLOW UP INFORMATION RECEIVED ON (B)(6) 2023, CLARIFIED THAT THE BACKFLOW MENTIONED IN THE SOURCE DOCUMENT WAS LEAKAGE FROM THE VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03400 | 2023-08-08 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | CUSTOMER REPORTED THAT THE PICC TUBE WAS PLACED, THE GUIDE WIRE WAS PLACED SMOOTHLY, AND THE VASCULAR SHEATH WAS READY TO BE PLACED. A SMALL SPLIT WAS FOUND AT THE FRONT END OF THE VASCULAR SHEATH, ABOUT 2MM IN LENGTH, AND A NEW VASCULAR SHEATH WAS IMMEDIATELY REPLACED. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03399 | 2023-08-08 | 2023-07-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN SOAKING, FLUSHING AND LOCKING THE SOLO CATHETER POST OPENING OF THE OUTER PACKAGE, THE CUSTOMER FOUND THAT LIQUIDS LEAKED FROM THE LOWER END OF THE VALVE. AN INSPECTION REVEALED CRACKS ON THE CATHETER. AFTER THE CATHETER WAS DISPOSED OF, ANOTHER NEW KIT OF CATHETER WAS UNPACKED IN A RUSH TO COMPLETE CATHETER PLACEMENT IN THIS PATIENT. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03398 | 2023-08-08 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN PRE-FLUSHING THE CATHETER, LEAKAGE WAS NOTICED. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03395 | 2023-08-08 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN PRE-FLUSHING THE CATHETER, LEAKAGE WAS NOTICED. NO OTHER INFORMATION PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03389 | 2023-08-08 | 2023-07-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | CUSTOMER REPORTED PATIENT COMPLAINED OF DISCOMFORT AROUND PICC SITE. PICC WAS BLOCKED, DIFFICULTY GETTING UROKINASE INSERTED INITIALLY UROKINASE EVENTUALLY SUCCESSFUL IN UNBLOCKING PICC. CHEST X-RAY CONFIRMED CORRECT POSITION. PATIENT CONTINUED TO COMPLAIN OF DISCOMFORT AROUND PICC SITE / NO OBVIOUS REASON FOR SAME LINOGRAM REQUESTED BY MEDICAL STAFF STATING QUERY FRACTURED LINE / LINOGRAM NOT COMPLETED AND ABOVE INFORMATION NOT COMMUNICATED TO NURSING STAFF OR RECORDED IN PATIENTS NOTES. PICC LINE REMOVED AS REQUESTED/TIP SENT FOR ANALYSIS. NO INJURY SUSTAINED / NO LEAKAGE FROM PICC. NO OTHER INFORMATION PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03387 | 2023-08-08 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER WITH SHERLOCK 3CG TPS STYLET FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1274108 |  | IT WAS REPORTED CUSTOMER RETURNED A LEAKING 3FR PICC. HAD TO CLEAN UP MESS AND REPLACE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03385 | 2023-08-08 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405J |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, LEAKAGE WAS FOUND DURING INFUSION. WHEN THE CATHETER WAS CHECKED, LEAKAGE WAS CONFIRMED FROM THE CONNECTION BETWEEN THE CATHETER AND THE OVERSLEEVE. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK IS CONFIRMED AND THE CAUSE APPEARS TO BE USE RELATED. THE PRODUCT RETURNED FOR EVALUATION WAS A 4 FR GROSHONG NXT CLEARVUE CONNECTOR ASSEMBLY. IT WAS RECEIVED ASSEMBLED. USE RESIDUE WAS EVIDENT THROUGHOUT THE DEVICE. THE CATHETER WAS NOT RETURNED. NO CATHETER FRAGMENT WAS VISIBLE WITHIN THE CONNECTOR. THE CONNECTOR WAS DISASSEMBLED AND FOLLOWING DISASSEMBLY, NO CATHETER FRAGMENTS WERE WITHIN THE CONNECTOR. FOLLOWING BISECTION OF THE DISTAL CONNECTOR SEGMENT, INSPECTION OF THE COMPRESSION SLEEVE REVEALED IT WAS NOT IN THE ADVANCED, LOCKED POSITION. INSPECTION OF THE PROXIMAL CONNECTOR SEGMENT CANNULA REVEALED MASS OF BIOLOGICAL RESIDUE ADHERED TO THE DISTAL TIP. IT APPEARED THAT THE MASS MAY HAVE WITHDRAWN THE COMPRESSION SLEEVE DURING DISASSEMBLY. DISASSEMBLY REVEALED THERE WAS NO CATHETER FRAGMENTS PRESENT WITHIN THE CONNECTOR, SUGGESTING THAT THE CATHETER WAS NOT PROPERLY INSERTED IN THE COMPRESSION SLEEVE DURING ASSEMBLY. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03383 | 2023-08-08 | 2023-07-14 | I, n, j, u, r, y | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540 |  | IT WAS REPORTED, THAT (B)(6) YEARS, (B)(6) MONTHS, AND (B)(6) DAYS POST PORT PLACEMENT. THE CATHETER WAS ALLEGEDLY BROKEN AND SEPARATED FROM THE PORT. IT WAS FURTHER REPORTED, THAT THE BROKEN CATHETER GOT DISLODGED IN THE HEART. REPORTEDLY, THE EXTERNAL PORT WAS REMOVED BY NORMAL PROCEDURE, BUT THE INTERNAL CANNULA WENT INTO THE RIGHT VENTRICLE, WHICH WAS REMOVED BY AN INTERVENTIONAL CARDIOLOGIST IN THE CARDIAC CATHETERIZATION LAB. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED. A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, AN IMAGE WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. D4 (EXPIRY DATE: 05/2023). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05916 | 2023-08-08 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE FOLEY CATHETER ON THE NEXT DAY AFTER PLACEMENT. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE EXACT LEAK LOCATION WAS NOT REPORTED BY THE CUSTOMER | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® I.C. SILVER FOLEY TRAY B SINGLE USE ONLY . [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05895 | 2023-08-08 | 2023-07-25 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 50614 |  | IT WAS REPORTED THAT CUSTOMER HAS BEEN USING PUREWICK FEMALE EXTERNAL CATHETER FOR 2 YEARS. ON SATURDAY, THEY WERE USING THE CATHETER BUT THIS TIME THE CATHETER CUT THE INSIDE OF THE PATIENT URETHRA. PATIENT WENT TO URGENT CARE AND WAS PUT ON ANTIBIOTIC. A NEW CATHETER WAS PUT IN FOR THE PATIENT. CUSTOMER SAID IT LOOKED LIKE THE END OF THE CATHETER WAS NOT FORMED PROPERLY AND THE TIP CUT THE URETHRA. PER FOLLOW UP VIA PHONE ON 25JUL2023, IT WAS REPORTED THAT THE CUSTOMER WAS USING THE ANTIBIOTIC, BACTRIM, TO HELP ALLEVIATE THE URINARY TRACT INFECTION. HE MADE IT CLEAR THAT THIS WAS AN INTERMITTENT CATHETER AND NOT A FEMALE EXTERNAL CATHETER. | THE REPORTED EVENT IS INCONCLUSIVE AS IT'S UNKNOWN HOW THE PRODUCTS REACTED IN THE PARTICULAR PATIENT'S BODY. VISUAL INSPECTION NOTED (1) MAGIC3 ISC IN OPENED PACKAGING. THE TIP OF THE ISC HAS A NICK. IT IS UNKNOWN HOW THE PRODUCT REACTED IN THE PARTICULAR PATIENT'S BODY. THEREFORE, IT IS UNKNOWN IF THE PRODUCT MEETS SPECIFICATIONS. THE POTENTIAL ROOT CAUSE COULD BE PATIENT SENSITIVITY TO MATERIALS. THE PRODUCT WAS USED FOR PATIENT DIAGNOSTIC OR TREATMENT. IT WAS UNKNOWN IF THE PRODUCT WAS INFLUENCED BY THE REPORTED EVENT. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS.WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED, | Y | N | LAY USER/PATIENT |  | 2023-08
1018233-2023-05917 | 2023-08-08 | 2023-07-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165SI18 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS WITHDRAWN SPONTANEOUSLY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE TORN RUBBERIZE WHICH COULD BE DUE TO INSUFFICIENT LATEX STRENGTH, LOW/NON UNIFORM RUBBERIZED THICKNESS ETC AND ALSO MIGHT BE CONTRIBUTED BY USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.], (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS, PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.]. [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.]. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.]. 2. APPLICABLE PATIENTS. (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX.. (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING. THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. SIZES OF CATHETERS. AVAILABLE IN SIZES 8 TO 24 EVERY 2FR. 1. FOLEY CATHETER. 2. ACCESSORIES. CLOSED DRAINAGE BAG. (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.). SYRINGE PREFILLED WITH STERILE WATER. WATER SOLUBLE LUBRICANT. ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION. TWEEZERS. GAUZE PADS. WATERPROOF SHEET. COTTON BALLS . GLOVES. STATLOCK® FOLEY. [INTENDED USE & EFFECT- EFFICACY]. THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE]. 1. METHOD OF USE. THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1). (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2). FIG.1. FIG.2. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3). (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4). FIG.3. FIG.4. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5). FIG.5. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6). (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7). FIG.6. FIG.7. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT. 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8). FIG. 8. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING. CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9). FIG. 9. 2. PRECAUTIONS FOR USE. (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ]. (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS]. 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS). (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS. (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.). B. BALLOON-RUPTURE METHOD. WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD. 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) . FIG. 10. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11). FIG. 11. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12). FIG. 12. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13). FIG. 13. B. BALLOON-RUPTURE METHOD. 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14). FIG. 14. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15). FIG. 15. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16). FIG. 16. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17). FIG. 17. 3. MALFUNCTION AND ADVERSE EVENTS. 1) MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE, DIFFICULTY OR FAILURE TO REMOVE THE DEVICE, OCCLUSION OF CATHETER INNER LUMENS, ENCRUSTATION, ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE, DEVICE DAMAGE DUE TO INAPPROPRIATE USE. 2) ADVERSE EVENTS: URINARY-TRACT INFECTION, HEMORRHAGE, HEMATURIA, ALLERGY REACTION TO THE DEVICE, CALCULUS FORMATION , EDEMA, PAIN, DISCOMFORT, INJURY OF BLADDER OR URETHRAL, URETHRITIS, URINARY INCONTINENCE, RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]. 1. STORAGE. STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE. INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03380 | 2023-08-08 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED THAT THE PATIENT LOST ACCESS IN THE OR IN THE MIDDLE OF A CASE AND HAD TO HAVE EMERGENT ACCESS PLACED PER THE CUSTOMER. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05914 | 2023-08-08 | 2023-07-25 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 51814 |  | IT WAS REPORTED, THAT PATIENT WAS HOSPITALIZED (B)(6) TIMES. AND HAD TO RECEIVE ANTIBIOTICS AND STEROIDS, DUE TO URINARY TRACT INFECTIONS. | THE REPORTED EVENT WAS INCONCLUSIVE, BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE, DUE TO "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN, WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER, AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR, WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS, PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE, WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM. 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-08
1018233-2023-05913 | 2023-08-08 | 2023-07-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123522A |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAS POPPED TWICE WHILE IN THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON BURST. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE PATIENT USE ONLY. DO NOT REUSE AND RESTERILIZE. FOR UROLOGICAL USE ONLY. USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PROTOCOL, THE MAY BE CUT OFF. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03372 | 2023-08-08 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA CATHETER W/SOLO VALVE TECHNOLOGY MAXIMAL BARRIER KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S3395108D |  | IT WAS REPORTED THAT THE STIFFENING WIRE IN THE BARD PICC BENDS AND CANNOT BE ADVANCED ONCE THE STIFFENING WIRE IS ALTERED. HAD TO USE 3 PICC LINE KITS ON THIS PATIENT. STIFFENING WIRE INSIDE PICC LINE BENT AT A 90 DEGREE APPROXIMATELY 15 CM FROM HUB. UNABLE TO FULLY THREAD PICC LINE AFTER WIRE BENT. HAD TO TRIM THE PICC LINE TO WHERE THE STIFFENING WIRE BENT TO USE PICC LINE. IT WAS REPORTED THAT THREE DIFFERENT PICC KITS WITH THREE DIFFERENT MATERIAL NUMBERS WERE USED. THIS REPORT ADDRESSES THE SECOND KIT AND MATERIAL NUMBER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03379 | 2023-08-08 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TPS STYLET 4F SINGLE-LUMEN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 91660417 |  | IT WAS REPORTED THAT AFTER INSERTION OF THE NEW PICC LINE, THE HUB WAS ATTACHED TO THE LINE AS PER PROTOCOL, AND THE LINE WOULD NOT FLUSH OR RETURN BLOOD. THIS HUB WAS CUT OFF AND THE LINE WAS REPAIRED USING A REPAIR KIT AS PER PROTOCOL. NO PATIENT AFFECTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF APPLICABLE INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF DIFFICULTY FLUSHING THE CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR SL GROSHONG CATHETER. INSPECTION OF THE RETURNED UNIT FOUND THAT THE CATHETER TUBING WAS BUNCHED UP AGAINST THE METAL CANNULA OF THE TWO-PIECE CONNECTOR. IT APPEARED THAT THE BUNCHED CATHETER MATERIAL FORCED THE COMPRESSION SLEEVE TOO FAR INTO THE DISTAL CONNECTOR SEGMENT DURING ASSEMBLY, CAUSING THE COMPRESSION SLEEVE TO PINCH THE CATHETER AND IMPEDE FLOW. SUCH DAMAGE CAN OCCUR IF THE CATHETER IS ADVANCED TOO FAR ONTO THE PROXIMAL CONNECTOR CANNULA AND IF EXCESS CATHETER MATERIAL IS FORCED INTO THE CONNECTOR BORE DURING CATHETER/CONNECTOR ASSEMBLY. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03377 | 2023-08-08 | 2023-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | CUSTOMER REPORTED THAT THE PATIENT IS HEMIPLEGIC DUE TO POOR VEINS AND LONG TREATMENT TIME, A "CENTRAL VENOUS CATHETER WITH PERIPHERAL INTUBATION" WAS PLACED ACCORDING TO THE DOCTOR'S ADVICE. THE MEDICAL STAFF OF OUR HOSPITAL SUCCESSFULLY PLACED THE TUBE FOR THE PATIENT, AND WHEN THE WITHDRAWAL WAS CONFIRMED, IT WAS FOUND THAT THE EXTRACTED BLOOD CONTAINED FOAM, AND SOMETIMES THERE WAS NO BLOOD RETURN (UNDER NORMAL CIRCUMSTANCES, THE BLOOD RETURN SHOULD BE SMOOTH, NO FOAM), WHICH WAS VERY CONFUSED, AND THE POSITION OF THE PIPELINE WAS ADJUSTED, AND THIS WAS THE CASE FOR REPEATED WITHDRAWAL AND INJECTION. LATER, IT WAS FOUND THAT THE END OF THE PIPE HAD BEEN DRIPPING WATER WHEN THE SALINE WAS INJECTED, SO IT WAS THOUGHT THAT THE INTERFACE WAS NOT TIGHTENED. AFTER RE-FIXING THE INTERFACE, THE LOWEST PART WAS CHANGED REPEATEDLY, BUT WATER WAS STILL LEAKING THERE. ACCORDING TO THE RESULTS OF CATHETERIZATION PREDICTION, FORTUNATELY, THIS SECTION OF PICC TUBE BELONGS TO THE CUTTING RANGE, AND IT DOES NOT NEED TO BE RE-CATHETERIZED OR REPLACED. AFTER CUTTING, BLOOD RETURN IS SMOOTH AND PROPERLY FIXED AFTER INSERTING THE KIT. NO PATIENT HARM REPORTED. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03375 | 2023-08-08 | 2023-07-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE OUTER WALL OF THE PICC LINE WAS DAMAGED. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED CATHETER WAS INCONCLUSIVE DUE TO THE SAMPLE CONDITION. A PHOTOGRAPH OF A POWERPICC CATHETER SHAFT, LIKELY FROM THE IMPLICATED 5FR DUAL-LUMEN CATHETER, WAS RETURNED FOR EVALUATION. A WHITISH/CLEAR RESIDUAL MATERIAL WAS NOTED ON THE CATHETER. THE RESIDUAL MATERIAL APPEARED CONSISTENT WITH ADHESIVE RESIDUE FROM THE CATHETER DRESSING. A SMALL AMOUNT OF REDDISH MATERIAL, WHICH MAY BE BLOOD RESIDUE, WAS ALSO SEEN ON THE CATHETER. AN ARROW WAS DRAWN ONTO THE PHOTOGRAPH POINTING TO THE CATHETER SHAFT. HOWEVER, THE IMPLICATED DAMAGE COULD NOT BE SEEN IN THE RETURNED PHOTOGRAPH. AS THE SUBMITTED PHOTOGRAPH DID NOT CONTAIN ENOUGH INFORMATION TO CONFIRM THE EVENT OR IDENTIFY A SPECIFIC ROOT CAUSE, THIS COMPLAINT WILL BE RECORDED AS INCONCLUSIVE AT THIS TIME. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03376 | 2023-08-08 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | CUSTOMER REPORTED THAT THE PATIENT WAS TO UNDERGO CHEMOTHERAPY AND WAS GIVEN AN INDURATING PICC CATHETER TO PROTECT THE BLOOD VESSELS. DURING THE CATHETERIZATION PROCESS, THE CATHETERIZATION WAS SMOOTH, AND THE PUNCTURE WAS SUCCESSFUL. THE APPEARANCE OF THE CANNULA SHEATH WAS NOT ABNORMAL, AND THERE WAS SLIGHT RESISTANCE DURING DELIVERY, AND THE CATHETER WAS SENT SMOOTHLY. AFTER WITHDRAWING THE CANNULA SHEATH, IT WAS FOUND THAT THE FRONT EDGE OF THE CANNULA SHEATH WAS CURLY AND BURR. AFTER WITHDRAWING THE CANNULA SHEATH, IT WAS FOUND THAT THE PUNCTURE POINT OOZED MORE BLOOD, AND THE ELASTIC BANDAGE WAS APPLIED TO HEMOSTASIS. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05905 | 2023-08-08 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175812 |  | IT WAS REPORTED THAT FOLEY CATHETER BALLOON BURST WHEN TESTED. AS PER PACKET INSTRUCTIONS, 10 ML OF STERILE WATER WAS INSERTED AND BALLOON BURST BEFORE IT REACHED CAPACITY. THE BALLOON BURST WHEN STERILE WATER WAS PUT IN. THE CUSTOMER TESTED THE BALLOON BEFORE INSERTING THE CATHETER INTO THE PATIENT. IF THEY HAD NOT TESTED THE BALLOON BEFORE INSERTION, IT COULD HAVE BURST INSIDE THE PATIENTS BLADDER. ANOTHER CATHETER FROM SAME BATCH WAS TESTED WITH NO ISSUES. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "HIGH MODULUS SILICONE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05915 | 2023-08-08 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT STERILE WATER WAS DIFFICULT TO REMOVE DURING A PRETEST. THE CATHETER WAS NOT USED. WHEN THE CATHETER WAS LEFT WITH WATER IN THE BALLOON, THE BALLOON DEFLATED AFTER ABOUT 3 HOURS. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 02AUG2023, THE EVENT OCCURRED DURING A PRETEST AND WAS NOT USED ON THE PATIENT. THE WATER COULD NOT BE REMOVED FROM THE BALLOON DURING PRETEST. HOWEVER, IT WAS REPORTED THAT UPON LEAVING IT, THE WATER (NOT THE CATHETER) SPONTANEOUSLY REMOVED FROM THE BALLOON IN 3 HOURS AFTER THE EVENT. UPON CHECKING THE RETURNED SAMPLE, THE WATER WAS NOT COMPLETELY REMOVED AND THE BALLOON REMAINED IN THE ASYMMETRICAL INFLATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05890 | 2023-08-08 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT IN THE OPERATING ROOM, THE TEMPERATURE SENSOR WAS DISPLACED AND STUCK IN THE FUNNEL AREA, AFFECTING EVEN THE DRAINAGE OF URINE, SO TEMPERATURE SENSING FOLEY CATHETER USE WAS DISCONTINUED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05912 | 2023-08-08 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE TWO OF FOLEY CATHETER BALLOON WOULD NOT DEFLATE. NURSE PLACING THE FOLEYS WAS PRIMING THE BALLOON PRIOR TO INSERTION SO THE FAULTY ITEMS DID NOT MAKE IT INTO A PATIENT. NURSE PREFILLED THE BALLOON WITH 5 ML TO TEST IT, AND IT WOULD NOT DEFLATE. THEY GOT ANOTHER TRAY AND HAD THE SAME ISSUE. FINALLY, WITH THE THIRD TRAY, THE BALLOON DEFLATED PROPERLY. THEY WAS NOT ABLE TO GET IT TO DEFLATE INITIALLY EITHER AND PUSHED THE SYRINGE INTO THE PORT REALLY HARD TO THE POINT THAT THE ORANGE RING POPPED OFF, AND THEN IT FINALLY DEFLATED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05889 | 2023-08-08 | 2023-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER HAD BEEN USING FOLEY CATHETER EVERYDAY FOR LAST TWO AND HALF YEARS. THE CUSTOMER WAS A PATIENT OF HYPOTONIC URINARY BLADDER AND REQUIRED SELF CATHETERIZATION AT LEAST TWICE A DAY. BUT THE CUSTOMER WAS OBSERVING THAT OVER LAST ONE TO TWO MONTHS, THE QUALITY OF THE CATHETERS WERE NOT UP TO THE MARK, REQUIRING REPLACEMENT AFTER USING 2 TO 3 TIMES ONLY UNLIKE PREVIOUSLY WHEN EACH CATHETER COULD BE USED 6 TO 7 TIMES. THE CATHETERS WERE OF SIZES 16 FR AND 18 FR. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE HIGH DRYING TEMPERATURE AND LONG DRYING TIME. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03366 | 2023-08-08 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED, THAT APPROXIMATELY (B)(6) WEEKS POST PORT PLACEMENT. THE PATIENT ALLEGEDLY EXPERIENCED SMALL BUMP ON SKIN NEAR THE PORT. REPORTEDLY, THE BUMP WAS DISAPPEARED. AND THE PATIENT CONTINUED TREATMENTS WITH THE PORT. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03361 | 2023-08-08 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT FOUR YEARS SEVEN MONTHS AND EIGHT DAYS POST A PORT PLACEMENT, LEAKAGE ALLEGEDLY OCCURRED FROM THE INSERTION SITE. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY BROKEN INSIDE THE INTERNAL JUGULAR VEIN. REPORTEDLY THE DISTAL CATHETER SEGMENT AND PORT BODY WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE X-PORT IMPLANTABLE PONE CATH-LOCK, AND ONE GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATION WERE PERFORMED. TWO PARTIAL CIRCUMFERENTIAL BREAKS WERE NOTED APPROXIMATELY 14.5CM AND 15.5CM FROM THE DISTAL END OF THE CATH-LOCK. FIRST PARTIAL CIRCUMFERENTIAL BREAK WAS SLIGHTLY INTACT AND THE SECOND PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED TO HAVE EVEN EDGES. THE EDGES OF THE COMPLETE CIRCUMFERENTIAL BREAK AT THE DISTAL END OF THE CATHETER WERE NOTED TO BE UNEVEN AND THE SURFACE WAS NOTED TO BE GRANULAR AND HAVE TO LONGITUDINAL SPLITS. UPON INFUSION, WATER AND THROMBUS WERE NOTED EXITING THE DISTAL END AND NO LEAKS WERE NOTED. ASPIRATION WAS UNSUCCESSFUL. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION AND IDENTIFIED WEAR AND DEFORMATION ISSUES. HOWEVER THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED LEAK ISSUE. THE IDENTIFIED BREAK IS TYPICAL OF FLEXURAL FATIGUE, WHICH IS DUE TO THE REPETITIVE, CYCLIC KINKING OF THE CATHETER. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
2020394-2023-00570 | 2023-08-08 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE OF A TARGET LESION BELOW THE KNEE VIA AN IPSILATERAL ANTEGRADE APPROACH, THE TIP OF THE CATHETER WAS ALLEGEDLY FOUND TO BE SEPARATED IN THE ANTERIOR TIBIAL ARTERY. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 09/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03362 | 2023-08-08 | 2023-07-13 | I, n, j, u, r, y | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540 |  | IT WAS REPORTED THAT ONE YEAR SEVEN MONTHS AND TWENTY-THREE DAYS POST PORT PLACEMENT, THE CATHETER ALLEGEDLY GOT BROKEN AND GOT SEPARATED FROM THE PORT. IT WAS FURTHER REPORTED THAT THE BROKEN CATHETER GOT DISLODGED IN THE HEART. REPORTEDLY, THE PORT WAS REMOVED BY THE SURGERY. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, AN IMAGE WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 05/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05872 | 2023-08-08 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165816 |  | IT WAS REPORTED THAT THERE WAS AN ISSUE OVER THE WEEKEND WITH 2 CERTAIN LOTS OF FOLEY BALLOONS THAT WERE NOT WORKING (LOT#NGHP0791) AND (LOT#NGHQ1143). PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT FOLEY CATHETER BALLOON WAS LEAKING AFTER IT WAS INSERTED INTO THE PATIENT (LOT#NGHP0791). IT WAS ALSO REPORTED THAT ANOTHER LOT WAS NOT DEFLATING PROPERLY (LOT#NGHQ1143). THE CUSTOMER WAS NOT ABLE TO CROSS-REFERENCE WHICH LOT HAD WHICH ISSUE, BUT LOT#NGHP079 WAS DISCARDED, AND 6-8 TRAYS FROM LOT#NGHQ1143 ARE AVAILABLE FOR A SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05887 | 2023-08-08 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS BLOCKED. THERE WAS NO URINE DISCHARGE IMMEDIATELY AFTER PLACEMENT IN THE PATIENT. THE USER FELT RESISTANCE WHEN INJECTING STERILE WATER INTO THE CATHETER. ALSO STATED THAT A LITTLE WHITE SUBSTANCE CAME OUT FROM THE CATHETER TIP WHEN STERILIZED WATER WAS INJECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05882 | 2023-08-08 | 2023-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER FELL OUT OF THE PATIENT IN A DAY AFTER PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05885 | 2023-08-08 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAD BURST. THE USER STARTED USING THE CATHETER ON JULY 19, AND IT WAS BEING USED NORMALLY. BUT, AFTER 1:00 P.M. ON JULY 21, IT WAS FOUND THAT THE CATHETER FELL OUT OF THE PATIENT AND THAT THE BALLOON WAS TORN. THE EVENT OCCURRED TWO DAYS AFTER PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05883 | 2023-08-08 | 2023-07-21 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53614G |  | IT WAS REPORTED THAT THE CATHETER WAS LOGGED AS A PRECAUTION AGAINST CATHETERS. A NURSE COMPLAINT HAS BEEN PREVIOUSLY LOGGED. IT WAS NOTED FROM THE LETTER THAT AFTER A VISIT FROM THE NURSE, THEY RECOMMENDED THE WRONG TYPE OF CATHETER. THE PATIENT HAD FIVE EMERGENCY ADMISSIONS IN THE MONTH OF APRIL, WHERE THE USER WAS DIAGNOSED WITH SEPSIS, PNEUMONIA AND DEVELOPED A SCROTAL ABSCESS THAT HAD TO BE REMOVED UNDER ANESTHETIC. AFTER COMING HOME, THE PATIENT COULD NOT CLIMB THE STAIRS. THE DISTRICT NURSE HAD TO DRESS THEIR WOUNDS, AND THEY COULD NO LONGER WALK UNAIDED. PER FOLLOW UP VIA EMAIL ON 15-FEB-2024, THE PATIENT¿S WIFE REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH PER DEATH CERTIFICATE WAS UROSEPSIS, MULTIPLE URINARY TRACT INFECTIONS (UTIS). ADDITIONAL INFORMATION WAS RECEIVED VIA EMAIL ON 27-FEB-2024. A BD (BECTON DICKENSON) REPRESENTATIVE MET WITH THE PATIENT AND HIS WIFE ON (B)96) 2023. THE PATIENT WAS ORIGINAL REFERRED TO THE BD NURSING SERVICE BY AN NHS CONTINENCE NURSE SPECIALIST ON (B)(6) 2023 FOR CONSIDERATION OF A MALE EXTERNAL CATHETER (MEC) WITH A VIEW TO MANAGING URINARY SYMPTOMS OVERNIGHT WHILST AWAITING AN APPOINTMENT WITH A UROLOGIST. HOWEVER, THE PATIENT WAS TAUGHT TO CATHETERIZE BY THE VISITING BD NURSE DURING A HOME VISIT ON (B)(6) 2023. PATIENT WIFE INFORMED BD THAT THE PATIENT SUBSEQUENTLY PRESENTED TO THE GENERAL PRACTITIONER (GP) AS HE FELT UNWELL ON (B)(6) 2023 AT WHICH TIME PATIENT CEASED INTERMITTENT SELF-CATHETERIZATION (ISC) ON THE ADVICE OF GP HAVING UNDERTAKEN THE PROCEDURE 2-3 TIMES. PATIENT HAD ALSO EXPERIENCED A SMALL AMOUNT OF BLEEDING WHEN SELF-CATHETERIZING. AS HE FELT NO BETTER THE FOLLOWING MORNING ((B)(6) 2023), PATIENT WAS ADMITTED TO HOSPITAL VIA A&E WHERE PATIENT WAS DIAGNOSED WITH A TESTICULAR ABSCESS WHICH REQUIRED ORCHIDECTOMY. WHILST IN HOSPITAL HE DEVELOPED PNEUMONIA AND ¿HEART PROBLEMS¿. FOLLOWING DISCHARGE, PATIENT FELT FRAIL AND WEAK, AND PATIENT WAS UNABLE TO CLIMB THE STAIRS. PATIENT WAS REFERRED TO THE DIETICIAN DUE TO WEIGHT LOSS. PATIENT HAD ANOTHER HOSPITAL ADMISSION LATER IN THE YEAR DUE TO A FURTHER TESTICULAR ABSCESS WHICH WAS TREATED WITH ORAL ANTIBIOTICS. PER FOLLOW UP VIA EMAIL ON 01-MAR-2024, A BD REPRESENTATIVE REPORTED THAT THE PATIENT ONLY USED HYDROSIL GO FOLLOWING THE INITIAL VISIT BY THE BD NURSE ¿ THIS APPEARS TO HAVE BEEN USED 2-3 TIMES SINCE THIS INITIAL VISIT. AT THE TIME BD REPRESENTATIVES MET WITH THE PATIENT IN (B)(6) 2023, THE PATIENT WAS NOT USING ANY PRODUCT. IT WAS UNKNOWN IF THE PATIENT USED A PRODUCT BEYOND THIS AS HIS CONDITION DETERIORATED. IF HE DID, IT WAS NOT SUPPLIED BY SCRIPT-EASY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PLEASE READ ALL INSTRUCTIONS BEFORE USING THIS DEVICE. CAUTION: FEDERAL (USA) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. DEVICE DESCRIPTION: THE HYDROSIL GO® INTERMITTENT URINARY CATHETER IS A SILICONE TUBE INSERTED INTO THE URETHRA FOR THE PURPOSE OF DRAINING THE BLADDER OF URINE. NOT MADE WITH NATURAL RUBBER LATEX DOES NOT CONTAIN DEHP INDICATIONS FOR USE: THE HYDROSIL GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. INSTRUCTIONS FOR USE: REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, USE THE ADHESIVE LABEL(S) TO HANG THE PACKAGE ON A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05886 | 2023-08-08 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE NURSE WAS UNABLE TO REMOVE FOLEY CATHETER FROM THE PATIENT. THE FOLEY CATHETER WAS PLACED IN THE HOSPITAL. IT APPEARED TO BE A TEMPERATURE SENSING. THEY PULLED THE PROBE FROM THE FOLEY, BUT THIS DID NOT CORRECT THE ISSUE. THEY EXPLAINED THAT THEY SHOULD DETERMINE IF THE INFLATION LUMEN WAS COLLAPSED AND THEY EXPLAINED HOW ASPIRATING WATER FROM THE BALLOON COULD DO THIS. THEY ADVANCED THE FOLEY SLIGHTLY TO DETERMINE IT WAS IN THE BLADDER AND REINSERTED 3ML OF WATER INTO THE INFLATION LUMEN. THEY ALLOWED IT TO PASSIVELY DEFLATED. THERE WAS SOME BLEEDING NOTED, BUT THEY WAS ABLE TO PULL THE FOLEY. THE INFLATION CUFF DID NOT CONTAIN ANY PRODUCT OR LOT NUMBERS THEY COULD READ ONLY SURE AND IT WAS A TEMPERATURE SENSING SILICONE FOLEY. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03364 | 2023-08-08 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM LOW-PROFILE IMPLANTABLE PORT WITH ATTACHABLE 8F GROSHONG SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0606150 |  | IT WAS REPORTED THAT THREE YEARS, FOUR MONTHS, AND SIX DAYS POST-PORT PLACEMENT, DURING INFUSION OF CYTOSTATIC, THE PATIENT EXPERIENCED OEDEMA AND ERYTHEMA AT THE IMPLANT SITE WITHOUT PAIN. IT WAS FURTHER REPORTED THAT AN X-RAY WAS PERFORMED, ALLEGEDLY SHOWING DISUNION OF THE CHAMBER AND CATHETER. CONTRAST OPACIFICATION ALLEGEDLY PRESENTED A LEAK AT THE CLAVICULAR LEVEL. FURTHERMORE, A SLIT WAS VISIBLE ON THE CATHETER AT THE SIXTEEN-CENTIMETER MARK. REPORTEDLY, CATHETER REMOVAL OF THE RIGHT-SIDE IMPLANTABLE CHAMBER WAS CONDUCTED, AND A NEW LEFT-SIDE CHAMBER WAS IMPLANTED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2023). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05966 | 2023-08-09 | 2023-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE TEMPERATURE DISPLAY OF THE FOLEY CATHETER WAS POOR. STATED THAT THERE WAS A DEFECTIVE TEMPERATURE DISPLAY IN THE HIGH CARE UNIT AND THE JACK WAS SUSPECTED TO BE DAMAGED. STATED THAT THERE WAS A REGULAR TEMPERATURE DISPLAY FAILURE. BOTH THE FUKUDA DENSHI MONITOR AND THE TEMPERATURE SENSOR CABLE (TCJ) HAD DISPLAY PROBLEMS LAST YEAR, SO THEY JUST SWITCHED THEM IN (B)(6) THIS YEAR. PER FOLLOW-UP INFORMATION RECEIVED FROM (B)(6) ON 02AUG2023, STATED THAT THE REGULAR TEMPERATURE DISPLAY FAILURE WERE ALREADY CAPTURED IN THE PREVIOUS RECORDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03415 | 2023-08-09 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED INCIDENT WITH A PICC LINE THAT WE DISCOVERED WAS LEAKING AND IT TURNED OUT TO BE A LEAK JUST AFTER THE 4 CM MARK. ADDITIONAL INFORMATION RECEIVED 26 JULY 2023: THE PATIENT IN QUESTION HAS RECEIVED CYTOSTATIC TREATMENT WITH FOLFOX, THE LEAKAGE WAS DISCOVERED WHEN HER CYTOSTATIC PUMP WAS TO BE DISCONNECTED ON (B)(6), FORTUNATELY, NO CYTOSTATIC HAD LEAKED OUT ON THE SKIN BUT WHEN I FLUSHED WITH NACL I SAW THAT IT WAS LEAKING, THE PATIENT THUS DID NOT HAVE ANY PROBLEMS. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03416 | 2023-08-09 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, THERE WAS A FRACTURED PICC, EXTERNAL LENGTH 20.5CM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03414 | 2023-08-09 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 3F SL MBPLUS (3CG, BIOPATCH, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1173108D5 |  | IT WAS REPORTED BY CUSTOMER, THAT GUIDEWIRE WOULD NOT ADVANCE. WHEN ATTEMPTING TO PULL OUT IT, GOT STUCK. A TISSUE SUBSTANCE APPEARS TO BE STUCK ON WIRE CLOSE TO TIP. PATIENT NEEDED GUIDEWIRE REMOVED BY TEAM OF PHYSICIANS. PATIENT NOT HARMED. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE, AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05961 | 2023-08-09 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123614A |  | IT WAS REPORTED THAT AFTER INDWELLING CATHETERIZATION, THE FOLEY CATHETER BALLOON RUPTURED AND THE URINARY TUBE FELL OFF. IT WAS STATED THAT THE CATHETER WAS RE-INJECTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05959 | 2023-08-09 | 2023-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT URINE OR STERILE WATER LEAKED FROM AROUND THE FOLEY CATHETER CONNECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03412 | 2023-08-09 | 2023-07-21 | I, n, j, u, r, y | Manufacturer report | POWERPICC SV CATHETER W/ SHERLOCK 3CG TIP POSITIONING SYSTEM, STYLET BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1274118 |  | IT WAS REPORTED VIA MEDWATCH "ULTRASOUND GUIDANCE WAS USED TO IDENTIFY VEIN. DYNAMIC, REAL-TIME IMAGING OF VESSEL CANNULATION PERFORMED. STERILE GEL AND PROBE COVER USED IN ULTRASOUND-GUIDED CENTRAL VENOUS CATHETER INSERTION. CHLORAPREP ANTISEPTIC USED DURING CENTRAL VENOUS CATHETER INSERTION. SKIN PREP AGENT COMPLETELY DRIED PRIOR TO PROCEDURE. HAND HYGIENE PERFORMED PRIOR TO CENTRAL VENOUS. CAP, MASK, STERILE GOWN, STERILE GLOVES AND STERILE DRAPE USED. LEFT BASILIC VEIN WAS LOCATED USING ULTRASOUND, NEEDLE WAS PLACED WITH BLOOD RETURN, GUIDEWIRE WAS ADVANCED 5-10 CM UNTIL THERE WAS RESISTANCE AND SO ADVANCEMENT WAS STOPPED AND GUIDEWIRE WAS PULLED BACK ALONG WITH THE NEEDLE 2-3 CM UNTIL RESISTANCE WAS MET AND STOPPED. ATTEMPTED TO REMOVE NEEDLE AND GUIDEWIRE TOGETHER BUT WAS UNABLE TO DO SO. NEEDLE WAS EASILY REMOVED OVER THE GUIDEWIRE. PROCEDURE WAS ABORTED. WITH THE GUIDEWIRE REMAINING IN PLACE, PRESSURE WAS APPLIED TO THE SITE WITH MINIMAL BLEEDING. LIDOCAINE WAS ADMINISTERED AT SITE AND PATIENT WAS GIVEN IV MORPHINE FOR PAIN CONTROL. SURGEONS WERE NOTIFIED, DOCTORS ALL RESPONDED TO BEDSIDE TO EVALUATE PATIENT. X-RAY WAS COMPLETED. DECISION WAS MADE TO TAKE PATIENT TO OPERATING ROOM FOR FOREIGN BODY REMOVAL. MOM AND PATIENT WERE UPDATED AND ALL QUESTIONS WERE ANSWERED." "PICC PLACEMENT WAS CONSULTED FOR PATIENT WHO REQUIRED LONG TERM IV ANTIBIOTICS." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03411 | 2023-08-09 | 2023-07-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001031 |  | IT WAS REPORTED BY THE CUSTOMER THAT THEY OPENED A #4 FRENCH DUAL LUMEN PICC KIT THAT HAD A HAIR INSIDE THE KIT. ADD INFO RECEIVED ON (B)(6) 2023 THE EVENT WAS DISCOVERED WHEN CUSTOMER OPENED THE KIT, SHE SAW A LARGE BLACK HAIR INSIDE THE KIT. SHE DISCARDED THIS KIT, SO IT WAS NOT USED ON THE PATIENT. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05956 | 2023-08-09 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 006173P |  | IT WAS REPORTED THAT THE TEMPERATURE PACING ELECTRODE CATHETER BALLOON WAS BURST. THE USER WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05954 | 2023-08-09 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT DURING CATHETER REMOVAL, RESISTANCE WAS FELT DUE TO A MUSHROOMED BALLOON. IT WAS STATED THAT WHEN THE USER TRIED TO REMOVE THE CATHETER ON THE DAY AFTER THE OPERATION, SINCE THE MUSHROOMED BALLOON WAS CAUGHT, IT WAS REMOVED BY USING A LUBRICANT. IT WAS ALSO STATED THAT THE 3 SIMILAR INCIDENTS OCCURRED IN THE LAST ABOUT 1 MONTH. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05955 | 2023-08-09 | 2023-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL AND DIAGNOSED WITH RIGHT URETERAL CALCULI. IT WAS STATED THAT ON (B)(6) 2023, AT 9:30 A.M., RIGHT URETERAL SOFT LENS HOLMIUM LASER LITHOTRIPSY, POSTOPERATIVE GENERAL ANESTHESIA TO PREVENT URINARY RETENTION AND RETROGRADE URINARY INFECTION, AND INDWELLING CATHETERIZATION AT 16:00 ON (B)(6) 2023, THE PATIENT'S URINARY TUBE CAME OUT SPONTANEOUSLY, AND THE RUPTURE OF THE FOLEY BALLOON WAS CHECKED. IT WAS ALSO STATED THAT CONSIDERING THE GENERAL POSTOPERATIVE SITUATION, THERE WERE NO CHILLS OR FEVERS, AND THE URINARY TUBE WAS NOT INDENTURED AGAIN, SO THE PATIENT'S VITAL SIGNS WERE STABLE AND NO SPECIAL DISCOMFORT WAS OBSERVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05950 | 2023-08-09 | 2023-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OFF. THE CATHETER WAS INSERTED TO THE PATIENT IN THE EMERGENCY ROOM. THEN, IT SPONTANEOUSLY FELL OUT OF THE PATIENT BODY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05949 | 2023-08-09 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE WATER WAS DIFFICULT TO REMOVE FROM THE FOLEY CATHETER. DURING A PRETEST IN THE OPERATING ROOM, THE STERILE DISTILLED WATER WAS NOT REMOVED FROM THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05946 | 2023-08-09 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE INFLATION VALVE OF THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05958 | 2023-08-09 | 2023-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAD RUPTURED AND WATER HAD LEAKED FROM THE BALLOON. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 02AUG2023, STATED THAT THE BALLOON HAD RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05962 | 2023-08-09 | 2023-07-18 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT PATIENT STATED THAT THEY WERE HAVING ISSUES WITH THE PUREWICK ACCESSORY REPLACEMENT KIT, THAT IT SMELLS REALLY BAD EVEN AFTER CLEANING THE WHOLE THING WITH THE INSTRUCTION PROVIDED. THEY ALSO SAID IT LEAKS A LOT. IT WAS NOTED THAT THE PUREWICK ACCESSORY KIT WAS ORIGINALLY PURCHASED WITH THE PUREWICK URINE COLLECTION SYSTEM ON (B)(6) 2023. PER FOLLOW UP VIA PHONE ON 18JUL2023, IT WAS REPORTED THAT THE CUSTOMER WAS WORKING WITH LMS TO REPLACE THE KIT. THE CUSTOMER STATED THAT PATIENT HAD DEVELOPED A RASH, BUT HAD ALREADY SOUGHT MEDICAL INTERVENTION, AND THAT THEY WERE STILL USING THE PUREWICK. PER FOLLOW UP ON 18JUL2023, THE PATIENT HAD SOUGHT MEDICAL INTERVENTION AND WAS CONTINUING TO USE THE WICKS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03418 | 2023-08-09 | 2023-07-16 | I, n, j, u, r, y | Manufacturer report | BASIC KIT 20G X 8 CM WITH GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120087T |  | CUSTOMER REPORTED: "ON (B)(6) 2023 RN ATTEMPTED TO ACCESS CEPHALIC VEIN WITH POWERGLIDE PRO 20G 8CM CATHETER USING ULTRASOUND. ON ASSESSMENT VEIN EASILY COMPRESSIBLE AND ABLE TO ACCOMMODATE SIZE OF CATHETER. CATHETER TIP APPEARED TO BE IN CENTER OF VESSEL, RN ADVANCED THE GUIDEWIRE WITH NO RESISTANCE. RN PROCEEDED TO ADVANCE CATHETER AND FELT RESISTANCE ALMOST IMMEDIATELY. BECAUSE GUIDEWIRE WAS ALREADY INSERTED, RN DEPLOYED CATHETER WITHOUT ADVANCING (MOVED HOUSING AWAY FROM PATIENT WHILE DEPLOYING CATHETER). RN KNEW CATHETER WAS NOT IN CORRECT PLACE THEREFORE ATTEMPTED TO REMOVE CATHETER, RESISTANCE WAS MET SO RN STOPPED AND WAITED A FEW MINUTES WHILE ANOTHER RN WENT TO GET HEAT. AFTER A FEW MINUTES RN WAS ABLE TO REMOVE CATHETER WITHOUT HEAT. TIP OF CATHETER DID NOT APPEAR INTACT AND ABOUT 1CM WAS MISSING. XRAY OBTAINED SHOWED ¿ ¿FINDINGS MOST COMPATIBLE WITH SMALL RETAINED FOREIGN BODY WITHIN THE ANTERIOR LATERAL SOFT TISSUES OF THE DISTAL UPPER ARM¿. INTERVENTIONAL RADIOLOGY ATTEMPTED REMOVAL VIA ULTRASOUND AND WAS UNSUCCESSFUL AND CLIPPED VEIN ABOVE CATHETER TIP. PER IR ¿ ¿BASED ON THE FOREIGN BODIES POSITION AND SMALL SIZE OF THE CEPHALIC VEIN AND ASSOCIATED THROMBOSIS I BELIEVE THAT ATTEMPTS AT REMOVAL OF THIS FOREIGN BODY USING A CUTDOWN MAY CAUSE GREATER DAMAGE THAN LEAVING THE FOREIGN BODY IN PLACE. SUSPECT THE CEPHALIC VEIN SURROUNDING THIS FOREIGN BODY WITH THROMBOSIS¿. VASCULAR SURGERY THEN PERFORMED A SUCCESSFUL LIGATION OF CEPHALIC VEIN, EXPLORATION OF CEPHALIC VEIN AND REMOVAL OF LEFT UPPER EXTREMITY FOREIGN BODY FROM A VASCULAR VENOUS STRUCTURE ON 7/18. THIS DID NOT CAUSE ANY DELAY IN TREATMENT." ADDITIONAL INFORMATION FROM CUSTOMER RESPONSE ON 07-26-2023. BREAK WAS SUBCUTANEOUS, ALL PIECES REMOVED FROM PATIENT AND ACCOUNTED FOR. NO IMAGES, FROM POST-OPERATIVE NOTE - ¿WE LIGATED THE VEIN PROXIMALLY AND DISTALLY WITH SEVERAL SILK LIGATURES AND IDENTIFIED THE FOREIGN BODY AND THE VEIN WAS REMOVED EN BLOC AND EXPLORED. THE FOREIGN BODY WAS EXPRESSED FROM THE VEIN AND FELT TO BE CONSISTENT WITH THE LINE TIP¿. THROMBUS HAS NOT RESOLVED, NO ADDITIONAL TREATMENT FOR THROMBUS. EVENT DID NOT INVOLVE AN URGENT/LIFE-THREATENING MEDICAL SITUATION, DID NOT CAUSE A CLINICALLY SIGNIFICANT DELAY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05967 | 2023-08-09 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER TUBE GOT STUCK,(DOE (B)(6) 2023) IT HAD HAPPENED TWICE IN THE LAST MONTH. (DOE UNK) ALSO STATED THAT THEY HAD BEEN USING FOR ALMOST 2 YEAR AND GOT STUCK. PATIENT WAS ELDER AND HARD OF HEARING. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, THE ISSUE HAD BEEN RESOLVED, AND CUSTOMER WAS ABLE TO DRAIN THE URINE PROPERLY. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03438 | 2023-08-09 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F SL CATHETER BASIC TRAY (W/ 70CM NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6175355 |  | IT WAS REPORTED DURING INSERTION WHEN THE NURSE TRIED TO INSERT THE GUIDE INTO THE NEEDLE, IT WOULD NOT GO UP, WITHOUT ANY BACK AND FORTH MOVEMENT. WHEN SHE TRIED TO REMOVE IT, SHE WAS UNABLE TO DO SO, SO SHE HAD TO MAKE A POINT ON THE SKIN WITH A SCALPEL TO REMOVE IT AND FIND THAT A KNOT HAD FORMED AT THE END OF THE GUIDE. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03437 | 2023-08-09 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/ SHERLOCK 3CG TPS STYLET W/ MICROINTRODUCER (4F) (SL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174118 |  | IT WAS REPORTED BY CUSTOMER THAT HOLE IN THE PICC. VAT NURSE ASSESSED AND FOUND A HOLE IN EXTERNAL PORTION OF PICC. PATIENT HAD TO RECEIVE A ULTRASOUND GUIDED PIV AND HAVE EXISTING PICC REMOVED WHILE VAT NURSE CONTACTED PHYSICIAN TO RECEIVE NEW ORDERS. NEW PICC WAS THEN INSERTED THE FOLLOWING DAY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03435 | 2023-08-09 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG 4F SL PICC PERIPHERALLY INSERTED CENTRAL VENOUS CATHETER BASIC | CATHETER, PERCUTANEOUS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7715405 |  | IT WAS REPORTED THAT DURING PLACEMENT OF THE PICC CATHETER, THE HEAD NURSE UNPACKED A NEW PICC CATHETER AND FOUND EVIDENT BENDING AT THE PROXIMAL END OF THE CATHETER. WHEN BEING PRE-FLUSHED, THE CATHETER LEAKED LIQUIDS AT THE SITE OF BENDING. THE PICC CATHETER WAS IMMEDIATELY REPLACED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAKING CATHETER WAS UNCONFIRMED BECAUSE THE PROBLEM COULD NOT BE REPRODUCED. THE PRODUCT RETURNED FOR EVALUATION WAS 4FR S/L GROSHONG CATHETER. THE SAMPLE WAS RECEIVED WITH AN INLAID STYLET. THE SAMPLE APPEARED FREE OF OBVIOUS USE RESIDUES. AN ATTEMPT TO INFUSE WATER THROUGH THE SAMPLE USING A 12ML SYRINGE REVEALED THE SAMPLE TO BE PATENT TO INFUSION AND ASPIRATION WITH NO OBSERVED LEAKS. NO LEAKS WERE OBSERVED DURING PRESSURIZATION. MICROSCOPIC INSPECTION OF THE SAMPLE DID NOT REVEAL EVIDENCE OF CATHETER/STYLET DEFORMATION OR LEAKAGE. NO DEFICIENCIES WERE DISCOVERED DURING EVALUATION OF THE RETURNED SAMPLE. CONSEQUENTLY THIS COMPLAINT IS UNCONFIRMED AT THIS TIME. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11 | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03434 | 2023-08-09 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER W/ SHERLOCK 3CG (TPS) STYLET FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1274108 |  | IT WAS REPORTED BY CUSTOMER THAT PICCS ARE LEAKING. HAD TO CLEAN UP MESS AND REPLACE. NO OTHER INFORMATION WAS PROVIDED. (B)(6) 2024 - THREE DEVICES WERE RETURNED FOR EVALUATION. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF LEAKING CATHETERS IS CONFIRMED AND WAS DETERMINED TO BE USE-RELATED. THREE 4 FR D/L POWERPICC SV CATHETERS WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION REVEALED ABUNDANT USE RESIDUES THROUGHOUT EACH OF THE THREE RETURNED CATHETERS. A SMALL SPLIT WAS OBSERVED ALONG SAMPLE 1, SAMPLE 2 AND SAMPLE 3 EACH ALONG THE CATHETER TUBING JUST DISTAL OF THE MOLDED SUTURE WING. A MICROSCOPIC OBSERVATION REVEALED THE SPLIT LOCATED JUST DISTAL OF THE MOLDED SUTURE JOINT FOR SAMPLE 1 WAS CIRCUMFERENTIALLY ALIGNED WITH ROUNDED, POLISHED EDGES AND ¿C¿ SHAPED IMPRESSIONS LEADING INTO THE SPLIT. THE SPLIT LOCATED JUST DISTAL OF THE MOLDED SUTURE JOINT FOR SAMPLE 2 WAS CIRCUMFERENTIALLY ALIGNED WITH ROUNDED, POLISHED EDGES. THE SPLIT LOCATED JUST DISTAL OF THE MOLDED SUTURE JOINT FOR SAMPLE 3 WAS CIRCUMFERENTIALLY ALIGNED WITH ROUNDED, POLISHED EDGES AND ¿C¿ SHAPED IMPRESSIONS LEADING INTO THE SPLIT. THE CHARACTERISTICS OBSERVED IN EACH SPLIT OF EACH OF THE RETURNED DEVICES COINCIDE WITH THE FEATURES TYPICAL OF FLEXURAL FATIGUE, OR CYCLIC KINKING OF THE CATHETER TUBING CAUSING THE MATERIAL TO WEAR. BASED ON THE MICROSCOPIC OBSERVATIONS OF THE SPLITS FOUND IN EACH OF THE THREE SAMPLES RETURNED FOR EVALUATION, THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED AND WAS DETERMINED TO BE DUE TO FLEXURAL FATIGUE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03433 | 2023-08-09 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 15CM (SHORT-TERM) (DUAL-LUMEN) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5593150 |  | IT WAS REPORTED THAT DURING CATHETER PLACEMENT BY THE OPERATOR ON (B)(6) 2023, THE CATHETER WAS PRE-FLUSHED TO INSPECT THE CATHETER FOR INTEGRITY. IT WAS FOUND THAT THE CATHETER LEAKED LIQUID AND WAS IMMEDIATELY REPLACED WITH A NEW CATHETER FOR OPERATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03432 | 2023-08-09 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER 4F BASIC TRAY (SL) (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6194118 |  | IT WAS REPORTED THE STAFF NOTICED THE LEAKAGE UNDER THE DRESSING OF THE CATHETER. CATHETER WAS REMOVED DUE TO LEAKAGE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03430 | 2023-08-09 | 2023-07-20 | I, n, j, u, r, y | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (18G) (8CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | PCK00005 |  | IT WAS REPORTED BY CUSTOMER THAT POWERGLIDE MIDLINE 18 GAUGE 8CM GUIDE WIRE CAME OFF IN PATIENT. PATIENT WAS SENT TO NEARBY HOSPITAL EMERGENCY ROOM TO HAVE IMAGING TO EVALUATE. ADDITIONAL INFO FROM CUSTOMER: ((B)(6) 2023). PT NEEDED TO BE TRANSFERRED TO ER FOR SCANS. DISCOVERED AFTER MIDLINE PLACEMENT WHEN NEEDLE WAS PULLED OUT OF PATIENT¿S ARM AND THE GUIDE WIRE WAS MISSING. UNKNOWN IF IT BROKEN IN MULTIPLE PIECES. IT WAS NEVER RECOVERED. IMAGES WERE TAKEN IN ER AND NOTHING SHOWED UP ON IMAGES. PATIENT JUST RECEIVED OUTSIDE REFERRAL FOR FURTHER SCANS AND IMAGES IN HER HOMETOWN. PATIENTS WAS ALREADY NERVOUS BEFORE MIDLINE, SO THE EVENT MADE IT EVEN WORSE. NO OTHER NEGATIVE OUTCOMES TO DATE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03428 | 2023-08-09 | 2023-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA CATHETER W/ SOLO VALVE SHERLOCK 3CG (TPS) STYLET (5F) (TL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1395108D |  | IT WAS REPORTED THAT INTRODUCER HAS MALFORMATION ON THE GRAY PART AND DOES NOT ALLOW WIRE TO THREAD. ADD INFO RECEIVED ON (B)(6) 2023. DIRECTLY INVOLVE PATIENT, THE INTRODUCER WOULDN¿T THREAD INTO PATIENT BECAUSE OF THE BURR ON THE END OF IT. THE ATTEMPTS MADE TO THREAD INTRODUCER FAILED AND TRAUMA IN TRYING TO USE IT AS WE NORMALLY WOULD MAY HAVE CAUSED SOME DAMAGE AND I¿M SURE INCREASED THE RISK OF DVT. WE WERE ABLE TO DROP A INTRODUCER KIT ONTO STERILE FIELD AND USE A NEW INTRODUCER WHICH THREAD EASILY INTO PATIENT AND ALLOWED US TO FINISH HIS PICC LINE INSERTION AT BEDSIDE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03429 | 2023-08-09 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | CUSTOMER REPORTED THAT THE PATIENT ATTENDED OP DEPARTMENT FOR REVIEW OF PICC AND ANTIBIOTIC ADMINISTRATION AS ARM SWELLING NOTED WITH PREVIOUS ADMINISTRATION. PICC ASSESSED AND TREATMENT COMMENCED. FLUID LEAKAGE FROM INSERTION SITE NOTED. FRACTURE SUSPECTED AND PICC REMOVED. SPLIT IN PICC AT 5.5CM. PICC SECURED WITH SUBCUTANEOUS ANCHOR ON THE TAPERED END. 4FR USED OVER LARGER TAPERED END OF CATHETER. ADDITIONAL INFO 07/25/2023: THE ONLY ISSUE FOR THE PATIENT WAS THAT THEIR TREATMENT WAS DELAYED UNTIL A NEW PICC WAS PLACED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED AND WAS DETERMINED TO BE USE-RELATED. ONE 4 FR S/L POWERPICC WAS RETURNED FOR EVALUATION. A MICROSCOPIC OBSERVATION REVEALED A LONGITUDINALLY ALIGNED SPLIT ALONG THE EXTRUSION LINE OF THE CATHETER BETWEEN THE 5 CM AND 6 CM DEPTH MARKER. A GRANULAR FRACTURE SURFACE WAS OBSERVED. BECAUSE THE FRACTURE WAS OBSERVED AT THE DEFINED KINK NEAR THE INSERTION SITE OF THE CATHETER, THE COMPLAINT IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03427 | 2023-08-09 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F DUAL-LUMEN CATHETER BASIC TRAY (W/ 70CM NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6275355 |  | IT WAS REPORTED THAT DURING VISIT TO VAT INCIDENTALLY MENTIONED A PICC WAS GIVEN TO THEM IN A BIO BAG. PICC HAD A LABEL ON ONE OF THE LUMENS SAYING HOLE WAS NOTED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03426 | 2023-08-09 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY CUSTOMER THAT THE CATHETER LEAKED FROM THE RUBBER-STOPPER WHEN PICC LINE WAS BEING INSERTED. NO PATIENT HARM. RESPONSE RECEIVED 26 JULY 2023. THE EVENT DID NOT INVOLVE AN URGENT OR LIFE THREATENING MEDICAL SITUATION, AND THERE WAS NO INTERRUPTION IN TREATMENT. THE DEVICE WAS SECURED WITH A STATLOCK. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03423 | 2023-08-09 | 2023-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED BY CUSTOMER THAT THE INTRODUCER WOULD NOT ADVANCE. WHEN IT WAS REMOVED AND INSPECTED, A NOTCH WAS NOTED WHERE THE INNER AND OUTER INTRODUCER SEGMENTS JOIN. THE CUSTOMER STATES THIS IS THE SECOND TIME THIS HAS HAPPENED. ADDITIONAL INFORMATION RECEIVED: IT WAS STATED THE EVENTS DIRECTLY AFFECTED PATIENTS, HAD TO FIND ALTERNATE INTRODUCER. THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION, ROUTINE INSERTION. THIS REPORT ADDRESSES THE SECOND EVENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03424 | 2023-08-09 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED," CALLER STATES SHE HAD A PICC PLACED FOR CANCER TREATMENTS IN 2022 AND NEEDS TO KNOW HOW LONG HER PICC CAN STAY IN. CALLER CONFIRMS HER PICC IS WORKING FINE AND DOES GET A BLOOD RETURN BEFORE SHE STARTS HER TPN. SHE STATES THE IR NURSE SAID THEY REPLACE PICS EVERY THREE MONTHS. SHE STATES IN THE PAST THE NURSES HAD TO USE TPA TO GET IT WORKING AGAIN BUT THEY WERE ABLE TO GET A BLOOD RETURN. RESPONSE: EXPLAINED THAT BD DOES NOT OFFER A DWELL TIME FOR PICCS. THE PICC SHOULD BE REMOVED WHEN CLINICALLY INDICATED, NOT BASED SOLELY ON LENGTH OF DWELL TIME BECAUSE THERE IS NO KNOWN OPTIMAL DWELL TIME. AS LONG AS THE PICC IS NEEDED FOR TREATMENT, BEING FLUSHED ROUTINELY, WORKING AS INTENDED AND SHOWS NO SIGNS OF INFECTION OF COMPLICATIONS, IT CAN REMAIN INDEFINITELY. CATHETER REMOVAL IS RECOMMENDED IF THE PATIENT DEVELOPS PHLEBITIS, INFECTION OR HAS A MALFUNCTIONING CATHETER." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05945 | 2023-08-09 | 2023-08-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON BURST, AND WATER LEAK OCCURRED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03420 | 2023-08-09 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | IT WAS REPORTED PICC WAS PLACED ON (B)(6) 2023 FOR AN SURGICAL PATIENT AND IV ANTIBIOTICS. THE CATHETER RUPTURED AT 10CM. PATIENT, IS NOW AT ANOTHER HOSPITAL AND RECEIVED TREATMENT THERE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03417 | 2023-08-09 | 2023-07-24 | I, n, j, u, r, y | Manufacturer report | BASIC KIT 22G X 8 CM WITH BIOPATCH | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F322087T |  | IT WAS REPORTED BY CUSTOMER THAT 22G X 8CM POWERGLIDE HAD BROKEN OFF IN A PATIENT. CATHETER WAS IDENTIFIED AND PATIENT IS SET TO HAVE IT REMOVED. ACCORDING TO CUSTOMER THE LINE WAS ATTEMPTED TO BE RE-CANULATED AFTER MEETING RESISTANCE WITH CATHETER ADVANCEMENT PASS THE WIRE. ADDITIONAL INFORMATION RECEIVED ON 27 JULY 2023: THE CATHETER BREAK WAS DISCOVERED DURING DEVICE REMOVAL. 2/3S OF THE CATHETER WAS RETAINED IN THE PATIENT. PATIENT REPORTED PAIN AT THE SITE. X-RAYS WERE TAKEN SHOWING THE RETAINED CATHETER AND AN EXTENSION LOOP FROM ANOTHER IV THAT WAS IN PLACE ABOVE THE CATHETER SITE. PATIENT HAD TO GO TO VASCULAR SURGERY TO HAVE THE DEVICE REMOVED, WHICH WAS SUCCESSFUL. PATIENT DID NOT MISS ANY INFUSIONS DUE TO HAVING ADDITIONAL ACCESS IN PLACE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03422 | 2023-08-09 | 2023-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED BY CUSTOMER THAT THE INTRODUCER WOULD NOT ADVANCE. WHEN IT WAS REMOVED AND INSPECTED, A NOTCH WAS NOTED WHERE THE INNER AND OUTER INTRODUCER SEGMENTS JOIN. ADDITIONAL INFORMATION RECEIVED: IT WAS STATED THE EVENT DIRECTLY AFFECTED PATIENT, HAD TO FIND ALTERNATE INTRODUCER. THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION, ROUTINE INSERTION. ADDITIONAL INFORMATION RECEIVED 17 AUGUST 2023: THE PATIENT WAS IN INTENSIVE CARE INTUBATED, SEDATED, INFUSING VASOPRESSORS AND IN NEED OF CENTRAL LINE ACCESS. THE DEFECT WAS DISCOVERED UPON TRYING TO INSERT INTRODUCER AND MEETING SIGNIFICANT RESISTANCE, ALMOST RECOILING. INTRODUCER ADVANCED THROUGH SKIN, BUT UNABLE TO ACCESS VEIN. NO HARM CAME TO PATIENT, THIS DID NOT DELAY CARE, NO NEGATIVE IMPACT NOTED. A NEW INTRODUCER KIT WAS UTILIZED AND FURTHER PICC PLACEMENT WAS SUCCESSFUL. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DIFFICULT INSERTION IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. THREE PHOTO SAMPLES OF A MICROEZ MICROINTRODUCER WERE PROVIDED FOR EVALUATION. THE FIRST IMAGE SHOWED THE DISTAL END OF THE DILATOR AND INTRODUCER SHEATH. DEFORMATION ON THE SHEATH WAS OBSERVED. THE SECOND PHOTO SHOWS ANOTHER ANGLE OF THE DAMAGED REGION. INWARD BUNCHING OF THE SHEATH TIP WAS PRESENT. THE THIRD IMAGE IS OF A PRODUCT STICKER LABEL INDICATING LOT: REHN3194. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE ADVANCEMENT AGAINST RESISTANCE, STEEP INSERTION ANGLE, AND INADEQUATE SKIN NICK. SINCE BUNCHING OF THE SHEATH AND EVIDENCE OF A DIFFICULT INSERTION WAS NOTED, THE COMPLAINT IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05939 | 2023-08-09 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT DURING THE PRETEST, THE STERILE WATER DID NOT COMPLETELY DRAIN OUT , SO IT WAS REFRAINED FROM USE. AFTER IT WAS DRAINED, THE BALLOON PORTION SEEMED TO HAVE HAD A CUFF ROLL (WRINKLED). 10 ML OF WATER WAS ASPIRATED BY SYRINGE, AND THE BALLOON WAS TRIED TO BE DEFLATED BY SPONTANEOUS OR GENTLE DEFLATION BY SYRINGE. USUALLY, THE BALLOON CAN BE DEFLATED IN AROUND 5 SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05940 | 2023-08-09 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT DURING THE PRETEST, THE STERILE WATER COULD NOT BE DRAINED WHEN THERE WAS ABOUT 2 TO 3 CC LEFT, WHICH WAS REMOVED LATER. 10 ML OF WATER WAS ASPIRATED BY SYRINGE, AND THE BALLOON WAS TRIED TO BE DEFLATED BY SPONTANEOUS OR GENTLE DEFLATION BY SYRINGE. USUALLY, A BALLOON CAN BE DEFLATED WITHIN A FEW SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05934 | 2023-08-09 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT AFTER SURGERY, A FOLEY CATHETER WAS PLACED FOR CONTINUOUS BLADDER IRRIGATION AND URINE DRAINAGE. AFTER OPENING THE CATHETER PACKAGE, IT WAS FOUND THAT THE CATHETER TIP WAS ABNORMAL OR DEFORMED (POSSIBLY DUE TO MOLD PROBLEMS). THE REPLACEMENT OF A NEW CATHETER AND THE INDWELLING DID NOT HARM THE PATIENT BUT AFFECTED THE OPERATION PROCESS. PER ADDITIONAL INFORMATION VIA EMAIL ON 06AUG2023, IT WAS STATED THAT THE TIP OF THE CATHETER DEFORMED. IT WAS STATED THAT THE THERE WAS NO IMPACT TO THE PATIENT AND THERE WAS NO FOREIGN MATERIAL PRESENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03408 | 2023-08-09 | 2022-05-10 | I, n, j, u, r, y | Manufacturer report | POWERPORT DUO M.R.I. IMPLANTABLE PORT, CHRONOFLEX DUAL-LUMEN, 9.5F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1829500 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT TWO MONTHS AND EIGHTEEN DAYS POST PORT PLACEMENT VIA RIGHT INTERNAL JUGULAR VEIN, THE PORT ALLEGEDLY RUPTURED AND CHEMOTHERAPY MEDICATION HAD BEEN LEAKED INTO PATIENT BODY. IT WAS FURTHER REPORTED THAT THE PATIENT HAD SWELLING AT THE PORT SITE AND THE DEVICE WAS UNABLE TO BE REMOVED FROM THE PATIENT BODY. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. MEDICAL RECORDS WERE PROVIDED AND REVIEWED. POST PORT PLACEMENT PATIENT REPORT BURNING ABOVE RIGHT CHEST PORT-A-CATH SITE WITH SOME SWELLING NOTED AT SITE. IT WAS INFORMED TO THE PATIENT THAT CATHETER LUMEN HAD RUPTURED. AROUND SOME DAYS LATER, REMOVAL OF EXISTING PORT WAS PERFORMED, AND PLACEMENT OF NEW PORT WAS INDICATED DUE TO MALFUNCTIONING PORT. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE. HOWEVER, THE INVESTIGATION IS KEPT INCONCLUSIVE FOR THE REPORTED FLUID LEAK AND DIFFICULT TO REMOVE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL CONDITION ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. BASED UPON THE PROVIDED INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: (12/2022), G3, H6 (METHOD) H11: B5, H6 (RESULT, CONCLUSION) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03402 | 2023-08-09 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT DUO M.R.I. IMPLANTABLE PORT, CHRONOFLEX DUAL-LUMEN, 9.5F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1829500 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE CHAMBER WAS ALLEGEDLY DISLODGED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
2020394-2023-00572 | 2023-08-09 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | RECON | SUPPORT CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | SCCTO110T |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE IN THE TIBIOPERONEAL TRUNK, THE SUPPORT CATHETER WAS ALLEGEDLY FOUND TO BE LODGED AND WOULD NOT ADVANCE TO REMOVE FROM THE PATIENT¿S PERONEAL ARTERY WITH THE CATHETER. IT WAS FURTHER REPORTED THAT THE TIP OF THE SUPPORT CATHETER ALLEGEDLY GOT FRAYED AND A PART GOT STUCK IN THE PATIENT'S VESSEL WALL. REPORTEDLY, THE CATHETER WAS REMOVED WITH DIFFICULTY AND A DISSECTION WAS OCCURRED TO REMOVE THE DEVICE AND ADDITIONAL PROCEDURE TO TAPER AND CLOSE THE LESION WAS REQUIRED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | 8.16 H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 09/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
2020394-2023-00574 | 2023-08-09 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CQF7586 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY FAILED TO DEFLATE. IT WAS FURTHER REPORTED THAT THE BALLOON ALLEGEDLY GOT STUCK INSIDE THE VESSEL AND WAS NOT EASILY REMOVED THROUGH THE SHEATH. REPORTEDLY, THE BALLOON WAS FORCED OUT THROUGH THE SHEATH WHILE BEING INFLATED AND REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS AND A VIDEO WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 08/2025). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05924 | 2023-08-09 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THERE WAS A DIFFICULTY IN REMOVAL OF THE FOLEY CATHETER HAS OCCURRED IN THE PAST. DETAILED INFORMATION AND THE DATE OF EVENT OCCURRENCE WERE UNKNOWN. THE HOSPITAL REQUESTED TO KNOW THE RECOMMENDED DIAMETER OF GW TO BE USED FOR TROUBLESHOOTING. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6) 2022, CLARIFIED THAT THE GW MENTIONED WAS GUIDEWIRE AND THE DEVICE WAS USED ON THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE THIN RUBBERIZE WALL THAT CAUSING COLLAPSE/PINCH INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT¿ AND MIGHT BE CONTRIBUTED BY USER RELATED, EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® I.C. SILVER FOLEY TRAY B SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.]. 2. APPLICABLE PATIENTS: (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX. (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING. THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR. 1. FOLEY CATHETER. 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.). SYRINGE PREFILLED WITH STERILE WATER, WATER SOLUBLE LUBRICANT, ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION, TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS, GLOVES, STATLOCK® FOLEY, [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE, THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE: (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD: 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE DIFFICULTY OR FAILURE TO REMOVE THE DEVICE OCCLUSION OF CATHETER INNER LUMENS ENCRUSTATION ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS: URINARY-TRACT INFECTION, HEMORRHAGE, HEMATURIA, ALLERGY REACTION TO THE DEVICE, CALCULUS FORMATION, EDEMA, PAIN, DISCOMFORT, INJURY OF BLADDER OR URETHRAL, URETHRITIS, URINARY INCONTINENCE, RETAINED BALLOON FRAGMENTS, [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05922 | 2023-08-09 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0174L18 |  | IT WAS REPORTED THAT A FOLEY CATHETER WAS INSERTED IN THE THEATER. ON RETURN TO THE WARD, THE PATIENT REPORTED THAT THE CATHETER FELL OUT WHILE IN THE BATHROOM. UPON INSPECTION, THE BALLOON WAS NOTED TO HAVE POPPED. THE USER WAS UNSURE IF RESIDUAL PRODUCT WAS STILL IN THE PATIENT'S BLADDER OR FLUSHED IN THE TOILET. THE TREATING TEAM WAS AWARE OF THE SITUATION. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 21JUL2023, STATED THAT THE PATIENT WAS OKAY AS PER REPRESENTATIVE'S UNDERSTANDING. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 26JUL2023, THE PRODUCT COMPLAINT STATED THAT THE CUSTOMER WAS UNSURE IF THE RESIDUAL PIECES OF BALLOON WERE STILL IN THE PATIENTS BLADDER OR IF THEY HAD PASSED INTO THE TOILET. THE TREATING TEAM WERE MADE AWARE AND AS PER THE REPRESENTATIVE'S UNDERSTANDING, NO MEDICAL INTERVENTION WAS MADE. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. BASED ON THE ATTACHED PHOTO, IT WAS OBSERVED THAT THE BALLOON WAS RUPTURE WITH MISSING PIECES. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. NO FURTHER TESTING COULD BE CONDUCTED DUE TO ONLY PHOTO SAMPLE IS PROVIDED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT/ EXPOSURE TO PETROLATUM BASED PRODUCTS MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05921 | 2023-08-09 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THERE WAS NO URINE DRAINED AFTER USING THE INDWELLING CATHETER, LEADING TO URINARY RETENTION IN THE PATIENT. IMMEDIATE EMERGENCY TREATMENT BY CLINICIANS WAS REQUIRED. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE PATIENT DID NOT HAD ANY IMMEDIATE EMERGENCY TREATMENT OR ANY MEDICAL INTERVENTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: SALT ACCUMULATION)/BLOCK DRAINAGE LUMEN/NO DRAINAGE EYE.) THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05925 | 2023-08-09 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE TEMPERATURE DISPLAY HAD FAILED ON TEMPERATURE SENSING FOLEY CATHETER. IT WAS STATED THAT DURING A LAPAROSCOPIC SURGERY, THE TEMPERATURE DISPLAY SHOWED A LOW OF 34°C OR DISAPPEARED. FOR THIS REASON, THEY WERE ASKED TO INVESTIGATE IF THERE WAS A PROBLEM ON THE CATHETER SIDE. THE RELAY CABLE COULD NOT BE RETRIEVED, BUT THEY HAD IT CHECKED BY THE CALIBRATOR AND FOUND NO RELAY CABLE PROBLEM. THE MONITOR WAS ALSO CHECKED, AND THEY FOUND NO PROBLEM. THE CLIENT REQUESTED TO CONFIRM THE PROBLEM BECAUSE THE SAME PROBLEM HAD BEEN OCCURRING SINCE LAST YEAR. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA EMAIL ON 24JUL2023, CLARIFIED THAT THE DEVICE WAS USED ON THE PATIENT AND THE DATE OF OCCURRENCE IN THE PAST CASES WAS NOT SPECIFIED BY THE CUSTOMER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ASSEMBLY TECHNIQUE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS : 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS : 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER SHAPE, CONFIGURATION AND PRINCIPLES: BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL : BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. SIZES OF CATHETER : AVAILABLE IN SIZES 14FR, 16FR 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY INTENDED USE & EFFECT- EFFICACY : THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. DIRECTIONS FOR USE : 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT : 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING : CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS : 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING : WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) MRI SAFETY : 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3.MALFUNCTION AND ADVERSE EVENTS : 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE : 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05932 | 2023-08-09 | 2023-07-26 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WAS HAVING ISSUES WITH WAKING UP WET FROM USING THE PUREWICK FEMALE EXTERNAL CATHETER. REPRESENTATIVE GAVE PLACEMENT TIPS AND THE MACHINE PASSED. PER THE FOLLOW UP VIA PHONE (B)(6) 2023, THE CUSTOMER HAD BEEN EXPERIENCING REOCCURRING URINARY TRACT INFECTIONS WITH USE OF THE PUREWICK. PATIENT STATED PATIENT TAKES AMOXICILLIN DAILY FOR THE URINARY TRACT INFECTIONS. CUSTOMER ALSO STATED THAT PATIENT WAS STILL EXPERIENCING SUCTION AND LEAKING ISSUES AFTER TROUBLESHOOTING WITH LIBERATOR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-08
1018233-2023-05931 | 2023-08-09 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT IMMEDIATELY AFTER CATHETER PLACEMENT TO THE PATIENT IN THE OPERATING ROOM, THE CATHETER FELL OUT DUE TO STERILE WATER LEAK WHICH COULD HAVE BEEN DUE TO A HOLE IN THE BALLOON. AS PER THE NOTIFICATION RECEIVED ON 10JUL2023, THERE WAS BALLOON BURST OBSERVED IN THE RETURNED SAMPLE. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED SOME SCRATCHED/ABRASIONS ON THE BALLOON, BUT NO SIGNS OF BALLOON BURST OR RUPTURE. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND IT IMMEDIATELY LEAKED OUT OF 0.012" PINHOLE FOUND ON THE BALLOON. THIS IS OUT OF SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER . WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS. GAUZE PADS. WATERPROOF SHEET. COTTON BALLS. GLOVES. STATLOCK FOLEY. [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05933 | 2023-08-09 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE WATER WAS DIFFICULT TO REMOVE FROM THE FOLEY CATHETER. DURING A PRETEST IN THE OPERATING ROOM, THE STERILE DISTILLED WATER WAS NOT REMOVED FROM THE BALLOON. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND SYRINGE. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05935 | 2023-08-09 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® PEDIATRIC ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165808 |  | IT WAS REPORTED THAT 3 OF THE CATHETERS SEEMED TO BE DEFECTIVE. ALSO STATED THAT IN 2 OF THE CATHETERS, THE BALLOON WOULD NOT STAY INFLATED, AND THE 3RD CATHETER SEEMED TO BE CLOGGED SOMEHOW BUT NOTHING WAS COMING OUT. THEY HAD SOME INTERMEDIATE CATHETERS SO THEY WERE STILL ABLE TO VOID, BUT THESE FOLEY CATHETERS WERE NOT WORKING LIKE THEY SHOULD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05936 | 2023-08-09 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE FOLEY CATHETER BALLOON DURING THE PRETEST. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. VISUAL INSPECTION NOTED IRREGULAR BURST WITHOUT MISSING PIECES. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE CONTRIBUTE BY "HIGH MECHANICAL STABILITY TIME (MST) OF INCOMING LATEX" OR "COAGULANT DIP 2- DIPPING SPEED TOO SLOW" AND ALSO MIGHT BE CONTRIBUTED BY USER RELATED ( EXAMPLE: MISHANDLING PRODUCT OR EXPOSED TO PETROLEUM OR SHARP OBJECT). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D,F,H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05928 | 2023-08-09 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE PATIENT HAD NOT BEEN PUTTING OUT ANY URINE. THEY ATTEMPTED TO MANIPULATE THE FOLEY CATHETER. IT WAS NOTICED THAT THE CATHETER WAS NO LONGER IN THE PATIENT. THE BALLOON WAS NOT INFLATED AND THE TOP ON THE CATHETER WAS PARTIALLY DETACHED. ALL PARTS OF THE CATHETER DID COME OUT OF THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿MOLDED IN PLACE, NOT GLUED". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED DIRECTIONS FOR USE WASH HANDS AND DON CLEAN GLOVES EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERIURETHRAL AREA 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL 5. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP 6. DON STERILE GLOVES 7. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE 8. POSITION FENESTRATED DRAPE ON PATIENT 9. SATURATE 3 FOAM SWAB STICKS IN POVIDONE IODINE 10.ATTACH THE WATER FILLED SYRINGE TO THE INFLATION PORT NOTE: IT IS NOT NECESSARY TO PRE-TEST THE FOLEY CATHETER BALLOON 11. REMOVE FOLEY CATHETER FROM WRAP AND LUBRICATE CATHETER 12. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA INORA MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD 13. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND A. WHEN CATHETER TIP HAS NTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE B. INSERT CATHETER TWO MORE INCHES AND INFLATE CATHETER BALLOON 14. INFLATE CATHETER BALLOON USING ENTIRE 10CC OF STERILE WATER PROVIDED IN THE PREFILLED SYRINGE NOTE: USE OF LESS THAN 10CC CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON 15. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK 16. SECURE THE FOLEY CATHETER TO THE PATIENT USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED (SEE STATLOCK® FOLEY STABILIZATION DEVICE IFU) NOTE: PLEASE MAKE SURE PATIENT IS APPROPRIATE FOR USE OF STATLOCK® STABILIZATION DEVICE 17. POSITION HANGER ON BED RAIL AT THE FOOT OF THE BED NOTE: EXERCISE CARE TO KEEP BAG OFF THE FLOOR 18. USE GREEN SHEETING CLIP TO SECURE DRAINAGE TUBE TO THE SHEET. MAKE SURE TUBE IS NOT KINKED 19. INDICATE TIME AND DATE OF CATHETER INSERTION ON PROVIDED LABELS. PLACE DESIGNATED LABELS ON PATIENT CHART AND DRAINAGE SYSTEM 20. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL FOLEY CATHETER REMOVAL 1. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE 2. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY 3. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION 4. IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL 5. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT¿. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
MW5122891 | 2023-08-09 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARDEX FOLEY CATHETER | CATHETER, UROLOGICAL (ANTIMICROBIAL) AND ACCESSORIES | BARD SDN. BHD. | 01658116 |  |  | FOLEY CATHETER SYRINGE USED TO INFLATE THE BALLOON BROKE WHILE ATTEMPTING TO REMOVE SYRINGE FROM THE PORT AND THE FOLEY CAME BACK OUT OF THE PATIENT'S BLADDER. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05972 | 2023-08-10 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE CATHETER FELL OUT SPONTANEOUSLY DURING RETENTION, AND IT WAS FOUND THAT THE CATHETER BALLOON HAD BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03396 | 2023-08-10 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC 4F S/L 3CG BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 22184118 |  | CUSTOMER REPORTED THERE WAS A FISSURE 8 CM FROM THE EXIT SITE ON THE PICC. FOLLOWING A MALFUNCTION OF THE DEVICE, THE OPERATOR DECIDES TO REMOVE IT AND FINDS A FISSURE OF THE CATHETER 8CM FROM THE EXIT SITE. THE OPERATOR EXPLAINS THE INCIDENT TO US AND SHOWS US THE AFFECTED CATHETER, SHOWING US THE DAMAGE. DEVICE REMOVED, NO PATIENT HARM REPORTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05993 | 2023-08-10 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119216M |  | IT WAS REPORTED THAT THE CUSTOMER EXPERIENCED AN ISSUE WITH A 16FR TEMP SENSING FOLEY CATHETER IN WHICH THE BALLOON DEFLATED IN THE PATIENT AND THE CATHETER FELL OUT. WHEN INSPECTING THE ITEM, THE BALLOON WOULD NOT STAY INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05970 | 2023-08-10 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 33422 |  | IT WAS REPORTED THAT THE CUSTOMER HAD MULTIPLE ISSUES WITH THE FOLEY CATHETER BALLOON. THE BALLOON HAD BURST INSIDE THE PATIENT SEVERAL TIMES. THE BALLOON ALSO DID NOT BLOW UP ON SOME OCCASIONS. THE BALLOON WAS ALSO LEAKING. THIS HAD HAPPENED WITH ONE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05984 | 2023-08-10 | 2023-07-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 998004 |  | IT WAS REPORTED THAT WHEN UNPACKING THE BALLOON DILATION CATHETER DURING OPERATION, THE PUMP WAS FOUND RUPTURED AND DAMAGED. THIS CAUSED NO ADVERSE EFFECT ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05977 | 2023-08-10 | 2023-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER HAD BEEN USING FOLEY CATHETER EVERYDAY FOR LAST TWO AND HALF YEARS. THE PATIENT WAS A PATIENT OF HYPOTONIC URINARY BLADDER AND REQUIRED SELF CATHETERIZATION AT LEAST TWICE A DAY. BUT THE CUSTOMER WAS OBSERVING THAT OVER LAST ONE TO TWO MONTHS, THE QUALITY OF THE CATHETERS WERE NOT UP TO THE MARK, REQUIRING REPLACEMENT AFTER USING 2 TO 3 TIMES ONLY UNLIKE PREVIOUSLY WHEN EACH CATHETER COULD BE USED 6 TO 7 TIMES. THE CATHETERS WERE OF SIZES 16 FR AND 18 FR. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE HIGH DRYING TEMPERATURE AND LONG DRYING TIME. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05978 | 2023-08-10 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 |  | A304916 |  | IT WAS REPORTED THAT 3CCS OF WATER DISTRIBUTED INSTEAD OF 10CCS FROM FOLEY CATHETER. PER FOLLOW UP VIA EMAIL ON (B)(6) 2023, CUSTOMER JUST WANTED TO LET YOU KNOW OUR FINDINGS WHEN THEY OPENED THE STERILE COUDE CATHETER KIT. CUSTOMER WAS TOLD FROM YOUR COMPANY THAT PERHAPS THE MANUFACTURER PLACED A PEDIATRIC SYRINGE OF 3ML STERILE WATER INSIDE THE ADULT KIT. CUSTOMER ASKED DO YOU FILL YOUR PEDIATRIC NS SYRINGES WITH JUST 3ML. THEY ALSO ASKED IS IT EVEN POSSIBLE THAT THE WATER COULD EVAPORATE WITHIN A CLOSED, STERILE SYRINGE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05979 | 2023-08-10 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE WATER WAS DIFFICULT TO REMOVE FROM THE FOLEY CATHETER. DURING A PRETEST IN THE OPERATING ROOM, THE STERILE DISTILLED WATER WAS NOT REMOVED FROM THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03468 | 2023-08-10 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8174118 |  | IT WAS REPORTED THAT "PATIENT HAD AN USG GUIDED PICC LINE (POWER PICC SINGLE LUMEN 4 FR) WAS INSERTED ON 22.06.23 IN RT BASILIC VEIN INSIDE WAS-39 CM AND OUTSIDE-3CM KEPT. X-RAY DONE ADVISED TO DO EVERY AFTER 7 DAYS DRESSING. PATIENT CAME IN OPD FOR PICC LINE DRESSING ON (B)(6) 2023 BEFORE REMOVING THE OLD DRESSING THERE IS BLOOD SOAKAGE IN THE GAUZE AND WHILE FLUSHING THE LINE THERE IS LEAKAGE FROM THE INSERTION SITE , INFORMED DOCTOR. ADVISED TO REMOVE THE LINE. SISTER WAS TRYING TO REMOVE THE LINE BUT THERE IS TIGHTNESS WHILE PULLING THE LINE. SO, INFORMED THE CATH LAB DOCTOR TO REMOVE THE LINE. USG GUIDED NEW PICC LINE INSERTED ON (B)(6) 2023." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05980 | 2023-08-10 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT DURING CATHETER REMOVAL, RESISTANCE WAS FELT DUE TO A MUSHROOMED BALLOON. IT WAS STATED THAT WHEN THE USER TRIED TO REMOVE THE CATHETER ON THE DAY AFTER THE OPERATION, SINCE THE MUSHROOMED BALLOON WAS CAUGHT, IT WAS REMOVED BY USING A LUBRICANT. IT WAS ALSO STATED THAT THE 3 SIMILAR INCIDENTS OCCURRED IN THE LAST ABOUT 1 MONTH. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05988 | 2023-08-10 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L20 |  | IT WAS REPORTED THAT THE CATHETER'S INFLATION VALVE CAME OFF WHEN INJECTING STERILE WATER. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, CLARIFIED THAT THE DEVICE WAS USED ON THE PATIENT AND IT WAS NOT A FITTING ISSUE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05975 | 2023-08-10 | 2023-07-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT WHEN THE BALLOON WAS INFLATED AFTER THE INSERTION OF THE CATHETER, STERILE WATER LEAKED FROM THE CONNECTION BETWEEN THE SYRINGE AND THE INFLATION VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05994 | 2023-08-10 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE CONNECTION BETWEEN THE VALVE AND SYRINGE WHEN THE CATHETER WAS PRETESTED TO BE PLACED FOR URINARY DRAINAGE IN AN OUTPATIENT DEPARTMENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE POOR SAMPLE CONDITION. THE RETURNED CATHETER WAS FUNCTIONALLY TESTED. RELEVANT MEASUREMENT WERE TAKEN, AND THE PRODUCT WAS FOUND TO BE WITHIN THE CURRENT SPECIFICATIONS. FURTHER FUNCTION TESTING COULD NOT BE PERFORMED DUE TO AFFECTED SYRINGE WAS NOT RETURNED FOR EVALUATION AND THE CONDITION OF THE RETURNED SAMPLE PREVENTED THE ABILITY TO DETERMINE CONFORMANCE TO APPLICABLE SPECIFICATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO CRACK/CRUSH/DEFORMED VALVE CAUSED BY INCORRECT AIR PRESSURE OR DUE TO DEFECTIVE COMPONENT FROM SUPPLIER. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05998 | 2023-08-10 | 2023-07-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 73018L |  | IT WAS REPORTED THAT CUSTOMER USED THREE FOLEY CATHETERS AND EACH HAD A SPLIT NEAR THE THREE-WAY CONNECTOR. STATED THAT THE PRODUCT DID NOT WORKED AS INTENDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03460 | 2023-08-10 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1618000 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, PIECES FROM THE PEEL-AWAY SHEATH WERE ALLEGEDLY BROKEN OFF. IT WAS FURTHER REPORTED THAT AS THE PHYSICIAN BEGAN TO PEEL THE SHEATH, LITTLE FRAGMENTS OF THE GRAY OR PLASTIC PART FROM THE PEEL-AWAY WERE LEFT INSIDE THE PATIENT'S BODY. REPORTEDLY, THE FRAGMENTS WERE REMOVED FROM THE PATIENT'S BODY USING SURGICAL INSTRUMENTS. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 11/2024) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03471 | 2023-08-10 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1608062 |  | IT WAS REPORTED THAT APPROXIMATELY TWO MONTHS POST PORT PLACEMENT ON THE LEFT SIDE, THE PATIENT ALLEGEDLY EXPERIENCED SORENESS WHEN SLEEPING, AND THE PORT DID NOT SEEM TO BE HEALING AS IT SHOULD. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05997 | 2023-08-10 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT DURING THE PRETEST, THE STERILE WATER DID NOT COMPLETELY DRAIN OUT, SO IT WAS REFRAINED FROM USE. AFTER IT WAS DRAINED, THE BALLOON PORTION SEEMED TO HAVE HAD A CUFF ROLL (WRINKLED). 10 ML OF WATER WAS ASPIRATED BY SYRINGE, AND THE BALLOON WAS TRIED TO BE DEFLATED BY SPONTANEOUS OR GENTLE DEFLATION BY SYRINGE. USUALLY, THE BALLOON CAN BE DEFLATED IN AROUND 5 SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05995 | 2023-08-10 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI12 |  | IT WAS REPORTED THAT THE PRE-TEST WAS FINE, BUT THE BALLOON RUPTURED AFTER CATHETER INSERTION. TWO CASES FOLLOWED THE SAME PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05999 | 2023-08-10 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THERE WAS URINE LEAKAGE NEAR THE CATHETER AND CONNECTING TUBE DURING INDWELLING. THERE WAS LEAKAGE OF URINE DUE TO A CRACK IN THE CATHETER. THERE WAS NO BALLOON RUPTURE OR TEAR. THE WATER LEAKED FROM THE BIFURCATION AREA. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06002 | 2023-08-10 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETERS WERE NORMALLY 26 CM IN LENGTH. HOWEVER, THE LAST 10 CATHETERS THE PATIENT HAD RECEIVED WERE SHORTED AND WERE ONLY ABOUT 20CM. AS PER ADDITIONAL INFORMATION STATED ON PHOTO SAMPLE ATTACHED, BOTH FEMALE ONE WAS 20CM AND THE OTHER WAS 26CM. THERE WAS ABOUT 2 INCH DIFFERENCE BOTH ON THE SAME ORDER CODE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-08
1018233-2023-06001 | 2023-08-10 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER TUBE GOT STUCK,(DOE 24JUL2023) IT HAD HAPPENED TWICE IN THE LAST MONTH. (DOE UNK) ALSO STATED THAT THEY HAD BEEN USING FOR ALMOST 2 YEAR AND GOT STUCK. PATIENT WAS ELDER AND HARD OF HEARING. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, THE ISSUE HAD BEEN RESOLVED, AND CUSTOMER WAS ABLE TO DRAIN THE URINE PROPERLY. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-08
1018233-2023-06003 | 2023-08-10 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE INFLATION VALVE OF THE FOLEY CATHETER WAS BROKEN. AFTER INSERTING THE CATHETER AND INJECTING DISTILLED WATER, THE GREEN VALVE CAP CAME OFF, STERILE WATER CAME OUT, CAUSING THE CATHETER TO COME OFF. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06004 | 2023-08-10 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE FOLEY CATHETER. THERE WAS A SCRATCH (A LITTLE CRACK) NEAR THE VALVE. STATED THAT THE WATER PROBABLY LEAKED FROM THERE. THERE WAS A SCRATCH FROM THE BEGINNING, EVEN THOUGH IT WAS NOT PULLED PARTICULARLY HARD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03473 | 2023-08-10 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC SOLO 5FR D/L, 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 22295118 |  | IT WAS REPORTED THAT DURING INSERTION OF PICC LINE THE PEEL AWAY SHEATH SPLIT AT THE DISTAL END DURING DILATION OF THE VESSEL. NO PATIENTS OR HEALTHCARE WORKERS WERE INJURED AS A RESULT OF THE ISSUE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03472 | 2023-08-10 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT WHILE HCP OPENED GROSHONG 4F CATHETER TO PLACE IN A PATIENT, HE SAW TRACTION IN CATHETER WHILE ASSEMBLING FOR PROCEDURE AND HENCE HE USED OTHER FRESH PACK OF GROSHONG 4F CATHETER AND COMPLETED PROCEDURE. 07/26/2023 IT WAS FOUND DURING AN INVESTIGATION THAT THERE IS A SHARP BEND IN THE STYLET EXTENDING PAST THE METAL CANNULA. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A BENT STYLET IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE 4 FR GROSHONG NXT CATHETER KIT WAS RETURNED IN OPEN PACKAGING. THE PRODUCT LABEL INDICATED LOT: REGN3971. THE CATHETER AND KIT COMPONENTS WERE PRESENT WITHIN THE OPEN TRAY. THE CATHETER WAS RETURNED CUT IN TWO SECTIONS AT THE 40 CM DEPTH MARKER. BLOOD RESIDUE AND EVIDENCE OF HANDING WAS NOTED. THE FLUSHING HUB STYLET WAS RETURNED OUTSIDE OF THE CATHETER AND CONTAINED A SHARP BEND IN THE STYLET EXTENDING PAST THE METAL CANNULA. THE BEND PRESENT IN THE STYLET WAS THE LIKELY CONTRIBUTING FACTOR TO THE RESISTANCE WITHIN THE CATHETER; THEREFORE, THE COMPLAINT IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. A REVIEW OF THE MANUFACTURING RECORDS DID NOT REVEAL EVIDENCE TO SUGGEST A MANUFACTURING RELATED ROOT CAUSE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05989 | 2023-08-10 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT A FOREIGN MATERIAL WAS ATTACHED TO THE FOLEY CATHETER AT ABOUT 5 CM TOWARD THE TIP FROM THE CATHETER BIFURCATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-05990 | 2023-08-10 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119110 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS DEFORMED DURING THE PRETEST. PER PHOTO SAMPLE RECEIVED FROM IBC ON (B)(6) 2023, IT WAS NOTED THAT THAT TWO LOT NUMBERS HAD THIS ISSUE (BATCH# NGGP3275 AND BATCH# NGGV5075) PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, CLARIFIED THAT THE BALLOON WRINKLED WHEN THE BALLOON WAS DEFLATED DURING PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06024 | 2023-08-11 | 2023-08-02 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53614G |  | IT WAS REPORTED THAT THESE INTERMITTENT CATHETERS WERE PAINFUL AND CAUSED PATIENT TO HAVE A URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. PER THE FOLLOW UP VIA CALL RECEIVED ON (B)(6) 2023, PATIENT WAS SENT INTERMITTENT CATHETERS, BUT IT WAS NOT THE CATHETER PATIENT NORMALLY USES. CUSTOMER STATED PATIENT EXPERIENCED A URINARY TRACT INFECTION WITH USE AND WAS SEEN AT AN URGENT CARE AND SENT HOME. CUSTOMER STATED THE DAY AFTER PATIENT WAS SENT HOME, PATIENT PASSED OUT AND WAS TAKEN TO EMERGENCY WHERE THEY ADMINISTERED ANTIBIOTIC BY IV. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-08
1018233-2023-06019 | 2023-08-11 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® CATHETER AND GLOVE SUCTION KIT | CATHETER AND GLOVE SUCTION KIT | C.R. BARD, INC. (COVINGTON) -1018233 |  | 140000 |  | IT WAS REPORTED THAT THERE WERE NO CONTENTS INSIDE TRAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06017 | 2023-08-11 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT THE CATHETER FELL OUT SPONTANEOUSLY AND THERE WAS NO MAJOR DAMAGE, BUT STERILE WATER LEAKED FROM THE BALLOON AREA. THE EVENT OCCURRED AFTER 6 DAYS OF PLACEMENT AND THERE WAS NO BALLOON RUPTURE OR TEAR. NO MATERIALS POSSIBLY CONTACTED TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06050 | 2023-08-11 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE CATHETER WAS PLACED PREOPERATIVELY AND WITHDRAWN SPONTANEOUSLY POSTOPERATIVELY. IT WAS STATED THAT THE WHEN WATER WAS RE-INJECTED, IT INFLATED NORMALLY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06022 | 2023-08-11 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT 4 TO 5 URETHRAL CATHETERS HAD THE PROBLEM OF OCCLUSION CONSECUTIVELY. THE TREATMENT WAS EVENTUALLY DONE WITH A SYRINGE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE POOR SAMPLE CONDITION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: SALT ACCUMULATION)/ BLOCK DRAINAGE LUMEN/ NO DRAINAGE EYE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06052 | 2023-08-11 | 2023-07-21 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER HAD CATHETER INSERTED IN THEIR TO ADMINISTER MEDICATIONS. THEY HAD PROBLEMS AND DOCTOR TOLD THEM TO GO TO EMERGENCY. AS PER UCC NOTIFICATION RECEIVED ON (B)(6) 2023, PER BD REPRESENTATIVE, (B)(4), THIS WAS MERELY AN INQUIRY REGARDING THE USE OF THE CATHETER. PRIMARILY SPANISH SPEAKING CALLER. CUSTOMER STATED THAT THEIR CAREGIVER HAD CATHETER INSERTED IN THEM TO ADMINISTER MEDICATIONS. THEY HAD PROBLEMS AND THE DOCTOR TOLD THEM TO GO TO THE ER. IT WAS EXPLAINED THAT OUR SIDE NOT INVOLVED IN THE ACTUAL CATHETER MANAGEMENT. THEY SHOULD CONSULT THE DOCTOR AND IF THE DOCTOR ADVISED TO GO TO THE ER, THEN THAT WAS THE APPROPRIATE ROUTE TO FOLLOW. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. NO POTENTIAL ROOT CAUSE COULD BE CONCLUDED DUE TO INSUFFICIENT INFORMATION. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03489 | 2023-08-11 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F MAXIMAL BARRIER TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED VIA MEDWATCH "PATIENT ARRIVED IN ER AT ANOTHER HOSPITAL WITH A LEAKING PICC (PERIPHERALLY INSERTED CENTRAL CATHETER) LINE, NO CAP IN PLACE. PATIENT STATES THAT IT WAS REMOVED BY THE HOME CARE SERVICES BECAUSE THE NEW SOLO PICCS ARE MADE TO NO LONGER REQUIRED A PRESSURIZED CAP. THE VAT (VASCULAR ACCESS TEAM) AT THIS FACILITY WAS NOTIFIED, AND ADVISED THEY COULD NOT GET THE PICC LINE TO STOP LEAKING BLOOD WITHOUT A CAP. PATIENT TRANSFERRED HERE FOR VAT TO PLACE A NEW PICC LINE. AFTER REVIEW, IT APPEARS THE VALVE ON THE PICC FAILED. THIS IS THE 2ND INCIDENT WITH THIS PARTICULAR PICC AND LOT# TODAY. THIS LINE WAS ORIGINALLY PLACED WHILE THE PATIENT WAS INPATIENT LAST MONTH." THE CUSTOMER REPORTED THIS EVENT OCCURRED A TOTAL OF 2 TIMES. THIS REPORT ADDRESSES THE FIRST EVENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE CUSTOMER REPORTED TWO EVENTS VIA MEDWATCH HOWEVER THERE WAS SEPARATE INFORMATION AND IT REQUIRED TWO FILES, THE FIRST INCIDENT WAS DOCUMENTED UNDER INTERNAL BD REFERENCE NUMBER PR (B)(4), AND THE SECOND INCIDENT WAS DOCUMENTED UNDER INTERNAL BD REFERENCE NUMBER PR (B)(4). | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06048 | 2023-08-11 | 2023-07-15 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THREE DAYS AFTER PLACEMENT, THE CATHETER RUPTURED DUE TO PATIENT SELF EXTRACTION. THE FOLEY CATHETER WAS PLACED DURING ORTHOPEDIC SURGERY. SELF-EXTRACTION OCCURRED THREE DAYS LATER. PATIENT HAD DEMENTIA. NO REINSERTION OF THE FOLEY CATHETER AFTER REMOVAL. NO ADVERSE EVENTS OCCURRED (ALMOST NO BLEEDING). THE REMAINING CATHETER FRAGMENT IN THE BLADDER WAS REMOVED BY CYSTOSCOPY BY UROLOGY ON (B)(6). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03488 | 2023-08-11 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F MAXIMAL BARRIER TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED VIA MEDWATCH "PATIENT ARRIVED IN ER AT ANOTHER HOSPITAL WITH A LEAKING PICC (PERIPHERALLY INSERTED CENTRAL CATHETER) LINE, NO CAP IN PLACE. PATIENT STATES THAT IT WAS REMOVED BY THE HOME CARE SERVICES BECAUSE THE NEW SOLO PICCS ARE MADE TO NO LONGER REQUIRED A PRESSURIZED CAP. THE VAT (VASCULAR ACCESS TEAM) AT THIS FACILITY WAS NOTIFIED, AND ADVISED THEY COULD NOT GET THE PICC LINE TO STOP LEAKING BLOOD WITHOUT A CAP. PATIENT TRANSFERRED HERE FOR VAT TO PLACE A NEW PICC LINE. AFTER REVIEW, IT APPEARS THE VALVE ON THE PICC FAILED. THIS IS THE 2ND INCIDENT WITH THIS PARTICULAR PICC AND LOT# TODAY. THIS LINE WAS ORIGINALLY PLACED WHILE THE PATIENT WAS INPATIENT LAST MONTH." THE CUSTOMER REPORTED THIS EVENT OCCURRED A TOTAL OF 2 TIMES. THIS REPORT ADDRESSES THE SECOND EVENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE CUSTOMER REPORTED TWO EVENTS VIA MEDWATCH HOWEVER THERE WAS SEPARATE INFORMATION AND IT REQUIRED TWO FILES, THE FIRST INCIDENT WAS DOCUMENTED UNDER INTERNAL BD REFERENCE NUMBER PR (B)(4), AND THE SECOND INCIDENT WAS DOCUMENTED UNDER INTERNAL BD REFERENCE NUMBER PR (B)(4). H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06037 | 2023-08-11 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® PEDIATRIC ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165808 |  | IT WAS REPORTED THAT 3 OF THE FOLEY CATHETERS SEEMED TO BE DEFECTIVE. ALSO STATED THAT IN 2 OF THE CATHETERS, THE BALLOON WOULD NOT STAY INFLATED AND THE 3RD CATHETER SEEMED TO BE CLOGGED SOMEHOW BUT NOTHING WAS COMING OUT. THEY HAD SOME INTERMEDIATE CATHETERS SO THEY WERE STILL ABLE TO VOID, BUT THESE FOLEY CATHETERS WERE NOT WORKING LIKE THEY SHOULD. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿WATER LOST VIA PERMEATION¿. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. ALLOW THE BALLOON TO DEFLATE SLOWLY ON ITS OWN. DO NOT ASPIRATE OR MANUALLY ACCELERATE THE DEFLATION OF THE BALLOON. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. RECOMMENDED INFLATION CAPACITIES 3CC BALLOON: USE 5ML STERILE WATER 5CC BALLOON: USE 10ML STERILE WATER 30CC BALLOON: USE 35ML STERILE WATER DO NOT EXCEED RECOMMENDED CAPACITIES." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06044 | 2023-08-11 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53614G |  | IT WAS REPORTED THAT THE WITH 8 INTERMITTENT CATHETERS, THE PACKAGING SEEMED TO HAVE BEEN FILLED WITH AIR AND THE WATER SATCHELS WERE ALSO EMPTY AND FILLED WITH AIR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03486 | 2023-08-11 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY W/NITINOL GUIDEWIRE 135CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174335 |  | IT WAS REPORTED THAT, 4FR SINGLE LUMEN PICC INSERTED BY EXPERIENCED PICC INSERTER ON THE (B)(6) 2023. PATIENT AND NURSES OBSERVED LEAKAGE FROM SITE AFTER RECEIVING MEDICATION. ONE NURSE OBSERVED WHITE POWDER LIKE SUBSTANCE UNDER THE DRESSING. PICC LINE REMOVED AND REPLACED ON THE (B)(6) 2023. AFTER REMOVAL THE PICC LINE WAS FLUSHED AND A SMALL HOLE WAS NOTICED CLOSE TO THE WINGS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03487 | 2023-08-11 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THAT "WHEN THE PICC LINE IS TO BE INSERTED, ALL PARTS OF THE SET ARE FLUSHED AND THE FUNCTION IS CHECKED. AT THE END OF THE PICC LINE CATHETER THERE IS A YELLOW STOPPER IN WHICH A CONDUCTOR IS INSERTED. WHEN THE PICC LINE IS FLUSHED, IT LEAKS THROUGH THIS YELLOW STOPPER, WHICH IT SHOULD NOT DO. THEREFORE, REPLACE THE SET WITH A NEW ONE AND COMPLETE THE INLAY. THE ERROR WAS DETECTED BEFORE USE AND THERE WAS NO PATIENT IMPACT." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06040 | 2023-08-11 | 2023-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER HAD BLACK DEBRIS AT THE END OF THE FOLEY TIP. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06035 | 2023-08-11 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 33422 |  | IT WAS REPORTED THAT THE CUSTOMER HAD MULTIPLE ISSUES WITH THE FOLEY CATHETER BALLOON. THE BALLOON HAD BURST INSIDE THE PATIENT SEVERAL TIMES. THE BALLOON ALSO DID NOT BLOW UP ON SOME OCCASIONS. THE BALLOON WAS ALSO LEAKING. THIS HAD HAPPENED WITH ONE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06032 | 2023-08-11 | 2023-07-17 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT PATIENT GOT A URINARY TRACT INFECTION FROM THE PUREWICK FEMALE EXTERNAL CATHETER. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE PATIENT WENT TO THE EMERGENCY ROOM AND WAS TOLD THAT THEY HAD A URINARY TRACT INFECTION. THEY WERE GIVEN IV ANTIBIOTICS IN THE EMERGENCY ROOM. THE PATIENT WAS ADMITTED AND HOSPITALIZED FOR FOUR DAYS. THE PATIENT WAS NOW BACK HOME AND FINISHED THEIR ORAL ANTIBIOTICS THAT THEY WERE PRESCRIBED ON YESTERDAY. CUSTOMER ALSO STATED THAT THE PATIENT WAS CONTINUING TO USE THE PUREWICK DEVICE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-08
1018233-2023-06031 | 2023-08-11 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT FOLEY CATHETER WAS NOT DRAINING WELL (BATCH# MYGR3749). ANOTHER FEEDBACK WAS RECEIVED FOR THE SAME QUALITY ISSUES OF OCCLUSION IN THE FAULTY CATHETER (BATCH# UNK). PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE USERS DID NOT KEEP ANY DETAILS ON PREVIOUS EVENTS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO DRAINAGE EYE OCCLUSION BLOCKED DRAINAGE LUMEN CAUSED BY PLC FAILURE AND ALSO MIGHT BE CONTRIBUTE BY NO DRAINAGE EYE OR USER RELATED (SALT ACCUMULATION). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03484 | 2023-08-11 | 2023-07-21 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO HF CATHETER 5F RADIOLOGY BASIC TRAY (TRIPLE-LUMEN)(70CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3395355Q |  | IT WAS REPORTED, "CALLER STATES HE WAS IN THE HOSPITAL FOR ABOUT 17 DAYS WITH A POWERPICC SOLO. CALLER STATES THEY REMOVED THE PICC FIVE DAYS AGO BUT HE CAN STILL FEEL TINGLING IN TWO FINGERS AND A THUMB. RESPONSE: EXPLAINED THAT IT IS POSSIBLE THAT THE CLINICIAN HIT A NERVE DURING PICC PLACEMENT. THIS MAY RESOLVE ITSELF IN A FEW DAYS BUT IT NEEDS TO BE REPORTED TO PHYSICIAN AS SOON AS POSSIBLE. THE DOCTOR MAY WANT TO FOLLOW UP WITH TREATMENTS. RESPONSE: EXPLAINED THAT IT IS POSSIBLE THAT THE CLINICIAN HIT A NERVE DURING PICC PLACEMENT. THIS MAY RESOLVE ITSELF IN A FEW DAYS BUT IT NEEDS TO BE REPORTED TO PHYSICIAN AS SOON AS POSSIBLE. THE DOCTOR MAY WANT TO FOLLOW UP WITH TREATMENTS." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06026 | 2023-08-11 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53614G |  | IT WAS REPORTED THAT THE WITH 8 INTERMITTENT CATHETERS, THE PACKAGING SEEMED TO HAVE BEEN FILLED WITH AIR AND THE WATER SATCHELS WERE ALSO EMPTY AND FILLED WITH AIR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-08
1018233-2023-06025 | 2023-08-11 | 2023-07-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2023, UNDERWENT THE OPERATION ON (B)(6) AND RETURNED TO THE WARD AT 10:20 AFTER THE OPERATION. THE PATIENT WAS PLACED IN A HORIZONTAL POSITION WITH A CLOUD PILLOW AND WAS GIVEN ECG MONITORING, OXYGEN INHALATION, AND BLOOD PRESSURE MEASUREMENTS ACCORDING TO THE DOCTOR'S ADVICE. IT WAS STATED THAT DURING THE OPERATION, ONE FOLEY CATHETER WAS USED, WHICH WAS PROPERLY FIXED, THE DRAINAGE WAS SMOOTH, THE BLADDER WAS CONTINUOUSLY FLUSHED WITH TIDE DRAINAGE, AND LIGHT BLOODY URINE WAS DRAINED. INFORM THE PATIENT AND FAMILY OF THE METHODS AND PRECAUTIONS TO PROTECT THE PIPES. AT 15:30, THE INSPECTION FOUND THE URETHRAL TUBE SLIP, THE INSPECTION FOUND THE BALLOON RUPTURE, AND THE DOCTOR IMMEDIATELY REINSERTED THE CATHETER AT THE BEDSIDE, PROPERLY FIXED IT, KEPT THE DRAINAGE SMOOTH, AND THE BLADDER DRAINAGE CONTINUED TO FLUSH. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
MW5144524 | 2023-08-11 | 2023-08-04 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | SINGLE LUMEN POWER PICC | CATHETER, INTRAVASCULAR, THERAPEUTIC, LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. |  |  |  | DURING PICC INSERTION, (B)(6) RN NOTED THAT THE STYLET IN THE PICC HAD AN UNUSUAL BEND NEAR THE END. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06020 | 2023-08-11 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT WAS TREATED ON ARCTIC SUN DEVICE. THERAPY WAS GOING WELL UNTIL THEY BATHED THE PATIENT. NOW THEY ADVISED NURSE THAT EITHER THE FOLEY PROBE OR THE TEMPERATURE-IN CABLE WAS THE PROBLEM. CONFIRMED CABLE WAS PLUGGED INTO TEMPERATURE PORT 1. NURSE OBTAINED A CABLE TO CONNECT THE FOLEY TO THE BEDSIDE MONITOR. THE PATIENT TEMPERATURE (PT) DID NOT DISPLAY ON THE BEDSIDE MONITOR. ADVISED THEY COULD REPLACE THE FOLEY OR UTILIZE AN ESOPHAGEAL OR RECTAL PROBE INSTEAD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03482 | 2023-08-11 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR POST PORT PLACEMENT, THE PATIENT NOTICED THAT THE AREA AROUND THE PORT BECAME VERY WARM AND HAD DISCOMFORT. IT WAS FURTHER REPORTED THAT THERE IS NO INFORMATION REGARDING THE ADDITIONAL INTERVENTION PERFORMED OR ADDITIONAL MEDICATION PRESCRIBED TO TREAT THE DISCOMFORT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06011 | 2023-08-11 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 33422 |  | IT WAS REPORTED THAT THE CUSTOMER HAD MULTIPLE ISSUES WITH THE FOLEY CATHETER BALLOON. THE BALLOON HAD BURST INSIDE THE PATIENT SEVERAL TIMES. THE BALLOON ALSO DID NOT BLOW UP ON SOME OCCASIONS. THE BALLOON WAS ALSO LEAKING. THIS HAD HAPPENED WITH ONE PATIENT. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE COULD BE "INCORRECT BALLOON DESIGN (INFLATION LENGTH UNDERSIZED)". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. DO NOT USE IF PACKAGE IS DAMAGED. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03479 | 2023-08-11 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP TI 8F GROSHW/OSH | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8708561 |  | IT WAS REPORTED THAT ONE MONTH AND SIX DAYS POST PORT PLACEMENT, THERE WAS A LEAK FROM THE PORT CATHETER DURING THE THIRD CHEMOTHERAPY INJECTION. REPORTEDLY, THE PATIENT HAD AN EXTRAVASATION AND REQUIRED NO TREATMENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03477 | 2023-08-11 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT APPROXIMATELY TWO DAYS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED HEADACHES AND NECK PAIN ON HER LEFT SIDE. IT WAS FURTHER REPORTED THAT THE HEADACHE AND NECK PAIN WAS DUE TO THE TIGHT CLOTHING. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06014 | 2023-08-11 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE FAMILY MEMBERS OF THE PATIENT FOUND THAT THE FOLEY CATHETER HAD FALLEN OFF AT 1:30 IN THE MORNING. THE NURSE ON DUTY IMMEDIATELY NOTIFIED THE DOCTOR ON DUTY TO ASSIST THE DOCTOR TO INSERT THE CATHETER AGAIN AND TO CONTINUE BLADDER IRRIGATION. AFTER EXAMINATION, THE FOLEY CATHETER BALLOON HAD RUPTURED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03485 | 2023-08-11 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716000J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, AFTER THE PUNCTURE NEEDLE WAS PUNCTURED AND THE GUIDEWIRE WAS INSERTED, THE MICRO INTRODUCER ALLEGEDLY COULD NOT BE INSERTED THROUGH THE PUNCTURE SITE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 6.5FR PEEL-APART SHEATH, ONE VESSEL DILATOR, AND ONE CATHETER WERE RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. THE PEEL-APART SHEATH AND VESSEL DILATOR WERE NOTED TO BE SLIGHTLY BENT AND THE DISTAL TIP OF THE VESSEL DILATOR WAS NOTED TO BE DEFORMED. AN ATTEMPT TO LOAD THE VESSEL DILATOR TO THE PEEL-APART SHEATH WAS PERFORMED AND WAS SUCCESSFUL AS THE DILATOR LOCKED INTO PLACE WITHOUT ISSUE; THE DILATOR THEN RETRACTED WITHOUT ISSUE. AN ATTEMPT TO LOAD THE CATHETER RETURNED INTO THE PEEL-APART SHEATH WAS ATTEMPTED AND WAS SUCCESSFUL. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED DEFORMATION ISSUES. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE TO ADVANCE ISSUES, AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED, AND THE REPORTED EVENT COULD NOT BE REPRODUCED IN THE LAB. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 08/2024) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03474 | 2023-08-11 | 2023-07-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 5608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED REDNESS FROM THE PORT ACROSS THE CHEST AND UP TO THE NECK. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 06/2021). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03480 | 2023-08-11 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP TI 8F GROSHW/OSH | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8708561 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THERE WAS A LEAK FROM THE PORT CATHETER DURING FLUSHING DURING THE THIRD CHEMOTHERAPY. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 05/2024). H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03476 | 2023-08-11 | 2023-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393420 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THERE WAS ALLEGEDLY LEAKAGE IN THE HUB OF THE CATHETER. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03478 | 2023-08-11 | 2023-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 5393230 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CLAMP OF THE DEVICE WAS ALLEGEDLY FOUND TO BE BROKEN. THE PROCEDURE WAS COMPLETED USING REPAIR KIT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 06/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03483 | 2023-08-11 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, AIR BUBBLES WAS ALLEGEDLY ASPIRATED FROM THE INTRODUCER NEEDLE AFTER CONNECTING A SYRINGE TO THE HUB OF THE INTRODUCER NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 09/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06010 | 2023-08-11 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT THE PATIENT EXPERIENCED A PROBLEM WITH THE SUPRAPUBIC CATHETER ON (B)(6) 2023. IT WAS STATED THAT ONE OF THE CONSULTANTS USED THE SUPRAPUBIC CATHETER ON TUESDAY, THERE WAS NO ISSUE AT THE TIME, BUT LATER THEY STATED THAT THE BALLOON OF THE CATHETER HAD RUPTURED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE THE LATEX. VISUALLY INSPECTED THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. -USE LUER TIP SYRINGE TO INFLATE WITH STATE 10ML OF STERILE WATER OR - FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. ALLOW THE PRESSURE WITHIN TO BALLOON FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
MW5137954 | 2023-08-12 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | POWER PORT | PORT AND CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | BARD |  |  |  | PATIENT STATES HER POWER PORT IS LEAKING. (B)(4). THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). |  | N | N |  |  | 2023-08
MW5143003 | 2023-08-12 |  | I, n, j, u, r, y | Voluntary report | MRI POWERPORT W/ATTACH 8FR CHRONO | PORT AND CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | BARD ACCESS SYSTEMS, INC. |  | 1808000 |  | SUMMARY OF INJURIES INCURRED DUE TO DEFECTIVE MEDICAL DEVICE: I ENDURED OVER 30 DAYS OF CHEMOTHERAPY DURING 5 YEARS OF FIGHTING AGENT ORANGE INDUCED LEUKEMIA WHICH RESULTED IN VEINS IN MY ARMS BECOMING HARDENED, PAINFUL, AND DANGEROUS. UPON INSISTENCE FROM MY NURSES AND DOCTORS, I HAD A INFUSION PORT IMPLANTED SUBDURAL ON (B)(6) 2011 BY THE SAME (B)(6) HOSPITAL WHERE I WAS BEING TREATED. THE PORT WAS A BARD POWER PORT, LOT: REUJ0764, REF: (B)(4), SERIAL#: (B)(6), MADE BY BARD ACCESS SYSTEMS. IT NEVER WORKED CORRECTLY. IT WAS DEFECTIVE IN EVERY WAY IMAGINABLE. THE INFUSION TECHNICIANS COULD NOT MAKE IT FUNCTION PROPERLY AND THEY REFUSED TO TRY TO ACCESS IT AGAIN AFTER SEVERAL FAILED ATTEMPTS. IT USUALLY COULD NOT BE ACCESSED NOR FLUSHED AND WOULD NOT RETURN BLOOD BUT THEN WHEN THE ACCESS LINE WAS REMOVED, IT LITERALLY SPRAYED BLOOD OVER THE ROOM. THIS BLEEDING WAS HARD TO STOP RESULTING IN REPEATED MASSIVE BRUISING AND SWELLING AROUND THE PORT. THE PORT WAS FREQUENTLY BLOCKED BY CLOTS. WHEN ACCESSED DURING TREATMENTS, THE FLOW STOPPED FREQUENTLY AND ALMOST ALWAYS STOPPED WHEN MY BODY MOVED. THE NURSES REFERRED TO THIS AS PINCH OFF SYNDROME. THIS RESULTED IN MY IV ALARM SOUNDING CONSTANTLY DURING TREATMENTS, CAUSING GREAT STRESS FOR ME AND ALL MEDICAL STAFF IN MY VICINITY, AND RESULTED IN SEVERAL UNCOMPLETED CRITICAL TREATMENTS. THE CATHETER WAS OBVIOUSLY DEFECTIVE SINCE IT KINKED IN THE MIDDLE OF AN ESSENTIAL AND GENTLE CURVE RESULTING IN 100% BLOCKAGE WHICH WAS CONFIRMED BY X-RAYS. BARD REFUSED REPEATEDLY TO DISCUSS THIS PROBLEM OR OFFER ANY ASSISTANCE. THEIR ONLY RESPONSE WAS TO INSIST REPEATEDLY THAT I RETURN THE PORT TO THEM. THAT WAS AN OUTRAGEOUS RESPONSE BY BARD. TWO OF MY DOCTORS OF ONCOLOGY AND SEVERAL RNS HIGHLY RECOMMENDED MY PORT BE REMOVED AND STATED IT WAS DEFECTIVE AND DANGEROUS. AFTER ANOTHER TRAUMATIC EXPERIENCE WITH THE PORT ON (B)(6) 2011, THE ATTENDING ONCOLOGIST RECOMMENDED AN IMMEDIATE TEST OF THE PORT BY INTERVENTIONAL RADIOLOGY AND STATED MY PORT WAS OBVIOUSLY TOTALLY DEFECTIVE AND DANGEROUS. ANOTHER ONCOLOGIST WROTE A RECOMMENDATION TO THE SURGICAL STAFF THAT MY PORT MUST BE REMOVED IMMEDIATELY AND WITHIN ONE WEEK BUT A CLERK (NON MEDICAL STAFF PERSON) FOR THE SURGICAL TEAM, ALSO IN THE SAME (B)(6) HOSPITAL, IGNORED THAT RECOMMENDATION, DOWNGRADED THE RECOMMENDATION, AND MY MANY APPEALS FOR IMMEDIATE TREATMENT WERE IGNORED BY THE ENTIRE (B)(6) STAFF ON (B)(6) 2011, MY PORT HAD TO BE REMOVED BY A DIFFERENT NON VA SURGEON AT MY EXPENSE. THE SURGEON WHO REMOVED IT STATED IN WRITING THAT HE TESTED IT AND FOUND A HOLE IN THE CATHETER AT MID LENGTH. I HAVE A COPY OF THAT REPORT. I ASKED THAT SURGEON IF ANY OF THE MANY LARGE IV BAGS OF CHEMOTHERAPY DRUGS WHICH I RECEIVED COULD BE FLOATING AROUND IN MY CHEST CAVITY. HE REPLIED: YOU NEED TO GET AN OPINION ON THAT. THAT RECOMMENDATION CAUSED GREAT ANXIETY. I CANNOT DETERMINE WHAT THE POSSIBLE MEDICAL CONSEQUENCES OF THE LEAKAGE AND BLEEDING THROUGH THAT LARGE HOLE ARE AND MIGHT BE IN THE FUTURE EXCEPT THAT IT IS CAUSING EXTREME MENTAL ANXIETY AND I AM EXPERIENCING UNUSUAL AND UNPRECEDENTED ABDOMINAL PAINS, AND SPORADIC BRUISING DIRECTLY BELOW WHERE THE PORT WAS IMPLANTED. THIS LEAKAGE FROM A MEDICAL DEVICE IS CLINICALLY REFERRED TO AS EXTRAVASATION INJURIES. I HAVE PHOTOS OF THESE INJURIES. I ASKED THE SURGEON WHO IMPLANTED IT, AND BARD, AND MY ONCOLOGIST, AND NURSES, ABOUT POSSIBLE SYMPTOMS AND INJURIES CAUSED BY LEAKAGE BUT THEY ALL BECAME DEFENSIVE AND NONRESPONSIVE. MY LAST CHEMOTHERAPY DRUG WAS RELATIVELY NEW TO THE VA AND FDA. THE CATHETER IS DISCOLORED IN THE VICINITY OF THE HOLE WHICH INDICATES POSSIBLE ADVERSE REACTION BETWEEN MY MEDICINES AND THE CATHETER WHERE THE CATHETER WAS BLOCKED AND APPARENTLY LEAKED. MY NEXT CRISIS WILL BE HOW TO TAKE FUTURE CHEMOTHERAPY WHEN REQUIRED BECAUSE, AFTER THIS CRISIS. HAVING ANOTHER PORT IMPLANTED WOULD BE EXTREMELY STRESSFUL. AS OF ON (B)(6) 2012, MY LEUKEMIA HAS REACTIVATED ENDING HOPES OF REMISSION. THIS RECURRENCE COULD POSSIBLY BE CAUSED BY THE PREVIOUSLY SCHEDULED TREATMENTS WHICH COULD NOT BE COMPLETED DUE TO THE DEFECTIVE PORT. I ESTABLISHED THE CHAIN OF CUSTODY OF THIS PORT AND THIS PORT IS PHYSICALLY IN MY POSSESSION. BARD DENIES ANY RECALLS BUT THE FDA AND SEVERAL BLOGGERS REPORT MANY RECALLS OF RELATED OR THE SAME BARD PRODUCT. OTHER BLOGGERS HAVE REPORTED SIMILAR LIFE THREATENING CRISES WITH BARD POWER PORTS. I HAVE A SEPARATE DESCRIPTION OF MY PERSONAL INJURIES CAUSED BY THIS DEFECTIVE PORT. (B)(4). THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). |  | N | N |  |  | 2023-08
3006260740-2023-03491 | 2023-08-12 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 9808560 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, AFTER THE CATHETER WAS CONNECTED TO THE PORT BODY, WHEN THE NON-CORING NEEDLE WAS PUNCTURED TO THE PORT BODY AND BLOOD RETURN WAS ATTEMPTED TO BE CONFIRMED, BLOOD COULD NOT BE ASPIRATED. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY DETACHED FROM THE PORT BODY AND THE CATHETER WAS CONNECTED TO A SYRINGE DIRECTLY, BLOOD RETURN COULD BE CONFIRMED. THEREFORE, THE PORT BODY WAS REMOVED AND ANOTHER NEW PORT BODY WAS REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
MW5143082 | 2023-08-12 |  | I, n, j, u, r, y | Voluntary report | BARD ATLAS | CATHETER, ANGIOPLASTY, PERIPHERAL, TRANSLUMINAL | BARD |  |  |  | MEDTRONIC RECEIVED INFORMATION THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER PULMONARY BIOPROSTHETIC VALVE, A BARD ATLAS PT A DILATATION CATHETER BALLOON RUPTURED ON A DISTAL CALCIFIC SPUR. THE BALLOON WAS REMOVED FROM THE PATIENT AND THE CONDUIT APPEARED INTACT. A CORDIS PALMAZ 4010 STENT WAS IMPLANTED AND NO EXTRAVASATION OF CONTRAST WAS NOTED ON THE INITIAL REVIEW. UPON SUBSEQUENT REVIEW, IT WAS REPORTED THAT THERE APPEARED TO BE AN AREA ALONG THE LATERAL BORDER WHERE CONTRAST EXTENDED SLIGHTLY BEYOND THE CONDUIT SUGGESTING A MILD CONTAINED DISRUPTION. THERE WERE NO REPORTED ADVERSE PATIENT EFFECTS AS A RESULT OF THE BALLOON RUPTURE OR SUBSEQUENT STENT IMPLANT. (B)(4). THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). |  | N | N |  |  | 2023-08
3006260740-2023-03490 | 2023-08-12 | 2023-03-03 | I, n, j, u, r, y | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE CATHETER ALLEGEDLY BROKE. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY GOT SEPARATED FROM THE PORT. FURTHERMORE, THE CATHETER ALLEGEDLY GOT DISLODGED IN THE HEART. REPORTEDLY, THE EXTERNAL PORT WAS REMOVED AS NORMAL PROCEDURE BUT THE INTERNAL CANNULA WENT INTO THE RIGHT VENTRICLE, WHICH WAS REMOVED BY INTERVENTIONAL CARDIOLOGIST. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 12/2021). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
MW5141734 | 2023-08-12 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | ATLAS GOLD BALLOON | CATHETER, ANGIOPLASTY, PERIPHERAL, TRANSLUMINAL | BARD |  |  |  | MEDTRONIC RECEIVED INFORMATION THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER PULMONARY BIOPROSTHETIC VALVE, INTO A NATIVE ANNULUS WITH HARD CALCIFICATION AROUND THE LANDING ZONE, A BARD ATLAS GOLD BALLOON BROKE DURING INFLATION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. (B)(4). THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). |  | N | N |  |  | 2023-08
MW5135680 | 2023-08-12 | 2014-08-22 | I, n, j, u, r, y | Voluntary report | BARD DIAGNOSTIC WOVEN QUAD CATHETER | CATHETER, ELECTRODE RECORDING, OR PROBE, ELECTRODE RECORDING | BARD MEDICAL DIVISION |  |  |  | THIS LETTER IS TO INFORM YOU OF THIS ADVERSE EVENT AS THE SUSPECT DEVICE (BARD MEDICAL DIVISION -BARD DIAGNOSTIC WOVEN QUAD CATHETER) IS NOT MANUFACTURED OR IMPORTED BY BIOSENSE WEBSTER, INC. IT WAS REPORTED THAT THE PATIENT SUFFERED A PERICARDIAL EFFUSION. DURING THE PROCEDURE THE PATIENT'S PRESSURE DROPPED AND THEY BECAME NAUSEOUS AND COMPLAINED OF CHEST PAIN. THE ICS CATHETER CONFIRMED A PERICARDIAL EFFUSION AND PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT IS STABLE AT THIS TIME AND NO FURTHER INTERVENTION IS ANTICIPATED. THE PHYSICIAN STATED THAT THE RV WAS PERFORATED BY THE RV DIAGNOSTIC CATHETER. THE CARTO 3 SYSTEM (11854), STOCKERT 70 (ST-0308), AND COOLFLOW PUMP (05047) WERE UTILIZED IN THE PROCEDURE. THE CALLER ALSO STATED THAT THE ICE CATHETER WAS FROM STRYKER, AND THE RV CATHETER WAS A BARD DIAGNOSTIC WOVEN QUAD CATHETER. (B)(4). THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). |  | N | N |  |  | 2023-08
MW5135287 | 2023-08-12 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | CS CATHETER EP XT OCTAPOLAR | CATHETER, ELECTRODE RECORDING, OR PROBE, ELECTRODE RECORDING | BARD |  |  |  | IT WAS REPORTED THAT THE CS CATHETER (BARD EP XT OCTAPOLAR. REPROCESSED BY LNNOVA HEALTH) BECAME "STUCK" IN THE CORONARY SINUS THIS LASTED FOR AT LEAST AN HOUR, WHILE THE PHYS1C1AN ATTEMPTED TO MANIPULATE THE CATHETER OUT OF THE CORONARY SINUS. THE PHYSICIAN BELIEVES THIS WAS DUE TO THE CATHETER BEING IN A DEFLECTED POSITION. THE DEFLECTION MECHANISM WAS NOT WORKING TO STRAIGL1TEN THE CATHETER. AND A SPASM IN THE VESSEL RESULTED FROM THE EXTENDED TIME OF CATHETER MANIPULATION. AFTER ADMINISTERING NITROGLYCERIN DIRECTLY TO THE CORONARY SINUS, THE CATHETER WAS ABLE TO BE REMOVED NO PATIENT CONSEQUENCES THE CALLER STATED THAT THE RA HAD BEEN MAPPED AT THIS POINT AFTER THE CATHETER WAS REMOVED FROM THE CORONARY SINUS. THE PROCEDURE WAS CONTINUED ABLATION WAS PERFORMED SUCCESSFULLY BWI PRODUCTS IN THE BODY? STSF ABLATION CATHETER AND A DIAGNOSTIC QUAD. (B)(4). THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2) |  | N | N |  |  | 2023-08
MW5131583 | 2023-08-12 | 2016-10-06 | I, n, j, u, r, y | Voluntary report | DYNAMIC DECAPOLAR CATHETER | CATHETER, ELECTRODE RECORDING, OR PROBE, ELECTRODE RECORDING | BARD |  |  |  | IT WAS REPORTED THAT AFTER THE PVC ABLATION THE PATIENT PRESENTED DIAPHORETIC AND THEIR PRESSURE STARTING DROPPING. THE PHYSICIAN USED THE SOUNDSTAR CATHETER TO EVALUATE FOR AN EFFUSION. A PERICARDIAL EFFUSION WAS SEEN ON US. THE PATIENT WAS STABLE AT THE TIME OF THE CALL AND THE PHYSICIAN WAS STILL DECIDING ON FURTHER TREATMENT. THE CASE WAS COMPLETE WHEN THE EFFUSION WAS FOUND. THE ABLATION CATHETER WAS ST DF #D132705. ADDITIONAL INFORMATION WAS RECEIVED ON 10/11/2016; PHYSICIAN'S OPINION REGARDING THE CAUSE OF THE ADVERSE EVENT WAS THAT IT WAS PROCEDURE-RELATED. PHYSICIAN BELIEVES THE INJURY WAS A RESULT OF ADVANCING THE WEBSTER CS CATHETER TOO FAR INTO THE CORONARY SINUS, RESULTING IN THE RUPTURE OF THE DISTAL PORTION OF THE CORONARY SINUS. THEREFORE, THIS EVENT WAS REPORTED TO THE FDA UNDER THE BWI WEBSTER CS CATHETER (MDR REPORT# 2029046-2016-00240) ADDITIONAL INFORMATION WAS LATER RECEIVED ON 11/09/2016: WEBSTER DECAPOLAR WAS NOT USED DURING THIS CASE. BARD DYNAMIC DECAPOLAR WAS USED. THE CATHETER AND THE BOX WAS DISPOSED OF AFTER THE CASE. NO ADDITIONAL INFORMATION IS AVAILABLE. PER THIS INFORMATION THE BWI CATHETER WAS NOT USED AND THEREFORE THIS EVENT IS NOT REPORTABLE TO BWI. THE SUSPECTED DEVICE IS THE BARD DYNAMIC DECAPOLAR CATHETER. PI1-157RXU5. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). |  | N | N |  |  | 2023-08
MW5133218 | 2023-08-12 | 2018-03-20 | I, n, j, u, r, y | Voluntary report | NRG | CATHETER, SEPTOSTOMY | BARD | NRG-E-HF-71- ; C1/NGFF110917; 406849/6253465 |  |  | IT WAS REPORTED THAT AFTER AN AFIB CASE, A PERICARDIA! EFFUSION WAS FOUND WHEN THE NEEDLE WAS NOTICED TO HAVE GONE TRANSEPTAL BEFORE THE ATTEMPT TO GO TRANSEPTAL HAD BEEN MADE. THE EFFUSION WAS CONFIRMED THROUGH ICE AND CARLO 3. CALLER REPORTED THAT THE MEDICAL INTERVENTION PROVIDED WAS PERICARDIOCENTESIS AND 2LITERS OF FLUID WERE REMOVED. THE PATIENT'S CONDITION IS STABLE. THE FOLLOWING BWI DEVICES WERE IN USE: CARLO 3 (R10041), GENERATOR (G4C-2365), PUMP (G4CP-2310), SMARTTOUCH CATHETER(D134805, 17755402L -NOT AVAILABLE), CS CATHETER (BD710FJ282CT, 17681365M), PENTARAY (0128210, 30006132L), SOUNSTAR (G90789, 10439236) ADDITIONAL INFORMATION RECEIVED STATED FACTORS CITED BY THE THORACIC SURGEON THAT MAY HAVE CONTRIBUTED TO THE ADVERSE EVENT INCLUDE THAT THE PATIENT'S HEART WAS VERY FIBROUS. PHYSICIAN'S OPINION REGARDING THE CAUSE OF THE ADVERSE EVENT IS THAT IT OCCURRED DURING TRANSSEPTAL PUNCTURE, WHICH WAS TOO ANTERIOR, CAUSING A PERFORATION IN THE LAA. (B)(4). THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). |  | N | N |  |  | 2023-08
3006260740-2023-03496 | 2023-08-13 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT DUO M.R.I. IMPLANTABLE PORT, CHRONOFLEX DUAL-LUMEN, 9.5F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1829500 |  | IT WAS REPORTED, THAT SOMETIME POST PORT PLACEMENT. THE PORT WITH A CHAMBER ALLEGEDLY POPPED OUT. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. D4: (EXPIRY DATE: 12/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03495 | 2023-08-13 | 2023-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1806060J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE PATIENT HAD A POWER PORT PLACED VIA THE INTERNAL JUGULAR VEIN BY CUTDOWN PROCEDURE, AND THE PORT WAS PLACED IN THE ANTERIOR CHEST. IT WAS FURTHER REPORTED THAT WHEN THE CATHETER WAS CONNECTED TO THE PORT BODY AND BLOOD RETURN WAS CONFIRMED BEFORE THE SKIN WAS SUTURED, MORE AIR THAN USUAL WAS ASPIRATED ALONG WITH THE BLOOD. REPORTEDLY, THE PORT SYSTEM WAS REMOVED AND ANOTHER NEW PORT SYSTEM WAS REPLACED. HOWEVER, WHEN THE PATENCY OF THE REMOVED PORT SYSTEM WAS CHECKED AGAIN, MORE AIR THAN USUAL WAS ASPIRATED. THE NON-CORING NEEDLE (STRAIGHT) AND THE NON-CORING NEEDLE (ANGLE) WERE USED, HOWEVER, BOTH NEEDLES HAD A LOT OF AIR ASPIRATION. THE CATHETER AND THE PORT BODY WERE PRIMED BEFORE PLACEMENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 11/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03521 | 2023-08-14 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THAT THE LINE FRACTURED. FOLLOWED GUIDELINES. LOCALIZE AND DRUG NEUTRALIZE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06075 | 2023-08-14 | 2023-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT SURGERY IN THE OPERATING ROOM ON THE MORNING OF (B)(6) AND RETURNED TO THE WARD AT 10:00. THE THREE-CAVITY CATHETER WAS SMOOTH IN PLACE AND FIXED SECURELY, AND THE USER WAS GIVEN CONTINUOUS BLADDER IRRIGATION. THE BLADDER IRRIGATION IN THE MORNING OF JULY 15TH ONLY ALLOWED IN BUT COULD NOT COME OUT, AND THEN THE THREE-CAVITY CATHETER CAME OUT ON ITS OWN, AND THE THREE-CAVITY CATHETER BALLOON WAS RUPTURED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO THIN SAC CAUSING BY SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03519 | 2023-08-14 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER (20G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202252 |  | IT WAS REPORTED BY CUSTOMER THAT SAFETY BUTTON WAS FULLY ENGAGED, HOWEVER NEEDLE TIP WOULD NOT SAFETY BACK INTO CLEAR HOUSING. CLINICIAN HAD TO REMOVE EXPOSED NEEDLE AND IMMEDIATELY SHARPS IT. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: CUSTOMER REPORTS THE EVENT DIRECTLY INVOLVED A PATIENT BUT DID NOT INVOLVE AN URGENT OR LIFE-THREATENING MEDICATION SITUATION. THERE WAS NO PATIENT HARM OR NEGATIVE OUTCOME, DELAY IN MEDICATION, AND NO UNPLANNED MEDICAL OR SURGICAL INTERVENTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03516 | 2023-08-14 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT DURING THE PROCESS OF PICC CATHETER PLACEMENT, THE CUSTOMER FOUND THAT THE GAP BETWEEN CONNECTING PARTS OF THE CONNECTOR FOR THE GROSHONG KIT WAS A LITTLE BIT WIDER. NO PROBLEM OCCURRED AFTER CONNECTION WAS COMPLETED ON SITE. THE NEXT DAY, THE CONNECTED PARTS WERE SEPARATED. THE CATHETER WAS TRIMMED IMMEDIATELY AND A NEW CONNECTING ACCESSORY WAS USED. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03517 | 2023-08-14 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT DURING THE PROCESS OF PICC CATHETER PLACEMENT, THE CUSTOMER FOUND THAT THE GAP BETWEEN CONNECTING PARTS OF THE CONNECTOR FOR THE GROSHONG KIT WAS A LITTLE BIT WIDER. NO PROBLEM OCCURRED AFTER CONNECTION WAS COMPLETED ON SITE. THE NEXT DAY, THE CONNECTED PARTS WERE SEPARATED. THE CATHETER WAS TRIMMED IMMEDIATELY AND A NEW CONNECTING ACCESSORY WAS USED. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03520 | 2023-08-14 | 2023-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174155 |  | IT WAS REPORTED THE DEVICE PRESENTS DIFFICULTY TO USE THE GUIDEWIRE. DELAY IN PROCEDURE. ADDITIONAL INFORMATION RECEIVED 04 AUGUST 2023: INITIALLY, IT WAS NOTED THAT THE GUIDEWIRE OF THE 4FR SINGLE LUMEN KIT WAS DIFFERENT FROM THE OTHERS USED. THE PUNCTURE WAS PERFORMED, BUT THE GUIDEWIRE FAILED TO ADVANCE, AS IT WAS NOTED A DIFFICULT TO SLIDE IT. UPON REMOVAL IT GOT OFF COMPLETELY TORTUOUS. IT WAS PUNCTURED ONCE AGAIN, BUT WITH THE GUIDEWIRE DAMAGED THERE WAS NO PROGRESSION, MAKING IMPOSSIBLE TO USE IT. A NEW CATHETER WAS TAKEN AND I USED THE GUIDEWIRE FROM A 3FR KIT IN THE 4FR CATHETER, PROCEDURE WAS SUCCESSFULLY COMPLETED AFTER THIS EXCHANGE. ADDITIONAL PUNCTURES NEEDED. PATIENT GOT DISCHARGED. NO MEDICAL INTERVENTION WAS NEEDED. IT WAS A PEDIATRIC PATIENT BEING TREATED WITH CHEMOTHERAPY. THERE WAS NOT A BREAK. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06081 | 2023-08-14 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175816 |  | IT WAS REPORTED THAT THE TWO OF FOLEY CATHETER BALLOON WOULD NOT DEFLATE. NURSE PLACING THE FOLEYS WAS PRIMING THE BALLOON PRIOR TO INSERTION SO THE FAULTY ITEMS DID NOT MAKE IT INTO A PATIENT. NURSE PREFILLED THE BALLOON WITH 5 ML TO TEST IT, AND IT WOULD NOT DEFLATE. THEY GOT ANOTHER TRAY AND HAD THE SAME ISSUE. FINALLY, WITH THE THIRD TRAY, THE BALLOON DEFLATED PROPERLY. THEY WAS NOT ABLE TO GET IT TO DEFLATE INITIALLY EITHER AND PUSHED THE SYRINGE INTO THE PORT REALLY HARD TO THE POINT THAT THE ORANGE RING POPPED OFF, AND THEN IT FINALLY DEFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06069 | 2023-08-14 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 |  | FOL0101 |  | IT WAS REPORTED THAT WITH 3 (THREE) SEPARATE FOLEY STATLOCK STABILIZATION DEVICES, THE CUSTOMER BELIEVED THE STABILIZATION DEVICE WAS CRIMPING THE CATHETER TUBING AND CREATING A PROBLEM FOR THE URINE FLOW. IT WAS NOTED THAT THERE WAS NO SEDIMENT IN THE TUBE AND NO ISSUE WITH THE CATHETER. THERE WAS A SNAP TO KEEP THE DEVICE IN PLACE, THUS THE CUSTOMER COULD NOT PUT THE DEVICE ON INCORRECTLY. MORE OFTEN THAN NOT, THE BUTTERFLY STABILIZATION DEVICE WAS IN PLACE AND SNAPPED, BUT NO URINE WOULD GO THROUGH THE FOLEY CATHETER TUBING INTO THE FOLEY BAG. IT WAS NOTED THAT THIS ISSUE WAS NOT DUE TO LACK OF GRAVITY. AS SOON AS THE FOLEY CATHETER TUBING WAS UNSNAPPED, THE BAG WOULD BE FILLED. THE CUSTOMER BELIEVED THE FOLEY WAS CUTTING OFF THE URINE FLOW FROM THE CATHETER. THERE WOULD BE 8-9 HOURS WITH NO URINE FLOW, BUT ONCE THE STABILIZATION DEVICE WAS REPLACED WITH TAPE OR JUST FREE HANGING, THE URINE FLOW WOULD BE INSTANT. TUBES CLEAR, PROPERLY INSTALLED. THE CUSTOMER FEARED PEOPLE COULD DIE FROM RUPTURED BLADDERS, URINARY TRACT INFECTIONS (UTIS), OR KIDNEY INFECTIONS. THE CUSTOMER ALSO STATED THIS WAS A CONTINUOUS ISSUE, NOT JUST SINGLE TIME. AT ONE POINT IN THE HOSPITAL, THERE WAS NO URINE MOVEMENT OVER 9-10 HOURS. UPON SEEING THIS, THE DEVICE WAS UNSNAPPED, AND WITHIN ABOUT 2 MINUTES THE LARGER BAG ALMOST FILLED TO CAPACITY. THE STABILIZATION DEVICE WAS SNAPPED BACK TO CHANGE THE BAG. AS SOON AS THE BAG WAS REPLACED AND THE DEVICE UNSNAPPED, URINE FLOW RESUMED. THE PATIENT WAS GIVEN PLENTY OF FLUIDS, SO THE CUSTOMER DID NOT BELIEVE THIS WAS A URINE ISSUE. THE CUSTOMER BELIEVED THE ISSUE WAS RELATED TO THE STATLOCK DEVICE. THE PATIENT WAS SEDENTARY WHILE FOLEY CATHETER AND THE STABILIZATION DEVICE WAS IN PLACE. THE PATIENT DID NOT HAVE THE ABILITY TO STAND OR WALK, SO THE PATIENT¿S FAMILY ASSISTED WITH MOVING THE PATIENT SIDE-TO-SIDE AND RAISING/LOWERING THE UPPER AND LOWER PART OF THE BED. THE PATIENT EXPERIENCED A URINARY TRACT INFECTION (UTI), BUT IT IS UNKNOWN WHETHER THE UTI WAS RELATED TO THE DEVICE. MEDICAL INTERVENTION WAS UNKNOWN. PER FOLLOW UP VIA EMAIL ON 03-AUG-2023, THE PATIENT¿S WIFE CONFIRMED THE STATLOCK STABILIZATION DEVICE COULD NOT BE CLOSED/SNAPPED INTO PLACE IF THE CATHETER TUBE WAS NOT IN PROPER POSITION. SHE REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2023. NO ADDITIONAL INFORMATION WAS PROVIDED. PER CLINICAL FOLLOW UP VIA PHONE ON 09-AUG-2023, THE PATIENT¿S WIFE REPORTED THE PATIENT HAD BRAIN CANCER. HE WAS RECEIVING RADIATION THERAPY AND CHEMOTHERAPY. HE WAS ADDITIONALLY PREPARING TO BEGIN IMMUNOTHERAPY. THE PATIENT WAS INITIALLY ADMITTED TO THE HOSPITAL DUE TO LOSS OF MOBILITY. WHILE THE PATIENT WAS IN THE HOSPITAL, HE EXPERIENCED DIFFICULTY URINATING. THE PATIENT WAS DIAGNOSED WITH A UTI, AND A FOLEY CATHETER WAS INSERTED. WHILE THE PATIENT WAS IN THE HOSPITAL, THE FOLEY CATHETER WAS SECURED IN PLACE WITH TAPE. THE PATIENT WAS DISCHARGED FROM THE HOSPITAL ON (B)(6) 2023 WITH THE FOLEY CATHETER IN PLACE. THE HOME HEALTH NURSE SUPPLIED THE STATLOCK STABILIZATION DEVICES. THE PATIENT¿S WIFE CONFIRMED THE STATLOCK WAS APPLIED CORRECTLY, AND THE CATHETER TUBING WAS IN CORRECT POSITION WHEN SHE NOTED THERE WAS NO URINE FLOW. WHEN SHE REMOVED THE FOLEY CATHETER TUBING FROM THE STATLOCK DEVICE, URINE BEGAN TO FLOW. THE STATLOCK WAS REPLACED, BUT THE SAME ISSUE OCCURRED. ON (B)(6) 2023, THE PATIENT EXPERIENCED RECTAL BLEEDING, AND HE WAS ADMITTED TO THE HOSPITAL DUE TO DIVERTICULITIS. THE PATIENT¿S WIFE STATED WHILE THE PATIENT WAS IN THE HOSPITAL THE SECOND TIME, THEY CONTINUED TO HAVE ISSUES WITH THE FOLEY CATHETER NOT DRAINING BECAUSE OF THE STATLOCK DEVICE. THE FOLEY CATHETER TUBING HAD TO BE REMOVED FROM THE STATLOCK IN ORDER FOR URINE TO DRAIN. THE 61-YEAR-OLD MALE PATIENT PASSED AWAY IN THE HOSPITAL ON (B)(6) 2023 DUE TO COMPLICATIONS FROM BRAIN CANCER. THE PATIENT¿S WIFE CONFIRMED THE PATIENT¿S DEATH WAS UNRELATED TO THE STATLOCK DEVICE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE UNDERSTANDING FOR CLOSING OF RETAINER LID. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: INDICATIONS FOR USE: THE STATLOCK DEVICE IS A STABILIZATION DEVICE FOR COMPATIBLE CATHETERS. CONTRAINDICATIONS: KNOWN TAPE OR ADHESIVE ALLERGIES. WARNINGS AND PRECAUTIONS: 1. DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, DIAPHORETIC OR NON-ADHERENT SKIN, OR WHEN THE ACCESS DEVICE IS NOT MONITORED DAILY. 2. OBSERVE UNIVERSAL BLOOD AND BODY FLUID PRECAUTIONS AND INFECTION CONTROL PROCEDURES, DURING APPLICATION AND REMOVAL OF THE STATLOCK DEVICE. 3. MINIMIZE CATHETER MANIPULATION DURING APPLICATION AND REMOVAL OF THE STATLOCK DEVICE. 4. DAILY MAINTENANCE: A. THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY INDICATED, AT LEAST EVERY SEVEN DAYS. B. IF PAD BECOMES SOILED, WASH WITH SOAP/WATER, SALINE OR HYDROGEN PEROXIDE. DO NOT USE ALCOHOL OR PREPACKAGED BATHING SYSTEMS, WHICH COULD LEAD TO EARLY LIFTING. C. IF SHOWERING/BATHING, COVER WITH PLASTIC WRAP OR WATERPROOF DRESSING. D. CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. E. USE CLINICAL JUDGMENT ON THE REMOVAL OF THE STATLOCK STABILIZATION DEVICE IF THE PATIENT EXPERIENCES ANY FLUID SHIFTS THAT MAY INTERFERE WITH SKIN INTEGRITY. APPLICATION TECHNIQUE: PREPARATION 1. PLACE FOLEY CATHETER INTO RETAINER. DIRECTIONAL ARROW SHOULD POINT TOWARDS CATHETER TIP, AND BALLOON INFLATION ARM SHOULD BE NEXT TO THE CLAMP HINGE. 2. CLOSE LID, BEING CAREFUL TO AVOID PINCHING THE CATHETER. 3. IDENTIFY SECUREMENT SITE BY LAYING THE DEVICE RETAINER ON THE FRONT OF THE THIGH, LEAVING 1 INCH OF CATHETER SLACK BETWEEN INSERTION SITE AND THE STATLOCK DEVICE RETAINER. 4. AFTER PLACING THE STATLOCK STABILIZATION DEVICE OFF TO THE SIDE, CLEANSE AND DEGREASE THE SECUREMENT SITE WITH ALCOHOL PER HOSPITAL POLICY. LET SKIN DRY. 5. APPLY SKIN PROTECTANT, IN DIRECTION OF HAIR GROWTH, TO AREA LARGER THAN SECUREMENT SITE. ALLOW TO DRY COMPLETELY (10-15 SECONDS). 6. USING PERMANENT MARKER, WRITE INITIALS AND DATE OF APPLICATION ON THE STATLOCK DEVICE ANCHOR PAD. NOTE: ALWAYS SECURE CATHETER INTO THE STATLOCK DEVICE RETAINER BEFORE APPLYING ADHESIVE PAD ON SKIN. PLACE AND PEEL: 7. ALIGN THE STATLOCK STABILIZATION DEVICE OVER SECUREMENT SITE LEAVING 1 INCH OF CATHETER SLACK. MAKE SURE LEG IS FULLY EXTENDED. 8. WHILE HOLDING THE RETAINER TO KEEP THE PAD IN PLACE, PEEL AWAY PAPER BACKING, ONE SIDE AT A TIME AND PLACE TENSION-FREE ON SKIN. REMOVAL TECHNIQUE: DISENGAGE 1. OPEN RETAINER BY PRESSING RELEASE BUTTON WITH THUMB, THEN LIFT TO OPEN. 2. REMOVE FOLEY CATHETER FROM THE STATLOCK DEVICE. DISSOLVE: 3. WIPE THE EDGE OF THE PAD USING AT LEAST 5-6 ALCOHOL PADS UNTIL A CORNER LIFTS. THEN CONTINUE TO STROKE UNDERSURFACE OF PAD WITH ALCOHOL TO DISSOLVE ADHESIVE PAD AWAY FROM SKIN. DO NOT PULL OR FORCE PAD TO REMOVE. THE STATLOCK STABILIZATION DEVICE SHOULD BE MONITORED DAILY AND REPLACED WHEN CLINICALLY INDICATED, AT LEAST EVERY 7 DAYS. THE CATHETER INSERTION SITE SHOULD BE TREATED PER LOCAL POLICY RECOMMENDATIONS. THE STATLOCK DEVICE IS CONTRAINDICATED ON PATIENTS WITH KNOWN TAPE AND ADHESIVE ALLERGIES. ALCOHOL AND ACETONE MAY WEAKEN THE ADHESIVE BOND BETWEEN THE STATLOCK DEVICE PAD AND THE SKIN. CARE SHOULD BE TAKEN WHEN USING THESE SOLUTIONS WHILE PERFORMING CATHETER SITE CARE. PLEASE CONSULT PRODUCT INSERTS AND LABELS FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. FOR ADDITIONAL INFORMATION ON THE STATLOCK FOLEY STABILIZATION DEVICE, PLEASE CALL YOUR LOCAL BARD REPRESENTATIVE. CORRECTION: B, D. UDI FORMAT UPDATED WITH AVAILABLE PRODUCT INFORMATION PER GUDID. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03526 | 2023-08-14 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK SLIDING POST TRICOT CRESCENT PAD INCLUDES SKIN PREP | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | PIC0220CE |  | IT WAS REPORTED PATIENT WITH PICCLINE, WHERE THE ATTACHMENT TO THE CATHETER IS BROKEN, THERE WAS A RISK THAT THE TUBE HAD GONE OUT. THE ATTACHMENT TO THE STATLOCK WAS THE PART OF THE HARD PLASTIC COMPLETELY RELEASED FROM THE DRESSING THAT WAS TAPED TO THE SKIN. THE HOSE HAD NO STOP, THE SAFETY MECHANISM WAS BROKEN. OCCURRED DURING USE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03525 | 2023-08-14 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945J |  | IT WAS REPORTED THAT THE TIP OF THE GUIDEWIRE WAS ROUNDED AND UNABLE TO PASS THROUGH THE TIP OF THE DILATOR. THE TIP OF THE DILATOR SPREAD OVER TIME AND WAS ABLE TO BE INSERTED. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03524 | 2023-08-14 | 2023-08-01 | I, n, j, u, r, y | Manufacturer report | POWERPICC CATHETER W/SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY CUSTOMER THAT AFTER ACCESS DURING A PICC PROCEDURE, NURSE FELT RESISTANCE WHEN ADVANCING THE GUIDEWIRE INTO THE PATIENT¿S VESSEL. THE NURSE WITHDREW THE GUIDEWIRE AND LEFT ACCESS NEEDLE IN PLACE AND REASSESSED VESSEL. AFTER VESSEL ASSESSMENT, NURSE REINSERTED GUIDEWIRE INTO PATIENT¿S VESSEL AND CONTINUED TO FEEL RESISTANCE AND WITHDREW AGAIN. ABOUT 2 CM OF THE TIP OF THE GUIDEWIRE WAS DISLODGED INTO THE PATIENT¿S ANATOMY. WHEN EVALUATING GUIDEWIRE AFTER ATTEMPTED INSERTION, NURSE NOTICED THE TIP OF THE GUIDE WIRE UNRAVELED. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03523 | 2023-08-14 | 2023-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT WHEN PLACING THE PICC CATHETER IN THIS PATIENT FROM THE HEMATOLOGY DEPARTMENT, THE PART BETWEEN THE HEPARIN CAP AND THE CATHETER PORT LEAKED LIQUIDS. NO OTHER INFORMATION WAS REPORTED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03522 | 2023-08-14 | 2023-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT WHEN PLACING THE PICC CATHETER IN THIS PATIENT FROM THE HEMATOLOGY DEPARTMENT, THE PART BETWEEN THE HEPARIN CAP AND THE CATHETER PORT LEAKED LIQUIDS. NO OTHER INFORMATION WAS REPORTED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06082 | 2023-08-14 | 2023-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT'S FOLEY CATHETER HAD SLIPPED WHEN THE WARD WAS INSPECTED ON (B)(6) 2023. DURING THE INSPECTION, IT WAS FOUND THAT THE URINARY CATHETER BALLOON WAS RUPTURED AND THE URINARY CATHETER WAS INTACT. THE PATIENT DID NOT COMPLAIN OF ANY SPECIAL DISCOMFORT, AND THE DOCTOR ON DUTY WAS IMMEDIATELY INFORMED TO RE-INDWELLING THE CATHETER BASED ON THE DOCTOR'S ADVICE. THE URINARY CATHETER DRAINAGE WAS SMOOTH AND LIGHT RED URINE WAS INDUCED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 06AUG2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. THERE WAS NO PATIENT INJURY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03515 | 2023-08-14 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT DURING THE PROCESS OF PICC CATHETER PLACEMENT, THE CUSTOMER FOUND THAT THE GAP BETWEEN CONNECTING PARTS OF THE CONNECTOR FOR THE GROSHONG KIT WAS A LITTLE BIT WIDER. NO PROBLEM OCCURRED AFTER CONNECTION WAS COMPLETED ON SITE. THE NEXT DAY, THE CONNECTED PARTS WERE SEPARATED. THE CATHETER WAS TRIMMED IMMEDIATELY AND A NEW CONNECTING ACCESSORY WAS USED. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF APPLICABLE INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LOOSE TWO-PIECE CONNECTOR WAS CONFIRMED. THE RETURNED SAMPLE WAS FOUND TO EXHIBIT THE SAME FEATURES AS A KNOWN ISSUE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03497 | 2023-08-14 | 2021-07-29 | D, e, a, t, h | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT TEN MONTHS AND TWENTY-SIX DAYS POST A PORT PLACEMENT VIA THE LEFT INTERNAL JUGULAR VEIN, THE CATHETER HAD ALLEGEDLY FRACTURED. IT WAS FURTHER REPORTED THAT THE FRACTURED CATHETER FRAGMENT HAD ALLEGEDLY MIGRATED INTO THE PATIENT¿S MAIN PULMONARY ARTERY UNDER COMPUTED TOMOGRAPHY EXAMINATION. FURTHERMORE, PATIENT ALLEGEDLY EXPERIENCED CHEST PAIN AND THE RETAINED PORTION OF THE FRACTURED CATHETER FRAGMENT WAS SUCCESSFULLY RETRIEVED WITH A SNARE DEVICE, AND THE PORT AND CATHETER WERE REMOVED INTACT. REPORTEDLY, THE PATIENT EXPIRED DUE TO METASTATIC TO MEDIASTINUM INVASIVE DUCTAL CARCINOMA OF RIGHT BREAST AND PNEUMONITIS. HOWEVER, THE RELATIONSHIP OF PATIENT'S DEATH WITH THE DEVICE WAS NOT PROVIDED, BUT THE DEVICE MAY OR MAY NOT HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. MEDICAL RECORDS WERE PROVIDED AND REVIEWED. THE MEDICAL RECORD SHOWED THAT THE CATHETER HAD FRACTURED AND THEN MIGRATED INTO THE PATIENT'S MAIN PULMONARY ARTERY. THE PORT AND CATHETER REMAINING WERE REMOVED ALONG WITH THE MIGRATED PORTION OF THE CATHETER. THE PATIENT LATER EXPIRED AND THE REASON FOR DEATH WAS METASTATIC TO MEDIASTINUM INVASIVE DUCTAL CARCINOMA OF RIGHT BREAST AND PNEUMONITIS. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED COMPLETE CATHETER FRACTURE WITH MATERIAL SEPARATION AND MIGRATION OF A CATHETER FRAGMENT. A RELATIONSHIP BETWEEN THE PATIENT DEATH AND THE DEVICE ISSUES CANNOT BE CONFIRMED AT THIS TIME. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10:G3, H6 (METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03518 | 2023-08-14 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2174108 |  | IT WAS REPORTED THEY HAVE NOTICED A EXTRAVASATION AND UPON REMOVAL, IT APPEARS THAT THERE IS A BREAK IN THE DEVICE. THEY EXPLAINED THAT AFTER SEEING THE EXTRAVASATION THAT WAS OCCURRING, THEY DECIDED TO REMOVE THE PICC AND REALIZED THAT THERE WAS A BREAKAGE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03507 | 2023-08-14 | 2022-04-18 | I, n, j, u, r, y | Manufacturer report | POWERPORT TI ISP, 8 FR CHRONOFLEX, SUTURE PLUG, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1708060 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT ONE DAY POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, THE PORT ALLEGEDLY DETACHED FROM THE CATHETER. IT WAS FURTHER REPORTED THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03503 | 2023-08-14 | 2020-01-21 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1808060 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SIX YEARS, SEVEN MONTHS, AND SEVENTEEN DAYS POST PORT PLACEMENT FOR THE PURPOSE OF ONGOING MULTIPLE MYELOMA TREATMENT, THE PATIENT WAS DIAGNOSED WITH A CANDIDEMIA INFECTION. IT WAS FURTHER REPORTED THAT THE PATIENT UNDERWENT SURGERY TO REMOVE THE DEFECTIVE DEVICE AND TO CORRECT COMPLICATIONS. HOWEVER, THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 05/2015). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03502 | 2023-08-14 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1859660 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, ONE HALF OF THE PLASTIC RING OF THE SHEATH ALLEGEDLY BROKE OFF OF THE SHEATH. IT WAS FURTHER REPORTED THAT IT WAS ALLEGEDLY DIFFICULT TO INSERT THE CATHETER INTO THE SHEATH AND VEIN. REPORTEDLY, THE CATHETER WAS ABLE TO BE GET INTO THE VEIN AND THE PROCEDURE WAS COMPLETED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE FULLY PEELED 10.0FR PEEL-APART SHEATH WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC AND FUNCTIONAL EVALUATIONS WERE PERFORMED. ONE VALVE CAP WAS RECEIVED DETACHED FROM THE T-HANDLE. THE OTHER HALF OF THE VALVE CAP WAS FULLY ATTACHED TO THE T-HANDLE. THEREFORE, THE INVESTIGATION HAS BEEN CONFIRMED FOR IDENTIFIED MATERIAL SEPARATION AND LOSS OF BOND ISSUES. THE INVESTIGATION IS UNCONFIRMED FOR FRACTURE ISSUE AS THE MORE APPROPRIATE FAILURE MATERIAL SEPARATION HAS BEEN CONFIRMED. ALSO, THE INVESTIGATION REMAINS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO INSERT AS THE EXACT CIRCUMSTANCES COULD NOT BE REPRODUCED AND THE CATHETER WAS NOT RETURNED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03514 | 2023-08-14 | 2023-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2022. ON (B)(6) 2023, A FEW HOURS BEFORE THE EVENT, THE BLOOD SAMPLING AND INFUSION COULD BE PERFORMED, HOWEVER, AFTER THAT THE CATHETER WAS OCCLUDED. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF AN OCCLUDED CATHETER WAS UNCONFIRMED BECAUSE THE PROBLEM COULD NOT BE REPRODUCED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR S/L GROSHONG CATHETER. THE SAMPLE WAS RECEIVED ASSEMBLED WITH A TWO-PIECE CONNECTOR. USAGE RESIDUES WERE EVIDENT THROUGHOUT THE SAMPLE. MICROSCOPIC INSPECTION OF THE SAMPLE DID NOT REVEAL ANY EVIDENCE OF DEVICE OCCLUSION. THE VALVE APPEARED WELL-FORMED. AN ATTEMPT TO INFUSE WATER THROUGH THE SAMPLE USING A 12ML SYRINGE REVEALED THE SAMPLE TO BE PATENT TO INFUSION AND ASPIRATION WITH NO OBSERVED LEAKS. THE EFFORT REQUIRED TO FLUSH AND ASPIRATE THE SAMPLE WAS COMPARABLE TO THAT REQUIRED FOR A SIMILAR, NON-COMPLAINANT DEVICE. NO DEFICIENCIES WERE DISCOVERED DURING EVALUATION OF THE RETURNED SAMPLE. CONSEQUENTLY THIS COMPLAINT IS UNCONFIRMED AT THIS TIME. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03499 | 2023-08-14 | 2023-05-18 | I, n, j, u, r, y | Manufacturer report | MRI POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK MRI POWER PORT |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT APPROXIMATELY FOUR MONTHS POST PORT PLACEMENT FOR THE PURPOSE OF CHEMOTHERAPY TREATMENT TO TREAT CANCER, THE PATIENT WAS DIAGNOSED WITH BACTERIAL CELLULITIS, WHICH WAS FOUND TO BE CAUSED BY THE PORT CATHETER AND CONTAINED METHICILLIN-SENSITIVE STAPHYLOCOCCUS AUREUS AND STAPHYLOCOCCUS BACTEREMIA. IT WAS FURTHER REPORTED THAT THE PORT CATHETER WAS REMOVED. REPORTEDLY, THE PATIENT ADDITIONALLY REQUIRED INTRAVENOUS ANTIBIOTICS AND WAS HOSPITALIZED. HOWEVER, THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03500 | 2023-08-14 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0602680 |  | IT WAS REPORTED THAT DURING THE PORT PLACEMENT PROCEDURE, THERE WAS ALLEGEDLY A FRACTURE BETWEEN THE RESERVOIR'S CONNECTOR WITH THE CATHETER. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE BARDPORT MRI IMPLANTABLE PORT KIT WAS RECEIVED FOR EVALUATION. THREE PHOTOS WERE PROVIDED FOR REVIEW. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. THE PORT STEM WAS NOTED TO BE MISSING FROM THE BARDPORT MRI IMPLANTABLE PORT. A CIRCUMFERENTIAL BREAK WAS NOTED ON THE PORT BODY. THE EDGES OF THE COMPLETE CIRCUMFERENTIAL BREAK ON THE PORT STEM AREA WERE NOTED TO BE UNEVEN AND THE SURFACE WAS NOTED TO BE GRANULAR. THE PORT STEM SEGMENT WAS REMOVED FROM INSIDE THE CATHETER. THE PHOTOS SHOWS THE COMPLETE BREAK OF THE PORT STEM AND SEPARATED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND THE IDENTIFIED MATERIAL SEPARATION ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: B5, D4 (EXPIRATION DATE: 03/2026), G3, H6 (DEVICE). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03501 | 2023-08-14 | 2023-07-21 | I, n, j, u, r, y | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0607540 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIMES POST PORT PLACEMENT FOR ONGOING TREATMENT OF BREAST CANCER, THE PORT ALLEGEDLY COMPRESS OR PINCH OFF. IT WAS FURTHER REPORTED THAT PATIENT WAS HOSPITALIZED FOR OCCLUSION OF THE CATHETER AND TO DECLOT AN EMBOLISM IN RELATION TO THE CATHETER AND WAS DIAGNOSED WITH STAPH AUREUS BACTEREMIA. OBSTRUCTION AT THE TIP OF THE CATHETER WAS NOTED, WHICH MAY BE DUE TO THROMBUS. THE PATIENT ALSO REPORTED OF HAVING GENERALIZED BODY PAIN AND EXTRAVASATION OF CONTRAST AROUND THE CATHETER EXTENDING PROXIMALLY WAS ALSO OBSERVED. HOWEVER, THE CURRENT STATUS OF THE PATENT IS UNKNOWN. | H11: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. MEDICAL RECORDS WERE PROVIDED AND REVIEWED. THE MEDICAL RECORD ALLEGE THAT THE RIGHT INTERNAL JUGULAR CENTRAL LINE CATHETER PLACEMENT WAS PERFORMED FOR CANCER TREATMENT INFUSION. APPROXIMATELY ONE YEAR LATER, PORT WAS GIVEN AN ATTEMPT TO ASPIRATE EVERY 30 MINUTES FOR TWO HOURS WITH NO BLOOD RETURN. OBSTRUCTION AT THE TIP OF THE CATHETER WAS NOTED, WHICH MAY BE DUE TO THROMBUS. PATIENT COMPLAINED OF GENERALIZED BONE PAIN. EXTRAVASATION OF CONTRAST AROUND THE CATHETER EXTENDING PROXIMALLY WAS ALSO OBSERVED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED OBSTRUCTION OF FLOW AND SUCTION ISSUE. FURTHERMORE, THE CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED FOR BACTERIAL INFECTION. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. B3, D4 (MEDICAL DEVICE EXPIRATION DATE, UNIQUE IDENTIFIER (UDI) #), G3, H6 (PATIENT, DEVICE, COMPONENT, RESULT). SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03498 | 2023-08-14 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT WITH ATTACHABLE 8F GROSHONG SINGLE-LUMEN | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1878070 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT ONE YEAR SEVEN MONTHS AND FOURTEEN DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN FOR ADMINISTRATION OF THYROID MEDICATION INFUSIONS, THE PORT ALLEGEDLY HAD DIFFICULTY IN ASPIRATING, OR DRAWING FLUIDS IN AND OUT. IT WAS FURTHER REPORTED THAT THE DEVICE HAD BROKEN AND WAS NEARLY COMPLETELY SEVERED IN THE MID-PORTION OF THE DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03506 | 2023-08-14 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT APPROXIMATELY A YEAR AND A HALF POST PORT PLACEMENT, THE PORT ALLEGEDLY FELT FLAT. IT WAS FURTHER REPORTED THAT THE PORT ALLEGEDLY FELT LIKE A HARD BALL. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03505 | 2023-08-14 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 5833690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE IN THE RIGHT INTERNAL JUGULAR VEIN, THE PUNCTURE NEEDLE IN THE TUBE PACK WAS ALLEGEDLY CRACKED AT THE BEHIND. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 11/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
2020394-2023-00582 | 2023-08-14 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER IQ ULTRASONIC | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | XCTO146 |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE TIP OF CATHETER WAS ALLEGEDLY SEPARATED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
1018233-2023-06059 | 2023-08-14 | 2023-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER WAS HAVING TROUBLE WITH THE FOLEY CATHETER, WHICH HAD CLOGGED UP AND WANTED TO KNOW IF THERE WAS ANYTHING THEY COULD DO TO HELP FIX IT. THE CUSTOMER HAD THIS FOLEY CATHETER FOR A YEAR, AND THEY CHANGED THE BAG EVERY THIRTY DAYS BUT NOT THE CATHETER ITSELF. PATIENTS HAVE HAD NIGHTS WHERE THEY WOKE UP AND HAD TO GO TO THE BATHROOM TO TRY TO GET IT UNCLOGGED, WHICH WAS VERY EMBARRASSING. THE PATIENT DID NOT FEEL ANY PAIN FROM IT. THEY CHANGE THE DRAINAGE BAG BUT NOT THE TUBE THAT GOES DIRECTLY INTO THEIR BODY. REPRESENTATIVE REPLIED THAT BD MAKES SIZES THAT GO UP TO 24 FR, BUT IT WOULD BE AT THE DISCRETION OF THEIR CARE TEAM IF THEY NEEDED TO SIZE UP OR WHAT TYPE OF CATHETER THE PATIENT WAS USING. ADVISED THE CUSTOMER THAT THEY WOULD NEED TO SPEAK DIRECTLY WITH THEIR PHYSICIAN REGARDING THIS MATTER AS THEY WERE UNABLE TO GIVE MEDICAL ADVICE ON THIS ISSUE. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, THE ISSUE HAD BEEN RESOLVED, AND CUSTOMER WAS ABLE TO DRAIN THE URINE PROPERLY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO DRAINAGE EYE OCCLUSION BLOCKED DRAINAGE LUMEN CAUSED BY PLC FAILURE AND ALSO MIGHT BE CONTRIBUTE BY NO DRAINAGE EYE OR USER RELATED (SALT ACCUMULATION). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-08
3006260740-2023-03511 | 2023-08-14 | 2023-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F SINGLE-LUMEN CATHETER BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6175118 |  | IT WAS REPORTED FAILED TO SEPARATE FROM THE PEEL-AWAY. ADDITIONAL INFORMATION RECEIVED ON 31 JULY 2023: CUSTOMER REPORTED THERE WAS NO HARM OR IMPACT TO THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF DIFFICULTY REMOVING THE DILATOR FROM THE SHEATH WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 5FR MICROINTRODUCER DILATOR/SHEATH. THE DILATOR WAS INLAID THROUGH THE PEEL-APART SHEATH; HOWEVER, THE DILATOR WAS PARTIALLY WITHDRAWN. NO OBVIOUS USE RESIDUES WERE OBSERVED. A SHARP KINK WAS OBSERVED IN THE DILATOR APPROXIMATELY 1CM DISTAL OF THE COLLAR. MICROSCOPIC INSPECTION OF THE VESSEL DILATOR CONFIRMED THE PRESENCE OF A SHARP KINK. MATERIAL DISCOLORATION WAS OBSERVED AT THE KINK SITE. INSPECTION OF THE DISTAL TIP WAS UNREMARKABLE. MICROSCOPIC INSPECTION OF THE PEEL-APART SHEATH REVEALED MATERIAL BUCKLING JUST DISTAL OF THE PULL TABS. THE TIP OF THE SHEATH APPEARED UNREMARKABLE. THE KINK IN THE DILATOR CORRESPONDED TO THE BUCKLED REGION OF THE SHEATH WHEN THE DILATOR WAS FULLY INSERTED. THE KINKS IN BOTH THE SHEATH AND DILATOR SUGGESTED THAT THE FULLY INSERTED ASSEMBLY WAS MANIPULATED IN A WAY THAT RESULTED IN PERMANENT DEFORMATION. THAT DEFORMATION LIKELY CAUSED THE REPORTED RESISTANCE WHEN ATTEMPTING TO SEPARATE THE TWO COMPONENTS. IT COULD NOT BE DETERMINED IF THE MANIPULATION OCCURRED DURING KIT ASSEMBLY OR AFTER THE KIT WAS OPENED. CONSEQUENTLY, THIS COMPLAINT IS CONFIRMED AT ¿CAUSE UNKNOWN¿ AT THIS TIME. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03513 | 2023-08-14 | 2023-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D3 |  | IT WAS REPORTED "CUSTOMER REPORTED A HAIR FOUND IN A 4 FR PICC LINE KIT. THERE WAS NO PATIENT HARM". NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03504 | 2023-08-14 | 2023-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 5733690 |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY CAME OUT BY ITSELF. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03539 | 2023-08-15 | 2023-08-01 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT PATIENT HAS BEEN INSERTED PICC LINE FOR CA COLON, AFTER FEW DAYS CAME BACK WITH PAIN IN UPPER ARM, DIAGNOSED AS DVT INVOLVING THE RIGHT UPPER LIMB VEIN THRU DOPPLER IMAGING. CHEMO REGIMEN(FOLFOX-6) HAS BEEN INTERRUPTED, DIFFICULT TO USE PICC LINE. CATHETER WAS NOT REMOVED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03540 | 2023-08-15 | 2023-08-01 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT PATIENT HAS BEEN INSERTED PICC LINE FOR NHL AFTER FEW DAYS CAME BACK WITH PAIN IN UPPER ARM, DIAGNOSED AS DVT INVOLVING THE RIGHT UPPER LIMB VEIN THROUGH DOPPLER IMAGING. CHEMO REGIMEN (DA-R-EPOCH REGIMEN) HAS BEEN INTERRUPTED AND ORAL ANTICOAGULANT STARTED, DIFFICULT TO USE PICC LINE. CATHETER WAS NOT REMOVED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03537 | 2023-08-15 | 2023-07-31 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SIL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED, THAT PATIENT HAS BEEN INSERTED PICC LINE FOR SARCOMA. AFTER FEW DAYS CAME BACK WITH PAIN IN UPPER ARM. DIAGNOSED AS DVT INVOLVING THE RIGHT UPPER LIMB VEIN WITH DOPPLER IMAGING. CHEMO REGIMEN (IA REGIMEN) HAS BEEN INTERRUPTED. AND ORAL ANTICOAGULANT STARTED. DIFFICULT TO USE PICC LINE. CATHETER WAS NOT REMOVED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03536 | 2023-08-15 | 2023-07-31 | I, n, j, u, r, y | Manufacturer report | ACCUCATH ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY CUSTOMER, THAT APPEARS PART OF THE CATHETER AND GUIDEWIRE BROKE OFF IN THE PATIENT WHEN REMOVAL WAS ATTEMPTED. PATIENT HAD AN US AND XRAY. AND WAS EVALUATED BY VASCULAR SURGERY. PATIENT WENT TO OR FOR REMOVAL OF FOREIGN BODY ON (B)(6) 2023. IT WAS NOT FOUND. SINCE PATIENT WAS SUCH A DIFFICULT ACCESS A TL RT IJ CVC WAS INSERTED BY IR ON (B)(6) 2023 FOR MED ADMINISTRATION. NO LIFE THREATENING SITUATION. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: INSERTED NEEDLE WITH US. LOOKED LIKE IT WAS IN THE VEIN. SLID GUIDEWIRE OUT SMOOTHLY AND ADVANCED CATHETER INTO THE VEIN. RETRACTED THE NEEDLE. INSERTER COULD NOT GET THE GUIDEWIRE OUT. HE DECIDED TO ABORT THE PROCEDURE AND WENT TO REMOVE THE ENTIRE CATHETER AND GUIDEWIRE. WHEN HE ATTEMPTED TO REMOVE, THE CATHETER BROKE OFF IN THE PATIENT'S ARM. APPROXIMATELY 1 CM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03533 | 2023-08-15 | 2023-07-31 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT. PATIENT HAS BEEN INSERTED PICC LINE FOR CA STOMACH, AFTER FEW DAYS CAME BACK WITH PAIN IN UPPER ARM, DIAGNOSED AS DVT INVOLVING THE RIGHT UPPER LIMB VEIN. CHEMO REGIMEN HAS BEEN INTERRUPTED AND ORAL ANTICOAGULANT STARTED, DIFFICULT TO USE PICC LINE. ADDITIONAL INFORMATION RECEIVED 08/02/2023: DEVICE INSERTED BASILIC VEIN ON UPPER ARM. PATIENT COMPLAINED OF PAIN ON SITE OF INSERTION AT UPPER ARM. CATHETER WAS NOT REMOVED. PATIENT WAS DIAGNOSED WITH CA STOMACH. ORAL ANTICOAGULANT WAS GIVEN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03534 | 2023-08-15 | 2023-07-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED CUSTOMER RETURNED A 5FR MICROINTRODUCER WITH A PLASTIC BURR ON THE END. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. H3 OTHER TEXT: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03528 | 2023-08-15 | 2023-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED BY CUSTOMER THAT THE CUSTOMER PLACED A PICC IN THE ICU SATURDAY AND THE T EXTENSION THAT THE GUIDE WIRE GOES INTO THE PICC HAD LEAKED. THEY FOUND THIS AFTER THEY HAD PLACED THE PICC AND WAS NEARLY DONE WITH EVERYTHING, THEY WENT TO CHECK FOR BLOOD RETURN AND GOT AIR BUBBLES. AFTER THEY SAW THAT, THEY FOUND THE MEMBRANE WAS LOOSE IN THE CONNECTOR AND THEY WAS ABLE TO PLUG IT WITH MY THUMB AND FLUSH AND GET BLOOD RETURN WITHOUT DIFFICULTY. THE CUSTOMER DIDN'T SAVE THE T CONNECTOR BECAUSE IT HAD BLOOD ON IT FROM TRYING TO FLUSH AND GET BLOOD RETURN. RESPONSE RECEIVED (B)(6) 2023. PATIENT NEEDED PICC LINE FOR PRESSORS, PLANNED TO GET LEVOPHED, POTASSIUM AND MAGNESIUM REPLACEMENT, CEFEPIME, AND HYDROCORTISONE. CAUSE A COUPLE MINUTE DELAY DURING INSERTION OF PICC LINE, NO DELAY IN CARE TO PATIENT. UPON INSERTION OF PICC WHEN ASPIRATING FOR BLOOD NOTICED PINK TINGED SALINE AND AIR BUBBLES TRAVELING INTO SYRINGE, INITIALLY WAS ABLE TO FLUSH LINE WITHOUT COMPLICATION. WHEN ATTEMPTING TO CHECK PLACEMENT AFTER ACCEPTABLE 3CG READING IS WHEN I NOTICED JUST BUBBLES AND NO BLOOD RETURN AND WHEN I FLUSHED THE LINE IT STARTED TO LEAK FROM THE MEMBRANE THAT THE WIRE TRAVELS THROUGH THAT IS CONNECTED TO THE T-CONNECTOR. I QUICKLY PLACED MY STERILE GLOVED HAND ON THE T-CONNECTOR AND SEALED IT WITH MY THUMB AND WAS ABLE TO GET PATENT BLOOD RETURN AND EASILY FLUSH THE LINE. NO ADVERSE EVENT OR INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03538 | 2023-08-15 | 2023-07-31 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT, PATIENT HAS BEEN INSERTED PICC LINE FOR CA COLON, AFTER FEW DAYS CAME BACK WITH PAIN IN UPPER ARM, DIAGNOSED AS DVT INVOLVING THE RIGHT UPPER LIMB VEIN THROUGH DOPPLER IMAGING. CHEMO REGIMEN(FOLFOX-6) HAS BEEN INTERRUPTED, DIFFICULT TO USE PICC LINE. CATHER WAS NOT REMOVED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-08
3006260740-2023-03535 | 2023-08-15 | 2023-07-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F TL MAX BARRIER PLUS KIT W/BIOPATCH | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1395108D4 |  | IT WAS REPORTED BY CUSTOMER THAT THEY FOUND A BUG IN 3 KITS. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: DID NOT USE ON THE PATIENT. LOOKS LIKE IT IS WITH THE STERILE SALINE FLUSHES. NO DIRECT PATIENT USE, NO LIFE THREATENING SITUATION, NO PATIENT HARM OR INTERRUPTION OF MEDS. THIS REPORT ADDRESSES THE FIRST KIT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-08
MW5144648 | 2023-08-15 | 2023-08-09 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PEDIATRIC CATH KIT 8FR | CATHETER, UROLOGICAL | C. R. BARD |  |  |  | ADDITIONAL INFORMATION RECEIVED ON OCTOBER 12, 20023 TO CHANGE MANUFACTURER'S NAME FOR REPORT NUMBER MW5144648. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-08

## Monthly Summary
Month | Count | Date
2023-02 | 266 | 2023-02-01
2023-03 | 756 | 2023-03-01
2023-04 | 812 | 2023-04-01
2023-05 | 898 | 2023-05-01
2023-06 | 882 | 2023-06-01
2023-07 | 746 | 2023-07-01
2023-08 | 400 | 2023-08-01

## Event Types
Event Type | Count
M, a, l, f, u, n, c, t, i, o, n | 3906
I, n, j, u, r, y | 841
D, e, a, t, h | 12
O, t, h, e, r | 1

## Brand Names
Brand Name | Count
BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | 262
BARDEX® LUBRICATH® FOLEY CATHETER | 257
BARDEX® I.C. FOLEY CATHETER | 211
BARDIA® FOLEY CATHETER SILICONE COATED | 171
BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | 151
BARDEX® FOLEY CATHETER SILICONE COATED | 122
MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | 96
LATEX FOLEY CATHETER | 94
PUREWICK FEMALE EXTERNAL CATHETER | 88
UNKNOWN POWER PORT | 83
BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | 79
MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | 78
POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | 73
MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | 72
MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | 67
UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | 66
MAGIC3 GO MALE CATHETER WITH SUREGRIP | 65
BARDEX® ALL-SILICONE FOLEY CATHETER | 49
BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | 49
GLIDEPATH | 46
POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | 41
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | 34
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | 32
UNKNOWN PORT | 27
X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | 27
UNKNOWN LATEX FOLEY CATHETER | 27
CROSSER | 26
BARD PICC | 26
HEMOSTAR | 24
ATLAS GOLD | 24
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | 24
UNKNOWN LATEX TEMP-SENSING CATHETER | 23
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | 22
POWERPORT ISP MRI 6F CHRONW/OS | 22
BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | 21
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | 21
POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | 20
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | 20
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | 19
BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | 18
POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | 18
APHERESIS PORT | 18
ULTRAVERSE RX | 17
POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | 17
BARD PICC CATHETER | 17
SURESTEP INTERMITTENT CATHETER TRAY | 17
BARD® X-FORCE® BALLOON DILATION CATHETER | 17
EQUISTREAM SPLIT TIP | 17
PERSONAL CATHETER¿ INTERMITTENT CATHETER | 17
POWERPICC PROVENA | 17
POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | 16
SHERLOCK 3CG+ SENSOR | 16
POWER PORT ISP M.R.I., 8FR GROSHONG | 16
GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | 16
BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | 16
ULTRAVERSE 035 | 15
MAGIC3¿ INTERMITTENT CATHETER | 15
BARD STATLOCK | 15
VACCESS | 14
SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | 14
CONQUEST | 14
BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | 14
BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | 14
BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | 13
BARD® SILASTIC® FOLEY CATHETER | 12
ACCUCATH, ACE 20GX2.25 | 12
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | 12
BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | 12
CATHETER STABILIZATION DEVICE | 12
POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | 12
BARDEX® SILICONE COATED FOLEY CATHETER | 11
POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | 11
POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | 11
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | 11
ALL SILICONE CATHETER | 11
POWERPICC PROVENA 4F DL MBPLUS (3CG) | 11
INTERMITTENT ROCHESTER CATHETERS | 10
POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | 10
POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | 10
POWERPORT ISP MRI 6CF INT W SP, ATT, SL | 10
SHERLOCK 3CG SENSOR (RADEL) | 10
BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | 10
M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | 10
CONQUEST 40 | 10
DOT PROVPPSOLO 4FDL MBP | 9
ACCUCATH ACE INTRAVASCULAR CATHETER (20G) | 9
POWERPICC SV CATHETER BASIC TRAY | 9
BARD POWERPICC | 9
HEMOSPLIT CATHETER | 9
POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | 9
ULTRASCORE 014 | 8
ATLAS | 8
BARDEX® LUBRICATH® HEMATURIA FOLEY CATHETER | 8
GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIG JAPAN HANG | 8
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/ MICROINTRODUCER) | 8
POWERPORT DUO M.R.I. IMPLANTABLE PORT, CHRONOFLEX DUAL-LUMEN, 9.5F | 8
BARDEX® LUBRICATH® PEDIATRIC FOLEY CATHETER | 8
DORADO | 8
ACCUCATH ACE INTRAVASCULAR KIT CATHETER (22G) | 8
BARDEX® LUBRICATH® FOLEY CATHETER COUDE TIP | 8
POWERPICC PROVENA 5F DL KIT W/BIOPATCH, PROBE COVER, & NEEDLE GUIDES | 8
BARDEX® LUBRICATH® FOLEY CATHETER CREEVY MODEL | 8
M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | 7
TEMPORARY PACING ELECTRODE CATHETER | 7
POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING | 7
POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | 7
HICKMAN PORT SINGLE LUMEN | 7
POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | 7
POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | 7
DOT,4FR POWERMIDLINE | 7
POWERPICC | 6
CROSSER IQ ULTRASONIC | 6
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F | 6
M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | 6
BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN, WHITE, 4.2F | 6
PPICC BASIC | 6
BARD POWERMIDLINE CATHETER BASIC KIT (4F) (SINGLE-LUMEN) (20CM) | 6
POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT | 6
FULL KIT 20G X 10 CM WITH BIOPATCH AND PROBE COVER | 6
GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (WITH MICROINTRODUCER) | 6
UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | 6
3FR SL PROVENA MIDLINE MAX BARRIER KIT W/BIOPATCH AND PROBE COVER | 6
POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | 6
FULL KIT 18G X 10 CM WITH GUARDIVA AND PROBE COVER | 6
POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) (MICROINTRODUCER) | 6
FOLEY CATHETER | 5
PRISTINE LONG-TERM HEMODIALYSIS CATHETER | 5
POWERPICC 4F SL CATHETER BASIC TRAY (W/ MICROINTRODUCER, NITINOL GUIDEWIRE) | 5
X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | 5
POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER, BELDAR NG) | 5
POWERPICC CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | 5
BARD DL POWERPICC | 5
TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | 5
POWERPICC PROVENA 3F SL MBPLUS (3CG, BIOPATCH, PROBE COVER) | 5
BARD POWERGLIDE | 5
BARD® TIEMANN MODEL UROLOGICAL CATHETER COUDE TIP | 5
BARD TRIALYSIS CATHETER | 5
POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING | 5
POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | 5
POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) | 5
GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) (WITH MICROINTRODUCER) | 5
GROSHONG 4F SINGLE-LUMEN PICC (60CM) | 5
FULL KIT 20G X 10 CM WITH PROBE COVER | 5
POWERGLIDE PRO MIDLINE CATHETER | 5
BARDEX® LUBRI-SIL® ALL-SILICONE TEMP SENSING FOLEY CATHETER | 5
UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH | 5
NIAGARA SLIM-CATH 15 CM PRE-CURVE KIT | 5
NAVARRE OPTI DRAIN | 5
POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC | 5
BARD MIDLINE | 5
UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 3.5F X 5CM | 4
POWERPORT CLEARVUE SLIM IMPLANTABLE PORT | 4
HEMOSPLIT LONG-TERM HEMODIALYSIS CATHETER, POLYURETHANE, KIT, 14.5F | 4
STATLOCK SLIDING POST TRICOT CRESCENT PAD INCLUDES SKIN PREP | 4
GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | 4
BARD® RUBBER UTILITY CATHETER | 4
UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH | 4
BARDEX® LUBRICATH® HYDROGEL-COATED FOLEY CATHETER KIT | 4
POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA | 4
POWERPICC SOLO CATHETER | 4
STATLOCK CV PLUS SLIDING POST, TRICOT HYDROCOLLOID CRESCENT PAD, PIGTAIL PAD | 4
BARD MAGIC 3 GO COUDE | 4
POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH) | 4
WIDEBAND¿ SELF-ADHERING MALE EXTERNAL CATHETER | 4
GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN | 4
POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY | 4
POWERGLIDE PRO MIDLINE CATHETER (20G) (10CM) | 4
DOT PROV 4F DL MAX 3CG | 4
POWERPORT TI L/P 6 CF INT WOSP ATT SL | 4
POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET 4F | 4
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | 4
POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX, 8F | 4
POWERPORT ISP IMPLANTABLE PORT, GROSHONG, 8F | 4
BARD® FLIP-FLO¿ CATHETER VALVE | 4
PUREWICK FEMALE EXTERNAL CATHETER LF HC | 4
BARDIA® FOLEY CATHETER INSERTION TRAY | 4
POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY | 4
UNIVERSAL MICROINTRODUCER KIT PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | 4
BARD TL PICC | 4
MAGIC 3 GO MALE COUDE CATHETER | 4
BARD POWERGLIDE MIDLINE CATHETER | 4
POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER) | 4
22 GAUGE POWERGLIDE PRO MIDLINE CATHETER 8CM BASIC KIT | 4
POWERPORT VUE M.R.I. IMPLANTABLE PORT | 4
POWER PORT MRI ISP, 8 FR. GROSHONG, INT W SP, ATTACHABLE SL | 4
DOT STC PPICC, 5FR TL, | 4
BARDIA® COUDE OLIVE TIP URETHRAL FOLEY CATHETER | 4
POWERPICC SOLO2 CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | 4
BARD® LUBRICATH® FOLEY CATHETER | 4
3FR SL PROVENA MIDLINE MAX BARRIER KIT | 4
FULL KIT 18G X 10 CM WITH BIOPATCH AND PROBE COVER | 4
POWERPORT ISP M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | 4
POWERPICC PROVENA CATHETER W/ SOLO VALVE TECHNOLOGY, SHERLOCK 3CG (TPS) | 4
POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | 4
BARDEX®LUBRICATH® FOLEY CATHETER ALCOCK MODEL | 4
PUREWICK MALE EXTERNAL CATHETER KIT | 3
POWERPICC 5F DUAL-LUMEN POLYURETHANE CATHETER RADIOLOGY BASIC TRAY | 3
BARDIA® URETHRAL CATHETER TRAY, CLEAR PLASTIC CATHETER | 3
POWER PORT MRI ISP, 8 FR GROSHONG, INT WO SP, SL | 3
POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS CATHETER | 3
POWERPICC 3F SL SOLO MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | 3
NIAGARA SLIM-CATH DIALYSIS CATHETER KIT 12F X 20CM | 3
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | 3
GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY W/MICROINTRODUCER | 3
POWERGLIDE PRO MIDLINE CATHETER (18G) (8CM) | 3
POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | 3
BARD POWERGLIDE MIDLINE | 3
FULL KIT 20G X 10 CM WITH GUARDIVA AND PROBE COVER | 3
PPICC, IR, 135 CM | 3
POWER PORT MRI ISP, IMPLANTABLE PORT, GROSHONG SINGLE LUMEN, 8F | 3
PUREWICK MALE EXTERNAL CATHETER | 3
MAGIC 3 GO 16 FR MALE COUDE CATHETER WITH SUREGRIP | 3
BARDEX® PEDIATRIC ALL-SILICONE FOLEY CATHETER | 3
REPLACEMENT CONNECTOR FOR 4F GROSHONG NXT CLEARVUE PICC WITH EXTENSION LEG | 3
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F MAXIMAL BARRIER TRAY | 3
POWERPORT TITANIUM ISP 6CF INT WSP ATT SL | 3
SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® LUBRICATH® FOLEY CATHETER | 3
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | 3
BARDEX® ALL-SILICONE TEMPERATURE-SENSING FOLEY CATHETER | 3
POWERPICC SOLO CATHETER 4F RADIOLOGY BASIC TRAY SL 70 CM NITINOL GUIDEWIRE | 3
ACCUCATH, ACE 18GX2.25 | 3
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC | 3
POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F FULL TRAY (DL) | 3
VACCESS CT POWER-INJECTABLE IMPLANTABLE PORT, CHRONOFLEX, 6F | 3
DELTEC POWER PORT-A-CATH II PORTS | 3
POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | 3
POWERPICC PROVENA SOLO 5F DL FT MAX BARRIER PLUS KIT | 3
STATLOCK IV ULTRA PED CATHETER RETAINER | 3
FLEXXICON DIALYSIS CATHETER KIT 11F X 20CM (SHORT-TERM) (DUAL-LUMEN) | 3
BROVIAC CV CATHETER, SINGLE-LUMEN, 4.2F | 3
DOT PROV SP 4F DL MAX | 3
DOT PPICC SOLO 4F SL 3CG | 3
DOT PROV 4F DL MAX TLS | 3
NIAGARA DIALYSIS CATHETER KIT 13.5F X 15CM (SHORT-TERM) (DUAL-LUMEN) | 3
NIAGARA DIALYSIS CATHETER KIT 13.5F X 20CM PRE-CURVED SHORT-TERM DL | 3
POWER PORT MRI ISP, 8FR GROSHONG, INT WO SP, ATTSL | 3
LOW ARTIFACT POWER PORT | 3
M.R.I. LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 7F | 3
UNKNOWN DIALYSIS | 3
POWERPICC SOLO CATHETER WITH (TPS) STYLET 5F FULL TRAY (DUAL-LUMEN) | 3
POWERGLIDE PRO MIDLINE CATHETER (22G) (8CM) | 3
POWERPORT TI ISP, 8 FR CHRONOFLEX, SUTURE PLUG, SL | 3
DECATHLON DF LONG-TERM HEMODIALYSIS CATHETER | 3
POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH, | 3
SHERLOCK 3CG* TCS SENSOR | 3
POWERPICC SOLO HF CATHETER WITH SHERLOCK 3CG TIP POSITIONING | 3
DOT PPICC SOLO 5F DL FT 3CG | 3
BARDEX LUBRI-SIL ALL-SILICONE TEMP SENSING FOLEY CATHETER | 3
TRUE DILATATION CATHETER | 3
LUBRI-SIL® I.C. COMPLETE CARE® ALL-SILICONE FOLEY CATHETER TRAY | 3
POWERGLIDE PRO MIDLINE CATHETER (20G) (8CM) | 3
UNIVERSAL MICROINTRODUCER PTFE PEEL-APART SHEATH 4.5F X 5CM | 3
POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH & PROBE COVER | 3
POWERPICC PROVENA SOLO 5F DL FT KIT W/GUARDIVA & PROBE COVER | 3
SHERLOCK 3CG+ SENSOR CE | 3
UNKNOWN FOLEY CATHETER | 3
ULTRASCORE 035 | 3
PPICC HF 5F TL | 3
POWERPICC CATHETER W/SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY W/MICROINTRODUCER | 3
POWERLOC SAFETY INFUSION SET 20G X 0.75IN WITH Y-SITE | 3
22 GAUGE POWERGLIDE PRO MIDLINE CATHETER 8CM INTERMEDIATE KIT | 3
POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER | 3
PPICC SV 4F DL BASIC | 3
X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | 3
POWERPICC SOLO CATHETER W/SHERLOCK 3CG TPS STYLET 4F | 2
GROSHONG NXT CLEARVUE CATHETER W/ SHERLOCK 3CG (TPS) STYLET FULL KIT (4F) (SL) | 2
FULL KIT 20G X 8 CM WITH BIOPATCH AND PROBE COVER | 2
M.R.I IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | 2
POWERPICC PROVENA CATHETER RADIOLOGY BASIC KIT (5F) (TRIPLE-LUMEN) (N135CM) | 2
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/MICROINTRODUCER) | 2
POWERPICC HF CATHETER WITH SHERLOCK 3CG TPS STYLET 5F | 2
POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | 2
POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH | 2
POWERPICC PROVENA 4F DL MBPLUS 3CG | 2
POWERPICC 5F DUAL-LUMEN CATHETER BASIC TRAY (NITINOL GUIDEWIRE 135CM) | 2
POWERPICC PROVENA 4F DL MBPLUS 3CG, GUARDIVA, PROBE COVER | 2
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 5F FULL TRAY DL | 2
POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | 2
BARD POWERMIDLINE CATHETER (4F) (SINGLE-LUMEN) (20CM) | 2
UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRO 4.5F X 5CM | 2
POWERPORT ISP IMPLANTABLE PORT | 2
SURESTEP INTERMITTENT CATHETER TRAY 16 FR STRAIGHT TIP | 2
NIAGARA DIALYSIS CATHETER KIT 13.5F X 20CM (SHORT-TERM) | 2
POWERPORT MRI ISP W/O SUTURE PLUGS, 9.6 FR SILICONE SL CATHETER | 2
BASIC KIT 20G X 8 CM WITH GUARDIVA | 2
NIAGARA DIALYSIS CATHETER KIT 13.5F X 24CM (SHORT-TERM) (DUAL-LUMEN) | 2
BARD® URETHRAL CATHETER | 2
POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH | 2
POWERPICC 6F RADIOLOGY BASIC TRAY (DUAL-LUMEN) (135 CM GUIDEWIRE) | 2
POWERPORT CLEARVUE ISP IMPLANTABLE PORT | 2
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F | 2
HICKMAN CV CATHETER, TRIPLE-LUMEN, 12.5F | 2
POWERPICC 3F SL SOLO MAX BARRIER PLUS KIT | 2
GROSHONG NXT CLEARVUE 4F SL BASIC TRAY WITH SAFETY INTRODUCER | 2
BALLOON BIPOLAR TEMPORARY PACING ELECTRODE CATHETER | 2
POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY W/MICROINTRODUCER 55CM | 2
PPICC 4F S/L 3CG BASIC | 2
PROVENA MIDLINE 3F SL | 2
SHERLOCK 3CG SENSOR | 2
DOT 4FR POWERMIDLINE | 2
GROSHONG NXT CATHETER WITH SHERLOCK 3CG (TPS) (5F) (DL) (55CM) | 2
POWER PORT | 2
POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | 2
BARD POWERPICC SOLO2 5F 3CG | 2
UROFORCE® BALLOON DILATION CATHETER | 2
POWERPORT VUE, ISP, MR, 8FR, MICRO | 2
POWERPICC SOLO CATHETER 4F BASIC TRAY (SL) (W/ MICROINTRODUCER) | 2
POWERPICC PROVENA 4F DL MBPLUS 3CG, GUARDIVA, PROBE COVER, BELDAR NG | 2
GROSHONG NXT CLEARVUE CATHETER (TPS) STYLET FULL KIT (4F) (SINGLE-LUMEN) | 2
POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX, 6F | 2
BARD ACCUCATH | 2
POWERPICC PROVENA 5F DL FT W/BIOPATCH, PROBE COVER, & BELDAR NEEDLE GUIDES | 2
SL CV FOR MAQUET IAB | 2
POWERPICC SOLO CATHETER 4F BASIC TRAY (SINGLE-LUMEN) | 2
GROSHONG NXT CATHETER WITH SHERLOCK 3CG | 2
POWERPICC SOLO 4F SINGLE-LUMEN CATHETER BASIC TRAY | 2
BARD POWERGLIDE PRO 18G | 2
GROSHONG MID LINE MINI TRAY 4F | 2
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP | 2
DOT PG PRO 22G 8CM RT | 2
BARD SINGLE LUMEN PICC | 2
ULTRAVERSE 018 | 2
PROV 4F DL IR 135CM NI | 2
POWERPORT IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | 2
POWERPICC 5 F POLYURETHANE CATHETER NURSING BASIC TRAY WITH MICROINTRODUCER | 2
POWERPICC SOLO | 2
BARD® WOVEN URETERAL CATHETER | 2
POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER | 2
POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F | 2
POWER PORT ISP TI 8F GROSHW/OSH | 2
HICKMAN HEMODIALYSIS/APHERESIS CATHETER, DUAL-LUMEN,13.5F | 2
POWERPICC SOLO CATHETER, 5F DL, 3CG | 2
POWERPORT MRI, 9.6 FR. SILICONE, INTER. KIT W/SUTURE PLUG, S/L | 2
BARD POWERGLIDE PRO MIDLINE CATHETER | 2
POWERPICC 5F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY | 2
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY (4F) SL | 2
M.R.I. HARD BASE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | 2
LATEX TEMP-SENSING CATHETER | 2
ACCUCATH, ACE 18G X 2.25 | 2
NBIH¿ TEMPORARY PACING ELECTRODE CATHETER | 2
GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) | 2
PRISTINE | 2
POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | 2
REFURB, SHERLOCK 3CG TCS | 2
GROSHONG 4F SINGLE-LUMEN PICC PERIPHERALLY INSERTED CVC BASIC TRAY | 2
DOT PROV SOLO 4F DL MAX | 2
POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY | 2
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC | 2
TITANIUM LOW-PROFILE IMPLANTABLE PORT WITH ATTACHABLE 8F GROSHONG SL | 2
SHERLOCK 3CG PLUS TCS | 2
BARDIA® ALL-SILICONE FOLEY CATHETER | 2
BARDEX LUBRI-SIL ALL-SILICONE FOLEY CATHETER | 2
DOT POWERPICC 4F SL 3CG | 2
POWER-TRIALYSIS SLIM-CATH STRAIGHT SHORT-TERM DIALYSIS CATHETER | 2
SEEKER | 2
X-PORT LOW-PROFILE IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | 2
POWERPICC 4F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY BASIC TRAY | 2
UNKNOWN PUREWICK DISPOSABLE | 2
POWERPICC PROVENA SOLO 5F TL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | 2
POWERPICC CATHETER W/ SHERLOCK 3CG TPS STYLET W/ MICROINTRODUCER (4F) (SL) | 2
POWERPICC 4F SINGLE-LUMEN CATHETER FULL TRAY (WITH MICROINTRODUCER) | 2
SURESTEP¿ FOLEY TRAY SYSTEM WITH BARD® LUBRI-SIL® TEMP SENSING FOLEY CATHETER | 2
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG | 2
BARDEX® LUBRICATH® TEMPERATURE-SENSING FOLEY CATHETER | 2
POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F BASIC W/ MICROINTRODUCER | 2
POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS CATHETER KIT | 1
POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT W/GUARDIVA | 1
POWERPICC SOLO2 CATHETER W/SHERLOCK 3CG TIP POSITIONING SYSTEM STYLET FULL TRAY | 1
POWERPICC SV CATHETER | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY 5F DUAL-LUMEN | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM, STYLET BASIC TRAY | 1
BARD POWERPORT SLIM IMPLANTABLE PORT | 1
2FR SL W/O GW BULK NON | 1
CROSSER IQ | 1
POWERPICC 5F SINGLE-LUMEN CATHETER BASIC TRAY (WITH MICROINTRODUCER) | 1
POWERPICC SV CATHETER WITH SHERLOCK 3CG TPS STYLET FULL TRAY | 1
BARD TRIALYSIS HD CATHETER 3F | 1
BARD TRIALYSIS TL CATHETER | 1
POWERPICC PROVENA CATHETER W/SOLO VALVE TECHNOLOGY MAXIMAL BARRIER KIT | 1
POWERPICC PROVENA 5F DL FT MAX KIT W/BIOPATCH, PROBE COVER, NEEDLE GUIDES | 1
POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVE, NEEDLE GUIDES | 1
GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TPS STYLET 4F SINGLE-LUMEN | 1
POWERPICC SV CATHETER W/ SHERLOCK 3CG TIP POSITIONING SYSTEM, STYLET BASIC TRAY | 1
BD BARD SET ROTAREX 6 F X 110CM | 1
POWERPORT M.R.I. IMPLANTABLE PORT WITH ATTACHABLE 8F GROSHONG SINGLE-LUMEN | 1
HICKMAN CATHETER | 1
HYDROSIL GO¿ INTERMITTENT URINARY CATHETER | 1
BARDEX 3-WAY URETHRAL CATHETER | 1
POWERPICC PROVENA 4F DL 3CG | 1
MAGIC 3 GO MALE COUDE CATHETER WITH SUREGRIP | 1
POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F W/ MICROINTRODUCER | 1
LUBRI-SIL FOLEY CATHETERS | 1
PPICC 4F S/L 3CG BASIC, | 1
POWERPICC PROVENA SOLO 5F TL MAX BARRIER PLUS KIT W/BIOPATCH | 1
POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY W/ MICROINTRODUCER | 1
POWERPICC FT CATHETER WITH SHERLOCK 3CG TIP POSITIONING | 1
BARD PICC INTRODUCER KIT | 1
POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS | 1
ACCUCATH ACE 20GX2.25 | 1
GROSHONG NXT CLEARVUE 4F SIL PICC BASIC TRAY (W/ MICROINTRODUCER) | 1
POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | 1
POWERPICC FT CATHETER WITH SHERLOCK 3CG TPS STYLET 5F | 1
POWERPICC SOLO 4F SL CATHETER BASIC TRAY (W/ 70CM NITINOL GUIDEWIRE) | 1
TITANIUM DOME IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | 1
18G POWERGLIDE PRO RT | 1
POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET (4F) (SL) | 1
BARD POWERPICC SOLO DL CATHETER | 1
MRI POWER PORT | 1
POWERPICC PROVENA SOLO | 1
MAGIC3 GO PEDIATRIC CATHETER WITH SUREGRIP | 1
MRI POWERPORT W/ATTACH 8FR CHRONO | 1
POWERPICC 4F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY | 1
GROSHONG NXT CLEARVUE CATHETER W/ SHERLOCK 3CG (TPS) (4F) (SL) | 1
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 5F | 1
POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY (MICROINTRODUCER) | 1
PPICC SOLO 5FR D/L, 3CG | 1
BARD® CATHETER AND GLOVE SUCTION KIT | 1
POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY W/NITINOL GUIDEWIRE 135CM | 1
SURESTEP¿ FOLEY TRAY SYSTEM BARD® LUBRI-SIL® FOLEY CATHETER TRAY | 1
POWERPICC SOLO HF CATHETER 5F RADIOLOGY BASIC TRAY (TRIPLE-LUMEN)(70CM) | 1
POWERPICC CATHETER 4F SL 3CG | 1
CS CATHETER EP XT OCTAPOLAR | 1
SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® I.C. COMPLETE CARE FOLEY CATHETER | 1
SEEKER CROSSING SUPPORT | 1
SINGLE LUMEN POWER PICC | 1
POWERPORT MRI ISP W/SP, ATTACH 9.6 FR. SIL S/L CATHETER | 1
FULL KIT 18G X 10CM WITH PROBE COVER | 1
BARD DIAGNOSTIC WOVEN QUAD CATHETER | 1
POWERPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | 1
POWERPICC PROVENA SHERLOCK 3CG (5F) (TRIPLE-LUMEN) (TPS) | 1
ATLAS GOLD BALLOON | 1
BARD ATLAS | 1
POWERPICC PROVENA SOLO 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH | 1
BARDEX FOLEY CATHETER | 1
POWERGLIDE PRO, REINFORCED TIP 18G 10 CM | 1
RECON | 1
POWERPICC 5F DUAL-LUMEN CATHETER BASIC TRAY (W/ 70CM NITINOL GUIDEWIRE) | 1
BARD PORT ACCESS DRESSING | 1
POWERPICC 5F SL CATHETER BASIC TRAY (W/ 70CM NITINOL GUIDEWIRE) | 1
SURESTEP¿ FOLEY TRAY SYSTEM LUBRI-SIL® IC COMPLETE CARE® | 1
GROSHONG 4F SL PICC PERIPHERALLY INSERTED CENTRAL VENOUS CATHETER BASIC | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 5F FULL TRAY DL | 1
POWERPICC SV CATHETER W/SHERLOCK TIP LOCATION SYSTEM (4F) DL (55CM) | 1
NIAGARA DIALYSIS CATHETER KIT 13.5F X 15CM PRE-CURVED SHORT-TERM DL | 1
POWERPICC SV CATHETER W/ SHERLOCK 3CG (TPS) STYLET FULL TRAY | 1
POWERPICC PROVENA CATHETER W/ SOLO VALVE SHERLOCK 3CG (TPS) STYLET (5F) (TL) | 1
GROSHONG MID LINE MINI TRAY 4F (EXPORT ONLY) | 1
NIAGARA DIALYSIS CATHETER KIT 13.5F X 20CM (SHORT-TERM) (DUAL-LUMEN) | 1
BARD DUAL LUMEN POWERPICC | 1
NRG | 1
POWERGLIDE PRO MIDLINE CATHETER (20G)(8CM) | 1
5F MICROEZ UNIVERSAL PTFE MICROINTRODUCER SAFETY KIT | 1
ALL SILICONE FOLEY CATHETER 16 | 1
DYNAMIC DECAPOLAR CATHETER | 1
DOT PPICC PROVENA 4F DL 3CG | 1
POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY W/MICROINTRODUCER | 1
BASIC KIT 22G X 8 CM WITH BIOPATCH | 1
GROSHONG 4F SINGLE-LUMEN PICC PERIPHERALLY INSERTED CENTRAL VENOUS CATHETER | 1
ULTRAVERSE 035 PTA DILATATION CATHETER | 1
VENUS CATHETER POWERPICC PROVENA | 1
BARD 18 FR FOLEY CATHETER | 1
POWERPICC PROVENA KIT SINGLE LUMEN | 1
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC (W/ MICROINTRODUCER) | 1
BARD POWERMIDLINE 4F SL BASIC TRAY | 1
POWERPICC SV CATHETER W/ SHERLOCK 3CG (TPS) STYLET MAXIMAL BARRIER TRAY | 1
GROSHONG 3F SINGLE-LUMEN PICC BASIC TRAY WITH MICROINTRODUCER | 1
GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG (TPS) STYLET (4F) (SL) | 1
POWERPICC CATHETER 5F BASIC TRAY (DUAL-LUMEN) (WITH MICROINTRODUCER) | 1
POWERPICC PROVENA CATHETER W/ SOLO VALVE & SHERLOCK 3CG (TPS) STYLET (5F) (TL) | 1
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F | 1
POWERPICC 5F SL CATHETER BASIC TRAY (WITH 135CM NITINOL GUIDEWIRE) | 1
POWERPICC SV CATHETER WITH SHERLOCK 3CG TPS STYLET (3F) (SINGLE-LUMEN) | 1
4 FR SL PROVENA MIDLINE BASIC KIT, CE | 1
BARD TRIPLE LUMEN POWERPICC | 1
3FR SL PROVENA MIDLINE MAX BARRIER KIT W/BIOPATCH | 1
POWERPORT MRI ISP, 8 FR GROSHONG, INT WO SP, SL | 1
PPICC SOLO 5FR D/L 3CG | 1
POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F TRAY W/ MICROINTRODUCER | 1
DOT PROV SOLO SP 3F SL | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) (4F) (SL) | 1
3FR PROVENA MIDLINE MAX BARRIER KIT | 1
POWERPICC SOLO CATHETER 4F RADIOLOGY BASIC TRAY | 1
GROSHONG NXT CLEARVUE CATHETER W/ SHERLOCK 3CG (TPS) STYLET (4F) (SL) | 1
BASIC KIT 20G X 10 CM WITH GUARDIVA | 1
3FR SL PROVENA MIDLINE FULL KIT | 1
POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER TRAY 13 F (15CM) (TL) | 1
GROSHONG 3F SINGLE-LUMEN PICC BASIC TRAY WITH SAFETY INTRODUCER | 1
DOT 3FR SL PROV ML | 1
POWERPICC SV CATHETER W/ SHERLOCK (TLS) STYLET (4F) (DUAL-LUMEN) (55CM) | 1
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET 4F | 1
POWERPICC HF CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F MAXIMAL BARRIER | 1
PROVEN A TM MIDLINE CATHETER 3F SL | 1
M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F | 1
X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | 1
POWERPICC 4F SINGLE-LUMEN CATHETER (W/ MICROINTRODUCER)(55CM) | 1
POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT W/BIOPATCH PROBE COVER & NEEDLE GUIDES | 1
NIAGARA SLIM-CATH STRAIGHT DUAL-LUMEN DIALYSIS KIT 12F X 15CM | 1
X-PORT ISP IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | 1
ACCUCATH ACE IV CATHETER WITH COILED GUIDEWIRE | 1
STATLOCK PICC PLUS FIXED POST, CRESCENT PAD | 1
POWERPICC HF CATHETER WITH SHERLOCK 3CG TIP POSITIONING | 1
BARDEX ALL-SILICONE FOLEY CATHETER | 1
BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED URETHRAL CATHETER KIT | 1
HEMOSPLIT | 1
BARDEX ANTI-INFECTIVE FOLEY CATHETER | 1
HICKMAN CV CATHETER, DUAL-LUMEN REPAIR KIT, RED/WHITE, 9F | 1
POWERLOC MAX POWER-INJECTIBLE INFUSION SET 20G X 0.75IN | 1
POWERPICC PROVENA 4F 55CM | 1
POWERPICC 6F TRIPLE-LUMEN CATHETER WITH MICROEZ MICROINTRODUCER | 1
POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG (TPS) STYLET FULL TRAY | 1
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 5F BASIC TRAY | 1
POWERGLIDE PRO MIDLINE CATHETER (18G) (10CM) | 1
REPLACEMENT CONNECTOR FOR 4F GROSHONG NXT CLEARVUE PICC W/ EXTENSION LEG | 1
POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | 1
POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY (WITH 70CM NITINOL GUIDEWIRE) | 1
SPECIALTY TUBES & CATHETER | 1
BARD POWERPICC PROVENA DL | 1
PER-Q-CATH PLUS 2F SINGLE-LUMEN PICC WITH STIFFENING STYLET | 1
DOT PPICC 4F SL 3CG | 1
POWERPICC PROVENA SOLO 5F DL FT MAX BARRIER KIT | 1
SURESTEP FOLEY TRAY SYSTEM BARDEX LUBRICATH FOLEY CATHETER | 1
GROSHONG 4F SL PICC PERIPHERALLY INSERTED CENTRAL VENOUS CATHETER | 1
SITE~RITE 8 ULTRASOUND SYSTEM | 1
SITE~RITE 8 ULTRASOUND SYSTEM CONSOLE WITH CONNECTOR | 1
POWERPICC SOLO CATHETER 4F BASIC TRAY (SINGLE-LUMEN) (WITH MICROINTRODUCER) | 1
POWERPICC CATHETER (4F) (SINGLE-LUMEN) (55CM) | 1
POWERPICC 5 F POLYURETHANE CATHETER NURSING BASIC TRAY W/MICROINTRODUCER 55CM | 1
POWERPICC SOLO2 HF CATHETER WITH SHERLOCK 3CG TIP POSITIONING | 1
BARD POWERPICC CATHETER | 1
POWERPICC 5F DUAL-LUMEN CATHETER BASIC TRAY (NITINOL GW 70CM) | 1
SURESTEP FOLEY TRAY SYSTEM LUBRI-SIL IC COMPLETE CARE | 1
POWERPICC SOLO2 CATHETER BASIC TRAY (4F) SINGLE-LUMEN MICROINTRODUCER | 1
POWERGROSHONG CATHETER WITH SHERLOCK 3CG TPS STYLET FULL KIT (5F) | 1
POWERPORT MRI 8GS, INT. W SP, ATT S/L | 1
GROSHONG NXT CLEARVUE 4F SL BASIC TRAY W/MICROINTRODUCER | 1
GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TPS STYLET | 1
BD TRIALYSIS/HD CATHETER | 1
POWERPICC FT CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F MAXIMAL BARRIER | 1
POWERPICC FT CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F MAXIMAL BARRIER | 1
POWERPICC HF CATHETER W/ SHERLOCK 3CG (TPS) STYLET (5F) (TL) | 1
BARDEX® URETHRAL CATHETER WHISTLE TIP | 1
BARD DL POWERPICCSOLO | 1
BD 20G 8CM POWERGLIDE | 1
POWERPICC 4F SL, 3CG | 1
BARD TRIPLE LUMEN PICC CATHETER | 1
HUBLESS SILICONE FLAT DRAIN 10MM | 1
3 FR. POLY MIDLINE | 1
FULL KIT 20G X 8 CM WITH PROBE COVER | 1
POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER FULL TRAY 13 F (15CM) (TL) | 1
STATLOCK PICC PLUS SLIDING POST, CRESCENT PAD, INCLUDES SKIN PREP | 1
POWER-TRIALYSIS SLIM-CATH STRAIGHT SHORT-TERM DIALYSIS (12F) (TL) (24CM) | 1
POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER KIT 13 F 15CM TL | 1
BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | 1
SURESTEP FOLEY TRAY SYSTEM BARDEX IC COMPLETE CARE | 1
BARD® FOLEY CATHETER SILICONE COATED | 1
BARD NIAGARA DIALYSIS CATHETER | 1
TITANIUM LOW-PROFILE IMPLANTABLE PORT, BROVIAC SINGLE-LUMEN, 6.6F | 1
POWERPICC SOLO CATHETER W/SHERLOCK 3CG TPS STYLET 4F MAXIMAL BARRIER TRAY | 1
GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) (W/MICROINTRODUCER) | 1
GLIDEPATH HEMODIALYSIS CATHETER | 1
POWER PORT ISP GROSHONG 8FR TITANIUM CUSTOM | 1
POWERPICC 6F TRIPLE-LUMEN CATHETER W/ MICROEZ MICROINTRODUCER NURSE BASIC TRAY | 1
20 GAUGE POWERGLIDE PRO MIDLINE CATHETER 10CM INTERMEDIATE KIT | 1
POWERPICC SOLO CATHETER W/ SHERLOCK 3CG TPS STYLET 4F MAXIMAL BARRIER | 1
POWERPICC SOLO2 CATHETER BASIC TRAY (4F)(SINGLE-LUMEN) | 1
POWERPICC SOLO 3CG 4F | 1
POWERPICC PROVENA SOLO 5F TL MAX BARRIER PLUS KIT W/GUARDIVA | 1
BARDEX® LUBRICATH® FOLEY CATHETER EDELMAN MODEL | 1
POWERPICC SOLO CATHETER W/ SHERLOCK 3CG TPS STYLET 4F MAXIMAL BARRIER TRAY | 1
MALE EXTERNAL CATHETER | 1
POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER KIT 13F 15CM TL | 1
POWERPORT MRI ISP 9.6 FR. SILILCONE SINGLE LUMEN CATHETER | 1
POWERPORT ISP M.R.I. IMPLANTABLE PORT | 1
POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TPS STYLET FULL TRAY | 1
POWER PORT MRI ISP, 8FR. GROSHONG, INT W SP, ATTACHABLE SL | 1
POWERPICC 5F DUAL-LUMEN POLYURETHANE CATHETER RADIOLOGY | 1
POWERPICC® CATHETER, 4F SL, 3CG | 1
BARD TRIALYSIS HD CATHETER | 1
STATLOCK PICC PLUS CRESCENT TRICOT FIX-POST | 1
DAVOL® FEMALE CATH KIT | 1
POWERPICC PROVENA SOLO (TLS) STYLET (5F) (TRIPLE-LUMEN) | 1
POWERPICC FT CATHETER WITH SHERLOCK 3CG TPS STYLET (5F) (DL) (TPS) | 1
POWERGLIDE PRO 20GA, 10CM | 1
POP-ON¿ SELF-ADHERING MALE EXTERNAL CATHETER | 1
BARD DL PICC CATHETER | 1
POWERPICC SOLO CATHETER, SHERLOCK 3CG, 5F MBT, MICROINTRODUCER | 1
BARD PICC 4F SINGLE LUMEN | 1
BARD DL TRIALYSIS CATHETER | 1
PowerPICC® Catheters | 1
BARD MIDLINE CATHETER | 1
M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | 1
BARD POWERMIDLINE CATHETER BASIC KIT 4F SL 20CM | 1
BARD SL POWERPICC PROVENA CATHETER | 1
MIDLINE CATHETER | 1
PICC | 1
POWERPICC 3F SL SOLO MAX BARRIER PLUS KIT, BIOPATCH, PROBE COVER & NEEDLE GUIDES | 1
POWERPORT ISP MRI 8CF INT. W/SP, ATT, SL | 1
PROVENA POWERPICC SOLO 4F DL 3CG | 1
BARD NEONATAL PICC | 1
GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG(TPS) STYLET(4F) (SINGLE-LUMEN) | 1
POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA & PROBE COVER | 1
REPLACEMENT CONNECTOR FOR 4F GROSHONG NXT CLEARVUE PICC W/EXTENSION LEG | 1
SIDEKICK | 1
GROSHONG NXT CLEARVUE 4F SL BASIC TRAY W/ SAFETY INTRODUCER | 1
BARDIA® URETHRAL CATHETER TRAY, RED RUBBER CATHETER | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY (4F) (SINGLE-LUMEN) | 1
UNKNOWN MRI POWERPORT ISP | 1
BARD 4FR S/L PER-Q-CATH | 1
POWERPICC SOLO CATHETER W/ 3CG (TPS) STYLET 4F FULL TRAY | 1
BARD PROVENA CATHETER | 1
BD PROVENA DL PICC | 1
POWERPICC FT CATHETER 5F RADIOLOGY BASIC TRAY DL (70CM NITINOL GUIDEWIRE) | 1
POWERPICC SOLO2 CATHETER BASIC TRAY (5F) DL (MICROINTRODUCER) | 1
SHERLOCK 3CG+ SMALL SENSOR | 1
HICKMAN CV CATHETER,TRIPLE-LUMEN REPAIR KIT, COMPLETE EXTERNAL SEGMENT, 12.5F | 1
ACCUCATH ACE 20G X 2.25IN | 1
BARD STRAIGHT CATHETER | 1
ARDEX® LUBRICATH® FOLEY CATHETER | 1
POWERPICC CATHETER WITH SHERLOCK TLS TECH 5F MAX | 1
SURESTEP FOLEY TRAY SYSTEM BARD® LUBRI-SIL FOLEY CATHETER TRAY | 1
PPICC SV 3F SL,BASIC, CE | 1
PPICC SOLO 4FR S/L, 3CG | 1
BARD PPICC 4F DL 3CG BASIC | 1
POWERPICC PROVENA SOLO 5F TL MAX BARRIER PLUS KIT | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET (4F) SL | 1
POWERPICC SOLO CATHETER W/SHERLOCK 3CG TPS STYLET 4F FULL TRAY | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 4F | 1
BARD POWERMIDLINE CATHETER BASIC KIT (3F) (SINGLE-LUMEN) (20CM) | 1
POWERPICC HF CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F MAX BARRIER | 1
NIAGARA SLIM-CATH DIALYSIS CATHETER 12F X 12.5CM (PRE-CURVED) | 1
3FR SL PROVENA MIDLINE MAX BARRIER KIT W/GUARDIVA AND PROBE COVER | 1
BROVIAC 4.2 F SINGLE-LUMEN CV CATHETER | 1
BARDIA® URETHRAL CATHETER | 1
COUDE CATHETER 12 FR. 5CC RIBBED BALLOON | 1
BARDEX COUDE SILVER HYDROGEL CATHETER, BARDEX 22FR 5CC | 1
ARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | 1
POWERPICC SOLO2 CATHETER BASIC TRAY (5F) DL (MICROINTRODUCER) | 1
BARD POWERPICC, 3CG | 1
POWERPORT TI L/P 6 CF INT WSP ATT SL | 1
POWERPICC CATHETER WITH SHERLOCK TIP LOCATION SYSTEM (TLS) STYLET 4F | 1
POWERGLIDE | 1
PERIPHERAL CATHETER INSERTION KIT | 1
POWERPICC PROVENA 4F DLMBPLUS (3CG) | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET (4F) (SINGLE-LUMEN) | 1
PPIC 4F S/L 3CG BASIC | 1
TITANIUM SLIMPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | 1
BARD® URINE METER FOLEY TRAY WITH LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | 1
DAVOL FEMALE CATH KIT | 1
BARD® FOLEY TRAY WITH BARDEX® ALL SILICONE FOLEY CATHETER | 1
CONQUEST 40 PTA DILATATION CATHETER | 1
18 GAUGE POWERGLIDE PRO MIDLINE CATHETER 10CM BASIC KIT | 1
POWERPICC HF CATHETER WITH SHERLOCK 3CG TIP TPS STYLET FULL TRAY (5F) TL | 1
POWERPICC SV CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | 1
POWERPICC PROVENA SOLO 5F DL FT MAX BARRIER PLUS KIT W/GUARDIVA | 1
POWERPICC SOLO 5F SINGLE-LUMEN CATHETER BASIC TRAY | 1
TEMPERATURE PACING ELECTRODE CATHETER | 1
POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/GUARDIVA | 1
NIAGARA SLIM-CATH DIALYSIS CATHETER KIT 12F X 15CM | 1
POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | 1
PEEL APART INTRODUCER KIT, 10F | 1
BARD® URETHRAL CATHETERIZATION TRAY WITH VINYL CATHETER | 1
GROSHONG 4F SL PICC (60 CM) | 1
BARD URINE METER FOLEY TRAY WITH LUBRI-SIL ALL-SILICONE FOLEY CATHETER | 1
INTERMITTENT ROCHESTER CATHETER | 1
POWERPORT | 1
GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TIP | 1
POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, | 1
BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN,WHITE, 4.2F | 1
ULTRAVERSE 018 PTA DILATATION CATHETER | 1
NATURAL¿ NON-ADHESIVE MALE EXTERNAL CATHETER | 1
BARD SURESTEP FOLEY TRAY SYSTEM 16 FR TEMPERATURE-SENSING | 1
SURE STEP¿ FOLEY TRAY SYSTEM BARDEX® LUBRICATH® FOLEY CATHETER TRAY | 1
BARD® FOLEY TRAY WITH BARDEX® LUBRICATH® FOLEY CATHETER | 1
PER-Q-CATH PLUS 2F SINGLE-LUMEN PICC WITH STIFFENING STYLET BASIC TRAY | 1
DOT POWER MIDLINE CATHETER 4F SL | 1
POWERPICC SOLO CATHETER 5F RADIOLOGY BASIC TRAY (DUAL-LUMEN) | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET MICROINTRODUCER 4F SL | 1
GROSHONG NXT CATHETER WITH SHERLOCK TLS STYLET 5F MAXIMAL BARRIER TRAY DL 45CM | 1
BARDEX® LUBRICATH® FOLEY CATHETER SHORT TIP 3 WAY CONTINUOUS IRRIGATION | 1
POWERPICC SOLO HF CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | 1
REPLACEMENT CONNECTOR FOR 4F GROSHONG NXT CLEARVUE PICC WITH EXTENSION | 1
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F WITH MICROINTRODUCER | 1
SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® LUBRICATH® TEMPERATURE SENSING CATHETER | 1
POWERPICC SOLO2 CATHETER WITH SHERLOCK TLS TECHNOLOGY 5F | 1
POWERPICC SOLO2 CATHETER BASIC TRAY (4F) SL (MICROINTRODUCER) | 1
POWERPICC SOLO 4F SINGLE-LUMEN CATHETER FULL TRAY | 1
SILICONE FOLEY CATHETER | 1
BARD POWERPICC DL | 1
AIRGUARD VALVED INTRODUCER, 16.5F | 1
BARD URINE METER FOLEY CATHETER TRAY WITH BARDEX LUBRICATH FOLEY CATHETER | 1
POWERPICC SV CATHETER WITH SHERLOCK 3CG TIP | 1
LUBRI-SIL I.C. COMPLETE CARE TEMP-SENSING FOLEY CATHETER TRAY AND URINE METER | 1
ACCUCATH ACE INTRAVASCULAR CATHETER 22GX1.25 | 1
POWERPORT M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, KIT, 9.6F | 1
LONG POLYURETHANE BARD MIDLINE CATHETER 4 FRENCH PICC LINE | 1
POWERPICC PROVENA CATHETER WITH SHERLOCK 3CG TPS STYLET 5F TRIPLE-LUMEN | 1
POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | 1
GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIC | 1
GROSHONG NXT CATHETER 5F BASIC TRAY DUAL-LUMEN W/MICROINTRODUCER | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY (6F) (TRIPLE-LUMEN) | 1
SPECIALTY TUBES & CATHETERS | 1
M.R.I. ULTRA SLIMPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | 1
BARD PICC 4F SL | 1
POWERPICC 4F SINGLE-LUMEN CATHETER (WITH 70CM NITINOL GUIDEWIRE) | 1
POWERGROSHONG CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | 1
DOT,4FR POEWRMIDLINE | 1
BARD POWERMIDLINE CATHETER FULL KIT (5F)(DUAL-LUMEN) | 1
NIAGARA SIM-CATH DIALYSIS CATHETER KIT 12F X 20CM (STRAIGHT) (SHORT TERM) | 1
POWERPICC PROVENA 5F TL MAX BARRIER PLUS | 1
POWERPICC SV CATHETER 4F RADIOLOGY BASIC TRAY (DUAL-LUMEN) | 1
DOT PPICC 5F DL RN | 1
BARD POWERPICC 3CG | 1
UNKNOWN MALE EXTERNAL CATHETER | 1
STATLOCK PICC PLUS SLIDING POST, CRESCENT PAD | 1
POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM, STYLET | 1
BARD® TIGERTAIL¿ FLEXIBLE TIP URETERAL CATHETERS | 1
BARD POWERPORT IMPLANTABLE PORT | 1
BARD® CLEAN-CATH® INTERMITTENT CATHETER | 1
PPICC BASIC 3CG | 1
PPICC SV 3FSL,3CG | 1
POWERPICC SOLO CATHETER, SHERLOCK 3CG TIP POSITIONING SYSTEM,STYLET 4F FULL TRAY | 1
POWER GLIDE PRO MIDLINE CATHETER | 1
BARD POWERPICC PROVENA CATHETER | 1
POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER | 1
POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY (W/ MICROINTRODUCER) (55CM) | 1
BARD SL PICC | 1
POWERPICC SOLO 4F SL 3CG | 1
BARD POWERGLIDE PRO | 1
POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER (13 F) (20CM) (TL) | 1
POWERPICC CATHETER W/ SHERLOCK 3CG TPS BASIC TRAY (4F) (SINGLE-LUMEN) | 1
4FR DL POWERMIDLINE FULL KIT | 1
POWERPICC SOLO 5F SINGLE-LUMEN CATHETER BASIC TRAY W/70CM NITINOL GW | 1
POWER-TRIALYSIS SHORT-TERM CURVED DIALYSIS CATHETER (13F) (15CM) (TL) | 1
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/ SAFETY INTRODUCER | 1
FOLEY SIL 2-WAY 16 FR 5CC | 1
POWERHICKMAN CV CATHETER, POLYURETHANE SINGLE-LUMEN, 8F | 1
ROTAREX | 1
ACCUCATH ACE 18G 2.25¿ INTERMEDIATE KIT | 1
POWERPICC SOLO 4F SINGLE-LUMEN CATHETER BASIC TRAY W/135CM NITINOL GW | 1
POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER & NEEDLE GUIDES | 1
ACCUCATH ACE INTRAVASCULAR CATHETER (18G) | 1
BARD FOLEY CATHETER | 1
M.R.I. HARD BASE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | 1
POWERPORT SLIM TI W 6FR CHRONOFLEX WO SP, INTER. KIT | 1
ATLAS GOLD ANGIOPLASTY BALLOON | 1
BARD | 1
POWER-TRIALYSIS SHORT-TERM EXTENSION DIALYSIS CATHETER | 1
ACCUCATH ACE | 1
VACCESS PTA DILATATION CATHETER | 1
BARD TRIPLE LUMEN PICC | 1
BARD DUAL LUMEN PICC | 1
CLEAR SILICONE SELF-ADHESIVE MALE CONTINENCE SHEATH | 1
PER-Q-CATH PLUS PICC 4F SINGLE-LUMEN PICC | 1
TITANIUM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | 1
POWERHICKMAN POLYURETHANE CATHETER WITH SURECUFF INGROWTH CUFF | 1
GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TPS | 1
POWERPICC SOLO CATHETER 5F RADIOLOGY BASIC TRAY (DL) (135CM GUIDEWIRE) | 1
SECURELOC SAFETY INTRODUCER NEEDLE 21G X 2.75 IN | 1
POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER DL | 1
POWER PORT MRI ISP, 8 FR CHRONOFLEX, INT WO SP, ATTACH SL | 1
4FR DL PROVENA MIDLINE FULL KIT W/GUARDIVA AND PROBE COVER | 1
4F DL PROVENA SOLO BASIC W/3CG ROW | 1
BARD POWER HOHN CATHETER | 1
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F MBT WITH MICROINTRODUCER | 1
POWERPICC SOLO CATHETER W/SHERLOCK 3CG(TPS) STYLET 5F MBT, W/MICROINTRODUCER | 1
BARD DUOGLIDE ACUTE DIALYSIS CATHETER | 1
DOT PWRMIDLINE CATHETER 4F SL | 1
DOT 4FR SL PWERML | 1
M.R.I. LOW-PROFILE IMPLANTABLE PORT,HICKMAN SINGLE-LUMEN, 6.6F | 1
STATLOCK CV PLUS SLIDING POST, TRICOT HYDROCOLLOID | 1
NIAGARA DIALYSIS CATHETER KIT 13.5F X 15CM (PRE-CURVED) (SHORT-TERM) (DL) | 1
POWERPICC CATHETER, 4F SL, 3CG | 1
UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER | 1
POWERPICC SOLO CATHETER 5F BASIC TRAY (DUAL-LUMEN) (WITH MICROINTRODUCER) | 1
DELTEC POWER PORT-A-CATH II VENOUS ACCESS | 1
BARD SURESTEP INTERMITTENT CATHETER TRAY PRECONNECTED DRAINAGE BAG | 1
POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET BASIC TRAY (4F) (SL) | 1
POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET BASIC TRAY (5F) (DL) | 1
BARD DL POWERPICC SOLO | 1
PER-Q-CATH PLUS PICC WITH STIFFENING STYLET WIRE 5F (DUAL-LUMEN)(65CM) | 1
NIAGARA DIALYSIS CATHETER KIT 13.5F X 15CM (PRE-CURVED) (SHORT-TERM) DL | 1
2.7F BROVIAC CV CATHETERS 0.5 MM LUMEN REPAIR KIT | 1
UNK BROVIAC REPAIR KIT | 1
BARDEX® I.C. COMPLETE CARE® FOLEY TRAY | 1
BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN, WHITE, 6.6F | 1
SOLO PICC LINE | 1
GROSHONG NXT CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC KIT 5F DL | 1
POWERPICC SOLO FT CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F | 1
BARDEX¿ LUBRICATH¿ FOLEY CATHETER | 1
POWERPICC SOLO FT CATHETER 5F RADIOLOGY BASIC TRAY DL (70 CM GUIDEWIRE) | 1
BARD GROSHONG PICC | 1
POWERPICC 4F SL MAX BARRIER PLUS KIT, BIOPATCH, PROBE COVER, BELDAR NEEDLE GUIDE | 1
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | 1
POWERPICC SOLO 5F DUAL-LUMEN CATHETER BASIC TRAY W/ 70CM NITINOL GUIDEWIRE | 1
POWERPICC 5F DUAL-LUMEN POLYURETHANE CATHETER RADIOLOGY 135 CM GUIDEWIRE | 1
POWERPICC SOLO FT CATHETER 5F RADIOLOGY BASIC TRAY (DUAL-LUMEN) | 1
POWERPICC PROVENA CATHETER W/ SOLO VALVE TECHNOLOGY RADIOLOGY BASIC KIT (5F) | 1
PUREWICK FEMALE EXTERNAL CATHETER, NOT MADE NATURAL RUBBER LATEX | 1
POWERPICC SOLO 5F DUAL-LUMEN CATHETER BASIC TRAY W/135CM NITINOL GUIDEWIRE | 1
POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM STYLET 4F | 1
POWERPICC 5F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY 70 CM GUIDEWIRE | 1
POWERPICC 5F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY BASIC TRAY | 1
POWERPICC SOLO 5F 55CM | 1
POWERHICKMAN POLYURETHANE CATHETER WITH SURECUFF INGROWTH | 1
BARDPORT, SLIMPORT, X-PORT | 1
POWERPICC HF CATHETER 5F RADIOLOGY BASIC TRAY | 1
DOT PROV 3F SL MAX 3CG | 1
POWERPICC PROVENA CATHETER W/ SHERLOCK 3CG (TPS) STYLET (5F) (TRIPLE-LUMEN) | 1
GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/ MICROINTRODUCER | 1
UNKNOWN INTERMITTENT CATHETER | 1
POWERPICC 5F DL POLYURETHANE CATHETER RADIOLOGY (135 CM NITINOL GUIDEWIRE) | 1
POWERPICC SV CATHETER 4F RADIOLOGY (DL) (W/ HYDRO-GLIDE INSERTION STYLET) | 1
LUTONIX® DRUG COATED BALLOON | 1
INDWELLING URINARY CATHETER | 1
SOLO POWER PICC-DOUBLE LUMEN | 1
POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY WITH MICROINTRODUCER | 1
POWERPICC SV CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET | 1
POWERPORT CLEARVUE SLIM IMPLANTABLE PORT SLIM MR 8F - LOG15341 | 1
POWER PORT TI ISP, 8 FR. GROSHONG, INT W SP, SINGLE LUMEN | 1
BARD POWERPICC SOLO | 1
GROSHONG NXT 5F DUAL-LUMEN REVERSE TAPER PICC | 1
PORT-A-CATH | 1
BARD POWERMIDLINE CATHETER | 1
SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® IC COMPLETE CARE® | 1
BARDEX LUBRICATH FOLEY CATHETER COUDE TIP | 1
POWERPICC PROVENA SOLO 5F TL 3CG | 1
4FR DL PROVENA MIDLINE MAX BARRIER KIT | 1
X-PORT ISP | 1
PER-Q-CATH PLUS PICC 3F WITH STIFFENING STYLET WIRE SL (65 CM) | 1
BARD SL POWERPICC SOLO | 1
POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY (70CM NITINOL GUIDEWIRE) | 1
POWERPORT ISP TI 8F CHRONWSH | 1
POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F | 1
POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT | 1
FOLEY URINARY CATHETER | 1
POWERPICC SOLO 5FR 55 CM DL FT 3CG | 1
POWERPICC SOLO FT CATHETER WITH SHERLOCK 3CG TPS STYLET 5F | 1
FOLEY BALLOON | 1
SURESTEP FOLEY TRAY SYSTEM | 1
PEDIATRIC CATH KIT 8FR | 1

## Case Metadata
Case Name | MDL Number | MDL Filing Date | Court | MAUDE Start Date | MAUDE End Date | Total Records | Date Range
Bard PowerPort | 3081 | 2023-08-15 | D. Arizona | 2023-02-15 | 2023-08-15 | 4760 | 2023-02-15 00:00:00 to 2023-08-15 00:00:00