3006260740-2023-01862 | 2023-05-12 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD DUOGLIDE ACUTE DIALYSIS CATHETER | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT THE DUO GLIDE DIALYSIS CATHETER HAS HUGE BUBBLE/ANEURYSM TYPE NEAR HUB. CUSTOMER DID SAY THEY POWER INJECTED THROUGH IT IN CAT SCAN." ADDITIONAL INFORMATION RECEIVED 5/06/2023: IT WAS REPORTED BY THE CUSTOMER "THE PRODUCT IS A DUO GLIDE ACUTE DIALYSIS CATHETER. THE PRODUCT WAS INSERTED IN A PATIENT WHO WENT TO CT (THEY POWER-INJECTED CONTRAST THROUGH THE CATHETER) AND THE CATHETER BUBBLED UP. THE PATIENT DIDN'T HAVE HARM BUT HAD TO HAVE THE CATHETER REPLACED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF BALLOONING OF THE EXTENSION TUBING OF A CATHETER IS CONFIRMED AND WAS DETERMINED TO BE RELATED TO DEVICE USE. ONE 24 CM DUOGLIDE DIALYSIS CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION REVEALED THAT THE PROXIMAL SEGMENT OF THE EXTENSION TUBING OF THE VENOUS SIDE OF THE CATHETER EXHIBITED BALLOONING. THE EVENT DESCRIPTION CLAIMS TO HAVE POWER INJECTED CONTRAST THROUGH THE CATHETER, BUT THE CATHETER IS NOT POWER INJECTABLE. BECAUSE THE CUSTOMER REVEALED ATTEMPTS OF POWER INJECTION WHILE USING THE DEVICE, THE RETURNED DUOGLIDE CATHETER BALLOONED AT THE PROXIMAL END OF THE EXTENSION TUBING OF THE VENOUS LUMEN AND THE COMPLAINT OF BALLOONING IS CONFIRMED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01859 | 2023-05-12 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/SHERLOCK 3CG(TPS) STYLET 5F MBT, W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1295108D |  | CUSTOMER REPORTED THAT THE RED PORTION OF THE DILATOR DID NOT PEEL LIKE IT SHOULD. CUSTOMER SAID PART OF THE DILATOR HAD TO BE CUT. THIN CIRCULAR RING. HAD TO STOP PROCEDURE TO CUT. PATIENT WAS URGENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01863 | 2023-05-12 | 2023-04-17 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | IT WAS REPORTED THAT THE PATIENT HAD THE GROSHONG CATHETER PLACED FROM THE RIGHT BRACHIAL VEIN AT THE HOSPITAL ON (B)(6) 2023. 38 CM OF THE CATHETER WAS INSERTED INTERNALLY, REACHING THE PRECAVA AND ANOTHER 3 CM WAS EXPOSED EXTERNALLY. AFTER DISCHARGE, THE PATIENT HAD THE CATHETER MAINTAINED AT THE HOSPITAL. TWO MONTHS LATER, THE CATHETER WAS FOUND RUPTURED DURING THE 8TH MAINTENANCE AT THE LOCAL HOSPITAL. THE INTERNAL PORTION OF THE CATHETER MOVED NEAR THE LUNG WITH BLOOD FLOW. THE PATIENT GOT BACK TO THE HOSPITAL, WHERE THE RESIDUAL CATHETER WAS REMOVED BY AN EMERGENCY SURGERY AT INTERVENTION DEPARTMENT. GROSHONG CATHETER RUPTURED; PATIENT AND FAMILY QUESTIONED ABOUT THE QUALITY OF THE PRODUCT AND ASKED FOR COMPENSATION. THE CATHETER RUPTURED IN THE PATIENT WITH A FLAT OPENING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01860 | 2023-05-12 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F MBT WITH MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1194108D |  | WAS REPORTED BY THE CUSTOMER THAT "WE WERE PLACING A PICC AT TCA AND WHEN THE NURSE PULLED BACK THE WIRE TO TRIM IT COILED UP." ADDITIONAL INFORMATION RECEIVED 4/27/2023: IT WAS REPORTED BY THE CUSTOMER "PART OF THE KIT WAS USED ON THE PATIENT HOWEVER WHEN THE NURSE WENT TO CLIP/TRIM THE CATHETER TO THE PROPER LENGTH, SHE PULLED THE WIRE BACK WHICH IS WHEN IT COILED. SHE ASKED FOR ANOTHER KIT TO BE OPENED AND UTILIZED A NEW PICC CATHETER. NO HARM OR DELAY TO THE PATIENT." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED STYLET IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. ONE 4 FR SINGLE LUMEN POWERPICC SOLO CATHETER WITH A 3CG STYLET INSERTED THROUGH A T-LOCK WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED A PORTION OF THE STYLET DISTAL TO THE T-LOCK WAS CURLED TIGHTLY AROUND ITSELF. THE STYLET WAS ALSO OBSERVED TO BE BENT NEAR ITS DISTAL END. NO BREAKS WERE OBSERVED IN THE RETURNED STYLET. THE SHAPE, LOCATION, AND NATURE OF THE DAMAGE OBSERVED ON THE STYLET, AS WELL AS THE EVENT DESCRIPTION PROVIDED BY THE COMPLAINANT AND SIMILARITIES OBSERVED DURING PAST ATTEMPTS TO REPLICATE THIS FAILURE MODE, SUGGEST THE STYLET WAS PULLED AGAINST THE LUER HUB OR ANOTHER HARD SURFACE DURING REMOVAL OF THE STYLET RESULTING IN THE OBSERVED DAMAGE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01866 | 2023-05-12 | 2023-04-25 | I, n, j, u, r, y | Manufacturer report | BARD POWER HOHN CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. DUAL LUMEN CATHETER PLACED IN THE CHEST. BD REP OBSERVED THAT THERE IS REDNESS OR SWELLING AT OR AROUND CENTRAL LINE INSERTION SITE. NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01865 | 2023-05-12 | 2023-04-25 | I, n, j, u, r, y | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. DUAL LUMEN PICC PLACED IN THE UPPER ARM. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01867 | 2023-05-12 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1275108FD5 |  | IT WAS REPORTED VIA MEDWATCH "BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET REF S1275108FD5, LOT # REGY0842 NOTED LEAKAGE AT RUBBER-LIKE STOPPER AND STYLET ONCE PLACED. STYLET ALSO NOTABLY DIFFICULT TO REMOVE FROM CATHETER. STYLET RETAINED. PICC SUCCESSFULLY PLACED." THIS REPORT ADDRESSES THE STYLET THAT WAS DIFFICULT TO REMOVE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01852 | 2023-05-12 | 2022-12-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | CUSTOMER REPORTED A 4FR DL POWERPICC PROVENA INSERTED IN THE LEFT UPPER ARM BASILIC VEIN ON (B)(6) 2022 WAS FOUND TO BE KINKED ON (B)(6) 2022 IN THE SUBCLAVIAN REGION AT THE STERNOCLAVICULAR JUNCTION. PICC DWELL TIME WAS 18 DAYS. PICC REMAINED IN PLACE FOR ANOTHER WEEK. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03344 | 2023-05-12 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON BURST. THE PATIENT HEARD A POPPING SOUND WHEN THEY TURNED OVER, AFTER THE CATHETER WAS PLACED IN THE OPERATING ROOM AND AFTER THEY WERE TRANSFERRED TO THE ICU. WHEN THE NURSE CHECKED LATER, IT WAS FOUND THAT THE CATHETER HAD BEEN REMOVED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE DETAILS REGARDING ANY MISSING PIECES OF BALLOON WEREN'T PROVIDED FROM CUSTOMER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03340 | 2023-05-12 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ALL SILICONE CATHETER | ALL SILICONE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT GOT UP AND OUT OF BED WITH FOLEY CATHETER, STRETCHED AND THE CATHETER SNAPPED IN HALF. ONE PORTION FOUND ON FLOOR, AND THE OTHER PORTION REMAINED IN PATIENT AND THEN CAME OUT PASSIVELY WITH NEXT VOID. THE DEVICE WAS IN USE FOR 5 DAYS. THE BALLOON CHECK WAS DONE PRIOR TO USE. A SECOND DEVICE WAS REQUIRED FOR THE PATIENT. THERE WAS NO REPORTED INJURY AND NO MEDICAL INTERVENTION REQUIRED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00331 | 2023-05-12 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 014 | PTA BALLOON DILATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | US14150530 | US14150530 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF CHRONIC TOTAL OCCLUSION IN THE SUPERFICIAL FEMORAL ARTERY, THE CONTRAST MEDIUM WAS ALLEGEDLY UNABLE TO BE ASPIRATED WHEN THE DEVICE WAS ATTEMPTED TO DEFLATE AFTER INFLATING THE BALLOON. REPORTEDLY, THE NEEDLE WAS PUNCTURED FROM THE PATIENT¿S BODY SURFACE TOWARD TO THE BALLOON AND THE DEVICE WAS REMOVED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01846 | 2023-05-12 | 2022-08-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT 4FR POWERMIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P4154108D |  | IT WAS REPORTED BY THE CUSTOMER, "A 4FR PICC INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2022 WAS FOUND TO BE KINKED THE SAME DAY NEAR THE INSERTION SITE. PICC DWELL TIME WAS ZERO (0) DAYS. TPA WAS NOT USED. PER CUSTOMER, NO INTERVENTION WAS NEEDED. PATIENT EXPIRED ON 8/11/2022." | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01847 | 2023-05-12 | 2022-09-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED BY THE CUSTOMER, "A 4FR DL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2022 WAS FOUND TO BE KINKED ON (B)(6) 2022 IN THE UPPER 1/3 OF THE UPPER EXTREMITY. PICC DWELL TIME WAS 11 DAYS. PICC REMAINED IN PATIENT AND CONTINUED TO BE USED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A KINKED CATHETER WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. AP AND LATERAL RADIOGRAPHS WERE RETURNED FOR EVALUATION OF THIS COMPLAINT. THE RADIOGRAPHS SHOWED THE RIGHT HUMERUS. A CATHETER, CONSISTENT WITH THE IMPLICATED 4FR D/L POWERPICC PROVENA, WAS SEEN IN THE IMAGES. A KINK WAS VISIBLE IN THE CATHETER SHAFT AT THE SKIN INSERTION SITE IN THE LATERAL VIEW BUT WAS NOT SEEN IN THE AP RADIOGRAPH. AS THE KINK WAS ONLY SEEN IN ONE OF THE RADIOGRAPHS, THE KINK WAS LIKELY POSITIONAL. HOWEVER, THE EXACT ROOT CAUSE OF THE KINK COULD NOT BE DETERMINED. POSSIBLE CONTRIBUTING FACTORS COULD INCLUDE CATHETER PLACEMENT OR SECUREMENT. THE IFU STATES, ¿AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT.¿ AND ¿AVOID SHARP OR ACUTE ANGLES DURING IMPLANTATION WHICH COULD COMPROMISE THE PATENCY OF THE CATHETER LUMEN¿. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE REPORT OF A KINKED CATHETER HAS BEEN DOCUMENTED AND WILL BE MONITORED AS PART OF COMPLAINT TRENDING. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01848 | 2023-05-12 | 2022-10-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1274108D5 |  | IT WAS REPORTED BY THE CUSTOMER, "A 4FR DL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BRACHIAL VEIN ON (B)(6)2022 WAS FOUND TO BE KINKED ON (B)(6)2022 IN THE SUBCLAVIAN REGION. PICC DWELL TIME WAS 14 DAYS. PICC WAS EXCHANGED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A KINKED CATHETER WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. A SINGLE AP CHEST RADIOGRAPH SUPINE WAS RETURNED FOR EVALUATION. A CATHETER, LIKELY THE IMPLICATED 4FR D/L POWERPICC, WAS SEEN IN THE IMAGE AS A RIGHT-SIDED PLACED PICC. THE SINGLE VIEW COULD NOT DEFINITIVELY CONFIRM A KINK. HOWEVER, IT DID APPEAR THAT A POTENTIAL KINK/COIL WAS PRESENT IN THE CATHETER SHAFT NEAR THE SUBCLAVIAN. IT IS POSSIBLE THAT A SMALL AMOUNT OF THE CATHETER WAS LOOPED INTO THE RIGHT IJ. CONTRIBUTING FACTORS FOR A KINK COULD INCLUDE ANATOMIC VARIABLES AND CATHETER PLACEMENT. THE IFU STATES, ¿AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT.¿ A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE REPORT OF A KINKED CATHETER HAS BEEN DOCUMENTED AND WILL BE MONITORED AS PART OF COMPLAINT TRENDING. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11 | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01841 | 2023-05-12 | 2022-10-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED BY THE CUSTOMER, "A POWERPICC PROVENA INSERTED IN THE LEFT UPPER ARM BRACHIAL VEIN ON (B)(6) 2022 WAS FOUND TO BE KINKED ON (B)(6) 2022 IN THE UPPER 1/3 OF THE UPPER EXTREMITY. PICC DWELL TIME WAS 55 DAYS. PICC WAS PULLED BACK AND KINK RESOLVED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01845 | 2023-05-12 | 2022-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED BY THE CUSTOMER, "A 4FR DL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BRACHIAL VEIN ON (B)(6)2022 WAS FOUND TO BE KINKED IN THE UPPER 1/3 OF THE UPPER EXTREMITY ON (B)(6)2022 (DWELL TIME 3 DAYS). PICC WAS INITIALLY PULLED BACK AND THEN REMOVED AND A NEW PICC WAS PLACED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A KINKED CATHETER WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. A SINGLE RADIOGRAPH SHOWING THE PATIENT¿S RIGHT ARM/HUMERUS WAS RETURNED FOR EVALUATION. A CATHETER, CONSISTENT WITH THE IMPLICATED 4FR D/L POWERPICC PROVENA CATHETER, WAS SEEN IN THE IMAGE. A KINK WAS SEEN IN THE CATHETER SHAFT AT THE SKIN INSERTION SITE. POSSIBLE CONTRIBUTING FACTORS FOR A KINK IN THE CATHETER COULD INCLUDE THE ANATOMIC VARIABLES, CATHETER PLACEMENT, OR CATHETER SECUREMENT. THE IFU STATES, ¿AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT.¿ AND ¿AVOID SHARP OR ACUTE ANGLES DURING IMPLANTATION WHICH COULD COMPROMISE THE PATENCY OF THE CATHETER LUMEN¿. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE REPORT OF A KINKED CATHETER HAS BEEN DOCUMENTED AND WILL BE MONITORED AS PART OF COMPLAINT TRENDING. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11 | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01842 | 2023-05-12 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC HF CATHETER WITH SHERLOCK 3CG TPS STYLET 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1385108QD |  | IT WAS REPORTED BY THE CUSTOMER, "PICC INSERTED IN THE RIGHT UPPER ARM BRACHIAL VEIN ON (B)(6) 2022 WAS FOUND TO BE KINKED ON (B)(6) 2022 (DWELL TIME 17 DAYS). A CHEST X-RAY WAS TAKEN ON (B)(6) 2022 THAT SHOWED THE CATHETER KINK BUT THE KINK WASN'T DISCOVERED UNTIL (B)(6) 2022. A NEW PICC WAS INSERTED IN THE LEFT UPPER ARM." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER KINK WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. AN AP RADIOGRAPH OF THE RIGHT ARM/HUMERUS WAS RETURNED FOR EVALUATION. A CATHETER, CONSISTENT WITH THE IMPLICATED 5FR T/L POWERPICC PROVENA CATHETER, WAS SEEN IN THE IMAGE. A KINK WAS VISIBLE IN THE CATHETER SHAFT. THE KINK APPEARED TO BE LOCATED EITHER AT THE SKIN INSERTION SITE OR AS THE CATHETER ENTERED THE MUSCLE. POSSIBLE CONTRIBUTING FACTORS FOR A KINK IN THE CATHETER COULD INCLUDE THE ANATOMIC VARIABLES, CATHETER PLACEMENT, OR CATHETER SECUREMENT. THE IFU STATES, ¿AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT.¿ AND ¿AVOID SHARP OR ACUTE ANGLES DURING IMPLANTATION WHICH COULD COMPROMISE THE PATENCY OF THE CATHETER LUMEN¿. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE REPORT OF A KINKED CATHETER HAS BEEN DOCUMENTED AND WILL BE MONITORED AS PART OF COMPLAINT TRENDING. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01844 | 2023-05-12 | 2022-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED BY THE CUSTOMER, "A 4FR DL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2022. IT WAS FOUND TO BE KINKED IN THE UPPER 1/3 OF THE UPPER EXTREMITY ON (B)(6) 2022 (DWELL TIME 10 DAYS). PICC WAS EXCHANGED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A KINK IN THE CATHETER WAS UNCONFIRMED AS THE REPORTED KINK COULD NOT BE CONCLUSIVELY SEEN IN THE RETURNED IMAGE. A SINGLE AP RADIOGRAPH OF THE PATIENT¿S RIGHT ARM/HUMERUS WAS RETURNED FOR EVALUATION. A CATHETER, CONSISTENT WITH THE IMPLICATED 4FR D/L POWERPICC PROVENA, WAS SEEN IN THE RETURNED IMAGE. THE ENTRANCE OF THE CATHETER INTO THE VEIN APPEARED TO BE STEEP. THE AP PERSPECTIVE COULD NOT BE USED TO EVALUATE WHETHER THERE WAS A KINK AT THE SKIN SURFACE OR VENOUS PUNCTURE SITE; HOWEVER, THE RADIOGRAPH DID NOT EXCLUDE THE POSSIBILITY OF KINKING. IT IS POSSIBLE THAT THE CATHETER WAS KINKED, BUT IT COULD NOT BE INDEPENDENTLY CONFIRMED FROM THE PROVIDED IMAGE. POSSIBLE CONTRIBUTING FACTORS FOR A KINK IN THE CATHETER COULD INCLUDE CATHETER PLACEMENT, PATIENT ANATOMY, OR CATHETER SECUREMENT. THE IFU STATES, ¿AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT.¿ AND ¿AVOID SHARP OR ACUTE ANGLES DURING IMPLANTATION WHICH COULD COMPROMISE THE PATENCY OF THE CATHETER LUMEN.¿ A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE REPORT OF A KINKED CATHETER HAS BEEN DOCUMENTED AND WILL BE MONITORED AS PART OF COMPLAINT TRENDING. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01843 | 2023-05-12 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC HF CATHETER WITH SHERLOCK 3CG TPS STYLET 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1385108QD |  | IT WAS REPORTED BY THE CUSTOMER, " A 5FR TL POWERPICC PROVENA INSERTED IN THE LEFT UPPER ARM BRACHIAL VEIN ON (B)(6) 2022 WAS FOUND TO BE KINKED IN THE UPPER 1/3 OF THE UPPER EXTREMITY ON (B)(6) 2022 (DWELL TIME 2 DAYS). PATIENT REFERRED TO IR FOR A PICC EXCHANGE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER KINK WAS CONFIRMED. A SINGLE AP RADIOGRAPH OF THE PATIENT¿S LEFT ARM/HUMERUS WAS RETURNED FOR EVALUATION. A CATHETER, CONSISTENT WITH THE IMPLICATED 5FR T/L POWERPICC HF CATHETER, WAS SEEN IN THE RETURNED IMAGE. THE PICC CONTAINED A LARGE COIL IN THE VEIN WITH A POSSIBLE KINK IN IN THE PROXIMAL ASPECT OF THE COILED CATHETER. THE COIL IN THE CATHETER WAS LIKELY DUE TO A PATULOUS VEIN WITH A PROMINENT VALVE. CONTRIBUTING FACTORS FOR THE KINK INCLUDED ANATOMIC VARIABLES AND CATHETER PLACEMENT. THE IFU STATES, ¿AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT.¿ A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE REPORT OF A KINKED CATHETER HAS BEEN DOCUMENTED AND WILL BE MONITORED AS PART OF COMPLAINT TRENDING. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01838 | 2023-05-12 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE DRIPS AROUND STYLET WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. CLINICIAN DISCARDS DEVICE AND OBTAINS A SECOND DEVICE FOR INSERTION, WHICH ALSO LEAKED AROUND THE STYLET. THE 3RD PICC LEAKED WHEN FLUSHING THE CATHETER AT THE END OF THE PROCEDURE, BUT DID NOT LEAK BEFORE THE PROCEDURE. NO PATIENT OR CLINICIAN HARM REPORTED. NO OTHER INFORMATION WAS PROVIDED. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01839 | 2023-05-12 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE DRIPS AROUND STYLET WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. CLINICIAN DISCARDS DEVICE AND OBTAINS A SECOND DEVICE FOR INSERTION, WHICH ALSO LEAKED AROUND THE STYLET. THE 3RD PICC LEAKED WHEN FLUSHING THE CATHETER AT THE END OF THE PROCEDURE, BUT DID NOT LEAK BEFORE THE PROCEDURE. NO PATIENT OR CLINICIAN HARM REPORTED. NO OTHER INFORMATION WAS PROVIDED. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01837 | 2023-05-12 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE DRIPS AROUND STYLET WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. CLINICIAN DISCARDS DEVICE AND OBTAINS A SECOND DEVICE FOR INSERTION, WHICH ALSO LEAKED AROUND THE STYLET. THE 3RD PICC LEAKED WHEN FLUSHING THE CATHETER AT THE END OF THE PROCEDURE, BUT DID NOT LEAK BEFORE THE PROCEDURE. NO PATIENT OR CLINICIAN HARM REPORTED. NO OTHER INFORMATION WAS PROVIDED. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01840 | 2023-05-12 | 2023-01-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 3F SL SOLO MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | S1193108D5 | S1193108D5 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE "ECG WIRE BENT WITHIN PICC LINE." ADDITIONAL INFORMATION 01/24/2023: IT WAS REPORTED BY THE CUSTOMER "HAD TO OPEN A NEW PROVENA PICC KIT DUE TO MALFUNCTIONS OF LINE. PICC LINE NEVER TOUCHED THE PATIENT." 4/20/2023: THE RETURNED DEVICE WAS CONFIRMED TO BE A BENT STYLET. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FAILURE WAS FOUND DURING EVALUATION OF A RETURNED SAMPLE. ADDITIONAL INVESTIGATION IS REQUIRED TO DETERMINE THE CAUSE OF THE FAILURE. RESULTS OF THE INVESTIGATION ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03374 | 2023-05-12 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT ON (B)(6), 2023, THE CATHETER WAS PLANNED TO BE REMOVED FROM THE PATIENT. HOWEVER, ONLY A LITTLE SALINE INJECTED IN THE BALLOON COULD BE DRAINED OUT BY THE SYRINGE AND REST OF THE SALINE COULD NOT BE WITHDRAWN, WHICH LED THE CATHETER UNABLE TO BE REMOVED FROM THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01833 | 2023-05-12 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN OF THE NEEDLE STICK WHEN THE PORT GETS ACCESSED. HOWEVER, THERE IS NO INFORMATION REGARDING TREATMENT OR MEDICATION PROVIDED FOR PAIN. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03372 | 2023-05-12 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PATIENT WAS PLACED 20F THREE-LUMEN CATHETER FOR URINARY CATHETERIZATION AT 10:30 ON (B)(6) 2023 AND WAS MAINTAINED WITH IRRIGATION. WHILE AT 23:55, THE IRRIGATION LIQUID WAS FOUND LEAKING FROM THE URETHRAL ORIFICE. UPON CHECKING, THE CATHETER HAD COMPLETELY FALLEN OUT OF THE PATIENT AND THE BALLOON WAS BURST WITH NO MISSING PIECES LEFT INSIDE THE PATIENT. THEN A NEW CATHETER WAS REPLACED AND CONTINUED IRRIGATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03376 | 2023-05-12 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED, THAT THE PATIENT COMPLETED THE OPERATION ON (B)(6). THE DOCTOR PERFORMED INDWELLING CATHETER AND INJECTED 40ML OF WATER BAG TO STOP THE BLEEDING. ON (B)(6), THE PATIENT FIRST COMPLAINED THAT THEY HAD THE DESIRE TO URINATE WHILE HOLDING BACK THEIR URINE. AND THE NURSE ON DUTY DID THE WARD ROUNDS WITHOUT ABNORMALITIES. ON THE EVENING OF APRIL 7TH, THE PATIENT LEAKED URINE DURING THE WARD ROUND, AND INFORMED THE DOCTOR TO PERFORM CATHETERIZATION IMMEDIATELY. DURING THE CATHETERIZATION PROCESS, THE NURSE SAW THAT THE CATHETER SLIPPED OUT A LOT AT 23:00. THE NURSE CHECKED THE CATHETER AGAIN AND FOUND THAT THE WATER SAC WAS BROKEN, SO THE CATHETER WAS REMOVED. THE INDWELLING CATHETER WAS PERFORMED AGAIN. ON PHYSICAL EXAMINATION, THE PATIENT'S VITAL SIGNS WERE STABLE AND HIS ABDOMEN WAS SOFT WITHOUT TENDERNESS. BOWEL SOUNDS WERE 3 TIMES/MIN. THE INCISION DRESSING WAS FIXED, AND FIXED WITHOUT EXUDATION. AS PER THE FOLLOW UP INFORMATION RECEIVED ON (B)(6) 2023, THE LEAKAGE HAPPENED, DUE TO BALLOON BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03370 | 2023-05-12 | 2023-04-20 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT ON MARCH 30, THE CATHETER WAS PLACED IN THE UROLOGY OUTPATIENT DEPARTMENT, AND ON APRIL 20, THE CATHETER WAS REMOVED AND REPLACED AT ANOTHER HOSPITAL DUE TO CATHETER RUPTURE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03364 | 2023-05-12 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER COULD NOT BE INJECTED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 26APR2023, CLARIFIED THAT THE STERILE WATER DID NOT ENTER. THE DETAILS REGARDING ANY DEFECT IN SYRINGE WEREN'T PROVIDED FROM CUSTOMER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01836 | 2023-05-12 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 4F DL PROVENA SOLO BASIC W/3CG ROW | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT NUMBERS ON PICC LINE WERE NOT CLEARLY DEFINED. IT APPEARS THE INK ON #40 IS SMEARED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03363 | 2023-05-12 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT WHEN WATER WAS INJECTED DURING PRETEST, THE FOLEY CATHETER BALLOON WAS INFLATED BUT WATER COULD NOT BE WITHDRAWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03362 | 2023-05-12 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 806318 | 806318 |  | IT WAS REPORTED THAT THE CUSTOMER STARTED GETTING THE 806318 FOLEY CATHETERS A COUPLE OF MONTHS AGO AND NOW THE URINE ALWAYS LEAKED AROUND THE CATHETER AND SOILED THE PATIENTS WHEELCHAIR AND BED. THEY WANTED TO GO BACK TO THE CATHETER THEY HAD BEFORE, BUT THE NUMBER THEY PROVIDED DID NOT MATCH ANYTHING BD HAD. THEY DESCRIBED THE CATHETER AS LOOKING LIKE IT WAS THE COLOR OF A TAN RUBBER BAND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03357 | 2023-05-12 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998006 | 998006 |  | IT WAS REPORTED THAT THE GUIDE WIRE COULD NOT BE INSERTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00330 | 2023-05-12 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQ7584 | CQ7584 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE BALLOON ALLEGEDLY CANNOT DEFLATE FULLY. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03356 | 2023-05-12 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER WAS HARD TO INSERT AND DRY. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03373 | 2023-05-12 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123614A |  | IT WAS REPORTED THAT THE PATIENT GOT CATHETERIZATION FOR URINE DRAINAGE AND WAS OBSERVED THE COLOR OF URINE. THEN THE BALLOON OF CATHETER WAS FOUND BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01869 | 2023-05-12 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 4FR DL PROVENA MIDLINE FULL KIT W/GUARDIVA AND PROBE COVER | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | S4254108DGP |  | CUSTOMER REPORTED THAT THE TEAR AWAY SHEATH DID NOT FULLY BREAK IN TWO. CATHETER STUCK IN ORANGE PART OF SHEATH. USER HAD TO TROUBLESHOOT SHEATH AND INSERT CATHETER BY BYPASSING ORANGE PORTION OF SHEATH. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF AN IMPROPER SHEATH PEEL WAS CONFIRMED AND THE CAUSE APPEARED TO BE RELATED TO DEVICE MANUFACTURE. THE PRODUCT RETURNED FOR EVALUATION WAS ONE PHOTOGRAPH WHICH DEPICTED A 4.5FR MICROINTRODUCER PEEL-APART SHEATH. THE DEVICE WAS HELD IN A GLOVED HAND. BLOOD RESIDUE WAS VISIBLE ON THE SAMPLE AND THE GLOVE. THE GREY SHEATH WAS COMPLETELY SPLIT AND ONE HALF, INCLUDING THE ORANGE PULL TAB WAS NOT VISIBLE. THE REMAINING PULL TAB WAS ATTACHED TO ONE OF THE SPLIT GREY TUBE HALVES. A RING OF ORANGE PLASTIC WAS VISIBLE AT THE PROXIMAL END OF THE PULL TAB AND A LENGTH OF POWERPICC CATHETER WAS VISIBLE INLAID THROUGH THAT RING. THE MATERIAL SURROUNDING THE CATHETER SHAFT WAS CONSISTENT WITH THE ORANGE PLASTIC OF THE PULL TAB AND DEMONSTRATED THAT THE PULL TABS HAD NOT SPLIT CORRECTLY DURING DEVICE USE. THE DEPICTED DEVICE HAD NOT SPLIT CORRECTLY, AS INDICATED IN THE EVENT DESCRIPTION. THE IMPROPER FRACTURE BETWEEN THE TWO ORANGE PULL TABS APPEARED TO BE CAUSED BY THE MOLDING/SKIVING PROCESS DURING DEVICE MANUFACTURE. PHOTOGRAPHS HAVE BEEN FORWARDED TO THE MANUFACTURING SITE FOR FURTHER EVALUATION. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01873 | 2023-05-13 | 2023-04-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT ALLEGEDLY MOVED AND WAS MUCH MORE SUPERFICIAL AT THE SURFACE OF THE SKIN AS THREE BUMPS WERE VISIBLE. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03387 | 2023-05-15 | 2022-11-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD UNCATEGORIZED URETHRAL DAMAGE ONLY . THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED, AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY (BATCH # KME21J1763 AND BATCH # KME22C3259). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03388 | 2023-05-15 | 2022-11-30 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THIS PATIENT HAD URETHRAL DAMAGE OF 1CM. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED, AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01883 | 2023-05-15 | 2023-03-13 | I, n, j, u, r, y | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602830 | 0602830 |  | IT WAS REPORTED THAT SOME TIME POST PORT PLACEMENT PROCEDURE, RUPTURE OF THE CATHETER WAS ALLEGEDLY LODGED IN THE INTRAVASCULAR REGION. THE PATIENT ALLEGEDLY EXPERIENCED HEART INJURY. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 12/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01878 | 2023-05-15 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5733690 | 5733690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, AIR WAS ALLEGEDLY SUCKED IN UNDER NEGATIVE PRESSURE. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01903 | 2023-05-15 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT MRI ISP, 8 FR CHRONOFLEX, INT WO SP, ATTACH SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808061 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE DEVICE ALLEGEDLY DID NOT PROGRESS WITH THE GUIDE WIRE IN THE FIRST PUNCTURE. IT WAS FURTHER REPORTED THAT THE GUIDEWIRE WAS ALLEGEDLY BENT DURING THE ATTEMPT TO PROGRESS AND WAS DIFFICULT TO INSERT. REPORTEDLY, THE PATIENT EXPERIENCED DISCOMFORT, PAIN AND BRUISE. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING ADDITIONAL INTERVENTION OR MEDICATION PRESCRIBED TO TREAT PAIN, DISCOMFORT AND BRUISE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT MRI ISP, 8 FR CHRONOFLEX, INT WO SP, ATTACH SL THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT MRI ISP, 8 FR CHRONOFLEX, INT WO SP, ATTACH SL ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01889 | 2023-05-15 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK POWER PORT |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03389 | 2023-05-15 | 2022-10-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD UNCATEGORIZED URETHRAL DAMAGE ONLY . THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00333 | 2023-05-15 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DORADO | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | DR4054 | DR4054 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, A HOLE WAS ALLEGEDLY FOUND AT THE SHAFT OF THE CATHETER NEAR THE ROOT OF THE BALLOON UPON INFLATING AT 8ATM. IT WAS FURTHER REPORTED THAT THE SHAFT WAS ALLEGEDLY FOUND TO BE BROKEN. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE DORADO PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING FUNCTIONAL EVALUATION, AN IN-HOUSE PRESTO INFLATION DEVICE WAS USED TO INFLATE THE BALLOON AND WATER WAS NOTED LEAKING FROM THE GLUE JOINT. UNDER MICROSCOPIC EXAMINATION, A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED TO THE GLUE JOINT. NO FURTHER TESTING WAS PERFORMED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE ALLEGED BREAK AS BREAK TO THE GLUE JOINT WAS NOTED UNDER MICROSCOPIC EVALUATION. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED BREAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 07/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03390 | 2023-05-15 | 2023-03-18 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THERE WAS NO URINE IN THE URINE BAG AFTER THE OPERATION. THE URINARY CATHETER HAD SLIPPED AND THE WATER POLO WAS FOUND TO BE BROKEN. THE PATIENT SAID THAT THEY DID NOT PULL THE URINARY CATHETER. THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD OR BODILY FLUID. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03396 | 2023-05-15 | 2022-12-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD UNCATEGORISED URETHRAL DAMAGE ONLY (PATIENT #15). THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01899 | 2023-05-15 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER DL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED BY THE CUSTOMER, "A TORTUOUS VESSEL MAKING IT IMPOSSIBLE TO PROGRESS THE MATERIAL, REQUIRING A NEW ATTEMPT, WHICH, WHEN REMOVING PREVIOUS MATERIAL, THE CATHETER WAS DEFORMED. A NEW PROCEDURE WAS SUCCESSFULLY PERFORMED IN A NEW VESSEL, THE MATERIAL WAS DISCARDED.". | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03426 | 2023-05-15 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER WAS UNABLE TO REMOVED FROM THE FOLEY CATHETER. THE BALLOON WAS INFLATED DURING A PRETEST BUT THE WATER WAS UNABLE TO REMOVED EVEN THOUGH IT WAS INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01902 | 2023-05-15 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6903230 |  | IT WAS REPORTED THAT FOUR DAYS POST DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY HAD CUT/CRACK IN ONE OF THE EXTENSIONS OF THE DEVICE. IT WAS FURTHER REPORTED THAT DEVICE ALLEGEDLY HAD BUBBLES IN ONE OF THE EXTENSIONS. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A VIDEO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03417 | 2023-05-15 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 997706 | 997706 |  | IT WAS REPORTED THAT THE URETERAL DILATION BALLOON WAS DAMAGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03422 | 2023-05-15 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT WHILE ATTACHING THE SYRINGE TO OBTAIN THE URINE FORM THE SAMPLE ATTACHED TO THE URINARY CATHETER CAME OFF AND URINE WAS LEAKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03407 | 2023-05-15 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175814 | 175814 |  | IT WAS REPORTED THAT ONCE THE FOLEY CATHETER WAS PUT IN PLACE, UPON INFLATION OF THE CATHETER BALLOON, A LEAK WAS NOTED AT THE Y-JUNCTION OF THE CATHETER (THE PART THAT THEY USUALLY SECURE IN THE STATLOCK). THE ITEM WAS DC'D AND A NEW CATHETER PLACED WITHOUT ISSUE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01897 | 2023-05-15 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED, "THEY WERE CARRYING OUT A PICC INSTALLATION PROCEDURE, WHEN OPENING THE PRODUCT THEY NOTICED THE EXISTENCE OF A HAIR IN THE STERILE MATERIAL." NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01901 | 2023-05-15 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F DUAL-LUMEN POLYURETHANE CATHETER RADIOLOGY BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275335 |  | IT WAS REPORTED THAT, "A PIECE OF HUMAN HAIR HAS BEEN FOUND IN THE PACKAGE AFTER OPENED SO THE PRODUCT WAS CONSIDERED AS CONTAMINATION AND CAN'T BE USED, THEN THE NEW ONE HAD TO BE REPLACED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF FOREIGN MATERIAL IS CONFIRMED; HOWEVER, THE ROOT CAUSE COULD NOT BE DETERMINED. THE PRODUCT RETURNED FOR EVALUATION WAS A 5FR DUAL LUMEN POWERPICC CATHETER, A 5FR MICROINTRODUCER, A 21GA INTRODUCER NEEDLE, AN IR GUIDEWIRE, AND TWO END CAPS. THE KIT TRAY WAS RECEIVED PARTIALLY OPENED. A HAIR WAS FOUND IN THE KIT TRAY. ALSO RECEIVED WAS A PHOTOGRAPH OF THE OPENED KIT WHICH ALSO SHOWED A HAIR WITHIN THE KIT TRAY. DUE TO THE OPEN STATE OF THE KIT, IT COULD NOT BE DETERMINED WHEN THE HAIR ENTERED¬ THE KIT TRAY. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01904 | 2023-05-15 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK APHERESIS PORT ACCESS SYS |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT , THE DEVICE ALLEGEDLY HAD OBSTRUCTION OF FLOW. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01874 | 2023-05-15 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT UPON RECEIVING THE DEVICE, THE PORT PACKAGE ALLEGEDLY HAD CUTS AND STERILITY OF THE DEVICE WAS COMPROMISED. THERE WAS NO PATIENT CONTACT. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE SEALED POWERPORT CLEARVUE ISP IMPLANTABLE PORT KIT WAS RECEIVED FOR EVALUATION AND TWO PHOTOS WERE PROVIDED FOR REVIEW. VISUAL EVALUATION WAS PERFORMED. A LONGITUDINAL SPLIT WAS NOTED ON THE TOP CENTER PORTION OF THE OUTER PACKAGING AND THE SPLIT APPEARED TO GO WITHIN THE PRODUCT TRAY LABEL. MANUFACTURING SITE EVALUATION OF THE SAMPLE FOUND A LONGITUDINAL CUT AT THE BOTTOM OF THE UNIT LABEL PLACEMENT, THE BACK OF THE TRAY AND THE INNER PACKAGING APPEARED TO BE FREE OF DAMAGE, THE CONDITION OF THE SHIPPING BOX IS UNKNOWN. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED PACKAGING OF THE DEVICE TORN ISSUE AND THE PHOTO REVIEW ALSO CONFIRMS THE SAME. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 11/2024), G3, H6 (METHOD). THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01895 | 2023-05-15 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1275108FD5 |  | IT WAS REPORTED BY THE CUSTOMER AND VIA MEDWATCH "DURING PICC PROCEDURE, BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLETS S1275108FD5, LOT# REHN2437 2 KITS NOTED LEAKAGE AT RUBBER-LIKE STOPPER AND STYLET BEFORE INSERTION AND REPLACED. 3RD ONLY LEAKED AFTER INSERTION AND RESOLVED WITH STYLET REMOVAL. PICC SUCCESSFULLY PLACED. 2 PICCS RETAINED." IT WAS REPORTED THIS OCCURRED WITH 3 DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE BEFORE INSERTION. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01875 | 2023-05-15 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT UPON RECEIVING THE DEVICE, THE PORT PACKAGE ALLEGEDLY HAD CUTS AND STERILITY OF THE DEVICE WAS COMPROMISED. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE SEALED POWERPORT CLEARVUE ISP IMPLANTABLE PORT KIT WAS RECEIVED FOR EVALUATION AND TWO PHOTOS WERE PROVIDED FOR REVIEW. VISUAL EVALUATION WAS PERFORMED. A LONGITUDINAL SPLIT WAS NOTED ON THE TOP CENTER PORTION OF THE OUTER PACKAGING AND THE SPLIT APPEARED TO GO WITHIN THE PRODUCT TRAY LABEL. MANUFACTURING SITE EVALUATION OF THE SAMPLE FOUND A LONGITUDINAL CUT WAS OBSERVED IN THE UPPER CENTRAL PART OF THE OUTER PACKAGING. THE CUT TORE PART OF THE UPPER CENTER SEAL AS WELL AS THE UNIT LABEL AND THE OUTER LID OF THE TRAY. THE OUTER PACKAGING SEEMED TO BE INTACT. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED PACKAGING OF THE DEVICE TORN ISSUE AND THE PHOTO REVIEW ALSO CONFIRMS THE SAME. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 11/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01879 | 2023-05-15 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE CATHETER ALLEGEDLY FRACTURED. IT WAS FURTHER REPORTED THAT THE PORT ALLEGEDLY HAD A LEAK. FURTHERMORE, PATIENT EXPERIENCED SWELLING OVER NECK AREA AFTER SALINE INJECTION. HOWEVER, THERE IS NO INFORMATION REGARDING MEDICAL INTERVENTION OR TREATMENT PROVIDED FOR SWELLING. REPORTEDLY, THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 06/2024). SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01881 | 2023-05-15 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 4808060 | 4808060 |  | IT WAS REPORTED THAT APPROXIMATELY NINE YEARS POST PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY HAD CLOTTING ISSUE WITH THE PORT. IT WAS FURTHER REPORTED THE PORT ALLEGEDLY HAD LOW FLOW RATE AND DIFFICULT TO FLUSH. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 05/2015). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01882 | 2023-05-15 | 2023-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR AND ELEVEN MONTHS POST PORT PLACEMENT, THE PORT ALLEGEDLY HAD NO BLOOD RETURN. IT WAS FURTHER REPORTED THAT THE PORT WAS ALLEGEDLY FOUND TO BE FRACTURED UNDER CHEST X-RAY EXAMINATION. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 03/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01908 | 2023-05-15 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118J |  | IT WAS REPORTED THAT PRIOR TO A PLACEMENT PROCEDURE, IT WAS NOTICED THAT THE ROUND TAPE WAS OPENED WITH A SHARP OBJECT, AND THE SELLOTAPE PASTED ON IT WAS ALSO SHARPLY CUT DIAGONALLY. IN ADDITION, THE BOX WAS SLIGHTLY BENT ON ONE SIDE WHERE YOU COULD PUT FINGER IN WHEN OPENING IT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01906 | 2023-05-15 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED VIA MEDWATCH "PICC LINES LEAKING AT THE RUBBER-LIKE STOPPER UPON ATTEMPTS TO FLUSH." THE EXACT QUANTITY OF LEAKING PICCS WAS NOT PROVIDED BY THE CUSTOMER TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. THE EXACT QUANTITY OF LEAKING PICCS WAS NOT PROVIDED BY THE CUSTOMER TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE HOWEVER BD HAS ADDRESSED 4 SIMILAR EVENTS UNDER THE FOLLOWING MEDWATCH REPORTS: 3006260740-2023-01578; 3006260740-2023-01579; 3006260740-2023-01580; 3006260740-2023-01619. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01909 | 2023-05-15 | 2023-04-28 | I, n, j, u, r, y | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (20G) (8CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120080T |  | IT WAS REPORTED BY THE CUSTOMER, THE POWERGLIDE CATHETER BROKE IN IN HALF UPON REMOVING AFTER HITTING OBSTRUCTION. ADDITIONAL INFORMATION 05/09/2023: PATIENT NEEDED TO HAVE ADDITIONAL SURGERY TO REMOVE RETAINED CATHETER. THEY WERE BEING TREATED FOR C-DIFF AND CRYPTOSPORIDIUM, ELECTROLYTE IMBALANCE. THEY DO HAVE AN INFECTION DISEASE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01910 | 2023-05-15 | 2023-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT GOT CALL ON MONDAY FROM THE IR CATH LAB IN CHARGE ABOUT THE PRODUCT COMPLAINT. ACCORDING TO THE USER FACE CHALLENGE DURING THE PROCEDURE OF PICC PLACEMENT. THEY PLACE THE PICC LINE BUT WHEN THEY CONNECT IT WITH THE CONNECTOR ASSEMBLY IT DID NOT PASS THROUGH THE CONNECTER SLEEVE AFTER TWO THREE ATTEMPT IT PASSES WITH THE SLEEVE AND CONNECTED. BUT AFTER CONNECTING IT WITH CONNECTOR, THEY ARE NOT ABLE TO ASPIRATE OR NOT ABLE TO GET THE BACK FLOW. HE TRIED 2-3 TIMES BUT FAIL TO START THE PICC LINE. THEREFORE, THEY DECIDE TO CHANGE THIS DEFECTED PICC LINE WITH OTHER PICC LINE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03416 | 2023-05-15 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER LEAKED FROM AROUND THE INFLATION VALVE WHEN STERILE WATER WAS INJECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01907 | 2023-05-15 | 2023-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SECURELOC SAFETY INTRODUCER NEEDLE 21G X 2.75 IN | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S03021-275 |  | IT WAS REPORTED BY THE CUSTOMER THAT "GREEN SAFETY MECHANISM BROKE OFF OF THE NEEDLE" NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01877 | 2023-05-15 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716000J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE PORT SYSTEM WAS ALLEGEDLY FAILED TO ASPIRATE. IT WAS FURTHER REPORTED THAT THE SALINE SOLUTION ALLEGEDLY FAILED TO INFUSE. THE PORT SYSTEM WAS REMOVED, AND ANOTHER NEW PORT SYSTEM WAS REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03434 | 2023-05-15 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 165812 |  | IT WAS REPORTED THAT A PHYSICIAN STATED IN AN ONLINE SURVEY THAT THEY WERE NOT ABLE TO SUCCESSFULLY PLACE THE FOLEY CATHETER IN A PATIENT BECAUSE IT WAS PLACED BY UROLOGY TEAM AFTER NON DRAINAGE OR RETENTION FROM PREVIOUS URETHRAL CATHETER AND URETHRAL TRAUMA TRYING TO INSERT A NEW CATHETER IN RELATION TO 2-WAY ALL-SILICONE UNCOATED FOLEY CATHETERS FEMALE LENGTH 10CC BALLOON, FRENCH SIZE 12. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01894 | 2023-05-15 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1275108FD5 |  | IT WAS REPORTED BY THE CUSTOMER AND VIA MEDWATCH "DURING PICC PROCEDURE, BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLETS S1275108FD5 LOT# REHN2437 2 KITS NOTED LEAKAGE AT RUBBER-LIKE STOPPER AND STYLET BEFORE INSERTION AND REPLACED. 3RD ONLY LEAKED AFTER INSERTION AND RESOLVED WITH STYLET REMOVAL. PICC SUCCESSFULLY PLACED. 2 PICCS RETAINED." IT WAS REPORTED THIS OCCURRED WITH 3 DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE BEFORE INSERTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03423 | 2023-05-15 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX®LUBRICATH® FOLEY CATHETER ALCOCK MODEL | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0134L18 | 0134L18 |  | IT WAS REPORTED THAT STERILE WATER DID NOT ENTER THE FOLEY CATHETER THROUGH THE INFLATION VALVE. INJECTING STERILE WATER FROM THE INFLATION VALVE AFTER THE FOLEY CATHETER INSERTION TO THE PATIENT, THE WATER WAS UNABLE TO INJECT. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE WERE CONFIRMED TO EXHIBIT THE REPORTED FAILURE. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), USED THREE-WAY LATEX CATHETER FOLEY. VISUAL INSPECTION OF THE SAMPLE NOTED NO VISIBLE DEFECTS WERE NOTED ON THE CATHETER. ATTEMPTED TO INFLATE THE BALLOON WITH 30CC DI WATER USING A LUER LOK SYRINGE, HOWEVER, ALL THE WATER BEGAN LEAKING FROM THE IRRIGATION EYE AND THE BALLOON COULD NOT BE INFLATED. THIS DOES NOT MEET SPECIFICATION WHICH STATES THAT "FUNCTIONAL - LEAK TESTING", "INFLATION OR DEFLATION DIFFICULTIES". A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE "OPERATOR ERROR¿. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE THE BALLOON TO BURST. CAUTION: DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. SINGLE PATIENT USE ONLY. DO NOT REUSE. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP TIP SYRINGE. DO NOT USE NEEDLE. RECOMMENDED INFLATION CAPACITIES; 3CC BALLOON: USE 5ML STERILE WATER, 5CC BALLOON: USE 10ML STERILE WATER, 15CC BALLOON: USE 20ML STERILE WATER, 20CC BALLOON: USE 25ML STERILE WATER, 30CC BALLOON: USE 35ML STERILE WATER, 40CC BALLOON: USE 45ML STERILE WATER, 75CC BALLOON: USE 80ML STERILE WATER. DO NOT EXCEED RECOMMENDED CAPACITIES. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE, AND FEDERAL LAWS AND REGULATIONS. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03406 | 2023-05-15 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 175814 | 175814 |  | IT WAS REPORTED, THAT THE FOLEY CATHETER WAS PLACED INSIDE THE PATIENT. AND NOTICED LEAKING COMING FROM THE FOLEY TUBING LOCATED NEAR THE END, WHERE THE FOLEY CATHETER SPLITS OFF TO THE BALLOON PORT, AND WHERE IT CONNECTED TO THE CLOSED BAG SYSTEM. FURTHER INSPECTION OF THE CATHETER TUBE REVEALED, THAT THERE APPEARED TO BE A PUNCTURE HOLE AT THE SITE OF THE LEAK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03403 | 2023-05-15 | 2023-01-11 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD HA URETHRAL DAMAGE OF 1.0 CM (PATIENT #19). THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03401 | 2023-05-15 | 2022-12-22 | I, n, j, u, r, y | Manufacturer report | ALL SILICONE CATHETER | UNKNOWN SILICONE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD HA URETHRAL DAMAGE OF 0.3 CM (PATIENT #18). THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03398 | 2023-05-15 | 2023-01-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLAEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD UNCATEGORIZED URETHRAL DAMAGE ONLY (PATIENT #16). THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED, AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03394 | 2023-05-15 | 2023-01-20 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD UNCATEGORISED URETHRAL DAMAGE ONLY (PATIENT # 14). THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03392 | 2023-05-15 | 2022-11-21 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD UNCATEGORISED URETHRAL DAMAGE ONLY (PATIENT # 6). THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03400 | 2023-05-15 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED, THAT CUSTOMER RECEIVED MAGIC 3, GO MALE COUDE INTERMITTENT CATHETER SAMPLES. WHICH APPEARS TO BE SLIGHTLY BIGGER IN DIAMETER. IN THE PHOTO PROVIDED, CUSTOMER STATES, THE CATHETER FELT LARGER WHEN USING. BUT DID NOT CAUSE SIGNIFICANT DISCOMFORT OR TRAUMA. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03409 | 2023-05-15 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER LEAKAGE OCCURRED DURING STERILE WATER INJECTION, AND WHEN CHECKED, THERE WAS A CRACK AT THE BASE OF THE INFLATION VALVE. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. IT IS UNKNOWN IF THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED MULTIPLE CRACKS/SCRATCHES UNDER THE INFLATION CAP. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND SOLUTION IMMEDIATELY LEAKED OUT A 0.975" TEAR FOUND UNDER INFLATION CAP. THIS IS OUT OF SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03405 | 2023-05-15 | 2023-01-12 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD OA URETHRAL DAMAGE OF 0.8 CM (PATIENT #20). THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01896 | 2023-05-15 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1275108FD5 |  | IT WAS REPORTED BY THE CUSTOMER AND VIA MEDWATCH "DURING PICC PROCEDURE, BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLETS S1275108FD5 LOT# REHN2437 2 KITS NOTED LEAKAGE AT RUBBER-LIKE STOPPER AND STYLET BEFORE INSERTION AND REPLACED. 3RD ONLY LEAKED AFTER INSERTION AND RESOLVED WITH STYLET REMOVAL. PICC SUCCESSFULLY PLACED. 2 PICCS RETAINED." IT WAS REPORTED THIS OCCURRED WITH 3 DEVICES. THIS REPORT ADDRESSES THE SUCCESSFULLY PLACED DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03393 | 2023-05-15 | 2023-01-16 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD UNCATEGORISED URETHRAL DAMAGE ONLY (PATIENT #13). THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03411 | 2023-05-15 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE EYELETS ON THE INTERMITTENT CATHETERS WERE NOT CUT DEEPLY ENOUGH WHERE THEY FALL OFF. CUSTOMER STATED THAT THEY HAD TO PICK THEM UP OFF OF THE FLOOR. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "IMPROPER/INADEQUATE MAINTENANCE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEM OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03424 | 2023-05-15 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DUE TO THE BALLOON BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03429 | 2023-05-15 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE FUNNEL UPON INFLATING THE FOLEY CATHETER BALLOON DURING A PRETEST. UPON CHECKING THE RETURNED SAMPLE, THERE WAS A PINHOLE AT THE ROOT OF THE FUNNEL. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND IT IMMEDIATELY LEAKED OUT OF A 0.013" PINHOLE FOUND AT THE TRIFURCATION. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE (B)(4), REVISION 0, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ CORRECTIONS: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03425 | 2023-05-15 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FIXED WATER WAS DIFFICULT TO DRAIN FROM THE FOLEY CATHETER. THIS EVENT WAS FOUND DURING A PRETEST AND THE CATHETER WAS NOT USED ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03430 | 2023-05-15 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT AFTER INSERTING THE CATHETER INTO THE PATIENT AND INFLATING THE BALLOON, NO URINE WAS OBSERVED TO FLOW OUT, SO THE BALLOON WAS ATTEMPTED TO DEFLATE, BUT WATER COULD NOT BE DRAINED. THE USER STATED THAT THE ISSUE MIGHT HAVE OCCURRED AS THE BALLOON WAS INFLATED WITHOUT CONFIRMING THE URINE OUTFLOW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03420 | 2023-05-15 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT GOT CATHETERIZATION ON THE NIGHT OF APRIL 20. THE POSITION AND DRAINAGE WAS GOOD WHILE THE YELLOW URINE CAN BE DRAINED. ON APRIL 21, AT 15:30, WHEN CLEANING PERINEUM, IT WAS FOUND THE CATHETER FELL OUT AND THE BALLOON WAS BURST. THEN A NEW CATHETER WAS REPLACED. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, THERE WAS NO MISSING PIECES INSIDE THE PATIENT AND NO HARM WAS EXPERIENCED BY THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03427 | 2023-05-15 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS COLLAPSED BEFORE PHYSICAL COLLECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01898 | 2023-05-15 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT DURING A CHRONIC CATHETER PLACEMENT PROCEDURE IN RIGHT INTERNAL JUGULAR VEIN WITH A POLYESTER CUFF FOR RENAL DIALYSIS, THE CONNECTION WITH THE INJECTION WAS ALLEGEDLY FOUND TO BE DAMAGED. IT WAS FURTHER REPORTED THAT THE PHYSICIAN PINCHED THE DAMAGED PART AND INSERTED THE NEEDLE AGAIN AND FED THE GUIDE WIRE THROUGH THE GUIDE WIRE HOLE OF THE PUNCTURE NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 01/2024). SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03414 | 2023-05-15 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT ONE OF THE INTERMITTENT CATHETERS FROM THE LAST ORDER HAD AN EXTRA PIECE OF PLASTIC ON IT THAT WAS STICKING OUT AND THE PATIENT DID NOT SEE WHEN INSERTING. PATIENT STATED THAT IT HURT THEM BUT NO BLEEDING AND THEY NOTICED THIS EXTRA PIECE OF PLASTIC. PATIENT STATED THAT THE DEFECT WAS ABOUT 4 AND 1/4 INCHES FROM THE INSERTION END OF THE CATHETER. PATIENT STATED THAT THEY STILL HAD THE CATHETER AT HOME. PATIENT HAD USED THIS PRODUCT BEFORE. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03408 | 2023-05-15 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT FROM THE BEGINNING, THERE WAS A HOLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01905 | 2023-05-15 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT AFTER PUNCTURE AND INSERTION OF A GUIDE WIRE IN THIS PATIENT, THE DOCTOR FOUND A BALL OF WIRE PROTRUDING BEYOND THE DISTAL END. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DEFORMATION ON THE END OF THE GUIDEWIRE WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. TWO PHOTOGRAPHS OF THE IMPLICATED 0.018¿ X 35CM GUIDEWIRE WERE RETURNED FOR EVALUATION. THE GUIDEWIRE WAS HELD WITHIN A GLOVED HAND AND CONTAINED EVIDENCE OF USE (E.G. RESIDUAL MATERIAL). THE END OF THE GUIDEWIRE APPEARED ROUGH AND IRREGULAR; WITH WHAT APPEARED TO BE A SLIGHTLY LARGER OUTER DIAMETER COMPARED TO THE REST OF THE WIRE. THE PHOTO QUALITY DID NOT ALLOW FOR A CLEAR VIEW OF THE DAMAGED WIRE SECTION; HOWEVER, THE IRREGULAR PROFILE OF THE END OF THE WIRE WAS LIKELY CAUSED BY THE OUTER WIRE BEING DISPLACED. ALTHOUGH THE DAMAGE TO THE WIRE COULD BE CONFIRMED, THE EXACT ROOT CAUSE OF THE DAMAGE COULD NOT BE DETERMINED FROM THE RETURNED PHOTOGRAPHS. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. POTENTIAL CONTRIBUTING FACTORS COULD INCLUDE ATTEMPTED INSERTION OF THE WIRE AGAINST RESISTANCE OR MANIPULATION OF THE WIRE CAUSING DISPLACEMENT OF THE OUTER WIRE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01927 | 2023-05-16 | 2023-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED "THE CLINICIAN WAS FEEDING INTRODUCER OVER GUIDEWIRE. SHE FELT RESISTANCE SO SHE WITHDREW GUIDEWIRE AND INTRODUCER AT THE SAME TIME. SHE NOTICED FRAYED WIRE." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GUIDEWIRE IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE 0.018 IN. NITINOL GUIDEWIRE WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLE. A MICROSCOPIC OBSERVATION REVEALED THE ADHESIVE FILLET NEAR THE PROXIMAL END OF THE COILED WIRE OF THE GUIDEWIRE APPEARED TO BE MISSING. THE COILED WIRE WAS OBSERVED TO BE UNRAVELED AND KINKED AT ITS PROXIMAL END WITH THE PROXIMAL END OF THE COILED WIRE PROTRUDING OUTWARDS. BASED ON THE EVIDENCE OBSERVED ON THE RETURNED SAMPLE, IT WAS NOT POSSIBLE TO DETERMINE WHEN OR WHERE THE GUIDEWIRE WAS DAMAGED OR DETERMINE IF A DEFECT EXISTED ON THE GUIDEWIRE PRIOR TO USE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03460 | 2023-05-16 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT AFTER INSERTING THE CATHETER INTO THE PATIENT AND INFLATING THE BALLOON, NO URINE WAS OBSERVED TO FLOW OUT , SO THE BALLOON WAS ATTEMPTED TO DEFLATE, BUT WATER COULD NOT BE DRAINED. THE USER STATED THAT THE ISSUE MIGHT HAVE OCCURRED AS THE BALLOON WAS INFLATED WITHOUT CONFIRMING THE URINE OUTFLOW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01928 | 2023-05-16 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER 5F RADIOLOGY BASIC TRAY (DL) (135CM GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3295135 |  | IT WAS REPORTED BY THE CUSTOMER PICC LINE PLACED APRIL 10TH ISSUES DEVELOPED WHERE PICC LINE PORT COULD NOT BE USED, UNABLE TO FLUSH AND ASPIRATE FROM RED PORT. ALTAPLASE WAS PUT IN LINE BUT IT DID NOT CLEAR THE LINE. NO APPARENT HARM, INCONVENIENT. ADDITIONAL INFORMATION 09-MAY-2023. THE PICCS WERE BEING USED FOR IV THERAPY INCLUDING FLUID, ANTIBIOTICS AND NUTRITION. THE STAFF WERE TRYING TO FLUSH THE PICCS AS PER PROTOCOL. WE ATTEMPTED TO DEAL WITH THE PICCS USING ALTEPLASE, ANTIBIOTICS AND NUTRITION HELD. CAUSED A DELAY IN ANTIBIOTIC THERAPY AND NUTRITION. THERE WERE NO FURTHER SURGERIES, TESTS EXCEPT FOR PICC LINES NEEDING TO BE RETHREADED AND/OR REPLACED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01929 | 2023-05-16 | 2023-05-02 | I, n, j, u, r, y | Manufacturer report | BARD TL PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. THIS EVENT INVOLVED 2 TL PICC IN THE UPPER ARM. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03473 | 2023-05-16 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT GOT CATHETERIZATION AFTER THE OPERATION AND WHILE CHECKING THE CATHETER, IT WAS FOUND THAT THE CATHETER WAS BLOCKED AND A NEW CATHETER WAS REPLACED IMMEDIATELY BEFORE INDWELLING CATHETERIZATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03474 | 2023-05-16 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT GOT CATHETERIZATION AND THE CATHETER BALLOON WAS FOUND BURST WHEN MONITORED AND THEN REPLACED WITH A NEW ONE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03471 | 2023-05-16 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT 2 DIFFERENT FOLEYS BECAME DISLODGED AT ST. VINCENT. THEY WERE BOTH IN THE SAME DEPARTMENT AND NEITHER OF THE FOLEYS WERE SAVED, NOR WAS THE PACKAGING. THEY BELIEVE THEM WERE A303416A, BUT HAD TROUBLE CONFIRMING. ALSO STATED THAT THE BALLOONS APPEARED TO NOT HAVE 10CC OF FLUID, AFTER THEY WERE FOUND DISLODGED IN THE PATIENT BEDS. THE BALLOONS WERE NOT FLAT, LEAKING OR RIPPED AND SUSPECTED THAT THEY WERE PERHAPS NOT FULLY FILLED WITH SALINE AT INSERTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03476 | 2023-05-16 | 2022-11-24 | I, n, j, u, r, y | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. THE PRODUCT REFERENCE WERE NOT KNOWN ALTHOUGH THE MAJORITY WERE BELIEVED TO BE EITHER BARDIA PTFE CATHETERS OR BARD LUBRI-SIL CATHETERS. THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01922 | 2023-05-16 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TPS | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED BY THE CUSTOMER "INNER CATHETER STYLET ATTACHED TO THE CATHETER WITHOUT POSSIBILITY OF REMOVAL." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01924 | 2023-05-16 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER " THEY USE IT AND FLUSHED SALINE AND THE CATHETER WAS LEAKING. IT WAS HOLE WITHIN THE HUB. THE CATHETER WAS PULLED OUT AND THEY NEEDED TO PLACE A NEW ONE." ADDITIONAL INFORMATION RECEIVED 4/28/2023: IT WAS REPORTED BY THE CUSTOMER "PICC HAD BEEN USED WITH FOLFIRINOX (IRINOTECAN, FLV OCH OXALIPLATIN) INSTILLERAT ALTEPLASE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER LEAK IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE 4 FR SL POWERPICC SOLO CATHETER WAS RETURNED FOR EVALUATION. THE CATHETER EXTENDED UP TO THE 29 CM DEPTH MARKER. THE DISTAL SEGMENT OF THE CATHETER WAS ALSO RETURNED AND EXTENDED FROM THE 30 CM DEPTH MARKER UP TO THE 46 CM DEPTH MARKER. DRIED BLOOD RESIDUES AND EVIDENCE OF USE WAS NOTED ON THE DEVICE. THE CATHETER WAS INFUSED WITH WATER AND A LEAK WAS NOTED AT THE PROXIMAL END OF THE CATHETER AT THE INTERFACE WITH THE MOLDED WINGED HUB. MICROSCOPIC INSPECTION OF THE BREAK LOCATION REVEALED A PARTIAL CIRCUMFERENTIAL SPLIT. THE BREAK SURFACE WAS FOUND TO BE ROUGH AND GRANULAR. THE CATHETER WAS CUT IN ORDER TO MEASURE THE LUMEN DIAMETER AND WALL THICKNESS. THE MEASUREMENTS WERE FOUND TO BE WITHIN MANUFACTURING SPECIFICATION. THE DURATION OF USAGE OF THE CATHETER IS UNKNOWN. BASED ON THE CHARACTERISTICS OF THE DAMAGE OBSERVED, POSSIBLE CONTRIBUTING FACTORS INCLUDE REPEATED KINKING OF THE CATHETER LEADING TO MATERIAL FATIGUE, DAMAGE DURING HANDLING, AND SECUREMENT TECHNIQUE. SINCE A PARTIAL BREAK IN THE CATHETER TUBING WAS FOUND TO BE THE SOURCE OF THE LEAK, THE COMPLAINT IS CONFIRMED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01925 | 2023-05-16 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT,4FR POWERMIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P4154108 |  | IT WAS REPORTED BY THE CUSTOMER "AFTER PLACING A BARD TRADITIONAL POWERMIDLINE CATHETER AND UPON REMOVING THE STYLET/WIRE, I NOTICED THE WIRE SEEMED VERY LONG AND FRAGILE/FLIMSY IN APPEARANCE. I PUT THE EQUIPMENT ASIDE AND COMPLETED THE MIDLINE INSERTION BY FLUSHING AND CHECKING BLOOD RETURN. AFTER SUCCESSFULLY COMPLETING MID INSERTION, I EXAMINED THE WIRE AND IT APPEARED FRAYED AND FRACTURED.". | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01926 | 2023-05-16 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THAT THE PICC LINE FRACTURED. PICC FLUSHING BUT NO VENOUS RETURN FROM PICC, PATIENT FEELING A "FUNNY SENSATION" IN CHEST AREA WHEN DEVICE FLUSHED. PICC REMOVED AS PER ADVICE FROM INFUSION SERVICES. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03468 | 2023-05-16 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0102L12 | 0102L12 |  | IT WAS REPORTED THAT THE FOLEY CATHETER HAD CRACKS AT Y CONNECTION. THEN THE BALLOON DEFLATES, AND THE CATHETER FALLS OUT. THIS HAD HAPPENED TO MY CLIENT 4 TIMES IN THE LAST FEW MONTHS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01930 | 2023-05-16 | 2023-05-02 | I, n, j, u, r, y | Manufacturer report | BARD TL PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. THIS EVENT INVOLVED 2 TL PICC IN THE UPPER ARM. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03437 | 2023-05-16 | 2022-11-17 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD URETHRAL DAMAGE 1.5CM, LONG TERM CATHETER ONLY . THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
16938409 | 2023-05-16 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERHICKMAN POLYURETHANE CATHETER WITH SURECUFF INGROWTH CUFF | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 0805810 | 0805810 |  | WHILE REMOVING THE DEVICE, THE CUFF BROKE OFF AND REMAINED IN PATIENT REQUIRED SURGICAL REMOVAL. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03446 | 2023-05-16 | 2022-12-06 | I, n, j, u, r, y | Manufacturer report | ALL SILICONE CATHETER | ALL SILICONE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD UNCATEGORIZED URETHRAL DAMAGE ONLY . THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03445 | 2023-05-16 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 165812 |  | IT WAS REPORTED THAT THE STAFF HAD EXPERIENCED DEFECTS ON THE FOLEY CATHETER AFTER INSERTION EPISODES TO THE SAME PATIENT. STATED THAT THERE WAS LEAKING FROM THE BALLOON. ALSO THE USER WAS UNABLE TO INFLATE THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03441 | 2023-05-16 | 2023-01-18 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165SI12 |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD UNCATEGORISED URETHRAL DAMAGE ONLY (PATIENT #17). THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03456 | 2023-05-16 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE PATIENT GOT CATHETERIZATION AND ON THE EARLY MORNING OF (B)(6), THE PATIENT DID NOT URINATE AND FOUND THAT THE CATHETER FELL OUT. WITH CHECK, THE BALLOON WAS BURST WITHOUT MISSING PIECES. REPORTED TO THE DOCTOR TO KEEP CLOSE OBSERVATION AND NOT REPLACE A NEW CATHETER. AT 8 AM, THE PATIENT WAS ABLE TO URINATE BY HIMSELF. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03436 | 2023-05-16 | 2022-11-17 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD URETHRAL DAMAGE AT LEAST 1CM, LONGTERM CATHETER . THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01912 | 2023-05-16 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808000 | 1808000 |  | IT WAS REPORTED THAT THREE YEARS THREE MONTHS AND TWENTY DAYS POST PORT PLACEMENT, THE DEVICE WAS ALLEGEDLY DIFFICULT TO FLUSH. IT WAS FURTHER REPORTED THAT, THE DEVICE WAS NOT FLUSHED REGULARLY. REPORTEDLY, THE DEVICE WAS STILL IMPLANTED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 11/2019). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01916 | 2023-05-16 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808060 | 1808060 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, CATHETER WAS ALLEGEDLY UNABLE TO BE REMOVED. IT WAS FURTHER REPORTED THAT THE CATHETER WAS SUTURED TO THE VESSEL AND LEFT IT IN THE PATIENT BODY DUE TO THE RESISTANCE. REPORTEDLY, THE PORT WAS NOT PURPLE WHEN THE PHYSICIAN REMOVED IT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01917 | 2023-05-16 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1678300 | 1678300 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, AIR WAS ALLEGEDLY ASPIRATED WHEN THE INTRODUCER NEEDLE WAS PUNCTURED INTO THE VEIN. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 04/2024). H11: SECTION A THROUGH F THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01921 | 2023-05-16 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1678300 | 1678300 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, AIR WAS ALLEGEDLY ASPIRATED WHEN THE INTRODUCER NEEDLE WAS PUNCTURED INTO THE VEIN. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 04/2024) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03439 | 2023-05-16 | 2022-10-29 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 20 SEPARATE DATIX INCIDENT REPORTS OF URETHRAL INJURY FROM BARD FOLEY CATHETERS HAD BEEN RECEIVED. THE INJURIES RANGE FROM UNCLASSIFIED TO 1.5 CM EROSION. IT WAS STATED THAT THIS PATIENT HAD UNCATEGORIZED URETHRAL DAMAGE ONLY. THE NUMBER OF PATIENTS INVOLVED WERE 20. IT WAS STATED THAT THE TISSUE VIABILITY TEAM WERE CONTACTED AND ALL PATIENTS WERE TREATED APPROPRIATELY FOR THEIR INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03440 | 2023-05-16 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT THESE FOLEY CATHETERS WERE LEAKED AND THEY BACKED UP. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01913 | 2023-05-16 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0605320 | 0605320 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, CATHETER WAS ALLEGEDLY UNABLE TO BE REMOVED. IT WAS FURTHER REPORTED THAT THE CATHETER WAS SUTURED TO THE VESSEL AND LEFT IT IN THE PATIENT BODY DUE TO THE RESISTANCE. REPORTEDLY, THE PORT WAS NOT PURPLE WHEN THE PHYSICIAN REMOVED IT. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03452 | 2023-05-16 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ALL SILICONE CATHETER | ALL SILICONE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE URINARY CATHETER BALLOON LEAKED/COLLAPSED WHILE ATTEMPTING TO INSERT THE FOLEY URINARY CATHETER. WHILE PLACING A NEW FOLEY CATHETER (18 FR NON-LATEX 30 ML BALLOON), THE NURSE USED 10 ML OF STERILE SALINE TO START INFLATING THE BALLOON. THERE WAS A MILD REPORT OF DISCOMFORT AND 5 ML MORE STERILE SALINE WAS INSTILLED WHEN PULLING BACK, BUT NO SALINE RETURNED. FOLEY REMOVED AND NOTED THE BALLOON WAS NOT INFLATED AND THAT THERE WERE S-SHAPED FLAPS WHERE THE BALLOON WAS SUPPOSED TO BE. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03453 | 2023-05-16 | 2023-04-21 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT AFTER INSERTION ON 20TH AND ATTEMPTED REMOVAL THE NEXT DAY ON 21ST, 10 ML OF STERILE WATER WAS WITHDRAWN, BUT THE TIP OF THE CATHETER WAS STUCK AND COULD NOT BE REMOVED. IT WAS LUBRICATED WITH GEL AND REMOVED WITH LIGHT PRESSURE. ON CHECKING THE CATHETER, THE LOWER PART WAS FOUND BENT. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, STATED THAT THE BALLOON WAS MUSHROOMED AT THE BOTTOM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01932 | 2023-05-17 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT FIVE YEARS THREE MONTHS AND THIRTEEN DAYS POST PORT PLACEMENT VIA THE RIGHT VEIN, WHEN THE PORT SYSTEM WAS ATTEMPTED TO BE REMOVED, THE CATHETER WAS ALLEGEDLY FOUND TO BE DAMAGED NEAR THE VESSEL INSERTION SITE. IT WAS FURTHER REPORTED THAT THE SALINE SOLUTION ALLEGEDLY LEAKED NEAR THE INSERTION SITE. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03538 | 2023-05-17 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998804 | 998804 |  | IT WAS REPORTED THAT THE OPERATOR FOUND THAT THE WORKING PRESSURE COULD NOT BE REACHED AND THAT THE PUMP LEAKED WATER WHEN THE USER TRIED TO INFLATE THE BALLOON USING A PUMP. THE BALLOON DILATION CATHETER WAS THEN REMOVED AND REPLACED WITH A NEW ONE TO CONTINUE THE PROCEDURE, CAUSING NO ADVERSE IMPACT ON THE PATIENT. | PER INVESTIGATIONAL FINDINGS, BD HAS DETERMINED THIS MDR AS NOT REPORTABLE AS THE REPORTED EVENT WAS CONFIRMED AS USER RELATED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
16945313 | 2023-05-17 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDEX LUBRI-SIL ALL-SILICONE FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | 175814 | 175814 |  | FOLEY DISLODGED AFTER INSERTION IN PT. CATHETER TIP SEEMED TO BE DEFECTIVE. WHEN TRIED TO INFLATE THE BALLOON, THE SALINE WAS COMING OUT FROM THE TIP OF THE CATHETER AND THE BALLOON DOES NOT INFLATE MUCH. |  | N | Y |  |  | 2023-05
2020394-2023-00342 | 2023-05-17 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8088 | VA8088 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY COULD NOT BE REMOVED BACK THROUGH SHEATH AFTER INFLATION. IT WAS FURTHER REPORTED THAT THE BALLOON BECAME ALLEGEDLY STUCK IN SHEATH. REPORTEDLY, SHEATH HAD TO BE REMOVED FROM PATIENT IN ORDER TO GET BALLOON OUT OF PATIENT ARM. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 11/2025). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01940 | 2023-05-17 | 2023-01-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PER-Q-CATH PLUS PICC 4F SINGLE-LUMEN PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4134105 |  | IT WAS REPORTED THAT UPON OPENING OF THE MIDLINE¿S PACKAGE, THE PROXIMAL END OF THE GUIDE WIRE INSIDE WAS FOUND BENT. IT WAS NOT SURE IF THE CATHETER WAS DAMAGED OR NOT. THE PRODUCT WAS THEREFORE UNUSABLE. (B)(6) 2023 - THE CUSTOMER REPORTED A BENT GUIDEWIRE HOWEVER A STYLET WAS RETURNED FOR EVALUATION AND EXHIBITED MULTIPLE BENDS. | BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BENT STYLET WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR S/L PER-Q-CATH CATHETER. THE SAMPLE WAS RECEIVED WITH AN INLAID STYLET AND ATTACHED T-LOCK ASSEMBLY. THE STYLET EXHIBITED MULTIPLE BENDS JUST PROXIMAL OF THE T-LOCK ASSEMBLY. ANOTHER BEND WAS OBSERVED BETWEEN THE 58CM AND 59CM DEPTH MARKINGS. BLOOD RESIDUE WAS OBSERVED ON THE OUTSIDE OF THE STYLET. MICROSCOPIC INSPECTION OF THE SAMPLE CONFIRMED BENDS IN THE STYLET AND THE PRESENCE OF BLOOD RESIDUE BUT WAS OTHERWISE UNREMARKABLE. THE DEFORMATION WAS CONSISTENT WITH STYLET MANIPULATION, INCLUDING POSSIBLE ATTEMPTS TO FEED THE STYLET THROUGH THE T-LOCK ASSEMBLY SEPTUM. THE BLOOD RESIDUE SUGGESTED THAT MANIPULATION OCCURRED DURING ATTEMPTED DEVICE USE. EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01938 | 2023-05-17 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER "WHEN THE SPECIALIST CATHETERIZATION NURSE PERFORMED PICC PUNCTURE AND CATHETERIZATION FOR THE PATIENT, IT WAS FOUND THAT THE GUIDE WIRE COULD NOT NORMALLY GUIDE THE CATHETER INTO THE BLOOD VESSEL. AFTER CHECKING THE GUIDE WIRE, IT WAS FOUND THAT THE FRONT END OF THE GUIDE WIRE WAS MALFORMED AND HYPERTROPHY, WHICH COULD NOT BE USED NORMALLY." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01936 | 2023-05-17 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 20CM PRE-CURVED SHORT-TERM DL | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5594200 |  | IT WAS REPORTED THAT, "WHEN INSERTING THE CATHETER AND REMOVING THE METALLIC GUIDE FROM THE CATHETER, IT GOT STUCK AND IT WAS HARD TO REMOVE THE CATHETER. CASE PRESENTED IN THE INTENSIVE CARE UNIT OF THE HOSPITAL." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01937 | 2023-05-17 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 20CM PRE-CURVED SHORT-TERM DL | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5594200 |  | IT WAS REPORTED THAT, "WHEN INSERTING THE CATHETER AND REMOVING THE METALLIC GUIDE FROM THE CATHETER, IT GOT STUCK AND IT WAS HARD TO REMOVE THE CATHETER." NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03535 | 2023-05-17 | 2023-04-25 | I, n, j, u, r, y | Manufacturer report | CLEAR SILICONE SELF-ADHESIVE MALE CONTINENCE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 786195 | 786195 |  | IT WAS REPORTED THAT THE PATIENT CALLED TO ADVISE PATIENT HAD TO BE ADMITTED TO A&E AS THE ADHESIVE ON THE FIRST MALE EXTERNAL CATHETER SHEATH OF A BRAND-NEW BOX WAS NOT COMING OFF, THE SURGEON HAD TO CUT IT OFF OF PATIENT. DUE TO THIS PATIENT HAD NOT USED ANY OF THE OTHER FROM THE BOX AS PATIENT DOES NOT WANT TO REPEAT THE EXPERIENCE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "VISCOMETER FAILURE". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE SELF-ADHERING MALE EXTERNAL. CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS. APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: 1) VERIFY CORRECT SIZE PRIOR TO USE. 2) TRIM PUBIC HAIR IF NECESSARY. 3) WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. 4) OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. 5) PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. 6) UNROLL THE CATHETER OVER PENIS. 7) GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. 8) CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: 1) ENSURE DRAINAGE BAG IS EMPTY. 2) DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. 3) GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03534 | 2023-05-17 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE CATHETER WAS UNABLE TO REMOVE. AFTER INSERTION, THERE WAS A LEAK OF STERILE WATER FROM THE SENSOR AREA. THINKING SOMETHING WAS WRONG, THE USER WITHDREW THE STERILE WATER FOR REINSERTING THE CATHETER, BUT ONLY 2 CC OF WATER WAS REMOVED. CHECKING IT, THE BALLOON WAS FOUND TO HAVE BEEN TRAPPED IN THE URETHRA. THE FUNNEL WAS CUT BUT COULD NOT BE DRAINED. THE PATIENT WAS TAKEN TO THE RADIOLOGY DEPARTMENT AND A GUIDE WIRE WAS PASSED THROUGH THE INFLATION LUMEN WHILE CHECKING WITH AN ECHO, AND THE WATER WAS REMOVED. ALLEGEDLY, AT THE TIME OF INSERTION, THE CATHETER WAS INSERTED UP TO THE ROOT AND THE BALLOON WAS INFLATED AFTER CONFIRMING THE DRAINAGE OF URINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03531 | 2023-05-17 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM NEAR THE INFLATION VALVE. WHEN THE USER WAS TRYING TO INJECT STERILE DISTILLED WATER AFTER INSERTING FOLEY CATHETER THROUGH THE URETHRA, STERILE DISTILLED WATER LEAKED FROM THE VICINITY OF THE VALVE WHERE THE SYRINGE WAS INSERTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03527 | 2023-05-17 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT STERILE WATER DID NOT ENTER THE CATHETER. SINCE RESISTANCE WAS STRONG, THE DEVICE WAS DISCONTINUED TO USE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO MANUFACTURING RELATED DUE TO COLLAPSED / PINCHED INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE 1. MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTIONS: E, F, H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03530 | 2023-05-17 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS DEFLATED AND AS SUCH IT WAS REMOVED FROM THE PATIENT. WHEN THE USER TRIED TO BLOW BALLOON UP, IT WOULD NOT INFLATE. WHEN THEY PUT WATER DOWN THE URINARY PORT, THE BALLOON INFLATED WHICH WAS NOT SUPPOSED TO HAPPEN. AS PER INVESTIGATOR NOTIFICATION RECEIVED ON 02JUN2023, STATED THAT THERE WAS TEAR AT INFLATION EYE OBSERVED IN THE EVALUATION WHICH CAUSED WATER UNABLE TO RETAIN IN BALLOON. THIS CAUSED DEFLATION DUE TO TRAUMA ISSUE. USER COULD NOT INFLATE BACK THE BALLOON SINCE WATER WAS UNABLE TO RETAIN IN CATHETER BALLOON. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND TO BE A CASCADING FAILURE OF AN EVENT PREVIOUSLY REPORTED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03521 | 2023-05-17 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE BALLOON OF FOLEY CATHETER WAS FOUND TO BE BURST DURING THE PROCESS OF DRAINING THE URINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03536 | 2023-05-17 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS DEFLATED AND AS SUCH IT WAS REMOVED FROM THE PATIENT. WHEN THE USER TRIED TO BLOW BALLOON UP, IT WOULD NOT INFLATE. WHEN THEY PUT WATER DOWN THE URINARY PORT, THE BALLOON INFLATED WHICH WAS NOT SUPPOSED TO HAPPEN. AS PER INVESTIGATOR NOTIFICATION RECEIVED ON 02JUN2023, STATED THAT THERE WAS TEAR AT INFLATION EYE OBSERVED IN THE EVALUATION WHICH CAUSED WATER UNABLE TO RETAIN IN BALLOON. THIS CAUSED DEFLATION DUE TO TRAUMA ISSUE. USER COULD NOT INFLATE BACK THE BALLOON SINCE WATER WAS UNABLE TO RETAIN IN CATHETER BALLOON. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE (EXAMPLE: USER RELATED/OPERATOR ERROR/ROUGH HANDLING OR EXPOSE TO SHARP OBJECT/MECHANICAL FORCE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03514 | 2023-05-17 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE DRAINAGE EYE WAS NOT OPEN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03487 | 2023-05-17 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT AFTER PLACING THE CATHETER IN A PATIENT, THE NURSE NOTICED THAT THE BALLOON HAD RUPTURED. AS PER THE ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023, IT WAS STATED THAT 0N (B)(6) 2023, A 14FR CATHETER WAS INSERTED FOR DISCHARGE. USED A DIB CAP AT HOME. ON (B)(6) 2023, THE PATIENT VISITED THE EMERGENCY DEPARTMENT AS THE BALLOON SPONTANEOUSLY FELL OUT. AT AROUND 20:00, THE PATIENT FELT A SHOCK WITH AN UNEXPLAINED POPPING SOUND IN THE LOWER ABDOMEN, AFTER WHICH THE PATIENT NOTICED THAT THE CATHETER HAD WITHDRAWN SPONTANEOUSLY. THE PATIENT WAS IN THE KOTATSU TABLE FOR A LONG TIME IN A SITTING POSITION WHEN THE PATIENT HAD A BREAK. THE USER DECIDED TO CHECK THE NEXT DAY, AS IF THERE WAS ANY DEFECT OR REMNANTS OF THE BROKEN PART, IT WOULD HAVE CAUSED INFLAMMATION. ON (B)(6) 2023, THE NURSE AND DOCTOR ON DUTY CHECKED IF THERE WAS ANY DAMAGE TO THE DAMAGED PART. NO OBVIOUS DEFECT WAS OBSERVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03520 | 2023-05-17 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123620A |  | IT WAS REPORTED THAT THE SALINE IN THE BALLOON LEAKED OUT AND THE FOLEY CATHETER FELL OUT DURING THE PROCESS OF BLADDER IRRITATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03488 | 2023-05-17 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER DID NOT HAVE ENOUGH LUBRICANT. IT WAS STATED THAT THE BURST PACKET FOR THE INTERMITTENT CATHETER WAS NOT WORKING FOR THE PATIENT. PER ADDITIONAL INFORMATION RECEIVED FROM LIBERATOR ON (B)(6) 2023,THE BURST PACKET FOR THE INTERMITTENT CATHETER WAS NOT WORKING FOR THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03491 | 2023-05-17 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THERE WAS NO LIQUID IN THE BAG TO LUBRICATE THE INTERMITTENT CATHETER. THE PATIENT STATED THIS HAD HAPPENED FOUR TIMES AND HAD CAUSED A LOT OF PAIN. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE. INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01933 | 2023-05-17 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0607540 | 0607540 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY DAMAGED. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 08/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03502 | 2023-05-17 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 165812 |  | IT WAS REPORTED THAT THE STAFF HAD EXPERIENCED DEFECTS ON THE FOLEY CATHETER AFTER INSERTION EPISODES TO THE SAME PATIENT. STATED THAT THERE WAS LEAKING FROM THE BALLOON. ALSO THE USER WAS UNABLE TO INFLATE THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03501 | 2023-05-17 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT ON THE SECOND DAY OF INDWELLING, THE CATHETER CAME OUT AND UPON CHECKING THE BALLOON, IT WAS FOUND RUPTURED. ACCORDING TO THE INQUIRY SHEET ATTACHED, NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER AND THE CATHETER WAS NOT PULLED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03503 | 2023-05-17 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE PATIENTS COMPLAINED OF URINATION WHEN AWAKENED AFTER THE SURGERY, EVEN THOUGH URINE WAS FLOWING THROUGH THE CATHETER. THE DOCTOR STATED THAT IT WAS LITTLE THICKER THAN A NORMAL BALLOON BECAUSE IT CONTAINED A TEMPERATURE SENSOR, SO IT COULD HAVE STIMULATED URINATION. THERE WAS NO DAMAGE DIRECTLY CAUSED TO THE PATIENT BY THE BALLOON. THERE HAVE BEEN CASES OF PEOPLE GETTING UP AND FALLING OFF THE BED TO URINATE DESPITE BEING UNDER ANESTHESIA. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03509 | 2023-05-17 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 123516A | 123516A |  | IT WAS REPORTED, THAT THE FOLEY CATHETER FELL OUT DURING USE. DUE TO A HOLE IN THE BALLOON. WHICH CAUSED DISCOMFORT, BUT NO SERIOUS INJURIES. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03507 | 2023-05-17 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175814 |  | IT WAS REPORTED THE PATIENT GOT WET ON THE GROIN AREA WHILE USING THE FOLEY CATHETER. THE TUBE WAS NOT SITTING PROPERLY INTERNALLY AND IT SEEMED LIKE SEEMS LIKE THERE WAS AN ISSUE WITH THE TUBE THAT GOES INTO THE PENIS. THE SIZING AND LENGTH SEEMED FINE. THEY WERE NOT SURE WHAT THE PROBLEM WAS, BUT THIS IS THE AREA WHERE THE LEAKAGE OCCURRED. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "BAD FIT WITH SHAFT ." IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE SILICONE. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LAWS AND REGULATIONS. UNITS: THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. DO NOT USE IF PACKAGE IS DAMAGED. BARD, BARDEX, BARDIA, BIOCATH AND LUBRICATH ARE TRADEMARKS AND/OR REGISTERED TRADEMARKS OF C. R. BARD, INC. SINGLE USE ONLY. DO NOT RE-STERILIZE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF Y.OU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. FEMALE USE ONLY SYRINGE OF STERILE WATER FOR BALLOON INFLATION. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03512 | 2023-05-17 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L18 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT IN THE WARD NEXT DAY AFTER SURGERY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01939 | 2023-05-18 | 2023-01-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT IT WAS DIFFICULT TO INSERT THE PROXIMAL END OF THE INTRODUCER SHEATH IN SELDINGER INTO THE VESSEL. CRACKS WERE FOUND WITH THE GREY SLEEVE, AFFECTING ITS USE. VIGOROUS FORCES WERE NOT APPLIED IN THE PROCESS OF INSERTION. INSTEAD, GENTLE FORCES WERE APPLIED AND CEASED IMMEDIATELY ONCE RESISTANCE WAS MET. THE DEVICE WAS REMOVED. NO OTHER INFORMATION WAS PROVIDED. 05/12/2023 - TWO INTRODUCERS WERE PROVIDED FOR EVALUATION. BOTH EXHIBITED DEFORMATION AND A SPLIT AT THE DISTAL TIP. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A SPLIT MICROINTRODUCER IS CONFIRMED AND WAS DETERMINED TO BE MANUFACTURING RELATED. ONE 4.5 FR X 5 CM MICROINTRODUCER AS WELL AS THREE PHOTOGRAPHS WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLE. THE DISTAL END OF THE INTRODUCER SHEATH WAS OBSERVED TO BE DAMAGED. A MICROSCOPIC OBSERVATION REVEALED PLASTIC DEFORMATION OF THE SHEATH MATERIAL AROUND THE ENTIRE CIRCUMFERENCE OF THE DISTAL TIP OF THE INTRODUCER SHEATH. A RELATIVELY LARGE, EVEN, AND STRAIGHT SPILT WAS ALSO OBSERVED IN THE DISTAL TIP OF THE INTRODUCER SHEATH. NO DEFORMATION WAS OBSERVED ON ONE SIDE OF THIS SPLIT. THE RETURNED PHOTOGRAPHS APPEAR TO SHOW ANOTHER SAMPLE WITH SIMILAR DAMAGE. THE SPLIT OBSERVED IN THE DISTAL TIP OF THE INTRODUCER SHEATH WITHIN THE RETURNED PHOTOGRAPHS WAS ALSO OBSERVED TO BE STRAIGHT, EVEN, AND RELATIVELY LARGE WITH LITTLE TO NO DAMAGE ADJACENT TO THE SPLIT. THE INVESTIGATION WAS FORWARDED TO THE MANUFACTURING FACILITY FOR FURTHER EVALUATION, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03590 | 2023-05-18 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER HAVING ISSUE FROM THE ORDER. THEY STATES THAT THE 7 OF THE CATHETERS WERE BENT IN THE BOX AND NOT LUBRICATED. RMA FOR 7 CATHETERS. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT DOES NOT APPEAR TO HAVE BEEN USED. HOWEVER, THE PRODUCT IS INTENDED TO BE USED FOR TREATMENT PURPOSE BUT IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE.THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01945 | 2023-05-18 | 2023-04-27 | I, n, j, u, r, y | Manufacturer report | BARD DUAL LUMEN PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. TWO EVENTS OCCURRED WITH PICCS INSERTED IN THE UPPER ARM. IT WAS REPORTED THIS OCCURRED WITH A SINGLE LUMEN PICC AND A DUAL LUMEN PICC. THIS REPORT ADDRESSES THE DUAL LUMEN PICC. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01947 | 2023-05-18 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD TRIPLE LUMEN PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. THE EVENT INVOLVED 2 PICCS, ONE SL INSERTED IN THE UPPER ARM AND ONE TL INSERTED MID-THIGH. THIS REPORT ADDRESSES THE TRIPLE LUMEN PICC. | SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03579 | 2023-05-18 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE BALLOON WAS BURST AND THE CATHETER FELL OUT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE LOW LATEX VISCOSITY, SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY'. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03591 | 2023-05-18 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50612 | 50612 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE TOO MESSY TO USED. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER. CAUTION: FEDERAL (USA) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01943 | 2023-05-18 | 2023-04-18 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6CF INT W SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806060 |  | IT WAS REPORTED THAT TEN MONTHS AND THIRTEEN DAYS POST PORT PLACEMENT, THE PATIENT FELT CHEST DISTRESS. IT WAS FURTHER REPORTED THAT UPON CHEST X-RAY THE CATHETER ALLEGEDLY FOUND TO BE RUPTURED AND MIGRATED INTO THE ATRIUM. IT WAS ALSO REPORTED THAT THE DETACHED CATHETER WAS REMOVED WITH THE PORT IN AN INTERVENTIONAL MANNER. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, ONE ELECTRONIC PHOTOS WAS PROVIDED FOR REVIEW. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND MATERIAL SEPARATION ISSUES AS THE PHOTO SHOWS ONE CATHETER SEGMENT. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MIGRATION ISSUE AS THERE WAS NO OBJECTIVE EVIDENCE PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 08/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03587 | 2023-05-18 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE WATER INSIDE THE FOLEY CATHETER BALLOON WAS UNABLE TO REMOVED BY THE SYRINGE UPON CHECKING THE BALLOON BEFORE USE ON THE PATIENT. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01941 | 2023-05-18 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616070 | 1616070 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE GUIDEWIRE WAS ALLEGEDLY UNABLE TO BE INSERTED THROUGH THE SHEATH. IT WAS FURTHER REPORTED THAT THE TIP OF THE SHEATH WAS ALLEGEDLY FOUND TO BE DAMAGED AND THE SHAPE OF THE SHEATH TIP SEEMED TO BE ABNORMAL. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE 6.5FR PEEL-APART SHEATH AND VESSEL DILATOR WERE RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC AND FUNCTIONAL EVALUATIONS WERE PERFORMED. THE DISTAL TIP OF THE VESSEL DILATOR WAS NOTED TO BE DEFORMED. BUNCHING WAS NOTED ON THE DISTAL END OF THE PEEL-APART SHEATH. THEREFORE, THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED MATERIAL INTEGRITY ISSUE AS THE SPECIFIC DAMAGE SUCH AS DEFORMATION DUE TO COMPRESSIVE STRESS ISSUE WAS IDENTIFIED DURING THE INVESTIGATION. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO INSERT ISSUE AS THE EXACT CIRCUMSTANCE AT THE TIME OF THE EVENT REPORTED WAS UNKNOWN. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 12/2023), G3. H11: H6 (DEVICE, METHOD, RESULT, CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01942 | 2023-05-18 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN AND BULGING AREA ON HER NECK WHEN SHE TURNS HER HEAD TO THE LEFT. IT WAS FURTHER REPORTED THAT THE PATIENT ALLEGEDLY EXPERIENCED ITCHING AT THE PORT BODY SITE. REPORTEDLY, AFTER DOCTOR EVALUATIONS WITH X-RAY THEY HAVE CONCLUDED THAT THE PORT IS IN A GOOD POSITION. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING ADDITIONAL INTERVENTION OR MEDICATION TO TREAT PAIN, BULGING AREA ON NECK AND ITCHING AT THE PORT BODY SITE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 05/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01954 | 2023-05-18 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR (RADEL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770352 |  | IT WAS REPORTED THE TRACKING IS NOT WORKING CORRECTLY. THE USER TRIED RECALIBRATING MULTIPLE TIMES STILL THE TRACKING IS OFF. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03585 | 2023-05-18 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE FOLEY CATHETER BALLOON DURING INJECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | OTHER |  | 2023-05
3006260740-2023-01944 | 2023-05-18 | 2023-04-27 | I, n, j, u, r, y | Manufacturer report | BARD SINGLE LUMEN PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. TWO EVENTS OCCURRED WITH PICCS INSERTED IN THE UPPER ARM. IT WAS REPORTED THIS OCCURRED WITH A SINGLE LUMEN PICC AND A DUAL LUMEN PICC. THIS REPORT ADDRESSES THE SINGLE LUMEN PICC. | SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01951 | 2023-05-18 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770201 |  | IT WAS REPORTED THE CUSTOMER TRIED TO PERFORM SOME ASSISTED PUNCTURES AND REALIZED THAT THE ISSUE OF DEPTH IS ALSO NOT RELIABLE (IN THE ULTRASOUND IT SHOWED A VEIN WITH A DEPTH OF 2 CM, BUT A 6 CM NEEDLE WAS FULLY INSERTED AND THE IMAGE SHOWED THAT IT WAS JUST PINCHING THE VEIN). 05/10/23 ADDITIONAL DETAILS: AFTER 3 DAYS, THE PATIENT PRESENTED ECCHYMOSIS, BUT DURING THE ATTEMPT TO PUNCTURE THE NEEDLE THAT WAS BEING USED, IT ONLY SHOWED A MICRO DROP OF BLOOD, IT ONLY PINCHED THE WALL OF THE VEIN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03570 | 2023-05-18 | 2023-05-01 | I, n, j, u, r, y | Manufacturer report | UROFORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 887406 | 887406 |  | IT WAS REPORTED VIA SURVEY RESPONSE THAT A MALE PEDIATRIC PATIENT (13 YEARS OLD) EXPERIENCED URINARY TRACT INFECTION AFTER PLACEMENT OF THE DEVICE, THIS COMPLICATION WAS DIRECTLY ATTRIBUTABLE TO USAGE OF BALLOON DILATION CATHETER WITH EAGLE INFLATION DEVICE. URINARY TRACT INFECTION WAS PRESENT PRIOR TO PLACEMENT OF THE DEVICE, AND DID NOT RESULT IN PATIENT INJURY REQUIRING MEDICAL OR SURGICAL INTERVENTION. MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE "COATING/TISSUE INCOMPATIBILITY". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DEVICE DESCRIPTION: THE UROFORCETM BALLOON DILATION CATHETER IS A DUAL-LUMEN CATHETER WITH A NYDEX* BALLOON MOUNTED ON ITS DISTAL TIP AND WITH TWO RADIOPAQUE MARKERS BENEATH THE BALLOON THAT DEFINE WORKING LENGTH. THE LUMEN LABELED ¿BALLOON¿ IS FOR BALLOON INFLATION. THE OTHER LUMEN ALLOWS THE CATHETER TO TRACK OVER A 0.038¿ (.97MM) DIAMETER GUIDEWIRE AND CAN BE USED FOR MONITORING OF PRESSURE OR INFUSION OF MEDICATION AND/OR CONTRAST MEDIUM. EACH BALLOON INFLATES TO A STATED DIAMETER AND LENGTH AT A SPECIFIC PRESSURE--TYPICALLY AT 10 ATM. THE BALLOON IS WRAPPED CLOCKWISE (WHEN VIEWED FROM THE PROXIMAL SHAFT TO THE DISTAL TIP) AROUND THE SHAFT AND IS PROTECTED BY A PROFILE REDUCING SHEATH THAT IS POSITIONED OVER THE BALLOON FOR PROTECTION BEFORE USE. A STYLET IS INSERTED IN THE DISTAL (GUIDEWIRE) LUMEN BENEATH THE BALLOON TO PREVENT KINKING DURING SHIPMENT AND STORAGE. INDICATIONS FOR USE: UROFORCETM BALLOON DILATION CATHETERS ARE RECOMMENDED FOR DILATION OF THE URINARY TRACT. CONTRAINDICATIONS: DO NOT USE THE UROFORCETM BALLOON DILATION CATHETER TO TREAT OBSTRUCTION THAT IS DUE TO EXTRINSIC FACTORS. WARNINGS: DO NOT EXCEED THE RECOMMENDED MAXIMUM BALLOON INFLATION PRESSURE THAT IS STATED ON THE LABEL. EXTREME CAUTION SHOULD BE USED WHEN CONSIDERING DILATION OF IRREGULAR, NON-COMPLIANT TISSUE OR IN THE PRESENCE OF CERTAIN CALCULI. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: THIS DEVICE SHOULD BE USED ONLY BY A PHYSICIAN WHO HAS AN UNDERSTANDING OF THE CLINICAL APPLICATIONS, TECHNICAL PRINCIPLES AND ASSOCIATED RISKS OF BALLOON DILATION WITHIN THE URINARY TRACT. DIRECTIONS FOR USE: CATHETER INSERTION 1. PRIOR TO USE, CAREFULLY INSPECT THE CATHETER FOR BENDS, KINKS, OR OTHER DAMAGE WHICH MAY HAVE OCCURRED DURING SHIPMENT. DO NOT USE THE PRODUCT IF DAMAGE IS EVIDENT. 2. DILATION PROCEDURES MAY BE CONDUCTED UNDER FLUOROSCOPIC GUIDANCE WITH APPROPRIATE X-RAY EQUIPMENT OR BY DIRECT VISION. 3. ENDOSCOPICALLY PLACE A 0.038¿ (.97MM) DIAMETER GUIDEWIRE WITH THE FLEXIBLE END IN THE RENAL PELVIS. 4. REMOVE THE PROTECTIVE BALLOON SHEATH AND STYLET FROM THE CATHETER. 5. ADVANCE THE CATHETER OVER THE GUIDEWIRE. 6. USE THE RADIOPAQUE MARKERS LOCATED UNDER THE BALLOON TO AID IN POSITIONING. CATHETER BALLOON INFLATION: 1. CONFIRM PROPER PLACEMENT OF THE BALLOON WITHIN THE URINARY TRACT. 2. INFLATE THE BALLOON WITH STANDARD INFLATION MEDIUM (E.G., DIATRIZOATE MEGLUMINE INJECTION U.S.P. 60% DILUTED AS APPROPRIATE) USING A 10CC OR LARGER SYRINGE OR INFLATION DEVICE. NOTE: DO NOT USE AIR OR ANY GASEOUS SUBSTANCE AS A BALLOON INFLATION MEDIUM. NOTE: THE USE OF AN INFLATION DEVICE TO MONITOR BALLOON PRESSURE IS STRONGLY RECOMMENDED TO DETERMINE THAT ADEQUATE PRESSURE IS APPLIED AND THAT MAXIMUM LIMITS OF THE BALLOON ARE NOT EXCEEDED. CAUTION: IF LOSS OF PRESSURE IN THE BALLOON OCCURS OR THE BALLOON RUPTURES, IMMEDIATELY STOP THE PROCEDURE, DEFLATE THE BALLOON AND REMOVE CAREFULLY. DO NOT ATTEMPT TO REINFLATE THE BALLOON. IF, UPON INSPECTION OF THE BALLOON, IT IS NOTED THAT PIECES OF THE BALLOON ARE MISSING, CHECK FOR AND ENDOSCOPICALLY RETRIEVE THE FRAGMENTS WHEN POSSIBLE. CATHETER BALLOON DEFLATION AND REMOVAL: 1. DEFLATE THE BALLOON USING A 10CC OR LARGER SYRINGE OR INFLATION DEVICE. SINCE DEFLATION TIMES VARY WITH BALLOON SIZES AND SHAFT LENGTHS, CHECK FLUOROSCOPICALLY TO CONFIRM DEFLATION BEFORE ATTEMPTING WITHDRAWAL. 2. REMOVE THE SYRINGE OR INFLATION DEVICE FROM THE CATHETER ALLOWING AMBIENT PRESSURE TO RELAX THE BALLOON. 3. GENTLY WITHDRAW THE CATHETER. AS THE BALLOON STARTS TO EXIT THE ENDOSCOPE USE A SMOOTH, GENTLE, STEADY COUNTERCLOCKWISE MOTION TO FACILITATE WITHDRAWAL. CAUTION: IF RESISTANCE IS FELT WHEN EITHER REMOVING THE GUIDEWIRE FROM THE CATHETER OR THE CATHETER FROM THE ENDOSCOPE, STOP AND CONSIDER REMOVING THEM AS A SINGLE UNIT TO PREVENT DAMAGE TO THE PRODUCT OR TISSUE TRAUMA. APPLYING EXCESSIVE FORCE TO THE CATHETER CAN RESULT IN TIP BREAKAGE OR BALLOON SEPARATION. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03561 | 2023-05-18 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED AT 13:00 ON (B)(6) 2023 AND WAS CONFIRMED URINE OUTFLOW UNTIL AROUND 6:00 NEXT DAY. LATER AROUND 8:00, THE FOLEY CATHETER WAS FOUND TO BE REMOVED AND THERE WAS A SMALL HOLE IN THE BALLOON. ACCORDING TO THE INQUIRY SHEET ATTACHED, THE EVENT OCCURRED AFTER 20 HOURS OF PLACEMENT AND THE WATER LEAKED FROM THE BALLOON. NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER AND THE CATHETER WAS NOT PULLED. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. A POTENTIAL ROOT CAUSE COULD BE LUMEN COMPROMISED BY A PUNCTURE OR CUT. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3-WAY TEMPERATURE SENSING CATHETER. VISUAL INSPECTION NOTED THAT THE TEMPERATURE WIRE WAS CUT UPON RECEIVAL OF SAMPLE. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE SOLUTION IMMEDIATELY WENT INTO THE DRAINAGE LUMEN AT BIFURCATION AND CREATED LUMEN TO LUMEN LEAK. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER¿ CORRECTION- D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03566 | 2023-05-18 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, THE STERILE WATER LEAKED. ACCORDING TO THE INQUIRY SHEET ATTACHED, THERE WAS NO BALLOON RUPTURE/TEAR. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. PATIENT'S URINE STONE WAS NONE. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. THE DEVICE HAD NOT MET SPECIFICATIONS. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND IT SLOWLY LEAKED OUT OF A 0.013" PINHOLE FOUND AT THE TRIFURCATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNINGS: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: METHOD FOR USE: DO NOT REUSE. DO NOT RE-STERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00347 | 2023-05-18 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CQ5064J |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY FAILED TO INFLATE AT THE TIP. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CONQUEST PTA BALLOON DILATATION CATHETER PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CONQUEST PTA BALLOON DILATATION CATHETER PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE CONQUEST PTA DILATATION CATHETER LOADED ALONG WITH AN UNKNOWN INFLATION DEVICE WAS RETURNED FOR EVALUATION. UNRAVELED FIBERS AND FIBER DISTURBANCE WERE ABLE TO BE NOTICED ON THE RETURNED BALLOON. NO OTHER ANOMALIES WERE NOTED DURING THE VISUAL EVALUATION. IN THE FUNCTIONAL TESTING, THE BALLOON WAS ATTEMPTED TO INFLATE WITH THE RETURNED INFLATION DEVICE BUT IT WAS UNSUCCESSFUL. FURTHERMORE, THE BALLOON WAS INFLATED WITH THE IN-HOUSE PRESTO INFLATION DEVICE. THE BALLOON WAS ABLE TO INFLATE TO 8ATM, MAINTAINING SHAPE AND PRESSURE. THEN THE INFLATED BALLOON WAS DEFLATED WITHOUT ANY PROBLEMS. ALL THE ANOMALIES UNDER MICROSCOPIC OBSERVATION. UNRAVELED FIBERS AND FIBER DISTURBANCE WERE ABLE TO BE NOTICED ON THE RETURNED BALLOON DURING THE VISUAL AND MICROSCOPIC OBSERVATION. THEREFORE, THE INVESTIGATION WAS CONFIRMED FOR THE IDENTIFIED UNRAVELED FIBER AND FIBER DISTURBANCE. THEN THE BALLOON WAS UNABLE TO INFLATE WITH THE RETURNED UNKNOWN INFLATION DEVICE. HOWEVER, UNDER A LABORATORY-CONTROLLED ENVIRONMENT, THE RETURNED BALLOON WAS ABLE TO INFLATE FULLY, MAINTAINING PRESSURE AND SHAPE, THEN DEFLATED WITHOUT ANY PROBLEMS. HENCE, THE INVESTIGATION FOR THE REPORTED ABNORMAL INFLATION ISSUE WAS UNCONFIRMED. A DEFINITIVE ROOT CAUSE FOR THE REPORTED ABNORMAL INFLATION ISSUE IDENTIFIED UNRAVELED FIBER COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 01/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00346 | 2023-05-18 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS PTA DILATATION CATHETER | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8094 | VA8094 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON WAS ALLEGEDLY PRESSURIZED ABOVE THE RATED BURST PRESSURE BY THE USER AND IT RUPTURED. IT WAS FURTHER REPORTED THE BALLOON ALLEGEDLY MUSHROOMED WHEN ATTEMPTING TO REMOVE IT FROM THE SHEATH. REPORTEDLY, DUE TO THE MUSHROOM, THE USER WAS FORCED TO REMOVE THE SHEATH AND UPSIZE THE NEW SHEATH TO THE REMOVE THE BALLOON WHICH CAUSED A BIGGER HOLE IN THE VESSEL. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING ADDITIONAL INTERVENTION OR MEDICATION TO TREAT THE HOLE IN THE VESSEL. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED INCORRECT USE OF THE DEVICE, AS THE USER STATED IN THE REPORTED EVENT DETAILS THEY INFLATED THE BALLOON ABOVE RATED BURST PRESSURE AND ARE AWARE OF THEIR INCORRECT USE OF THE DEVICE. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED RUPTURE, DIFFICULTY TO REMOVE, AND DEFORMED BALLOON AS THE DEVICE WAS NOT RETURNED AND NO OBJECTIVE EVIDENCE WAS PROVIDED TO CONFIRM THESE FAILURES. AS PER THE REPORTED EVENT DETAILS, THE USER KNOWINGLY OVER INFLATED THE BALLOON CAUSING THE BALLOON RUPTURE. THEREFORE THE LIKELY CAUSE OF THE REPORTED BALLOON RUPTURE IS THE INCORRECT USE OF THE DEVICE. IT IS ALSO POSSIBLE THE REPORTED RUPTURE MAY HAVE LED TO THE ISSUES WITH THE DEFORMED BALLOON MATERIAL AND DIFFICULTY REMOVING FROM THE SHEATH. HOWEVER THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED AS THE DEVICE AND NO EVIDENCE WERE PROVIDED FOR REVIEW. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. PER THE REPORTED INFORMATION, THE USER INFLATED THE BALLOON ABOVE THE RATED BURST PRESSURE. THIS CONTRADICTS THE INSTRUCTIONS FOR USE IN THE VACCESS INSTRUCTION FOR USE UNDER WARNINGS, STEP 4: "DO NOT EXCEED THE RBP RECOMMENDED FOR THIS DEVICE. BALLOON RUPTURE MAY OCCUR IF THE RBP RATING IS EXCEEDED. TO PREVENT OVER PRESSURIZATION, USE OF A PRESSURE MONITORING DEVICE IS RECOMMENDED." H10: B5, D4 (EXPIRATION DATE: 01/2026), G3, H6 (DEVICE, METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03443 | 2023-05-18 | 2023-04-21 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE BALLOON OF THE TEMPORARY PACING ELECTRODE CATHETER AT THE TIP WAS UNABLE TO BE INFLATED. STATED THAT THE USER WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. 1SAMPLE WERE CONFIRMED TO EXHIBIT THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE PRODUCT CAUSED THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITH ORIGINAL PACKAGING), USED TEMPERATURE PACING ELECTRODE CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED SOME RESIDUE ON THE TIP OF THE CATHETER. THE BALLOON WAS NOTED TO BE BURST WITH NO APPARENT MISSING PIECES, PREVENTING IT FROM BEING INFLATED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE ¿CONTAMINANTS, LATEX DEFECTS OR ABRASION OF BALLOON AFTER ATTACHMENT TO SHAFT, WINDOWS OR INCLUSIONS OVER STRETCHED BALLOON. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "EXCESSIVE BENDING, TORQUING, OR KINKING OF THE ELECTRODE CATHETER MAY CAUSE DAMAGE TO THE CATHETER INCLUDING DAMAGE TO INTERNAL WIRES. INSPECT THE STERILE PACKAGE CAREFULLY FOR DAMAGE DURING TRANSIT OR STORAGE. DO NOT USE THE CATHETER IF THE PACKAGE IS DAMAGED. VISUALLY INSPECT THE CATHETER, UNDER STERILE CONDITIONS, FOR KINKS IN THE CATHETER SHAFT, INTEGRITY OF THE CONNECTOR, CONDITION OF ELECTRODES, AND ANY OTHER DAMAGE." THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03541 | 2023-05-18 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OFF AND THE BALLOON WAS DEFLATED ON (B)(6) 2023 AT 5:00 A.M. AND THE CUFF WATER WAS DRAINED. THERE WERE NO ISSUES AT THE TIME OF PLACEMENT TO THE PATIENT ON (B)(6) 2023 FOR A SURGERY. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE DETAILS REGARDING ANY DAMAGE TO THE BALLOON WASN'T PROVIDED FROM CUSTOMER. | THE REPORTED EVENT IS INCONCLUSIVE DUE TO THE CONDITION OF THE SAMPLE RECEIVED. VISUAL: RECEIVED A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED THAT THE CATHETER WAS CUT OFF AT THE BIFURCATION. THE SHAFT AND BALLOON PART OF THE CATHETER WERE NOT RETURNED. INJECTED 10 ML OF SOLUTION INTO CATHETER AND SOLUTION FLOWED THROUGH THE FUNNEL AND OUT THE BIFURCATION WITH NO LEAKS NOTED. FLUSHED THE DRAINAGE LUMEN AND BIFURCATION WITH D.I. WATER AND THERE WERE NO LEAKS NOTED. ALSO RECEIVED 2 PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWS OVERVIEW OF INFLATION FUNNEL AND CAP, TAMPER EVIDENT SEAL, AND SAMPLE PORT WITH CATHETER CUT OFF AT BIFURCATION SITE. SECOND PHOTO SAMPLE ZOOMS IN BIFURCATION SITE WITH CATHETER NOT PRESENT. BASED ON PHYSICAL SAMPLES RECEIVED IT IS UNKNOWN IF THE PRODUCT MEETS SPECIFICATIONS. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED POTENTIAL ROOT CAUSES INCLUDE: ¿ POOR BALLOON DESIGN ¿ INADEQUATE MATERIAL SELECTION THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ CORRECTIONS: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00345 | 2023-05-18 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRUE DILATATION CATHETER | VALVULOPLASTY BALLOON CATHETER | BARD PERIPHERAL VASCULAR, INC. | 0244512 | 0244512 |  | IT WAS REPORTED THAT DURING PREPARATION OF A VALVULOPLASTY PROCEDURE, THERE WAS ALLEGEDLY A TWENTY-TWO MILLIMETER BALLOON CATHETER IN THE CARTON INSTEAD OF TWENTY-FOUR MILLIMETER BALLOON CATHETER. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE TRUE PTV OUTER PACKAGING BOX AND ONE TRUE PTV DEVICE WERE RETURNED FOR EVALUATION. THE OUTER PACKAGING BOX WAS OPENED AND NO INNER PACKAGING WAS RETURNED. NO SEALED PACKAGING WAS PROVIDED. THE OUTER PACKAGING BOX LABELING MATCHED THE INFORMATION REPORTED IN TRACKWISE, AND SPECIFIED A 24 MM BALLOON. THE DEVICE RETURNED IN THE OPEN PACKAGING BOX HAD A 22 MM BALLOON SPECIFICATION PRINTED ON THE HUB. THE RETURNED SAMPLE WAS RETURNED TO THE MANUFACTURING FACILITY AND THE PRODUCT WAS MEASURED AND MET THE SPECIFICATIONS MATCHING THE 22 MM BALLOON SPECIFICATION PRINTED ON THE HUB. ALTHOUGH THE RETURNED DEVICE DID NOT MATCH THE SPECIFICATIONS OF THE RETURNED LABELING, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED PRODUCT MIX UP. PER THE REPORTED EVENTS, THE CUSTOMER PURCHASES BOTH 24MM AND 22MM PRODUCTS. SALES INFORMATION WAS REQUESTED BUT UNABLE TO BE OBTAINED AS THE PRODUCT WAS PURCHASED THROUGH A DISTRIBUTOR. THE PRODUCT WAS RECEIVED IN AN OPEN OUTER PACKAGING BOX WITH NO INNER PACKAGING RETURNED. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE AS NO EVIDENCE WAS FOUND TO CONFIRM THE PRODUCT IN THE SEALED PACKAGING WAS INCORRECT. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INSTRUCTIONS FOR USE STATES: PRECAUTIONS: 1. CAREFULLY INSPECT THE CATHETER PRIOR TO USE TO VERIFY THAT CATHETER HAS NOT BEEN DAMAGED DURING SHIPMENT AND THAT ITS SIZE, SHAPE AND CONDITION ARE SUITABLE FOR THE PROCEDURE FOR WHICH IT IS TO BE USED. DO NOT USE IF PRODUCT DAMAGE IS EVIDENT. 2. THE PRODUCT LABELING CONTAINED VALID INFORMATION GIVEN FOR THE LOT NUMBER AND PRODUCT CATALOG NUMBER, BUT DID NOT MATCH THE PRODUCT INSIDE THE PACKAGING BOX. H10: D4 (EXPIRATION DATE: 02/2026), G3 . H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03554 | 2023-05-18 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT STERILE WATER WAS NOT DRAINED DURING PRETEST. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR EVENT IS NOT REPORTABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03547 | 2023-05-18 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI12 |  | IT WAS REPORTED THAT TWO BALLOON RUPTURES OCCURRED IN THE SAME PATIENT AND IN THE SAME LOT. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED BALLOON BURST MEASURING 0.2995" WITH NO MISSING PIECES NOTED. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, "BALLOON MUST NOT BE TORN". A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINHOLE IN BALLOON OR CUFF. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿BARD® SILVER LUBRI-SIL® FOLEY TRAY SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY.] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
16952480 | 2023-05-18 | 2022-07-16 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | ACCUCATH ACE | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | AC1202252 | AC1202252 |  | ICU RN CALLED BY SUPERVISOR TO PLACE ACCUCATH FOR PATIENT IN ROOM (B)(6) WITH DIFFICULT IV ACCESS. RN ARRIVED TO ROOM TO PLACE ACCUCATH. PATIENT'S RIGHT UPPER EXTREMITY WAS UTILIZED, AS NO LIMB ALERT PRESENT. ACCUCATH PLACED UNDER ULTRASOUND GUIDANCE, BRACHIAL VEIN ACCESSED, AT SUPERFICIAL LOCATION NEAR ANTECUBITAL FOSSA. GUIDEWIRE THREADED WITHOUT RESISTANCE, CATHETER ADVANCED OFF GUIDEWIRE. VENOUS BLOOD RETURN WAS PRESENT AT THIS TIME. GUIDEWIRE AND NEEDLE WERE SUBSEQUENTLY RETRACTED. RESISTANCE WAS MET TO GUIDEWIRE WITHDRAWAL. ATTEMPT TO GENTLY REMOVE CATHETER REVEALED ONCE AGAIN RESISTANCE AND THE DEVICE WAS STUCK IN THE PATIENT'S ARM. PATIENT DISCOMFORT NOTED WITH THIS ATTEMPT. THE DEVICE WAS SECURED, AND THE MD/SUPERVISOR WERE IMMEDIATELY NOTIFIED. DEVICE REMOVED BY MD WITHOUT IMMEDIATE COMPLICATION. UPON REMOVAL THE DEVICE, CATHETER, AND GUIDEWIRE WERE INTACT, HOWEVER THE GUIDEWIRE WAS CAUGHT AT THE DISTAL END OF THE CATHETER AND UPON RETRACTION CURLED THE CATHETER BACK ONTO ITSELF, KINKED AT MULTIPLE LOCATIONS, WHICH OCCURRED INSIDE THE PATIENT¿S VESSEL. GUIDEWIRE WAS REMOVED FROM CATHETER FOR CLOSER INSPECTION AND REMAINED INTACT. PRODUCT CATALOG NUMBER: AC1202252. 20GX2.25 IN ACCUCATH ACE INTRAVASCULAR CATHETER WITH 0.008 IN. NITINOL GUIDEWIRE TIP. GUIDEWIRE FOR PLACEMENT OF ACCUCATH BECAME BENT AND DIFFICULT TO REMOVE FROM PATIENTS VEIN PATIENT COMPLAINED OF SOME DISCOMFORT WHEN TENSION WAS MET UPON ATTEMPTED REMOVAL OF GUIDEWIRE. ICU RN NOTIFIED MHO AND PHYSICIAN IMMEDIATELY AFTER SECURING THE LINE. PHYSICIAN WAS ABLE TO REMOVE GUIDEWIRE AND THERE WAS NO PATIENT HARM. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03559 | 2023-05-18 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER BALLOON BURST. THE CATHETER FELL OUT OF THE PATIENT IN THE WARD AFTER SURGERY. UPON REMOVING AND CHECKING THE CATHETER, IT WAS FOUND BURST. ACCORDING TO THE INQUIRY SHEET ATTACHED, THE EVENT HAS OCCURRED AFTER 3 HOURS OF PLACEMENT AND HAND (GLOVES) POSSIBLY CAME IN CONTACT WITH THE CATHETER. THE CATHETER WAS PULLED RARELY. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03562 | 2023-05-18 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE CATHETER WAS WITHDRAWN IMMEDIATELY AFTER PLACEMENT. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND IT IMMEDIATELY LEAKED OUT OF A 0.016" PINHOLE FOUND ON THE BALLOON. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE (B)(4), REVISION 0, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE.A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿WARNINGS 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA.THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03564 | 2023-05-18 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT THE SYRINGE WAS DAMAGED, SO STERILE WATER COULD NOT BE INJECTED. | THE REPORTED EVENT WAS CONFIRMED AS SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. THE AFFECTED SYRINGE WAS MANUFACTURED BY HANSCENT. FURTHER INVESTIGATION IS DOCUMENTED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. PACKAGING LOTS NUMBER MANUFACTURED WITH AFFECTED SYRINGE BATCHES IDENTIFIED AND DOCUMENTED. THE DEFORMATION AT THE TIP OF THE SYRINGE IS THE ROOT CAUSE THAT CONTRIBUTED TO THE LEAK; HOWEVER, THE ROOT CAUSE OF THE DEFORMATION IS UNKNOWN. NO FAILURES WERE FOUND WITHIN THE SUPPLIER INVESTIGATIONS. BOTH FLEX AND HANSCENT SHOWED NO CLEAR EVIDENCE THAT THE FAILURES OCCURRED DURING THEIR PROCESSES. HANSCENT DID NOT FIND ANY ROOT CAUSE OF THE LEAKAGE. THIS INVESTIGATION DOES NOT REQUIRE A DHR REVIEW DUE TO A CLOSURE LETTER NOT REQUIRED FOR THIS COMPLAINT AND THIS FAILURE WAS CONFIRMED. THE REPORTED EVENT IS UNDERSTOOD, CHARACTERIZED, AND CONFIRMED AS UNRELATED TO THE LABELING ISSUE. IT WAS CONFIRMED RELATED TO A SUPPLIER, THEREFORE LABELING REVIEW IS NOT REQUIRED FOR THIS REPORTED EVENT. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03563 | 2023-05-18 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER CAME OUT WITHIN A FEW HOURS AFTER PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03551 | 2023-05-18 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED THAT THE CUSTOMER WAS USING ONE OF THE INTERMITTENT CATHETERS AND EXPERIENCED PAIN WHEN THEY TRIED TO INSERT IT. THE CUSTOMER THEN WITHDREW THE CATHETER AND INSPECTED IT. THEY SAW THAT IT HAD A FAULT, INSTEAD OF THE USUAL ROUNDED END, IT HAD A SHARP POINT WITH BLOOD ON IT. IT WAS AS IF SOMEONE HAD CUT THE END WITH A KNIFE AT AN ANGLE. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT IS CONFIRMED - CAUSE UNKNOWN. IT WAS UNKNOWN HOW THE CATHETER HAD BROKEN TIP WITH A SHARP POINT. VISUAL INSPECTION NOTED FOUR (4) PHOTO SAMPLES. FIRST AND SECOND PHOTO SAMPLE SHOWCASES SHOWS ISC WITH A BROKEN TIP WITH A SHARP POINT. BASED ON THE SAMPLES RECEIVED, THE SHARP BROKEN TIP OF THE ISC MAY HAVE EXPERIENCED PAIN WHEN USER TRIED TO INSERT. A POTENTIAL ROOT CAUSE COULD BE INSUFFICIENT EYELET AREA. THEREFORE, THE PRODUCT DOES NOT MEET SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿FOR UROLOGICAL USE ONLY. INTENDED FOR USE BY PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICES ARE PASSED TO THE URINARY BLADDER VIA THE URETHRA. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. 1. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. 2. PRIOR TO OPENING THE SEALED CATHETER POUCH, APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. 3. WET THE CATHETER BY HOLDING THE PACKAGE WITH THE PRINTED SIDE UP AND TIP THE PACKAGE END-TO-END THREE TO SIX TIMES TO WET THE CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. 4. PEEL OPEN THE PACK AT THE FUNNEL/HANDLE END JUST ENOUGH TO EXPOSE 5 CM (2¿) OF THE CATHETER, OR FOR PRODUCTS WITH AN INSERTION SLEEVE, EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 5. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. 6. WASH YOUR HANDS AGAIN. 7. USING THE CATHETER WITH INSERTION SLEEVE (GRIPPER): A. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. B. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. C. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. WITHOUT INSERTION SLEEVE (GRIPPER): A. HOLD THE HANDLE/FUNNEL END AND REMOVE THE CATHETER FROM THE PACKAGING. B. GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA AND ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. 8. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. 9. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01946 | 2023-05-19 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD SINGLE LUMEN PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. THE EVENT INVOLVED 2 PICCS, ONE SL INSERTED IN THE UPPER ARM AND ONE TL INSERTED MID-THIGH. THIS REPORT ADDRESSES THE SINGLE LUMEN PICC. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03620 | 2023-05-19 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER CREEVY MODEL | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 01267520 | 01267520 |  | IT WAS REPORTED THAT THE VALVE GOT DISCONNECTED AND FELL OUT. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. THE USED THREE-WAY LATEX FOLEY CATHETER WAS RETURNED WITHOUT THE ORIGINAL PACKAGING. VISUAL EVALUATION: VISUAL EVALUATION NOTED DRIED RESIDUE WITH THE DRAINAGE EYES AND FUNNEL, AS WELL AS WITHIN THE IRRIGATION FUNNEL. THE INFLATION CAP/VALVE COMPLETELY SEPARATED FROM THE FROM THE FUNNEL. FURTHER EVALUATION NOTED THE INFLATION CAP WAS PLACED UPSIDE DOWN ON THE VALVE. THE CAP/VALVE WAS REATTACHED TO THE FUNNEL AND AN ATTEMPT WAS MADE TO INFLATE THE CATHETER, HOWEVER, THE VALVE DETACHED ONCE THE WATER WAS INSTILLED. WHEN ATTEMPTING TO REMOVE THE SYRINGE, THE CAP DISCONNECTED AND LEFT THE VALVE PORTION STILL ATTACHED TO THE LUER LOCK SYRINGE; IT IS UNKNOWN IF THIS OCCURRED DURING USE AND THE USER REPLACED THE CAP/VALVE IN THE WRONG ORIENTATION. THIS DOES NOT MEET SPECIFICATION " DETACHED VALVES/CAPS". A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03619 | 2023-05-19 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER FELL OFF FROM A PATIENT THE NEXT DAY AFTER PLACEMENT. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 09MAY2023,THE BALLOON WAS BURST | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE CONTRIBUTE BY "HIGH MECHANICAL STABILITY TIME (MST) OF INCOMING LATEX" OR "COAGULANT DIP 2- DIPPING SPEED TOO SLOW" AND ALSO MIGHT BE CONTRIBUTED BY USER RELATED ( EXAMPLE: MISHANDLING PRODUCT OR EXPOSED TO PETROLEUM OR SHARP OBJECT). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03616 | 2023-05-19 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER HAVING ISSUE FROM THE ORDER. THEY STATES THAT THE 7 OF THE CATHETERS WERE BENT IN THE BOX AND NOT LUBRICATED RMA FOR 7 CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. BASED ON PATIENT CODE 2645 THE PRODUCT WAS NOT USED ON THE PATIENT. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN, THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03610 | 2023-05-19 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT THE USER FACED DIFFICULTY WHILE INJECTING THE FIXED WATER INTO THE FOLEY CATHETER. WHEN THE FIXATION WATER WAS INJECTED INTO THE FOLEY CATHETER, THE WATER LEAKED FROM THE INFLATION VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03606 | 2023-05-19 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM BOTTOM OF THE BALLOON UPON INJECTING STERILE WATER INTO THE FOLEY CATHETER. THE EVENT OCCURRED JUST AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BALLOON. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01966 | 2023-05-19 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT TWO YEARS THREE MONTHS AND FIFTEEN DAYS POST PORT PLACEMENT, WHEN THE PHYSICIAN INTENDED TO REMOVE THE PORT FROM THE PATIENT, THE PHYSICIAN ONLY GOT THE PORT BODY WHILE THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN AND DISCONNECTED COMPLETELY AT ABOUT TWENTY-TWO CENTIMETERS SITE. IT WAS FURTHER REPORTED THAT IMMEDIATE X-RAY EXAMINATION REVEALED THAT THE BROKEN CATHETER WAS MIGRATED TO THE PULMONARY ARTERY. REPORTEDLY, THE BROKEN CATHETER WAS REMOVED FROM THE ARTERY AND SOME REMAINED CATHETER ATTACHED IN THE REMOVED PORT BODY. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE POWERPORT MRI IMPLANTABLE PORT ATTACHED TO A CATHETER IN TWO SEGMENTS WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATION WERE PERFORMED. A COMPLETE COMPOUND BREAK WAS NOTED ON THE DISTAL END OF THE ATTACHED CATHETER. A COMPLETE COMPOUND BREAK AND A LONGITUDINAL SPLIT WAS NOTED ON THE PROXIMAL END OF THE DISTAL CATHETER SEGMENT. THE EDGES OF THE COMPLETE COMPOUND BREAK ON THE DISTAL END OF THE ATTACHED CATHETER AND THE PROXIMAL END OF THE DISTAL CATHETER SEGMENT WERE NOTED TO BE UNEVEN. THE SURFACE WAS NOTED TO BE GRANULAR. IN ADDITION, THREE PHOTOS WERE PROVIDED FOR REVIEW WHICH ALSO CONFIRMS THE REPORTED FRACTURE AND MATERIAL SEPARATION ISSUES. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND MATERIAL SEPARATION ISSUES. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MIGRATION ISSUE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 08/2022). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03601 | 2023-05-19 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT ON THE NEXT DAY AFTER PLACEMENT, URINE OR STERILE WATER LEAKED FROM THE CATHETER SHAFT. AS PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 09MAY2023,VISUAL INSPECTION REVEALED A PINHOLE IN THE CATHETER SHAFT, BUT NO CONFIRMATION AS TO WHETHER THE LEAK WAS WATER OR URINE. | THE REPORTED EVENT WAS CONFIRMED-MANUFACTURING RELATED. SAMPLE HAS BEEN EVALUATED AND OBSERVED PINHOLE ON SHAFT OF CATHETER. THE ORIGIN OF PINHOLE ARE COME FROM BUBBLE. A POTENTIAL ROOT CAUSE FOR THIS POTENTIAL FAILURE MODE COULD BE DUE TO BUBBLE PRESENT DURING DIPPING. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE(S) THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX SHAPE, CONFIGURATION AND PRINCIPLES: BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY INTENDED USE & EFFECT- EFFICACY: THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. DIRECTIONS FOR USE: 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR. (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01969 | 2023-05-19 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT DURING THE PREPARATION OF THE PORT PLACEMENT PROCEDURE VIA THE RIGHT INTERNAL JUGULAR VEIN, THE PORT ALLEGEDLY HAD DIFFICULTY FLUSHING. IT WAS FURTHER REPORTED THAT AN OCCLUSION WAS FOUND DURING FLUSHING. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 12/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01968 | 2023-05-19 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT ONE YEAR THREE MONTHS AND THIRTEEN DAYS POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN UPON REMOVAL. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. A FRACTURE WAS NOTED NEAR THE PROXIMAL END OF THE CATHETER. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 02/2023), G3, H6 (METHOD). H11: B5, H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01967 | 2023-05-19 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT THREE YEARS ELEVEN MONTHS AND TWENTY-FIVE DAYS POST PORT PLACEMENT, A SMALL BREAKAGE WAS ALLEGEDLY NOTED ON THE CATHETER UPON REMOVAL. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT, ONE CATH-LOCK AND ONE GROSHONG CATHETER WAS RETURNED FOR EVALUATION. IN ADDITION TO THE RETURNED SAMPLE ANALYSIS, ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS A COMPOUND BREAK ON THE CATHETER BODY. FUNCTIONAL, GROSS VISUAL, TACTILE, MICROSCOPIC VISUAL AND DIMENSIONAL EVALUATIONS WERE PERFORMED. A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED APPROXIMATELY 5.8CM FROM THE PROXIMAL END OF THE CATHETER RETURNED. THE RADIALLY ADJACENT SPLITS WERE NOTED ON THE PROXIMAL PORTION OF THE CATHETER RETURNED. UPON INFUSION, A LEAK WAS OBSERVED FROM THE SPLITS AND THE PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED. ASPIRATION WAS ATTEMPTED AND WAS UNSUCCESSFUL. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. BASED ON THE MEASUREMENTS RECORDED DURING SAMPLE EVALUATION AND APPLICABLE DRAWINGS, NO MEASUREMENTS RECORDED COULD BE CONFIRMED TO BE OUT OF TOLERANCE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 11/2021). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03607 | 2023-05-19 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED THROUGH PINHOLE. WHEN THE USER TRIED TO INJECT STERILE DISTILLED WATER TO INFLATE THE BALLOON, THE STERILE DISTILLED WATER LEAKED DUE TO A PINHOLE. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE OVER DIPPING CAUSED STRIPPING DIFFICULTY, TORN FUNNELS AND PREMATURE DEFLATORS AND MIGHT BE CONTRIBUTED BY USER (EXAMPLE: CONTACT WITH SHARP OBJECT/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARDEX® SILVER LUBRICATH® FOLEY CATHETER SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03623 | 2023-05-19 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS DID NOT HAVE ENOUGH LUBRICANT ON THEM, MADE IT DIFFICULT FOR PATIENTS TO INSERT THEM. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03613 | 2023-05-19 | 2023-04-28 | I, n, j, u, r, y | Manufacturer report | PUREWICK MALE EXTERNAL CATHETER | PUREWICK MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWM030 |  | IT WAS REPORTED THAT PATIENTS AND STAFF REPORTED THAT PATIENTS SKIN TRYING TO HOLD THE PUREWICK IN PLACE. THEY USE DIAPERS, CHUCKS AND A LOT OF TAPE THAT IS DAMAGING PATIENTS SKIN AND CAUSING INFECTIONS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03597 | 2023-05-19 | 2023-04-25 | I, n, j, u, r, y | Manufacturer report | WIDEBAND¿ SELF-ADHERING MALE EXTERNAL CATHETER | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 36103 | 36103 |  | IT WAS REPORTED THAT THE CUSTOMER WAS HAVING ISSUES WITH THE ADHESIVE ON THE MALE EXTERNAL CATHETERS. THE CUSTOMER HAD TWO BOXES OF CATHETERS (BATCH# JUGT1701 AND BATCH# JUGU9119). ON (B)(6) 2022, THE CUSTOMER STATED THAT THE BOX OF CATHETERS WERE NOT WORKING AND THEY KEPT COMING OFF. THE CUSTOMER STATED THAT THEY HAD TO CHANGE THE CATHETERS FOUR TIMES BECAUSE THEY KEPT COMING OFF. SO THE REPRESENTATIVE RECOMMENDED TO USE THE PRODUCT FROM THE BOTTOM OF THE BOX AS THE TOP PRODUCT MIGHT BE DEFECTIVE ON (B)(6) 2022. ON (B)(6) 2022, THE CUSTOMER STATED THAT THEY PICKED THE CATHETERS FROM THE BOTTOM AS RECOMMENDED BY REPRESENTATIVE. HOWEVER, THE CUSTOMER HAD TO CHANGE 2 CATHETERS AS THEY WERE STILL COMING OFF AND THE PATIENT WAS GETTING WET. ON (B)(6) 2022, THEY CUSTOMER STATED THAT THE PATIENT WAS NOT ABLE TO GET INTO THEIR CHAIR BECAUSE THEY KEPT GETTING WET AND HAD TO STAY IN BED. THE CUSTOMER COULD NOT GO OUT BECAUSE THEY HAD TO MAKE SURE TO KEEP CHANGING THE PATIENT'S CATHETER AND CLEAN THEM SO THAT THEY DO NOT GET INFECTIONS. THEY COULD NOT GO OUT BECAUSE THE CATHETERS KEPT LEAKING. PER FOLLOW-UP INFORMATION RECEIVED ON (B)(6) 2023, STATED THAT BOTH OF THE BOXES OF CONDOMS LEAKED. THE PATIENT ENDED UP WITH AN INJURY AND HAD AN OPEN WOUND IN THEIR BUTT DUE TO ALL THE WETNESS GETTING IN FROM THE URINE. IT WAS HARD TO LEAVE THE HOUSE AS THE PATIENT WAS GETTING WET ALL THE TIME. THEY WERE CONFINED TO STAYING INSIDE DURING THE HOLIDAYS BECAUSE OF THE BIG LEAKS. THE GLUE WAS NOT STICKING WELL. THE CUSTOMER HAD A FEW THAT DID WORK OUT OF THE 2 BOXES, BUT MOSTLY THEY HAD TO CHANGE AT LEAST 3 TIMES A DAY. THE CUSTOMER WAS NOT HAPPY WITH THE PRODUCT DUE TO THE AMOUNT OF TIME THEY SPENT TO KEEP THE PATIENT DRY, WAKING UP AT NIGHT TO CHANGE THE PATIENT AND ALSO DUE TO THE INJURY THAT THE PATIENT HAD SUSTAINED FROM THE PRODUCT. PER FOLLOW-UP INFORMATION RECEIVED ON (B)(6) 2023, STATED THAT THE WOUND CARE NURSE LOOKED AT THE WOUND AND WAS TREATING IT. PER ADDITIONAL INFORMATION RECEIVED FROM CUSTOMER ON (B)(6) 2023, STATED THAT THE MALE EXTERNAL CATHETERS HAD STARTED TO LEAK AGAIN. IN 5 DAYS, THE PATIENT HAD TO CHANGED IT 5 TIMES. THE ADHESIVE WAS NOT STICKING AND THEY WERE GETTING SKIN BREAK DOWN (BATCH# JUGW0577). PER FOLLOW-UP RECEIVED FROM CUSTOMER ON (B)(6) 2023, STATED THAT THEY HAD ALREADY RETURNED SAMPLES WITH 2 DIFFERENT LOT NUMBERS. THEY HAD ALREADY REPORTED SKIN BREAK DOWN AND MEDICAL INTERVENTION WAS THAT A WOUND CARE NURSE WAS TREATING IT. THE NURSE WAS TREATING THE SKIN WITH INER DRY AND COLOPLAST PASTE. THE LOT NUMBER FOR THIS BATCH WAS JUGW0577, THE ADHESIVE GLUE DID NOT STICK WELL. IT HAD THE SAME PROBLEM. PER FOLLOW-UP RECEIVED FROM CUSTOMER ON (B)(6) 2023, STATED THAT THEY HAD BREAK DOWN OF THE SKIN ON BOTH SIDES OF THE GROIN FROM THE CONDOMS LEAKING AND GETTING MOISTURE FROM THE URINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | LAY USER/PATIENT |  | 2023-05
16960607 | 2023-05-19 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWER-TRIALYSIS SHORT-TERM EXTENSION DIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | BARD ACCESS SYSTEMS, INC. | 5613150 | 5613150 |  | DURING INSERTION OF POWER TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS CATHETER, THE CATHETER GOT HUNG UP ON THE GUIDEWIRE DURING THE NORMAL EXCHANGE. INITIALLY, PROVIDER THOUGHT THERE WAS A PROBLEM WITH THE CATHETER, LEFT THE GUIDEWIRE IN PLACE AND REQUESTED A SECOND CATHETER. THE SECOND CATHETER ALSO HUNG UP ON THE GUIDEWIRE IN THE SAME WAY AS THE FIRST CATHETER. UPON FURTHER INSPECTION, THE GUIDEWIRE IS FRAYED AT THE DISTAL END. NO IMMEDIATE HARM TO PATIENT DETECTED BUT FRAYED WIRE COULD HAVE BROKEN OFF AS AN EMBOLIC STRUCTURE IN PATIENT'S VASCULATURE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03596 | 2023-05-19 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER WAS DIFFICULT TO DRAIN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01974 | 2023-05-19 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PPICC SOLO 5F DL FT 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | IT WAS REPORTED BY THE CUSTOMER THAT "INTRODUCER WILL NOT PENETRATE SKIN WITHOUT BENDING OR BREAKING." CLINICIAN OPENED A NEW INSERTION KIT. NO OTHER INFORMATION WAS PROVIDED. 05/16/2023 - TWO INTRODUCERS WERE RETURNED FOR EVALUATION AND BOTH EXHIBITED KINKS. THIS REPORT ADDRESSES THE SECOND DEVICE. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF INTRODUCERS NOT PENETRATING THE SKIN WITHOUT BENDING IS CONFIRMED. THREE 5.0 FR MICROINTRODUCERS WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED BLOOD USE RESIDUES ON ALL THREE OF THE RETURNED DEVICES. SAMPLE ONE HAD A KINK TOWARDS THE DISTAL END OF THE GRAY SHEATH, SAMPLE TWO HAD TWO KINKS ALONG THE GRAY MICROINTRODUCER SHEATH, AND SAMPLE THREE HAD A SLIGHT BEND THROUGHOUT THE ENTIRE DEVICE. A MICROSCOPIC OBSERVATION REVEALED SAMPLE ONE TO HAVE A LONGITUDINALLY ALIGNED SPLIT ON THE DISTAL END OF THE GRAY SHEATH THAT APPEARS TO BE MANUFACTURING RELATED. SAMPLES TWO AND THREE HAD NO SPLITS IN THE GRAY SHEATHS AND NO OBVIOUS DEFORMATION TO THE DISTAL ENDS OF THE DEVICES. SAMPLES TWO AND THREE HAD SMOOTH TRANSITIONS FROM THE DISTAL END OF THE GRAY SHEATH TO THE INTRODUCER. THE SMOOTH TRANSITIONS OF THE MICROINTRODUCER SHEATHS FOR SAMPLES TWO AND THREE ARE WITHIN MANUFACTURING SPECIFICATIONS. THE COMPLAINT OF THE MICROINTRODUCERS NOT PENETRATING THE SKIN WELL IS CONFIRMED. POTENTIAL FAILURE MODES FOR MICROINTRODUCERS COULD BE THE ANGLE OF PENETRATION INTO THE SKIN DEPENDENT ON THE SUBCUTANEOUS THICKNESS OF THE PATIENT¿S SKIN. ANGLES THAT ARE TOO SHARP COULD CAUSE DIFFICULTY DURING INSERTION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01965 | 2023-05-19 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN FOR CHEMOTHERAPY, SUBCUTANEOUS LEAKAGE WAS ALLEGEDLY FOUND. IT WAS FURTHER REPORTED THAT PATIENT ALLEGEDLY EXPERIENCED SWELLING NEAR THE PORT BODY. HOWEVER, IT WAS CONFIRMED THAT NO ADDITIONAL MEDICATION OR INTERVENTION PROVIDED TO TREAT PATIENT SWELLING. REPORTEDLY, THE PORT BODY WAS REMOVED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE X-PORT IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATION WERE PERFORMED. THE PORT BODY WAS PATENT TO BOTH INFUSION AND ASPIRATION WITHOUT ISSUE AND NO LEAKS WERE OBSERVED THROUGHOUT TESTS. THEREFORE, THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED LEAK ISSUES AS NO EVIDENCE OF THE REPORTED EVENT WAS NOTED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, G3 H11: H6 (METHOD, RESULT, CONCLUSION) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00348 | 2023-05-19 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | AT75144 | AT75144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE BALLOON ALLEGEDLY WOULDN'T DEFLATE. IT WAS FURTHER REPORTED THAT PHYSICIAN NEED TO UPSIZE THE SHEATH TO GET THE BALLOON OUT. REPORTEDLY, THE BALLOON WAS SEPARATING FROM THE PLASTIC. THE PHYSICIAN HAD TO TO ACCESS ANOTHER AREA IN THE FISTULA TO FINISH THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ATLAS PTA DILATATION CATHETER WAS RETURNED FOR EVALUATION. THE BALLOON WAS LOADED IN AN UNKNOWN SHEATH. THE BALLOON HAD A BREAK FROM THE CATHETER LEAVING THE PORTHOLES AND SOME PART OF THE INNER GUIDE WIRE LUMEN WAS EXPOSED. THE BALLOON WAS STILL ATTACHED TO THE CATHETER, NOT COMPLETELY DETACHED AND WAS COMPLETELY BUNCHED NEAR TO THE DISTAL END. NO OTHER ANOMALIES WERE NOTED DURING THE VISUAL EVALUATION. DURING THE FUNCTIONAL TESTING, AN ATTEMPT TO REMOVE THE LOADED SHEATH WAS PERFORMED BUT WAS UNSUCCESSFUL. HOWEVER, THE BALLOON WAS COMPLETELY ADVANCED THROUGHOUT THE SHEATH, EXPOSING THE INNER GUIDEWIRE LUMEN AND GLUE BULLET SEATED INCORRECTLY. ALSO, TWO PHOTOS WERE PROVIDED AND REVIEWED. THE FIRST PHOTO SHOWS THE BALLOON LOADED WITHIN THE INTRODUCER SHEATH AND THE BALLOON WAS NOT TO BE BUNCHED NEAR THE DISTAL END OF THE DISTAL TIP. THE SECOND PHOTO SHOWS THE BUNCHED BALLOON MATERIAL. NO OTHER ANOMALIES ARE NOTED IN THE SUBMITTED PHOTO. BASED ON THE RETURNED SAMPLE AND PHOTOS, THE RETURNED CATHETER WAS LOADED WITHIN AN UNKNOWN SHEATH AND HAD A BREAK EXPOSING THE PORT HOLES AND GLUE BULLET. THEREFORE, THE INVESTIGATION WAS CONFIRMED FOR THE REPORTED SHEATH REMOVAL ISSUE AND CATHETER BREAK. HOWEVER, THE FUNCTIONAL TESTING COULDN¿T BE PERFORMED FOR THE DEFLATION ISSUE DUE TO THE DEVICE'S CONDITION BEING RETURNED FOR EVALUATION. HENCE REMAIN THE INVESTIGATION WILL REMAIN INCONCLUSIVE FOR REPORTED DEFLATION ISSUE. A DEFINITIVE ROOT CAUSE FOR THE REPORTED SHEATH REMOVAL ISSUE AND CATHETER BREAK AND REPORTED DEFLATION ISSUE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 12/2025), G3, H6 (DEVICE, METHOD) H11: B5, H6 (RESULT, CONCLUSION) H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10 | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01963 | 2023-05-19 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT DURING DIALYSIS CATHETER IMPLANT VIA RIGHT INTERNAL JUGULAR VEIN FOR DIALYSIS, THE GUIDEWIRE ALLEGEDLY GOT STUCK INSIDE THE NEEDLE AND WOULD NOT ALLOW THE NEEDLE TO BE REMOVED OFF. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE GUIDEWIRE LOADED TO AN INTRODUCER NEEDLE WERE RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC, DIMENSIONAL AND FUNCTIONAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED PHYSICAL RESISTANCE AND DIFFICULT TO REMOVE ISSUES AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED AND THE REPORTED EVENT COULD NOT BE REPRODUCED IN THE LAB. HOWEVER, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED STRETCHED AND MATERIAL PROTRUSION ISSUES AS THE GUIDEWIRE WAS NOTED TO BE STRETCHED AND UNCOILED AND THE CORE WIRE WAS PROTRUDING THE COILS OF THE GUIDEWIRE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRY DATE: 05/2024). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03599 | 2023-05-19 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE ARCTIC SUN DEVICE HAD BEEN WARMING PATIENT AND TRYING TO REWARM SINCE 5:15. INTENSIVIST CAME IN AND ADJUST THE SETTINGS. THE DEVICE WAS SET TO TARGET TEMPERATURE WAS 37C AT RATE OF 0.5C/HR, BUT INTENSIVIST CHANGED TO REWARM FROM 37C AT MAXIMUM RATE TO TARGET TEMPERATURE OF 37C. THERAPY CURRENTLY STOPPED. 4 PADS CONNECTED. EVENT LOG SHOWS 105 (COOL PATIENT END) AND ALERT 113 (REDUCED WATER TEMPERATURE CONTROL). MIS HAD NURSE TO RESTART THERAPY AND AFTER A COUPLE MINUTES CHECKED SYSTEM DIAGNOSTICS SHOWED OUTLET MONITOR TEMPERATURE (T1) WAS 27.7C, OUTLET CONTROL TEMPERATURE (T2) WAS 27.7C, INLET TEMPERATURE (T3) WAS 30.5C, CHILLER TEMPERATURE (T4) WAS 7.5C, WATER FLOW RATE WAS 2.6 L/M, INLET PRESSURE WAS -7PSI, CIRCULATION PUMP COMMAND WAS 81 PERCENTAGE, MIXING PUMP COMMAND WAS 0, HEATER SHOWED 12 PERCENTAGE, WATER RESERVOIR LEVEL SHOWED 5, SYSTEM HOURS WERE 6858.9 AND PUMP HOURS WERE 6002.5. THE WATER TEMPERATURE WAS 29.4C. MIS WALKED NURSE THROUGH ADJUSTING SETTINGS TO REWARM FROM PATIENT CURRENT TEMPERATURE WAS 34.9C TO TARGET TEMPERATURE OF 37C. MIS HAD NURSE TO STOP THERAPY, DRAIN PADS AND DRAIN 16 OUNCES FROM RIGHT DRAIN PORT. NURSE RESTARTED THERAPY. MIS ENCOURAGED ADDING BAIR HUGGERS OR WARM BLANKETS TO HANDS, HEAD AND FEET IF OK WITH PROTOCOL. MIS WOULD CALL BACK IN 20 MINUTES TO CHECK WATER TEMPERATURE. NURSE STATED THAT THEY RECEIVED ALERT ABOUT PATIENT PROBE. NURSE SWAPPED OUT TO ESOPHAGEAL PROBE. THE PATIENT TEMPERATURE WAS NOW REGISTERING HIGHER AT 35.4C, WATER TEMPERATURE WAS 29.8C AND THERAPY RESTARTED ABOUT 20 MINUTES AGO. NURSE CONFIRMS PATIENT WAS GENERATING HEAT. DISCUSSED WAYS TO ADDRESS HEAT GENERATION. NURSE NOTED DEVICE WOULD NOT MAKE WARM WATER IF IT WAS TRYING TO COUNTERACT TEMPERATURE INCREASE TOO QUICKLY. PATIENT TREND SHOWING 1 BAR TRENDING UPWARD. MIS ADVISED TO CALL BACK WITH ANY QUESTIONS OR CONCERNS. AS PER FOLLOW UP INFORMATION RECEIVED VIA EMAIL ON 01JUN2023, IT WAS REPORTED THAT THERAPY WAS COMPLETED AND THERE WAS NO IMPACT TO THE PATIENT. THE PATIENT WAS MALE AND NO OTHER IDENTIFYING INFORMATION WAS PROVIDED. DID TROUBLESHOOTING WITH MIS AND WAS WALKED THRU THE ERROR CORRECTING PROCESS. DEVICE WORKS AND WAS NOT SENT TO BIOMED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "SENSOR WIRE TOO WEAK." THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03598 | 2023-05-19 | 2023-05-01 | I, n, j, u, r, y | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 996081 | 996081 |  | IT WAS REPORTED THAT IN AN ONLINE SURVEY, A PHYSICIAN STATED A TISSUE TRAUMA (HEMATURIA, BLEEDING) ADVERSE EVENT, DIRECTLY ATTRIBUTABLE TO THE DEVICE 996081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, W/ INFLATION DEVICE, STERILE. ALSO STATED THAT THE PATIENT EXPERIENCED UTI PRIOR TO DEVICE 996081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, W/ INFLATION DEVICE, STERILE PLACEMENT AND DIDN'T RESULT IN PATIENT INJURY REQUIRING MEDICAL OR SURGICAL INTERVENTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE COATING/TISSUE INCOMPATIBILITY. A DHR REVIEW IS NOT REQUIRED AS NO LOT NUMBER WAS PROVIDED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." CORRECTION: G,F,H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03595 | 2023-05-19 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTERMITTENT ROCHESTER CATHETERS | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE BOTH INTERMITTENT CATHETERS WERE HURT AND WERE DIFFICULT TO PUT IN. THE PATIENT STATED ITEM INTERMITTENT CATHETER WAS MESSY AFTER POPPING THE BURST PACKET. THEY SENT THE PATIENT SAMPLES OF DIFFERENT CATHETERS SO THEY COULD FIND A PRODUCT THEY COULD USE. NO MEDICAL INTERVENTION WAS REPORTED. PER CLARIFICATION MAIL RECEIVED ON (B)(6)2023, THESES CATHETERS HURT, WERE DIFFICULT TO INSET, AND MESSY AFTER POPPING THE BURST PACKET. PER CLARIFICATION MAIL RECEIVED ON (B)(6)2023, THE PATIENT REPORTED THAT THE RTU14M WAS MESSY AFTER POPPING THE BURST PACK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
16960987 | 2023-05-19 | 2023-02-19 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARD | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. |  |  |  | BALLOON DEFLATED WITHIN ABOUT 30 SECONDS. |  | N | Y |  |  | 2023-05
MW5117731 | 2023-05-19 | 2023-03-13 | I, n, j, u, r, y | Voluntary report | ATLAS GOLD ANGIOPLASTY BALLOON | CATHETER, PERCUTANEOUS | BARD PERIPHERAL VASCULAR, INC. |  | ATG80144 |  | AN ESRD PATIENT THAT CAME TO THE OFFICE FOR ROUTINE ANGIOPLASTY FOR RECURRENT STENOSIS. A BARD, ATLAS GOLD 14MM X 4 CM BALLOON WAS USED. DURING BALLOON INFLATION, THE BALLOON COMPLETELY SEPARATED FROM THE SHAFT AND STARTED MIGRATING, SO IT HAD TO BE STENTED IN PLACE. THE BALLOON WAS INFLATED TO 14 ATM OF PRESSURE AND NOT PASSED THE STATED BURST PRESSURE OF 18 ATM. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01979 | 2023-05-21 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNK LOW ARTIFACT POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT CLEARVUE SLIM IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017¿ MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THIS EVENT IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01981 | 2023-05-21 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE DEHISCENCE REPORTED DURING THE USE OF THE POWERPORT ISP IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01978 | 2023-05-21 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY (B)(6) TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWERPORT CLEARVUE ISP IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017¿ MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THIS EVENT IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01980 | 2023-05-21 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNK LOW ARTIFACT POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE HEMATOMA REPORTED DURING THE USE OF THE POWERPORT CLEARVUE SLIM IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017¿ MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THIS EVENT IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02012 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017- (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02011 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02026 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE PHLEBITIS REPORTED DURING THE USE OF THE POWERPORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01984 | 2023-05-22 | 2023-04-09 | D, e, a, t, h | Manufacturer report | POWER PORT MRI ISP, 8 FR GROSHONG, INT WO SP, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808561 |  | IT WAS REPORTED THAT DURING PORT INFUSION TREATMENT, RESISTANCE WAS ALLEGEDLY EXPERIENCED. IT WAS FURTHER REPORTED THAT AN X-RAY IDENTIFIED THAT THE TIP OF THE CATHETER WAS BLOCKED. REPORTEDLY, TWO ATTEMPTS WERE MADE TO FLUSH THE CATHETER WITH UROKINASE; HOWEVER, THE HCP WAS UNABLE TO REMOVE THE BLOCKAGE, SO THE DEVICE WAS REMOVED FROM THE PATIENT. THE PATIENT EXPIRED AFTER BEING DISCHARGED FROM THE HOSPITAL. THE PATIENT'S DEATH WAS DUE TO MULTIPLE ORGAN METASTASIS OF ADVANCED LIVER CANCER. THE RELATIONSHIP BETWEEN THE DEVICE AND THE PATIENT DEATH IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 02/2024) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01985 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: H6 (CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01983 | 2023-05-22 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE DEVICE WAS ALLEGEDLY DAMAGED DURING USE. IT WAS FURTHER REPORTED THAT, THE DEVICE WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02009 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02016 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02014 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02029 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02008 | 2023-05-22 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02033 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02035 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02015 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02032 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02013 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01986 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01987 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01990 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01989 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: H6 (CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02028 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02018 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE HEMATOMA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02025 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE THROMBOSIS REPORTED DURING THE USE OF THE POWERPORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02021 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE PNEUMOTHORAX REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02023 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE PROCEDURE-INDUCED SEPSIS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02020 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE HEMOTHORAX REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6)2017-(B)(6)2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02004 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02007 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY (B)(6) TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02006 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02005 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY (B)(6) TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01996 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01988 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01999 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02002 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02000 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01998 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6), 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3, H11: D1, G4, H6 (CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02003 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01997 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY (B)(6) CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01994 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01995 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017- (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01993 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: D1, G4, H6 (CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01992 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01991 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE INFECTION REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02036 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02001 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02017 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02040 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02037 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03649 | 2023-05-22 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENTS MOTHER STATED THAT THE PATIENT GOT A URINARY TRACT INFECTION. PATIENT'S MOTHER WAS UNSURE WHERE IT CAME FROM. PATIENTS' MOTHER WAS STOPPING PATIENT USE FOR NOW. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. PER FOLLOW UP VIA PHONE ON 17MAY2023, IT WAS REPORTED THAT THE CUSTOMER WAS PARALYZED FROM THE WAIST DOWN AND HAD A URINARY TRACT INFECTION THAT THEY WERE CURRENTLY TAKING ANTIBIOTICS FOR. THE CUSTOMER WAS STILL ACTIVELY USING THE DEVICE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE " DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. " " REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER." "ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS." BAW2456229 REV 0 IS FOUND TO BE ADEQUATE TO FULFILL S03FA211 REVISION 18 AS IT STATES: PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE:- EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE. "DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS." "6. REPLACE THE PUREWICK¿ FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICK¿ FEMALE EXTERNAL CATHETER." "ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03650 | 2023-05-22 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MAGIC3 GO MALE CATHETER WITH SUREGRIP | C.R. BARD INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATED THAT HE HAD GOTTEN A URINARY TRACT INFECTION WHEN USING THE INTERMITTENT CATHETERS ON THE RECALL LIST. PATIENT SAID THAT PATIENT HAD TO GO TO CRITICAL CARE WAS TREATED FOR A URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. PER FOLLOW UP VIA PHONE ON 20MAY2023, IT WAS REPORTED THAT THE URINARY TRACT INFECTION WAS CONFIRMED TO BE FROM THE CATHETER AND THE CUSTOMER WAS PRESCRIBED ANTIBIOTICS. ALSO STATED THAT THE CUSTOMER WAS USING A DIFFERENT CATHETER AND WAS UNSURE OF THE NAME OR BRAND, BUT STATED IT WAS WORKING WELL. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01982 | 2023-05-22 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808000 | 1808000 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE DEVICE WAS ALLEGEDLY DIFFICULT TO FLUSH. IT WAS FURTHER REPORTED THAT, THE DEVICE WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 11/2019). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02024 | 2023-05-22 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808000 | 1808000 |  | IT WAS REPORTED THAT THREE YEARS FOUR MONTHS AND EIGHTEEN DAYS POST PORT PLACEMENT PROCEDURE, THE DEVICE WAS ALLEGEDLY DIFFICULT TO FLUSH. IT WAS FURTHER REPORTED THAT, THE DEVICE WAS NOT FLUSHED REGULARLY. REPORTEDLY, THE DEVICE WAS STILL IMPLANTED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 11/2019). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02041 | 2023-05-22 | 2023-04-25 | I, n, j, u, r, y | Manufacturer report | POWERPORT SLIM TI W 6FR CHRONOFLEX WO SP, INTER. KIT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1716001 | 1716001 |  | IT WAS REPORTED THAT ONE YEAR POST PORT REMOVAL, THE CATHETER ALLEGEDLY FRACTURED AND MIGRATED AND REMAINED OVER THE PULMONARY ARTERY. IT WAS FURTHER REPORTED THAT THE REMAINED CATHETER WAS REMOVED. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED EVENT COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 05/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02043 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN MRI POWERPORT ISP | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK MRI POWERPORT ISP |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE PNEUMOTHORAX REPORTED DURING THE USE OF THE POWERPORT VUE M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017- (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02046 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY (B)(6) TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017- (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02049 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | POWERPORT VUE M.R.I. IMPLANTABLE PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT VUE M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02051 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | POWERPORT VUE M.R.I. IMPLANTABLE PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT VUE M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02052 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | POWERPORT VUE M.R.I. IMPLANTABLE PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY (B)(6) TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWERPORT VUE M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02053 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | POWERPORT VUE M.R.I. IMPLANTABLE PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT VUE M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6), 2017 (B)(6) , 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3 H11: H6 (CONCLUSION) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03646 | 2023-05-22 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER SEEMED TO BE THICKER, CAUSING PATIENT ISSUES WHEN INSERTING IT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "UN CALIBRATED VISCOMETER ,OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03648 | 2023-05-22 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT DID GET A URINARY TRACT INFECTION AND WAS GIVEN MEDICATION FROM THEY DOCTOR ALSO THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03647 | 2023-05-22 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAD URINARY TRACT INFECTION BUT COULD NOT DETERMINE IF IT WAS CAUSED BY THE MAGIC 3 GO MALE INTERMITTENT. CATHETER PATIENT STATES THEY HAD BEEN TAKING MEDICINE FOR 2 YEARS. (LOT#JUGT0289 , LOT#JUGU1969 , LOT#JUGU0783 , LOT#JUGV0665). | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03639 | 2023-05-22 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CATHETER WAS LEAKING. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO WRONG PRODUCT SIZE SELECTED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03630 | 2023-05-22 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE GRIPPER WAS ALL THE WAY DOWN AT THE BOTTOM OF THE CATHETER. PATIENT WAS ASKED FOR THE LOT # AND PATIENT STATED PATIENT DID NOT WANT TO GET THE INFORMATION BECAUSE PATIENT DID NOT WANT TO BE BOTHERED. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "MOLDING FLASH IN PROCESS - INADEQUATE AIR PRESSURE". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03628 | 2023-05-22 | 2023-05-01 | I, n, j, u, r, y | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 997006 | 997006 |  | IT WAS REPORTED THAT A MALE PEDIATRIC PATIENT (14 YEARS OLD) EXPERIENCED PERFORATION AND URINARY TRACT INFECTION AFTER PLACEMENT OF THE DEVICE, THESE COMPLICATIONS WERE DIRECTLY ATTRIBUTABLE TO USAGE OF BALLOON DILATION CATHETER. PERFORATION WAS NOT PRESENT PRIOR TO PLACEMENT OF THE DEVICE AND DID NOT RESULT IN PATIENT INJURY REQUIRING MEDICAL OR SURGICAL INTERVENTION. URINARY TRACT INFECTION WAS PRESENT PRIOR TO PLACEMENT OF THE DEVICE, AND DID RESULT IN PATIENT INJURY REQUIRING MEDICAL OR SURGICAL INTERVENTION SUCH AS URETERAL SCARRING. MEDICAL INTERVENTION WAS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03633 | 2023-05-22 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123520A | 123520A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS LEAKING. PER FOLLOW UP VIA PHONE ON (B)(6)2023, IT WAS REPORTED CUSTOMER WERE CURRENTLY GETTING THEIR SUPPLIES THROUGH ADAPTHEALTH. THE BALLOON DEFLATED ON THE CATHETER CONSTANTLY WHICH MAKES IT DIFFICULT TO COLLECT ANY URINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03629 | 2023-05-22 | 2023-04-27 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS REMOVED SPONTANEOUSLY AND THERE WAS NO BALLOON PART. AN ECHOCARDIOGRAM WAS DONE BUT COULD NOT CONFIRM A RESIDUE IN THE BLADDER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03631 | 2023-05-22 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 129416M | 129416M |  | IT WAS REPORTED THAT NURSE TRIED TO FLUSH TEMPERATURE SENSING FOLEY CATHETER TO CLEAR SEDIMENT. PORT BROKE OFF AND FOLEY STARTED LEAKING. FOLEY WAS REMOVED AND NEW PRODUCT WAS REINSERTED. NO PATIENT HARM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03626 | 2023-05-22 | 2022-12-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT REGISTERED NURSE PLACED THE FOLEY AND THEN THE CHEMO DRUG, GEMZAR, WAS INSTILLED IN THE BLADDER AT 1132. THERE WAS A CLAMP PLACED ON THE FOLEY TUBING AND REGISTERED NURSE RETURNED FROM LUNCH AT 1230, ANOTHER REGISTERED NURSE REPORTED THAT THE CLAMP NEEDED TO BE REMOVED FROM THE FOLEY TUBING AT 1232 SO THAT REGISTERED NURSE REMOVED THE CLAMP FROM THE FOLEY AND SAW URINE PASSING THROUGH THE FOLEY TUBING. LATER, THE RESIDENT ASKED THE REGISTERED NURSE WHERE THE FOLEY WAS. REGISTERED NURSE FOUND THE FOLEY BY THE HEAD OF THE BED AND SHOWED RESIDENT AND NOTED THAT THERE WAS ABOUT 175 IN THE FOLEY BAG. REGISTERED NURSE REMINDED THEM THAT 50CC OF GEMZAR HAD BEEN INSTILLED IN THE BLADDER, SO THEY DOCUMENTED 125 OF URINE OUTPUT. AROUND 1500 THE ROBOT WAS UNDOCKED, AND THE PATIENT WAS REPOSITIONED, AND AT THAT TIME IT WAS NOTED THAT THE FOLEY WAS ON THE GROUND. AT 1505, RESIDENT REPLACED THE FOLEY AND IMMEDIATELY 500CC OF URINE CAME OUT OF THE BLADDER INTO THE NEWLY PLACED BAG. SURGEON WAS PRESENT AND DID NOT STATE THAT ANYTHING ELSE NEEDED TO BE DONE. REGISTERED NURSE ASKED RESIDENT HOW MUCH URINE OUTPUT THEY HAD DOCUMENTED IN TOTAL AND THEY HAD CHARTED THE PREVIOUS 125 FOR THE 1130 OUTPUT. BECAUSE OF THE CHANCE THAT GEMZAR MIGHT HAVE BEEN IN THE BLADDER LONGER THAN THE RECOMMENDED ONE-HOUR INSTILLATION. | PER INVESTIGATOR EVALUATION, BD HAS DETERMINED THAT THIS MDR EVENT IS NOT REPORTABLE. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03644 | 2023-05-22 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53812G | 53812G |  | IT WAS REPORTED THAT THE PATIENT MIGHT HAD A INFECTION FROM THE INTERMITTENT CATHETER BUT THEY WERE NOT SURE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR INFECTION. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND THAT THERE WAS NO ALLEGATION AGAINST PRODUCT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02030 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6)2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00350 | 2023-05-22 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG120144 | ATG120144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE FOR AN ABDOMINAL AORTIC ANEURYSM GRAFT VIA FEMORAL ACCESS, THE PTA BALLOON ALLEGEDLY RUPTURED AFTER POST DILATION OF STENT. IT WAS FURTHER REPORTED THAT UPON REMOVAL OF THE BALLOON THE INTERNAL SHAFT WIRE SEPARATED. THE BALLOON WAS REMOVED SUCCESSFULLY. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 02/2026). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03656 | 2023-05-22 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE PACKED IMPROPERLY WHEN THEY OPENED THE BOX. THE END TIPS, OPPOSITE OF THE INSERTION END, ARE JAGGED AND NOT CUT PROPERLY. THE PATIENT STATED THAT THEY APPEAR TO NOT BE LUBRICATING CORRECTLY AND HAVE CAUSED MINOR BLEEDING. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿BLUNT KNIVES ¿. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02034 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02038 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6)2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02039 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02022 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE PNEUMOTHORAX REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02027 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY (B)(6) CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02019 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE HEMATOMA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017 - (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02044 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE PNEUMOTHORAX REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL EVENT ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02048 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017- (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02045 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE SEPSIS REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02047 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWERPORT ISP M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017- (B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02042 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE OTHER COMPLICATIONS REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM JUNE 17, 2017-MAY 11, 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENTS REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03668 | 2023-05-22 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0102L16 | 0102L16 |  | IT WAS REPORTED, A FOLEY CATHETER INSERTED ON (B)(6). AND THE CUSTOMER NOTICED WETNESS AT BALLOON PART OF CATHETER THAT EVENING. IN THE MORNING, THE CATHETER FELL OUT WHEN THE CUSTOMER WAS USING WASHROOM AND THE BALLOON WAS COMPLETELY DEFLATED. THE CUSTOMER WAS PRESENTED TO HOMECARE, AND A PINHOLE WAS NOTICED IN CATHETER. STATED, THAT THERE WAS AN POTENTIAL ADVERSE OUTCOME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03667 | 2023-05-22 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT PARTICULAR RESIDENT WHO USED FOLEY TRAY WAS EXPERIENCING THE BALLOON DEFLATING. CUSTOMER WANTED TO KNOW IF THIS WAS COMMON WITH SILICONE FOLEYS AND IF THIS WAS PATIENT SPECIFIC. REPRESENTATIVE RESPONDED THIS SOMETIMES HAPPENS WITH FOLEY CATHETERS, PARTICULARLY WITH SILICONE BALLOONS. THIS WAS NOT DUE TO THE DWELL TIME, BUT RATHER DUE TO OSMOSIS. OSMOSIS WAS DEFINED AS THE PASSAGE OF WATER FROM A REGION OF HIGH CONCENTRATION THROUGH A SEMI-PERMEABLE MEMBRANE TO A REGION OF LOWER CONCENTRATION. THE RATE OF FLUID MOVEMENTS WAS DEPENDENT ON THE CONCENTRATION OF THE FLUID ON THE OTHER SIDE. IF A PATIENT HAD A HIGH URINE SPECIFIC GRAVITY, THEN FLUID WOULD MOVE OUT OF THE BALLOON FASTER. IF THE PATIENT¿S URINE LOWER SPECIFIC GRAVITY, OSMOSIS WOULD BE AT A SLOWER RATE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PINHOLE IN BALLOON OR CUFF WALL THICKNESS TOO THIN". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿ BARD® LATEX-FREE FOLEY TRAY WITH BARDEX¿ ALL-SILICONE FOLEY CATHETER, BARD® EZ-LOK¿ SAMPLING PORT CONTENTS: ¿ BARDEX¿ 5CC ALL-SILICONE FOLEY CATHETER 2000ML (APPROX. VOL.) CAPACITY DRAINAGE BAG WITH ANTI-REFLUX CHAMBER BARD® EZ-LOK¿ SAMPLING PORT STRING HANGER ATTACHED SHEETING CLIP CONTROL-FIT¿ OUTLET DEVICE ¿ TAMPER-EVIDENT SEAL ¿ URO-PREP¿ TRAY WITH: CATHETER LUBRICATING JELLY SYRINGE LATEX-FREE, POWDER-FREE GLOVES (2) UNDERPAD (WATERPROOF), DRAPE POVIDONE-IODINE SWAB PACKET 10CC SYRINGE (PREFILLED WITH STERILE WATER) TO INFLATE FOLEY CATHETER ONLY. ¿ SPECIMEN CONTAINER, CAP AND LABEL STERILE: STERILE UNLESS PACKAGE IS OPENED OR DAMAGED, EXCEPT FOR ANY INDIVIDUALLY PACKAGED COMPONENTS WITHIN THE TRAY WHICH ARE NOT LABELED AS STERILE. THESE COMPONENTS ARE NOT TERMINALLY STERILIZED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE SILICONE AND MAY CAUSE BALLOON TO BURST. DIRECTIONS FOR USE: VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. 1. OPEN CSR WRAP TO FORM STERILE FIELD AND PLACE UNDERPAD BENEATH PATIENT, PLASTIC SIDE DOWN. 2. PUT ON CUFFED GLOVES AND POSITION DRAPE ON PATIENT. 3. OPEN CATHETER LUBRICATING JELLY. 4. REMOVE TOP TRAY AND OPEN PLASTIC POUCH SURROUNDING CATHETER. 5. LUBRICATE CATHETER. 6. PREPARE PATIENT WITH POVIDONE-IODINE SWAB PACKET. 7. PROCEED WITH CATHETERIZATION IN USUAL MANNER. TO INFLATE CATHETER, SIMPLY INSERT TIP OF WATER-FILLED SYRINGE GENTLY INTO VALVE (DO NOT OVERPENETRATE) AND DEPRESS PLUNGER. INSTILL ENTIRE AMOUNT OF STERILE WATER (10CC). 8. POSITION HANGER ON BEDSIDE RAIL NEAR FOOT OF THE BED AND USE SHEETING CLIP TO SECURE DRAINAGE TUBE TO DRAW SHEET. 9. TO EMPTY BAG: A. REMOVE OUTLET TUBE FROM HOUSING; GENTLY SQUEEZE CONNECTOR ARMS AND PULL TUBE FROM HOUSING. B. RELEASE CLAMP AND EMPTY BAG. C. AFTER EMPTYING, RECLAMP OUTLET TUBE AND SLIDE CONNECTOR INTO HOUSING UNTIL CONNECTOR ARMS ENGAGE. 10. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. ALLOW THE BALLOON TO DEFLATE SLOWLY ON ITS OWN. DO NOT ASPIRATE OR MANUALLY ACCELERATE THE DEFLATION OF THE BALLOON. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. 11. PERIODIC OBSERVATIONS OF THIS SYSTEM SHOULD BE MADE TO ENSURE THAT URINE IS FLOWING FREELY. IF A STANDING COLUMN OF URINE IS OBSERVED, CHECK FOR CORRECT POSITIONING OF BAG AND THEN FOR A PHYSICAL OBSTRUCTION. IF CORRECT POSITIONING OR REMOVAL OF PHYSICAL OBSTRUCTION DOES NOT ALLOW FREE FLOW, THE BAG MAY HAVE TO BE CHANGED. TAMPER-EVIDENT SEAL DISCONNECT INSTRUCTIONS: IF NECESSARY TO DISCONNECT CATHETER AND DRAINAGE TUBE CONNECTOR, FLEX CATHETER AS ILLUSTRATED, GRASP TAB AND SLOWLY PULL ALONG PERFORATIONS UNTIL SECTION IS REMOVED. DIRECTIONS FOR USING BARD® EZ-LOK¿ SAMPLING PORT: BARD® EZ-LOK¿ SAMPLING PORT ACCEPTS A LUER-LOCK OR SLIP TIP SYRINGE. 1. KINK DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. 2. SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT (AT APPROXIMATELY 80-100 DEGREE ANGLE). PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. NOTE: IMPROPER PENETRATION TECHNIQUE COULD CAUSE FORMATION OF A DROP OF URINE ON THE SURFACE OF THE SAMPLING PORT. PERIODIC OBSERVATION OF THE SAMPLING PORT IS RECOMMENDED. 4. ASPIRATE DESIRED VOLUME OF URINE. 5. UNKINK TUBING AND SEND SPECIMEN TO LABORATORY NOTE: AGGRESSIVE TRACTION, PARTICULARLY IN THE PRESENCE OF SUTURING, IS NOT RECOMMENDED FOR 100% SILICONE FOLEY CATHETERS.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03665 | 2023-05-22 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT PATIENT WAS HOSPITALIZED WITH A URINARY TRACT INFECTION IN (B)(6) OF THIS YEAR. STATED THAT THEY WERE STILL USING THE INTERMITTENT CATHETERS WHICH THEY PREFERRED AND SAID THAT THEY STILL HAD 23 CATHETERS LEFT AS PATIENT WAS USING UP WHAT THEY HAD GOT FROM THEIR DOCTOR AND LIBERATOR MEDICAL SUPPLY. REPRESENTATIVE ADVISED TO CUT UP AND TOSS DUE TO THE RECALL. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03664 | 2023-05-22 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® WOVEN URETERAL CATHETER | URETERAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 36304 | 36304 |  | IT WAS REPORTED THAT THE PATIENT WAS HAVING ISSUES WITH THE URINE FLOWS SLOWLY DOWN TUBING IN MALE EXTERNAL CATHETER. PATIENT SAID THAT IT COULD BE DUE TO BENT OR KINK IN TUBING, SEDIMENT OR HOW THE BODY WAS RELEASING URINE SLOWER BUT NOT SURE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INCORRECT DIAMETER ON FORM". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: 1) VERIFY CORRECT SIZE PRIOR TO USE. 2) TRIM PUBIC HAIR IF NECESSARY. 3) WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. 4) OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. 5) PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. 6) UNROLL THE CATHETER OVER PENIS. 7) GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. 8) CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: 1) ENSURE DRAINAGE BAG IS EMPTY. 2) DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. 3) GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER.¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03661 | 2023-05-22 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53812G | 53812G |  | IT WAS REPORTED THAT THE PATIENT HAD 3 UTI (URINARY TRACT INFECTION) FROM (B)(6) 2022. PATIENT WAS UNSURE IF THE CATHETERS WERE CAUSE OF INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION PROVIDED. AS PER FOLLOW UP VIA PHONE ON 24MAY2023, IT WAS REPORTED THAT THE CUSTOMER DID INDEED DEVELOP 3 URINARY TRACT INFECTIONS BACK-TO-BACK WHILE USING THE CATHETERS. THE CUSTOMER WAS TREATED EACH TIME WITH ANTIBIOTICS AND STOPPED USE AFTER THE THIRD TIME. THE CUSTOMER STATED THAT LIBERATOR WAS ABLE TO SEND HIM A DIFFERENT TYPE OF CATHETER AND THEY HAD NOT HAD ANY ISSUES. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "CAUTION: FEDERAL (USA) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. DEVICE DESCRIPTION: THE MAGIC3 GO® INTERMITTENT URINARY CATHETER IS A SILICONE TUBE INSERTED INTO THE URETHRA FOR THE PURPOSE OF DRAINING THE BLADDER OF URINE. NOT MADE WITH NATURAL RUBBER LATEX DOES NOT CONTAIN DEHP. INDICATIONS FOR USE: THE MAGIC3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, USE THE ADHESIVE LABEL(S) TO HANG THE PACKAGE ON A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03663 | 2023-05-22 | 2023-05-05 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION FROM USING THE PUREWICK FEMALE EXTERNAL CATHETER. PATIENT WAS ON A FOLEY CATHETER AT THIS TIME AND NOT SURE WHEN THEY WOULD TAKE OFF THE FOLEY. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. PER CLARIFICATION MAIL RECEIVED FROM LIBERATOR ON 12MAY2023, NO ADDITIONAL INFORMATION PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. WARNINGS: ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. ¿ PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03660 | 2023-05-22 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD A UTI (URINARY TRACT INFECTION), AND THEY WENT TO THE DOCTOR AND GOT TREATMENT. PATIENT STATES THAT IT HAD CAME ACROSS SEVERAL OPENING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03659 | 2023-05-22 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT WAS DIAGNOSED WITH UTI( URINARY TRACT INFECTION). BUT THE PATIENT DIDN'T SPECIFY THE TIMEFRAME. PATIENT STATES THAT THEY GOT THE NOTICE OF THE SUPPLIES BEING PART OF RECALL. PATIENT HADN'T NOTICED ANY BROKEN SEALS IN THE PACKAGING. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA AND IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03655 | 2023-05-22 | 2023-04-27 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE CATHETER WAS UNABLE TO REMOVE. AFTER INSERTION, THERE WAS A LEAK OF STERILE WATER FROM THE SENSOR AREA. THINKING SOMETHING WAS WRONG, THE USER WITHDREW THE STERILE WATER FOR REINSERTING THE CATHETER, BUT ONLY 2 CC OF WATER WAS REMOVED. CHECKING IT, THE BALLOON WAS FOUND TO HAVE BEEN TRAPPED IN THE URETHRA. THE FUNNEL WAS CUT BUT COULD NOT BE DRAINED. THE PATIENT WAS TAKEN TO THE RADIOLOGY DEPARTMENT AND A GUIDE WIRE WAS PASSED THROUGH THE INFLATION LUMEN WHILE CHECKING WITH AN ECHO, AND THE WATER WAS REMOVED. ALLEGEDLY, AT THE TIME OF INSERTION, THE CATHETER WAS INSERTED UP TO THE ROOT AND THE BALLOON WAS INFLATED AFTER CONFIRMING THE DRAINAGE OF URINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03653 | 2023-05-22 | 2023-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119316M |  | IT WAS REPORTED THAT THE FOLEY CATHETER FELL OUT OF PATIENT AND BALLOON WAS NOT INFLATED . CUSTOMER KEPT OLD FOLEY AND TRIED TO INFLATE AND NOTED A LEAK AT THE Y SITE WHERE INJECTION PORT MEETS THE TUBING. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED WITH NO CUFFING OR LEAKS NOTED FROM THE UNDER THE INFLATION CAP OR THE CATHETER ITSELF. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE IP7603444, REVISION 0, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02031 | 2023-05-22 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5733690 | 5733690 |  | IT WAS REPORTED THAT FOURTEEN DAYS POST DIALYSIS CATHETER PLACEMENT, THE CATHETER WAS ALLEGEDLY FRACTURED. IT WAS FURTHER REPORTED THAT BLOOD WAS LEAKING FROM THE RUPTURED PART. REPORTEDLY, THE DEVICE WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE 19CM HEMOSPLIT D/L CATHETER WAS RETURNED FOR EVALUATION. IN ADDITION TO THE RETURNED SAMPLE, ONE ELECTRONIC PHOTO AND TWO ELECTRONIC VIDEOS WERE PROVIDED FOR REVIEW. THE PHOTO SHOWS ONE HEMOSPLIT CATHETER IMPLANTED IN A PATIENT AND A SPLIT WAS NOTED ON THE RED LUER EXTENSION LEG. THE VIDEOS SHOWS THE CATHETER IMPLANTED IN THE PATIENT BODY AND A SPLIT WAS NOTED ON THE RED LUER. GROSS, MICROSCOPIC VISUAL AND FUNCTIONAL TESTING WERE PERFORMED. TWO LONGITUDINAL SPLITS, IN-PARALLEL, WERE OBSERVED ON THE RED LUER EXTENSION LEG AND UPON INFUSION OF THE RED LUER, LEAKING WAS NOTED FROM THE PARALLEL SPLITS ON THE EXTENSION LEG. ASPIRATION OF WATER INTO THE SYRINGE WAS ATTEMPTED BUT WAS UNSUCCESSFUL; ONLY AIR WAS ASPIRATED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND FLUID LEAK ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02010 | 2023-05-22 | 2023-04-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | IMPLANTABLE PORT | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | THE FOLLOWING ADVERSE EVENTS WERE IDENTIFIED IN SEVERAL STUDIES UTILIZING ELECTRONIC HEALTH RECORD DATA PROVIDED BY MAYO CLINIC TO BD FOR THE PURPOSE OF EVALUATING CLINICAL OUTCOMES FOR BARD¿ LONG-TERM IMPLANTABLE PORTS (POWER AND NON-POWER) USED IN A CONTEMPORARY REAL-WORLD COHORT OF PATIENTS. THIS MDR WILL CAPTURE THE ARRHYTHMIA REPORTED DURING THE USE OF THE POWER PORT M.R.I. IMPLANTABLE PORT PRODUCT (UNKNOWN PRODUCT CODE). EVENTS RANGE IN DATES FROM (B)(6) 2017-(B)(6) 2022. POTENTIAL DEVICE-RELATED ADVERSE EVENTS WERE DETECTED USING THE PRESENCE OF DIAGNOSIS AND PROCEDURE CODES OCCURRING ON THE DATE OF PORT-IMPLANT OR UP TO THE DATE OF A TERMINAL EVENT. TO ENSURE PRIVACY, THE DATA HAVE UNDERGONE STATISTICAL ANONYMIZATION TO PRECLUDE REIDENTIFICATION WITH HIGH PROBABILITY. THIS DE-IDENTIFICATION PROCESS INCLUDES ALL DATES AND TIMES, WHICH ARE RANDOMLY ADJUSTED. THEREFORE, ANY INDIVIDUAL PROCEDURE, PATIENT OR EVENT INFORMATION IS NOT AVAILABLE. NO PART OR LOT NUMBERS ARE AVAILABLE, AND NO DEVICES WERE RETURNED FOR EVALUATION AT THE TIME OF THE STUDY. NO ADDITIONAL INFORMATION REGARDING THESE EVENTS IS AVAILABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLES WERE NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATIONS OF THE REPORTED EVENTS ARE INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THESE EVENTS IS UNKNOWN. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02074 | 2023-05-23 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE 4F SL BASIC TRAY | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "A COLLEAGUE TOLD ME ABOUT THE DISCOVERY OF A NON-EXTERNALIZED MANDREL JUST BEFORE THE CONNECTION OF THE CATHETER TO THE INFUSION, REMOVED WITH DIFFICULTY IN THE ABSENCE OF A "SWOLLEN" PROXIMAL END." NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02072 | 2023-05-23 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5618150 |  | IT WAS REPORTED THE GUIDE-WIRE GOT LOOSE SUCH AS CRACK. IT WAS REPORTED THAT THE USER HAD FAILED, BECAUSE THE GUIDE WIRE WAS LOOSE WHEN USING THE CATHETER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF DAMAGED GUIDEWIRE WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE GUIDEWIRE AND ONE 15CM POWER-TRIALYSIS CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE CAUSED BY RETRACTION OF THE GUIDEWIRE AGAINST THE BEVEL OF THE INTRODUCER NEEDLE. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND THE GUIDEWIRE WAS CONFIRMED TO BE BROKEN WHICH ALLOWED THE OUTER COIL WIRE TO BECOME UNRAVELED. MICROSCOPIC EXAMINATION OF THE FRACTURE SITES REVEALED THE FOLLOWING: ¿ NARROWING OF THE WIRE CROSS-SECTION NEAR THE FRACTURE SITE, WHICH IS A CHARACTERISTIC FEATURE OF A STRONG PULL ON THE WIRE. NORMAL MOVEMENT OF THE GUIDEWIRE IS AWAY FROM THE SHARPENED BEVEL AND WILL NOT DAMAGE THE WIRE. IF THE GUIDEWIRE DIRECTION IS REVERSED, THE GUIDEWIRE IS THEN PULLED AGAINST THE SHARPENED EDGE OF THE NEEDLE BEVEL AND CAN CAUSE SHEARING DAMAGE OF THE WIRE AND/OR CAUSE THE WIRE TO BECOME STUCK WITHIN THE NEEDLE. AN EXAMINATION OF THE WIRE STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03703 | 2023-05-23 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | URETHRAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5108 | 4A5108 |  | IT WAS REPORTED THAT THE CATHETER TUBING KINKED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03700 | 2023-05-23 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998806 | 998806 |  | IT WAS REPORTED THAT PRIOR TO USE IN THIS PATIENT WITH URETEROSTENOSIS, THE DOCTOR INSPECTED AND TESTED THE DEVICE AND FOUND THAT THE BALLOON LEAKED WATER. AND A NEW ONE WAS USED AND NO PROBLEM OCCURRED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03689 | 2023-05-23 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | MAGIC 3 GO 16 FR MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT MALE INTERMITTENT CATHETER AT THE TIME OF USING THE M3G THEY GOTTEN TO UTI. THEY DID SEEK MEDICAL ATTENTION WAS TREATED WITH ANTIBIOTICS LAST WEEK. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03692 | 2023-05-23 | 2022-10-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION ON (B)(6)2022 (LOT#LUGR2051) AND THEY HAD SECOND URINARY TRACT INFECTION ON (B)(6) 2022 (LOT# JURG0224). THE PATIENT STATED THAT THE PATIENT ONLY HAD ONE URINARY TRACT INFECTION. THEY CONTRIBUTED TO THE PATIENT WAS NOT USING THE INTERMITTENT CATHETER. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03690 | 2023-05-23 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | MAGIC 3 GO 16 FR MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAD AN UTI( URINARY TRACT INFECTION). IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN; WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA, REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03688 | 2023-05-23 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT STATED THAT PATIENT HAD A URINARY TRACT INFECTION AND WAS TREATED BY PATIENT'S PRIMARY CARE PHYSICIAN. PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE SYMPTOMS OR TREATMENT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03687 | 2023-05-23 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123520A | 123520A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS LEAKING. PER FOLLOW UP VIA PHONE ON (B)(6)2023, IT WAS REPORTED CUSTOMER WERE CURRENTLY GETTING THEIR SUPPLIES THROUGH ADAPTHEALTH AND THE FOLEY CATHETER BALLOON DEFLATED CONSTANTLY WHICH MAKES IT DIFFICULT TO COLLECT ANY URINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03675 | 2023-05-23 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAD URINARY TRACT INFECTION BUT COULD NOT DETERMINE IF IT WAS CAUSED BY THE MAGIC 3 GO MALE INTERMITTENT. CATHETER PATIENT STATES THEY HAD BEEN TAKING MEDICINE FOR 2 YEARS. FOR LOT#JUGT0289, FOR LOT#JUGU1969, FOR LOT#JUGU0783, FOR LOT#JUGV0665. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03676 | 2023-05-23 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50814G |  | IT WAS REPORTED THAT PATIENT HAD URINARY TRACT INFECTION BUT COULD NOT DETERMINE IF IT WAS CAUSED BY THE MAGIC 3 GO MALE INTERMITTENT. CATHETER PATIENT STATES THEY HAD BEEN TAKING MEDICINE FOR 2 YEARS. FOR LOT #JUGT0289, FOR LOT #JUGU1969, FOR LOT #JUGU0783, FOR LOT #JUGV0665. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER: WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03671 | 2023-05-23 | 2023-05-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT STOPPED USING THESE INTERMITTENT CATHETER BECAUSE THEY GAVE THEM URINARY TRACT INFECTION IN THE PAST. IT WAS STATED THAT THE PATIENT HAD TWO URINARY TRACT INFECTION ONE IN (B)(6) 2021 AND ONE (B)(6) 2022.IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. PER FOLLOW UP VIA PHONE (B)(6) 2023, CUSTOMER STATED THE PROBLEM HAS BEEN CORRECTED AND THEY REACHED OUT TO LIBERATOR TO REQUEST DIFFERENT CATHETERS. THEY DID NOT INTEND TO SEND ANYTHING BACK FOR EVALUATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
2020394-2023-00354 | 2023-05-23 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 018 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U875310 | U875310 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED AT 8 ATM AT THE FIRST INFLATION ATTEMPT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 02/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00353 | 2023-05-23 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG120124 | ATG120124 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE BILATERAL ILIAC VEIN VIA THE BILATERAL FEMORAL VEIN, THE PTA BALLOON ALLEGEDLY RUPTURED AT 8 ATM. IT WAS FURTHER REPORTED THAT THE BALLOON BODY DETACHED FROM THE CATHETER DURING REMOVAL. REPORTEDLY, SNARE WAS USED TO REMOVE THE BALLOON SEGMENT AND WITHDRAW IT WITH THE VASCULAR SHEATH TOGETHER FROM THE BODY. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ATLAS GOLD PTA DILATATION CATHETER WAS RETURNED FOR EVALUATION. THE BALLOON WAS NOTED TO HAVE DETACHED COMPLETELY FROM THE CATHETER AND NOTED TO BE PROLAPSED AND INVERTED. GLUE BULLET WAS NOTED TO BE INCORRECTLY SEATED. NO OTHER SPECIFIC ANOMALIES DURING THE VISUAL EVALUATION. NO FUNCTIONAL TESTING WAS PERFORMED DUE TO THE CONDITION OF THE DEVICE. ONE RADIO GRAPHIC VIDEO REVIEWED. THE VIDEO LOOP DEMONSTRATES BALLOON BEING DILATED AT THE END OF A ILIAC VEIN STENT. THE BALLOON RUPTURES TOWARD THE END OF MAXIMUM DILATION. THE BALLOON IS BEING DILATED IN A SELF-EXPANDING STENT AND AT THE END OF THE STENT. THE STENT STRUTS CAN DEFORM AND PUNCTURE THE BALLOON AS WELL AS OVER EXPANSION CAN CAUSE RUPTURE. AFTER RUPTURE, THE BALLOON CAN GET CAUGHT ON THE STRUTS REQUIRING RETRIEVAL. THE SUBMITTED RADIOGRAPHIC VIDEO CONFIRMS THE EVIDENCE OF BALLOON RUPTURE AS THE BALLOON RUPTURES TOWARD THE END OF MAXIMUM DILATION COULD BE OBSERVED AND THE RETURNED DEVICE FOR ANALYSIS ALSO CONFIRMS THE BALLOON IS NOTED TO BE DETACHED AND THE RETURNED BALLOON NOTED TO BE INVERTED AND PROLAPSED. HENCE, THE INVESTIGATION CONFIRMS ALL THE REPORTED EVENTS SUCH AS BALLOON RUPTURE, BALLOON DETACHMENT AND DIFFICULTY WITH REMOVAL OF THE DETACHED BALLOON RESPECTIVELY. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BALLOON RUPTURE, BALLOON DETACHMENT AND DIFFICULTY WITH REMOVAL OF THE DETACHED BALLOON COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 06/2025), G3. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02054 | 2023-05-23 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM DOME IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602870 | 0602870 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, NEEDLE ALLEGEDLY DID NOT ASPIRATE. IT WAS FURTHER REPORTED THAT SERUM ALLEGEDLY CAME OUT THROUGH A HOLE ON THE SIDE OF THE NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO AND VIDEOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 10/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02062 | 2023-05-23 | 2023-04-26 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1606052 | 1606052 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA THE INTERNAL JUGULAR VEIN, THE TIP OF THE GUIDEWIRE WAS ALLEGEDLY BROKE. IT WAS FURTHER REPORTED THAT THE BROKEN GUIDEWIRE TIP ALLEGEDLY REMAINED INSIDE THE VESSEL. THE CURRENT PATIENT STATUS IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02063 | 2023-05-23 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY FIVE YEARS POST PORT PLACEMENT VIA THE SUBCLAVIAN VEIN, THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE DISTAL CATHETER SEGMENT AND THE PORT BODY WERE REMOVED. REPORTEDLY, THE PORT WAS REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03702 | 2023-05-23 | 2023-05-05 | I, n, j, u, r, y | Manufacturer report | HYDROSIL GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 71814 | 71814 |  | IT WAS REPORTED THAT THE PATIENT HAS CONFIRMED THAT DURING THE MONTH OF APRIL, THEY CONTRACTED A UTI. IT HAS BEEN CONFIRMED THAT THE PATIENT WAS USING THE CATHETER DURING THIS MONTH. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "FILTER FAILURE". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03704 | 2023-05-23 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE CATHETERIZATION WITH INFLATED 30ML WAS PERFORMED FOR THE PATIENT. AFTER FIVE DAYS, THE CATHETER SUDDENLY FELL OUT AND THE BALLOON WAS DEFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03706 | 2023-05-23 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT PATIENT HAD BEEN EXPERIENCED PAIN AND LEAKAGE WITH THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03699 | 2023-05-23 | 2023-05-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT STOPPED USING THESE INTERMITTENT CATHETER BECAUSE THEY GAVE THEM URINARY TRACT INFECTION IN THE PAST. IT WAS STATED THAT THE PATIENT HAD TWO URINARY TRACT INFECTION ONE IN (B)(6) 2021 AND ONE (B)(6) 2022.IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. PER FOLLOW UP VIA PHONE (B)(6) 2023, CUSTOMER STATED THE PROBLEM HAS BEEN CORRECTED AND THEY REACHED OUT TO LIBERATOR TO REQUEST DIFFERENT CATHETERS. THEY DID NOT INTEND TO SEND ANYTHING BACK FOR EVALUATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02073 | 2023-05-23 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER BASIC KIT (4F) (SINGLE-LUMEN) (20CM) | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P6154118 |  | IT WAS REPORTED BY THE CUSTOMER "PLACEMENT OF THE MIDLINE CATHETER FOR WHICH THE MANDREL (USUALLY ALREADY IN PLACE IN THE CATHETER WITH A THREE-WAY CONNECTOR AND A "SWOLLEN" PROXIMAL END) IS ABSENT - MANY CHANGES OF BRANDS AND MODELS DUE TO SUPPLY DIFFICULTIES. THE CATHETER IS INSERTED USING THE CLASSIC TECHNIQUE THROUGH A PEELABLE INTRODUCER. THE CATHETER WAS RINSED. THE NEXT DAY, INJECTION DIFFICULTIES PERSISTED AFTER RINSING. AT D2 OR D3 THE CATHETER WAS REMOVED: DISCOVERY OF A MANDREL INSIDE THE CATHETER, BEING EXPELLED BY THE DISTAL END. SIGNIFICANT RISK OF INTRAVASCULAR METALLIC FOREIGN BODY. CURRENT PATIENT STATUS: NO CONSEQUENCES EXCEPT FOR LOSS OF VENOUS LINE. ACTIONS TAKEN IN THE HEALTH CARE INSTITUTION FOR THE MANAGEMENT OF THE PATIENT: WITHDRAWAL OF THE MEDICAL DEVICE, OBSERVATION OF THE DEFECT. NO PARTICULAR MEASURE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03715 | 2023-05-23 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE THE FOLEY CATHETER WAS PUT ON PATIENT. AFTER INJECTING 10 ML OF STERILE WATER AT THE CUFF, THE ENTIRE VALVE RELEASED FROM THE CATHETER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE RETURNED DEVICE WAS FUNCTIONALLY TESTED AND THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03716 | 2023-05-23 | 2023-04-26 | I, n, j, u, r, y | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED ON (B)(6). ON (B)(6), THE CATHETER WAS FOUND FALLING OUT AND THE BALLOON WAS BURST AND A PIECE MISSING, ABOUT 2*1CM. KEEPING OBSERVATION THAT THE URINATION WAS NORMAL AND NO HARM CAUSED. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 21MAY2023, IT WAS STATED THAT THE MISSING PARTS FOUND IN BLADDER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03712 | 2023-05-23 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 995081 | 995081 |  | IT WAS REPORTED VIA SURVEY RESPONSE THAT THE PHYSICIAN WAS NOT ABLE TO USE THE REFOLDING TOOL TO REFOLD THE BALLOON DILATION CATHETER BECAUSE AFTER FURTHER CASE UPDATE, THE FOLDING BALLOONS WERE NOT SUBJECTIVE FOR THE INNER LINING TRACE CAN REQUIRE ASSERTIONS AND ATTENTIVENESS TO HEAL THE DIGESTIVE TRACT OR URINALYSIS OF PREVENTION FROM PREVIOUS ADVERSE REACTIONS FROM THE PAST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03710 | 2023-05-23 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK MALE EXTERNAL CATHETER KIT | PUREWICK MALE EXTERNAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWM030 |  | IT WAS REPORTED THAT THE ADHESIVE ON THE PUREWICK MALE EXTERNAL CATHETER CAUSED A TEAR TO THE PATIENT'S TESTICLE. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03711 | 2023-05-23 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR 2 WEEKS DUE TO POOR URINATION. ON THE MORNING OF (B)(6) THE CATHETERIZATION WAS OPERATED WELL FOR THE PATIENT, FOR WHICH THE FIXATION AND DRAINAGE WAS GOOD. AT 8:00PM OF THE NIGHT , THE PATIENT COMPLAINED HEARING A SOUND OF BURST, THEN THE CATHETER FELL OUT FROM THE PATIENT AND THE BALLOON WAS BROKEN. THE CATHETER WAS REPLACED AFTER REPORTING TO THE DOCTOR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02060 | 2023-05-23 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT ALLEGEDLY SHIFTING AND TWISTING SO THEY COULDN'T BE ACCESSED. REPORTEDLY, PORT WAS REMOVED FROM PATIENT AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02075 | 2023-05-23 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2174108 |  | IT WAS REPORTED BY THE CUSTOMER "DURING USE, THE CATHETER WAS NOT REFLUXING WELL AND WHEN THEY WERE ABOUT TO REMOVE IT, THERE WAS A PARTIAL RUPTURE AT 13CM. THEY HAD TO REMOVE THE CATHETER." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02061 | 2023-05-23 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT ALLEGEDLY SHIFTING AND TWISTING SO THEY COULDN'T BE ACCESSED. REPORTEDLY, PORT WAS REMOVED FROM PATIENT AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03708 | 2023-05-23 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT AFTER OPENING THE PACKAGE, THE VALVE WAS DISCONNECTED BEFORE INSERTION INTO THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02057 | 2023-05-23 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER (20G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202252 |  | IT WAS REPORTED THAT THE CLINICIAN TOOK THE CATHETER OUT OF THE PACKAGE, AND WHEN SHE TOOK THE CAP OFF, THE NEEDLE WAS BENT WITH THE CATHETER ON IT. ADDITIONAL INFORMATION RECEIVED 2/24: CUSTOMER STATES: "I DID NOT USE IT ON A PATIENT, I JUST PULLED ANOTHER KIT AND USED THE CATHETER FROM THE SECOND ACCUCATH PACKAGE. I HAVE A FEELING THAT THE SHIPPING BOX THE PACKAGES ARE SENT IN IS BEING EXTERNALLY DAMAGED (SOME BOXES ARE DAMAGED WHEN WE RECEIVE THEM)AND IMPACTING THE CATHETERS, BUT I AM UNSURE AS NOTHING ELSE IN THE PACKAGE SHOWS DAMAGE.". IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BENT ACCUCATH NEEDLE WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 20GA X 2.25¿ ACCUCATH ACE PERIPHERAL IV CATHETER ASSEMBLY. THE SAMPLE WAS RECEIVED FULLY ASSEMBLED, WITH THE CATHETER OVERLAYING THE NEEDLE SHAFT. NO OBVIOUS USE RESIDUES WERE OBSERVED. THE WIRE EXTENDED APPROXIMATELY 1CM PAST THE NEEDLE TIP. THE SAFETY BUTTON WAS DEPRESSED AND THE NEEDLE WAS PARTIALLY WITHDRAWN INTO THE HOUSING. A SHARP BEND WAS OBSERVED IN THE NEEDLE SHAFT APPROXIMATELY 2CM FROM THE NEEDLE CARRIER. ATTEMPTS TO ADVANCE AND RETRACT THE GUIDEWIRE WERE UNSUCCESSFUL. THE WIRE WAS IMMOBILE AS IT WAS APPARENTLY CONSTRAINED BY THE BENT NEEDLE, PREVENTING FULLY SAFETY ACTIVATION. MICROSCOPIC INSPECTION OF THE NEEDLE TIP WAS UNREMARKABLE. MICROSCOPIC INSPECTION OF THE CATHETER WAS UNREMARKABLE. MICROSCOPIC INSPECTION OF THE GUIDEWIRE CONFIRMED THAT IT WAS INTACT BUT WAS OTHERWISE UNREMARKABLE. NEEDLE DEFORMATION WAS OBSERVED; HOWEVER, NO EVIDENCE WAS DISCOVERED THAT INFORMED THE CAUSE OF THE DEFORMATION. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE DEVICE MANIPULATION PRIOR TO OR DURING ATTEMPTED USE. H3: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02056 | 2023-05-23 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER (20G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202252 |  | IT WAS REPORTED THAT THE CLINICIAN TOOK THE CATHETER OUT OF THE PACKAGE, AND WHEN SHE TOOK THE CAP OFF, THE NEEDLE WAS BENT WITH THE CATHETER ON IT. ADDITIONAL INFORMATION RECEIVED 2/24: CUSTOMER STATES: "I DID NOT USE IT ON A PATIENT, I JUST PULLED ANOTHER KIT AND USED THE CATHETER FROM THE SECOND ACCUCATH PACKAGE. I HAVE A FEELING THAT THE SHIPPING BOX THE PACKAGES ARE SENT IN IS BEING EXTERNALLY DAMAGED (SOME BOXES ARE DAMAGED WHEN WE RECEIVE THEM)AND IMPACTING THE CATHETERS, BUT I AM UNSURE AS NOTHING ELSE IN THE PACKAGE SHOWS DAMAGE." IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BENT ACCUCATH NEEDLE WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 20GA X 2.25¿ ACCUCATH ACE PERIPHERAL IV CATHETER ASSEMBLY. THE SAMPLE WAS RECEIVED FULLY ASSEMBLED, WITH THE CATHETER OVERLAYING THE NEEDLE SHAFT. NO OBVIOUS USE RESIDUES WERE OBSERVED. THE WIRE EXTENDED APPROXIMATELY 1CM PAST THE NEEDLE TIP. THE SAFETY BUTTON WAS DEPRESSED AND THE NEEDLE WAS PARTIALLY WITHDRAWN INTO THE HOUSING. A SHARP BEND WAS OBSERVED IN THE NEEDLE SHAFT APPROXIMATELY 2CM FROM THE NEEDLE CARRIER. ATTEMPTS TO ADVANCE AND RETRACT THE GUIDEWIRE WERE UNSUCCESSFUL. THE WIRE WAS IMMOBILE AS IT WAS APPARENTLY CONSTRAINED BY THE BENT NEEDLE, PREVENTING FULLY SAFETY ACTIVATION. MICROSCOPIC INSPECTION OF THE NEEDLE TIP WAS UNREMARKABLE. MICROSCOPIC INSPECTION OF THE CATHETER WAS UNREMARKABLE. MICROSCOPIC INSPECTION OF THE GUIDEWIRE CONFIRMED THAT IT WAS INTACT BUT WAS OTHERWISE UNREMARKABLE. NEEDLE DEFORMATION WAS OBSERVED; HOWEVER, NO EVIDENCE WAS DISCOVERED THAT INFORMED THE CAUSE OF THE DEFORMATION. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE DEVICE MANIPULATION PRIOR TO OR DURING ATTEMPTED USE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02058 | 2023-05-23 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT ALLEGEDLY SHIFTING AND TWISTING SO THEY COULDN'T BE ACCESSED. REPORTEDLY, PORT WAS REMOVED FROM PATIENT AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03697 | 2023-05-23 | 2023-05-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT SAID THEY HAD A FEW URINARY TRACT INFECTIONS AND SAID IT WAS CAUSED BY THE AFFECTED MAGIC 3 GO INTERMITTENT CATHETER. REPRESENTATIVE ADVISED TO REPLACE THE CATHETER WITH A DIFFERENT PRODUCT. PATIENT ASKED ABOUT WHAT COURSE THEY SHOULD TAKE AND THE REPRESENTATIVE EXPLAINED IF THEY EXPERIENCING SYMPTOMS OF AN INFECTION TO GO AND SPEAK WITH THE PHYSICIAN. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03707 | 2023-05-23 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS URETHRAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE CUSTOMER HAD BEEN USING URETHRAL CATHETERS FOR ABOUT 2 YEARS. IT HAD BEEN NOTED THAT THERE WAS A RATHER INCONSISTENT AMOUNT OF LUBRICANT APPLIED TO THE CATHETER INSIDE THE COLLECTION BAG. THE MAJORITY OF THEM WERE PERFECTLY FINE, BUT SOMETIMES THERE WAS ALMOST NO LUBE AT ALL, WHICH MADE USING THE CATHETERS UNCOMFORTABLE AND DIFFICULT TO USE. ALSO STATED THAT A COUPLE TIMES PER MONTH THEY HAD EVEN BEEN NOTICING BLOOD VISIBLE IN THE CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿CHANNEL WITH INSUFFICIENT LUBRICANT¿. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® TOUCHLESS¿ PLUS UNISEX PRE-LUBRICATED URETHRAL CATHETER KIT CONTENTS: VINYL URETHRAL CATHETER 1100CC COLLECTION BAG 3 POVIDONE-IODINE SWABS 2 AMBIDEXTROUS GLOVES WATERPROOF UNDERPAD INSTRUCTIONS FOR USE: PREPARING FOR CATHETERIZATION: 1. OPEN KIT AND REMOVE CONTENTS. 2. THE GLOVES ARE NOT NECESSARY TO MAINTAIN ASEPTIC CONDITIONS DURING THE PROCEDURE. THE GLOVES ARE FOR THE OPERATOR'S PROTECTION. MALE/FEMALE CATHETERIZATION: 1. REMOVE CAP FROM INSERTION TIP WHILE SQUEEZING TABS. SAVE CAP FOR CLOSING THE BAG. 2. SLIDE THE CATHETER HALFWAY INTO THE INSERTION TIP. 3. MALE - PREPARE THE MALE URETHRAL MEATUS AND THE SURROUNDING AREA WITH POVIDONE-IODINE SWABS PROVIDED. FEMALE - PREPARE THE FEMALE URETHRAL MEATUS BY HOLDING THE OUTER LABIA APART AND PREP THE URETHRAL MEATUS AND SURROUNDING AREA WITH POVIDONE IODINE SWABS PROVIDED. 4. MALE - HOLDING THE PENIS, ADVANCE THE INSERTION TIP INTO THE URETHRA NO FURTHER THAN THE FLANGE BASE. FEMALE - SPREAD THE INNER LABIA, ADVANCE THE INSERTION TIP INTO THE URETHRA NO FURTHER THAN THE FLANGE BASE. RELEASE THE INNER LABIA. 5. PLACE YOUR NONDOMINANT HAND ON THE FINGER CONTROL GUIDE TO STABILIZE CATHETER IN URETHRA. WITH YOUR DOMINANT HAND, GRASP CATHETER THROUGH BAG APPROXIMATELY 1" BELOW THE FINGER CONTROL GUIDE AND PUSH CATHETER INTO URETHRA. THE CATHETER SHOULD BE INTRODUCED BY SHORT, REPETITIVE PUSHING MOTIONS. REPEAT MOTIONS UNTIL CATHETER REACHES BLADDER AND URINE STARTS TO FLOW. 6. ALLOW URINE TO FLOW FREELY, MAKING CERTAIN THE CATHETER GUIDE IS ELEVATED AT LEAST 4" ABOVE LOWER PORTION OF THE BAG. ALLOW FLOW UNTIL BLADDER IS EMPTY OR UNTIL BAG IS FILLED. PRECAUTION: IT IS RECOMMENDED THAT THE COLLECTION BAG BE HELD. THE CATHETER COULD POSSIBLY SEPARATE FROM THE COLLECTION BAG WHEN URINE INCREASES WEIGHT OF THE BAG. 7. WITHDRAW THE CATHETER FROM THE URETHRA. REMOVE THE REMAINING PORTION OF THE CATHETER FROM THE COLLECTION BAG BY PULLING IT THROUGH THE INSERTION TIP. NOTE: THE CATHETER IS DESIGNED TO PASS THROUGH THE INSERTION TIP. 8. THE FILLED COLLECTION BAG MAY BE CLOSED BY REPLACING THE CAP OVER THE INSERTION TIP. MAKE SURE THE CAP SNAPS ON SECURELY. SPECIMEN COLLECTION: TO COLLECT A SPECIMEN, OBTAIN A SAMPLE CUP/TUBE AND ALCOHOL WIPE. REMOVE GUIDE TIP BY PULLING TAB(S) UPWARD FROM BAG. WIPE PORT WITH ALCOHOL. POUR SPECIMEN THROUGH PORT AT TOP OF BAG INTO CUP/TUBE. DRAINING BAG: REMOVE INSERTION TIP BY PULLING TAB(S) UPWARD FROM BAG. DRAIN URINE THROUGH PORT AT TOP OF BAG. CAUTION: FEDERAL (USA) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03693 | 2023-05-23 | 2023-05-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT PATIENT STATES THAT THEY HAD FOUR UTIS( URINARY TRACT INFECTION) IN THE PAST 12 MONTHS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02059 | 2023-05-23 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT ALLEGEDLY SHIFTING AND TWISTING SO THEY COULDN'T BE ACCESSED. REPORTEDLY, PORT WAS REMOVED FROM PATIENT AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03673 | 2023-05-23 | 2023-05-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT EXPERIENCED A CHRONIC URINARY TRACT INFECTION. THEY WERE TREATED AT THE DOCTOR'S OFFICE IN MARCH. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02083 | 2023-05-24 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN USING THE MST-0668945 DURING PLACEMENT OF A THREE-WAY VALVE CATHETER IN THIS PATIENT WITH LYMPHOMA, THE HEMATOLOGY DEPARTMENT FOUND THAT THE PEEL-AWAY SHEATH HAD A JAGGED TIP. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02080 | 2023-05-24 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808000 | 1808000 |  | IT WAS REPORTED THAT APPROXIMATELY EIGHT YEARS POST PORT PLACEMENT, THE TIP OF THE CATHETER HAD ALLEGEDLY ADHERED TO THE SUPERIOR VENA CAVA VESSEL WALL. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 07/2017). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
MW5117879 | 2023-05-24 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARDEX 3-WAY URETHRAL CATHETER | CATHETER, UROLOGICAL (ANTIMICROBIAL) AND ACCESSORIES | C.R. BARD, INC. | 0167SI24 |  |  | CATHETER STOPPED DRAINING URINE - NURSE WENT TO DEFLATE BALLOON TO CHANGE THE CATHETER AT WHICH TIME NO LIQUID WAS REMOVED. THE CATHETER WAS EASILY REMOVED - NURSE THEN NOTED THE PORT TO ATTACH SYRINGE WAS DAMAGED. |  | N | Y |  |  | 2023-05
MW5117877 | 2023-05-24 | 2023-05-06 | I, n, j, u, r, y | Voluntary report | HICKMAN CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. |  |  |  | PATIENT INITIALS: (B)(6). DATE OF AWARENESS: (B)(6) 2023. PROVIDER REMS ID: (B)(6). REPORTER: PROVIDER HOSPITALIZATION START DATE: (B)(6) 2023, HOSPITALIZATION END DATE: (B)(6) 2023. 38YOF WITH PMH ON INFUSION WITH R SIDED HICKMANN CATHETER ADMITTED FOR CATHETER LINE INFECTION AND CELLULITIS. PATIENT UNDERWENT CATHETER REMOVAL IN OR AND TREATED WITH IV ANTIBIOTICS. |  | Y | N |  |  | 2023-05
MW5117867 | 2023-05-24 | 2023-05-08 | I, n, j, u, r, y | Voluntary report | BD BARD SET ROTAREX 6 F X 110CM | CATHETER, PERIPHERAL, ATHERECTOMY | BD / BARD PERIPHERAL VASCULAR, INC. |  |  |  | THE FOLLOWING IS A DESCRIPTION OF A BD BARD ROTAREX 6FR X 110CM DEVICE ERROR THAT OCCURRED DURING A PROCEDURE, INJURED THE PATIENT, AND REQUIRED ADDITIONAL INTERVENTION TO PREVENT HARMFUL DAMAGE. PATIENT SCHEDULED FOR LEFT LOWER EXTREMITY ANGIOGRAM WITH RIGHT GROIN AND LEFT DISTAL ACCESS; ROTAREX ATHERECTOMY LEFT SFA; LEFT SFA COVERED STENTING; BALLOON ANGIOPLASTY LEFT SFA; PERCLOSE CLOSURE. WHILE DOING THE ATHERECTOMY THE WIRE CAME OUT OVER THE DEVICE ON ITS OWN EXTREMELY FAST AS IF THE DEVICE SUCTIONS THE WIRE AND THEN AN ANGIOGRAM WAS DONE AT THIS STAGE WAS NOTICED THAT THERE IS DAMAGE TO THE POSTERIOR TIBIAL ARTERY WHERE THE WIRE WAS PRESENT AND DAMAGE TO THE SUPERFICIAL FEMORAL ARTERY. DIFFERENT STENTING WAS REQUIRED TO FIX THE DAMAGE CAUSED BY THE DEVICE AS WELL AS ACCESSING THE ARTERY IN TWO DIFFERENT LOCATIONS. |  | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02079 | 2023-05-24 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT MRI, 9.6 FR. SILICONE, INTER. KIT W/SUTURE PLUG, S/L | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1859600 | 1859600 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CLEAR PLASTIC PIECE ALLEGEDLY BROKE. IT WAS FURTHER REPORTED THAT THE GRAY SHEATH WAS ALLEGEDLY SPLIT HALFWAY DOWN WITHOUT THE DOCTOR PULLING ON IT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 10/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02088 | 2023-05-24 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PORT ACCESS DRESSING | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT "HAVING ISSUE WITH THE SENTRINEX DRESSING WITH PEDIATRIC PATIENTS." NO OTHER INFORMATION WAS PROVIDED. A SPECIFIC NUMBER OF AFFECTED DEVICES OR PATIENTS WAS NOT PROVIDED BY THE COMPLAINANT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02087 | 2023-05-24 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED, "ISSUE WITH THE STATLOCK TODAY. THE PIECE AT THE END DID NOT BREAK OFF, INSTEAD THE TUBING CRACKED. " ADDITIONAL INFORMATION RECEIVED ON MAY 9, 2023: TUBING CRACKED AND BLOOD BEGAN LEAKING OUT. RESULTED IN BLEEDING FROM THE ARTERIAL INSERTION. INCIDENT WAS IDENTIFIED AND TREATED QUICKLY. ARTERIAL LINES ARE COMMONLY USED IN THE INTENSIVE CARE UNIT TO MONITOR THE HEMODYNAMIC STATUS OF CRITICALLY ILL PATIENTS. THE INCIDENT CAUSED THE INTERRUPTION OF THIS MONITORING. NO DELAY IN THE DELIVERY OF MEDICATION TO THE PATIENT. THE TUBING WAS IMMEDIATELY REMOVED AND THE CONNECTION WAS MADE FROM THE PRESSURE TUBING DIRECTLY TO THE HUB OF THE ALINE CATHETER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02085 | 2023-05-24 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER W/SHERLOCK TIP LOCATION SYSTEM (4F) DL (55CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9274108 |  | IT WAS REPORTED BY CUSTOMER THAT PICC BROKE AT THE SUTURE WING. HAD TO INSERT ANOTHER PICC. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03745 | 2023-05-24 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50818G | 50818G |  | IT WAS REPORTED THAT THE PATIENT HAD SOME URINARY TRACT INFECTIONS WHEN USING THESE MAGIC 3 GO MALE COUDE CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03742 | 2023-05-24 | 2023-05-06 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MALE COUDE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT PATIENT CURRENTLY HAD A URINARY TRACT INFECTION THAT WAS DISCOVERED THREE DAYS AGO. ALSO SAID THEY HAD TO GO TO THE HOSPITAL IN EARLY (B)(6) OF 2023 FOR A URINARY TRACT INFECTION AND HAD ANOTHER URINARY TRACT INFECTION IN (B)(6) OF 2023. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03747 | 2023-05-24 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119108 | 119108 |  | IT WAS REPORTED THAT WATER COULD NOT BE REMOVED FROM THE BALLOON OF THE FOLEY CATHETER. UPON TRYING TO REMOVE THE CATHETER, THE CATHETER COULD NOT BE REMOVED BECAUSE A SPECIFIED AMOUNT OF STERILE DISTILLED WATER COULD NOT BE DRAINED. THEN THE CATHETER WAS CUT OFF AND WATER WAS DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03746 | 2023-05-24 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION, SEPSIS AND HAD TO GET THE GALL BLADDER REMOVED. IT WAS STATED THAT THE PATIENT ALMOST DIED AND DIDN'T SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. IT WAS UNKNOWN IF THE MAGIC 3 INTERMITTENT CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION, SEPSIS AND MEDICAL INTERVENTION WAS UNKNOWN AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03738 | 2023-05-24 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | MAGIC 3 GO MALE COUDE CATHETER | MAGIC 3 GO MALE COUDE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD A MILD RASH PRIOR TO RECEIVING THE RECALL NOTICE. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THEIR DOCTOR PRESCRIBED MEDICATED OINTMENT TO ALLEVIATE THE RASH THEY DEVELOPED. STATED THAT THEY DID NOT FEEL THE INTERMITTENT CATHETER CONTRIBUTED TO THAT HAPPENING AND DID NOT SPECIFY WHAT THEY WAS USING, AND DID NOT NEED TO SEND ANY SAMPLES IN FOR EVALUATION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN. WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03737 | 2023-05-24 | 2023-05-06 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT PATIENT CURRENTLY HAD A URINARY TRACT INFECTION THAT WAS DISCOVERED THREE DAYS AGO. ALSO SAID THEY HAD TO GO TO THE HOSPITAL IN (B)(6) 2023 FOR A URINARY TRACT INFECTION AND HAD ANOTHER URINARY TRACT INFECTION IN (B)(6) 2023. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03736 | 2023-05-24 | 2023-05-06 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION AND BECAME SEPTIC WHICH LED THEM TO GO TO THE HOSPITAL AND HAD TO DO REHABILITATION. REPRESENTATIVE CONFIRMED THAT THE LOT# OF THE CURRENT BOX OF CATHETERS WERE NOT LISTED ON THE RECALL LETTER. MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03732 | 2023-05-24 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998606 | 998606 |  | IT WAS REPORTED THAT THE SURGICAL NURSE FOUND THAT THE BALLOON OF THE BALLOON DILATION CATHETER LEAKED WATER PRIOR TO USE FOLLOWING UNPACKING. | THE REPORTED EVENT WAS CONFIRMED AND CAUSE WAS UNKNOWN. PHOTO SAMPLES WERE SUBMITTED. EVALUATION OF THE FIRST PHOTO SAMPLE NOTED ONLY THE PACKAGING WITH THE PRODUCT IDENTIFIERS PRESENT. THE SECOND PHOTO NOTED THE EAGLE INFLATION DEVICE, HOWEVER, COULD NOT BE EVALUATED FOR THE REPORTED FAILURE. VISUAL EVALUATION NOTED, NO DAMAGES WERE OBSERVED IN THE BALLOON, HOWEVER IT WAS IDENTIFIED A CRACK IN THE BALLOON HUB CONNECTOR. FUNCTIONAL EVALUATION NOTED A 0.038" GUIDEWIRE WAS INTRODUCED THROUGH THE WIRE HUB AND IT WAS NOTICED THAT THE INNER TUBE IS COLLAPSED IN THE BODY OF THE BALLOON, NOT ALLOWING THE FULL PASSAGE OF THE GUIDEWIRE. THEN THE CATHETER WAS INFLATED WITH DISTILLED WATER USING AN INFLATION DEVICE, THE BALLOON WAS ABLE TO BE INFLATED NEVERTHELESS IT WAS NOT POSSIBLE TO PRESSURIZE IT TO OPERATIONAL PRESSURE DUE TO THE LEAK IN THE BALLOON HUB. THE BALLOON HUB IS COMPLETE ONLY WITH A SMALL CRACK CLOSE TO THE WINGS AND THERE IS NO EVIDENCE THAT IT HAS BEEN DEFORMED BY A CRUSH, THEREFORE IS LIKELY THAT THE CONNECTOR HAS BEEN OVER TORQUED. MANUFACTURING PROCESS WAS REVIEWED, AND IT WAS CONFIRMED THAT ONLY PLASTIC CONNECTORS ARE USED TO CONNECT BALLOON HUB FOR LEAK TEST AND FOLDING. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE LEAKS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INSPECTION PRIOR TO USE: THE X-FORCE® BALLOON DILATION CATHETER IS A STERILE, SINGLE USE DEVICE. CAREFULLY INSPECT THE CATHETER AND THE STERILE PACKAGING FOR SIGNS OF DAMAGE THAT MAY HAVE OCCURRED DURING SHIPMENT. DO NOT USE THE PRODUCT IF DAMAGE IS EVIDENT. INFLATING THE BALLOON CATHETER: WARNING: DO NOT EXCEED THE RECOMMENDED RATED BURST PRESSURE (RBP) FOR THIS DEVICE. BALLOON RUPTURE MAY OCCUR IF THE RBP RATING IS EXCEEDED. PLEASE REFER TO THE DEVICE LABEL FOR THE RECOMMENDED MAXIMUM RATED BURST PRESSURE. EXTREME CAUTION SHOULD BE USED WHEN CONSIDERING DILATION OF IRREGULAR, NOT-COMPLIANT TISSUE OR IN THE PRESENCE OF CERTAIN CALCULI.¿ H11: THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02078 | 2023-05-24 | 2023-03-06 | I, n, j, u, r, y | Manufacturer report | POWERPORT DUO M.R.I. IMPLANTABLE PORT, CHRONOFLEX DUAL-LUMEN, 9.5F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1829500 | 1829500 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED SKIN REACTION. IT WAS FURTHER REPORTED THAT ANTIBIOTICS WERE PRESCRIBED FOR TREATMENT. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 07/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03730 | 2023-05-24 | 2023-05-06 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MALE COUDE INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT PATIENT CURRENTLY HAD A URINARY TRACT INFECTION THAT WAS DISCOVERED THREE DAYS AGO. ALSO SAID THEY HAD TO GO TO THE HOSPITAL IN EARLY (B)(6) OF 2023 FOR A URINARY TRACT INFECTION AND HAD ANOTHER URINARY TRACT INFECTION IN (B)(6) OF 2023. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN. WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02076 | 2023-05-24 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833350 | 5833350 |  | IT WAS REPORTED THAT DURING THE PREPARATION OF DIALYSIS CATHETER PLACEMENT PROCEDURE, UPON UNPACKING OF THE CATHETER, A PROBLEM WAS ALLEGEDLY FOUND WITH THE SHEATH WHICH HAD UNEVEN RUBBER. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE UNSEALED HEMOSTAR D/L KIT WAS RETURNED FOR EVALUATION AND ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. GROSS VISUAL AND FUNCTIONAL TESTING WERE PERFORMED. DURING VISUAL TESTING BUNCHING WAS NOTED THROUGHOUT SHEATH SHAFT AND MATERIAL DEFORMATION WAS ALSO NOTED AT THE DISTAL END OF THE PEEL-APART SHEATH AND THE EDGES OF THE INTRODUCER SHEATH WAS NOTED TO BE UNEVEN FROM THE PROVIDED PHOTO. MANUFACTURING REVIEW WAS PERFORMED, AND IT CONFIRMS TOP CAP DETACHED. THEREFORE, THE WHOLE INVESTIGATION IS CONFIRMED FOR THE REPORTED DEFORMATION AND IDENTIFIED DETACHMENT ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 12/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00359 | 2023-05-24 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U35130410 | U35130410 |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN UPON OPENING THE PACKAGE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ULTRAVERSE 35 DILATATION CATHETER WAS RETURNED FOR EVALUATION. ON THE VISUAL EVALUATION, A KINK WAS OBSERVED ON THE CATHETER SHAFT AND NO OTHER ANOMALIES NOTED, NO CATHETER BREAK WAS NOTED. NO FUNCTIONAL PERFORMED DUE TO CONDITION OF DEVICE. ALSO, TWO PHOTOS WERE REVIEWED. THE FIRST PHOTO SHOWS THE CATHETER WITHIN A PLASTIC POUCH, AND THE LABEL WAS OF THE DEVICE, IN WHICH THE PRODUCT DETAILS COULD BE VERIFIED. THE SECOND PHOTO SHOWS THE CATHETER MARKED AT A SPECIFIC LOCATION WHICH SHOWS THE CATHETER KINK. NO EVIDENCE OF A BREAK WAS NOTED IN THE CATHETER; NO OTHER ANOMALIES WERE NOTED. THERE IS NO OBJECTIVE EVIDENCE OF A CATHETER BREAK NOTED ON THE CATHETER, BUT A CATHETER KINK CAN BE OBSERVED ON THE SUBMITTED PHOTO AND THE RETURNED SAMPLE AS CONFIRMS THE SAME. HENCE, THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED CATHETER BREAK AND CONFIRMED FOR THE IDENTIFIED CATHETER KINK. A DEFINITIVE ROOT CAUSE FOR THE REPORTED CATHETER BREAK AND IDENTIFIED CATHETER KINK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. EXPIRATION DATE: 05/2025. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02077 | 2023-05-24 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK APHERESIS PORT ACCESS SYS |  | IT WAS REPORTED THAT APPROXIMATELY FOUR YEARS POST PORT PLACEMENT, THE PORT WAS ALLEGEDLY UNABLE TO BE ACCESSED. IT WAS FURTHER REPORTED THAT THE PORT WAS ALLEGEDLY UNABLE TO BE CANNULATED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03729 | 2023-05-24 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS URETHRAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE CUSTOMER HAD BEEN USING URETHRAL CATHETERS FOR ABOUT 2 YEARS. IT HAD BEEN NOTED THAT THERE WAS A RATHER INCONSISTENT AMOUNT OF LUBRICANT APPLIED TO THE CATHETER INSIDE THE COLLECTION BAG. THE MAJORITY OF THEM WERE PERFECTLY FINE, BUT SOMETIMES THERE WAS ALMOST NO LUBE AT ALL, WHICH MADE USING THE CATHETERS UNCOMFORTABLE AND DIFFICULT TO USE. ALSO STATED THAT A COUPLE TIMES PER MONTH THEY HAD EVEN BEEN NOTICING BLOOD VISIBLE IN THE CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿CHANNEL WITH INSUFFICIENT LUBRICANT¿. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03717 | 2023-05-24 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166V14S |  | IT WAS REPORTED THAT ON (B)(6), 2023, THE CATHETERIZATION WAS PLANNED TO BE PERFORMED BEFORE THE OPERATION. DURING THE PROCESS OF INJECTING LIQUID INTO THE BALLOON, LEAKAGE WAS FOUND AND A RESISTANCE WAS FELT, SO THE NURSE COULD NOT PUSH THE SYRINGE. ACCORDING TO THE EVENT CAUSE ANALYSIS PROVIDED THE BALLOON PORT WAS FLAWED AND THE SYRINGE NIPPLE COULDN'T BE INSERTED INTO PLACE. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 06MAY2023, THE CUSTOMER STATED THAT THE DESIGN OF THE BALLOON PORT WAS NOT PERFECT AND SO THE SYRINGE TIP WAS NOT EASY TO INSERT INTO THE APPROPRIATE POSITION OF THE VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00363 | 2023-05-24 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 014 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | US14150510 | US14150510 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE SUPERFICIAL FEMORAL ARTERY VIA A CROSS OVER APPROACH, THE GUIDEWIRE WAS ALLEGEDLY STUCK AROUND THE COMMON FEMORAL ARTERY. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03719 | 2023-05-24 | 2023-05-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION FROM THESE INTERMITTENT CATHETER. THE PATIENT DID NOT INDICATE THE TIMEFRAME BUT THE PATIENT SAID DID SEEK TREATMENT AND WAS PRESCRIBED MEDICATION. THERE ARE NO SAMPLES AVAILABLE FOR RETURN AS THE PATIENT DISCARDED THE CATHETERS IN ACCORDANCE WITH THE RECALL NOTICE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
2020394-2023-00361 | 2023-05-24 | 2023-04-12 | I, n, j, u, r, y | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG120146 | ATG120146 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE ILIAC VEIN, THE PTA BALLOON ALLEGEDLY RUPTURED AND THE RUPTURED BALLOON FELL INTO THE FEMORAL VEIN. IT WAS FURTHER REPORTED THAT THE INCISION OF THE FEMORAL VEIN WAS MADE TO REMOVE THE RUPTURED BALLOON. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 01/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03721 | 2023-05-24 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT WHEN THE NURSE WAS DISCONTINUING THE FOLEY, THEY ATTACHED THE SYRINGE TO EMPTY THE BALLOON AND IT DID NOT EMPTY UNTIL THEY APPLIED VERY LIGHT MINIMAL PRESSURE TO IT AND ONLY 8 ML CAME OUT WITH THE GENTLE PRESSURE ON THE SYRINGE. FOLEY DID NOT PULL OUT EASILY. THEY PULLED ON THE SYRINGE LIGHTLY AND WAS ABLE TO GET 9 ML OUT BUT NOT THE 10 THAT IS STATED IN THE LDA. THEY THEN GENTLY PULLED ON THE FOLEY AND IT SLID OUT WITH NO RESISTANCE UNTIL THE MEATUS WHERE IT STOPPED. THEY PUT MINIMAL PRESSURE SUPPORT ON THE MEATUS AT THE OPENING AND WAS ABLE TO ROTATE THE FOLEY AROUND AND OUT. THEY NOTED A RIDGE ON THE END OF THE BALLOON, IT DID NOT COMPLETELY EMPTY FLATLY. PATIENT COMPLAINED OF PAIN BUT IT RESOLVED WITHIN A FEW MINUTES. PER FOLLOW UP VIA EMAIL ON 20MAY2023, THE PATIENT WAS DISCHARGED FROM THE HOSPITAL WITH NO APPARENT COMPLICATIONS FROM EVENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03723 | 2023-05-24 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION AFTER USING THE FIRST BATCH OF THE INTERMITTENT CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS ¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS; NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03725 | 2023-05-24 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE MESSY, HARD TO OPEN, NOT ENOUGH LUBRICANT, AND WERE PAINFUL WHEN REMOVING. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03726 | 2023-05-24 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 265728 | 0165L28 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS LEAKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03727 | 2023-05-24 | 2023-05-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. THEY STATED THAT THEY HAD THE INTERMITTENT CATHETER AND HAD BEEN GETTING URINARY TRACT INFECTIONS. IT WAS UNKNOWN IF THE INTERMITTENT CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03728 | 2023-05-24 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® COUDE OLIVE TIP URETHRAL FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 802520 | 802520 |  | IT WAS REPORTED THAT THE PATIENT HAD BEEN USING CATHETERS AND HAVING PROBLEMS SINCE LAST YEAR, TIP OF CATHETER VARY IN SIZE, THE SKINNY TIP WAS SUCCESSFUL, THE BIGGER TIP DID NOT FIT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03720 | 2023-05-24 | 2022-09-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION ON (B)(6) 2022 (LOT#LUGR2051) AND THEY HAD SECOND URINARY TRACT INFECTION ON (B)(6) 2022 (LOT# JURG0224). THE PATIENT STATED THAT THE PATIENT ONLY HAD ONE URINARY TRACT INFECTION. THEY CONTRIBUTED TO THE PATIENT WAS NOT USING THE INTERMITTENT CATHETER. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02108 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F FULL TRAY (DL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2275108 |  | IT WAS REPORTED BY THE CUSTOMER THAT "WE ARE HAVING ISSUES WITH A LEAKING STOPPER ON THE DL PICC LINES. THE SIDE THAT HAS THE STYLET IN IT LEAKS WHEN IT IS FLUSHED. TODAY WE HAD TO OPEN A SL KIT, REPLACE THE STYLET WITH A NON LEAKING STOPPER AND THEN INSERT THE DL PICC LINE". ADDITIONAL INFORMATION RECEIVED ON 15-MAY-2023. THIS INVOLVED MULTIPLE PATIENTS WITH DIFFERENT DEMOGRAPHICS. THE EVENTS DIRECTLY INVOLVED A PATIENT AND USER AND DID NOT INVOLVE AN URGENT MEDICAL SITUATION. PATIENT HARM OR INJURY DID NOT OCCUR AS WE WERE ABLE TO INTERVENE BEFORE INSERTING THE PICC LINES. THE EVENTS DID NOT INTERRUPT THE ADMINISTRATION OF MEDICATION. A CHEST X-RAY WAS ORDERED TO CONFIRM PLACEMENT AS WE WERE NOT SURE IF THE SL WIRE WOULD WORK IN DL PICC. WE ALSO HAS TO OPEN A SL AND DL PICC LINE. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02103 | 2023-05-25 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108D |  | IT WAS REPORTED BY THE CUSTOMER THAT "THE 3FR INTRODUCER IN THE PROVENA MIDLINE KIT BUNCHED UP WHEN TRYING TO ADVANCE." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02102 | 2023-05-25 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED, "ISSUE WITH THE STATLOCK. THE HARD PLASTIC END DETACHED COMPLETELY FROM THE TUBING.". ADDITIONAL INFORMATION RECEIVED ON 9-MAY-2023. "RESULTED IN BLEEDING FROM THE ARTERIAL INSERTION. FORTUNATELY, WAS IDENTIFIED AND TREATED QUICKLY. THE FAILURE OF THE TUBING WHICH WAS CONNECTED TO THE PATIENT¿S ARTERIAL LINE CAUSED BLEEDING FROM THE ARTERIAL SITE. THIS CAUSED EACH PATIENT TO BLEED FROM THEIR ARTERY. FORTUNATELY, THE ISSUE WAS IDENTIFIED BEFORE HARM CAME TO THE PATIENT. ARTERIAL LINES ARE COMMONLY USED IN THE INTENSIVE CARE UNIT TO MONITOR THE HEMODYNAMIC STATUS OF CRITICALLY ILL PATIENTS. THE INCIDENT CAUSED THE INTERRUPTION OF THIS MONITORING. NO DELAY IN THE DELIVERY OF MEDICATION TO THE PATIENT WAS AFFECTED TO THE BEST OF MY KNOWLEDGE. THE TUBING WAS IMMEDIATELY REMOVED AND THE CONNECTION WAS MADE FROM THE PRESSURE TUBING DIRECTLY TO THE HUB OF THE ALINE CATHETER.". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02101 | 2023-05-25 | 2023-05-01 | I, n, j, u, r, y | Manufacturer report | FULL KIT 20G X 10 CM WITH PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | F220108PT |  | IT WAS REPORTED THAT DURING THE PLACEMENT OF A POWERGLIDE PRO RT 20G 10CM THE USER EXPERIENCES A SIGNIFICANT EVENT. THE USER "SUCCESSFULLY ACCESSED THE VESSEL AND SMOOTHLY ADVANCED THE WIRE WITH NO RESISTANCE, THE USER THEN BEGAN TO ADVANCE THE CATHETER BUT FELT RESISTANCE SHORTLY AFTER. THE USER THEN RETRACTED THE CATHETER BACK ONTO THE NEEDLE AND REATTEMPTED. THE USER EVENTUALLY HAD TO REMOVE THE ENTIRE DEVICE AND UPON DOING SO NOTICED THE LOWER 1-2 CM OF THE CATHETER WERE NO LONGER ATTACHED AND REMAINED IN THE PATIENT. THE LOWER PORTION OF THE CATHETER IS NOW LODGED IN THE PATIENT TISSUE NOT VESSEL SO THE SURGICAL TEAM DECIDED NOT TO REMOVE THE FRAGMENT AS IT POSED NO ADDITIONAL HARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02106 | 2023-05-25 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD DL POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT "NURSE CALLED IN REGARDING THE CLAMP BROKE ON THE DUAL LUMEN POWER PICC. THIS OCCURRED ABOUT A MONTH AGO.". ADDITIONAL INFORMATION RECEIVED MAY 11, 2023: "THE CLAMP WAS UNCLAMPED TO CHANGE NEEDLESS CONNECTOR AND THE CLAMP BROKE. THE PRODUCT WAS A DOUBLE LUMEN POWER PICC IN R CHEST. NO HARM DONE OR EXPERIENCED BY PATIENT.". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02109 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT,4FR POWERMIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P4154108 |  | IT WAS REPORTED THAT WIRE SHEARED BEFORE PLACING IN PATIENT. ADDITIONAL INFORMATION RECEIVED THE EVENT WAS DISCOVERED BY THE PERSON PERFORMING THE PROCEDURES WITNESSING THEM IN REAL TIME. NO REPORTED NEGATIVE OUTCOMES, JUST HAD TO GET NEW KITS AND START AGAIN. THERE WAS NOT A SERIOUS DELAY IN TREATMENT AS THEY WERE NOT EMERGENT LINE PLACEMENTS. THE ONE WAS FOR DE-ESCALATING PICCS PER OUR PROTOCOL SO THE PATIENT HAD AN ALTERNATE LINE. I CANNOT RECALL THE OTHER PATIENT BUT THE ONLY DELAY WAS GOING TO GET NEW KITS. THIS FILE ADDRESSES THE DE-ESCALATING PICC PER FACILITY PROTOCOL. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02107 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F FULL TRAY (DL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2275108 |  | IT WAS REPORTED BY THE CUSTOMER THAT "WE ARE HAVING ISSUES WITH A LEAKING STOPPER ON THE DL PICC LINES. THE SIDE THAT HAS THE STYLET IN IT LEAKS WHEN IT IS FLUSHED. TODAY WE HAD TO OPEN A SL KIT, REPLACE THE STYLET WITH A NON LEAKING STOPPER AND THEN INSERT THE DL PICC LINE". ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023 THIS INVOLVED MULTIPLE PATIENTS WITH DIFFERENT DEMOGRAPHICS. THE EVENTS DIRECTLY INVOLVED A PATIENT AND USER AND DID NOT INVOLVE AN URGENT MEDICAL SITUATION. PATIENT HARM OR INJURY DID NOT OCCUR AS WE WERE ABLE TO INTERVENE BEFORE INSERTING THE PICC LINES. THE EVENTS DID NOT INTERRUPT THE ADMINISTRATION OF MEDICATION. A CHEST X-RAY WAS ORDERED TO CONFIRM PLACEMENT AS WE WERE NOT SURE IF THE SL WIRE WOULD WORK IN DL PICC. WE ALSO HAS TO OPEN A SL AND DL PICC LINE. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02105 | 2023-05-25 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1295108D |  | IT WAS REPORTED, PICC WAS BROKEN IN HALF AT THE SUTURE WING. THE REST OF THE PICC WAS INTACT, BUT THE EXTENSION LEGS AND LUMEN WERE LYING ON THE FLOOR. THEY BELIEVE THE PATIENT COULD HAVE INTERFERED WITH THE PICC. HAD TO INSERT ANOTHER PICC." IT WAS STATED THERE WAS NO HARM TO THE PATIENT. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02117 | 2023-05-25 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | CUSTOMER REPORTED THAT THE TIP OF THE STYLET INSIDE THE POWERPICC CATHETER WAS FOUND BIFURCATED IN THE LAST STEP OF STYLET REMOVAL DURING CATHETER PLACEMENT. NO OBVIOUS RESISTANCE WAS FELT DURING REMOVAL OF THE GUIDE WIRE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BREAK IN THE STYLET IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. ONE STYLET WAS RETURNED FOR EVALUATION. A MICROSCOPIC OBSERVATION REVEALED THE BREAK IN THE CORE WIRE OF THE RETURNED STYLET TO HAVE A FLAT APPEARANCE WHERE THE BREAK OCCURRED, AND THE BREAK SITE OF THE CORE WIRE APPEARED GRANULAR. A PROFILE VIEW OF THE CORE WIRE REVEALED THE WIRE TO HAVE A TAPER. THE COMPLAINT OF A BREAK IN THE STYLET IS CONFIRMED AND WAS DETERMINED TO BE DUE TO A TENSILE FAILURE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02112 | 2023-05-25 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | POWERPICC CATHETER 4F SL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000775A |  | IT WAS REPORTED, "SITUATION WHERE THE PATIENT COMPLAINED OF FEELING HOT AND THEN DEVELOPED RED BLOTCHY SKIN AND HIVES IMMEDIATELY FOLLOWING THE PICC PLACEMENT. THE PICC POSITION WAS CONFIRMED VIA THE SHERLOCK 3CG TIP LOCATION SYSTEM. THE PATIENT THEN BECAME UNRESPONSIVE AND A CODE BLUE WAS INITIATED. ONCE ROSC WAS ACHIEVED THE PICC WAS REMOVED IMMEDIATELY AND THE PATIENT'S CONDITION BEGAN IMPROVING." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02111 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED THAT FRACTURED PICC AT HUB. ADDITIONAL INFORMATION: THIS EVENT DID NOT INVOLVE AN URGENT MEDICAL SITUATION. IT WAS DISCOVERED DURING A DRESSING CHANGE, AN EXTERNAL FRACTURE WAS FOUND TO BE LEAKING, CAUSING THE NEED TO CHANGE DRESSINGS. THE PICC WAS REMOVED AND REPLACED WITH A MIDLINE. SURGERY WAS NOT REQUIRED TO REMOVE THE BROKEN PIECES, AND ALL WERE ACCOUNTED FOR AND REMOVED SUCCESSFULLY. THE EVENT DID NOT INTERRUPT OR DELAY TREATMENT. LSL STERILE CENTRAL LINE KIT WAS USED WITH PICC/CVC SECUREMENT TEGADERM CHG. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A FRACTURE IN THE CATHETER IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. ONE 4FR SL POWERPICC CATHETER WAS RETURNED FOR EVALUATION. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED IN THE CATHETER BODY ADJACENT TO THE MOLDED SUTURE WING. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED. FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR. 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER. OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING) . AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02114 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIG JAPAN HANG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405J |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS IMPLANTED, RESISTANCE WAS FELT DURING INFUSION, AND THE MEDICINE COULD NOT BE INFUSED. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER UNABLE TO INFUSE IS CONFIRMED AND WAS DETERMINED TO BE MANUFACTURING RELATED. ONE 4 FR NXT CLEARVUE GROSHONG CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED NO BLOOD USE RESIDUES THROUGHOUT THE RETURNED CATHETER, AND THE STYLET WAS RETURNED IN THE CATHETER. NOTHING REMARKABLE WAS OBSERVED DURING AN INITIAL VISUAL OBSERVATION. A FUNCTIONAL TEST OF ATTEMPTING TO INFUSE WATER INTO THE CATHETER USING A 12 ML SYRINGE REVEALED THE CATHETER WAS NOT PATENT TO INFUSION, AND THE WATER DID NOT MOVE THROUGH THE CATHETER BEYOND THE WHITE LUER. A MICROSCOPIC OBSERVATION REVEALED THE LUER WAS BLOCKED BY EXCESS MOLDED PRODUCT OBSERVED WITHIN THE DEVICE. THE CANNULA OF THE WHITE LUER APPEARED TO BE BLOCKED BY A SMALL, CYLINDRICAL PIECE OF MOLDED PLASTIC. THE COMPLAINT OF AN OCCLUDED CATHETER IS CONFIRMED AND WAS DETERMINED TO BE RELATED TO A MANUFACTURING DEFECT. THE INVESTIGATION WAS FORWARDED TO THE MANUFACTURING FACILITY FOR FURTHER EVALUATION, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02115 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED THAT DURING PLACEMENT OF THE PICC CATHETER, THE CUSTOMER OPENED THE MST AND FOUND THAT THE TIP OF THE VASCULAR SHEATH WAS RUPTURED, UNABLE TO BE INSERTED PROPERLY. ANOTHER NEW MST WAS OPENED FOR PLACEMENT OF THE CATHETER. THE PATIENT DID NOT EXPERIENCE ANY SYMPTOMS. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A SPLIT MICROINTRODUCER SHEATH WAS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE 4.5 FR X 5 CM MICROEZ MICROINTRODUCER WAS RETURNED FOR EVALUATION. ONE PHOTO SAMPLE OF A MICROINTRODUCER WAS ALSO PROVIDED AND THE CONDITION OF THE SHEATH CORRESPONDED WITH THE RETURNED PHYSICAL SAMPLE. THE SHAFT WAS OBSERVED TO BE CURVED AND THE DILATOR CONTAINED A SHARP BEND. BLOOD RESIDUE AND EVIDENCE OF USE WAS NOTED. MICROSCOPIC INSPECTION OF THE DISTAL END OF THE SHEATH REVEALED A LONGITUDINAL SPLIT. THE SPLIT EDGES WERE OBSERVED TO BE SLIGHTLY UNEVEN. FIBERS OF THE GREY SHEATH MATERIAL WERE STILL ATTACHED BETWEEN THE TWO EDGES THROUGHOUT THE LENGTH OF THE SPLIT. NO APPARENT ISSUES WITH THE SHEATH TIP TAPER WERE NOTED. THE PRESENCE OF THE CURVED DILATOR SHEATH AND DAMAGE TO THE SHEATH SUGGEST THAT A DIFFICULT INSERTION MAY HAVE BEEN EXPERIENCED. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE STEEP INSERTION ANGLE, ADVANCEMENT AGAINST RESISTANCE, AND DAMAGE TO MATERIAL PRIOR TO USE. SINCE A SPLIT IN THE SHEATH WAS FOUND ON THE DEVICE, THE COMPLAINT IS CONFIRMED BUT THE EXACT CONTRIBUTING FACTORS REMAIN UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02104 | 2023-05-25 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 18G X 10CM WITH PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F218108PT |  | IT WAS REPORTED, "EXTENSION SET IS LEAKING, HAD TO CLEAN UP MESS AND REPLACE." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02116 | 2023-05-25 | 2023-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 18G X 10 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | F318108PT |  | IT WAS REPORTED BY THE CUSTOMER THAT "A MIDLINE WAS PLACED WITH SMOOTH INSERTION, UPON REMOVAL OF THE NEEDLE AND GUIDEWIRE IT WAS NOTED THAT THE ENTIRE GUIDEWIRE WAS NOT INTACT, HOWEVER THERE WAS NO EVIDENCE OF THE BROKEN WIRE. MD NOTIFIED AND X-RAY TAKEN, NOTHING FOUND IN X-RAY." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02118 | 2023-05-25 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER WHEN CONDUCTING PUNCTURE FOR THE GROSHONG CATHETER, THE SPECIALIST NURSE OPENED THE SELDINGER KIT AND PULLED OUT THE PROXIMAL END OF THE GUIDE WIRE AND FOUND EVIDENT SLITS. THE SELDINGER KIT WAS REPLACED TO CONTINUE THE PUNCTURE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02120 | 2023-05-25 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REPLACEMENT CONNECTOR FOR 4F GROSHONG NXT CLEARVUE PICC WITH EXTENSION LEG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7812400 |  | CUSTOMER REPORTED THAT THE PATIENT WAS ADMITTED TO HOSPITAL FOR "APLASTIC ANEMIA". WHEN ADMITTED, THE PATIENT HAD INDENTURED PICC CATHETER, AND THE PIPELINE WAS UNOBSTRUCTED. THE INFUSION TREATMENT WAS PERFORMED AT 9:00 ON (B)(6) 2023. DURING THE CATHETERIZATION BEFORE THE USE OF THE CATHETER, NURSE INSPECTED THE PERIPHERAL INTUBATION CENTRAL VENOUS CATHETER: THE APPEARANCE WAS INTACT AND THERE WAS NO DAMAGE. WHEN THE PULSE WAS PUSHED IN, IT WAS FOUND THAT THE TRANSPARENT TUBE OF THE ACCESSORY PART OF THE PICC CATHETER BULKED UP AFTER SALINE INJECTION. THE CATHETER WAS EASILY BROKEN, SO THE USE WAS STOPPED IMMEDIATELY, AND THE NURSE REPLACED IT WITH A NEW ONE, WHICH DID NOT AFFECT THE PATIENT. NO PATIENT HARM REPORTED, NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02123 | 2023-05-25 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | BASIC KIT 20G X 8 CM WITH GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120087T |  | IT WAS REPORTED, WRITER WAS PLACING A MIDLINE ON PATIENT. WRITER WAS USING POWERGLIDE PRO 20G 8CM. REF F120087T LOT REGW1332 WRITER PLACED MIDLINE IN VESSEL, ADVANCED THE GUIDEWIRE WITH NO RESISTANCE, THEN PROCEEDED TO ADVANCE THE CATHETER BUT MET RESISTANCE ONCE ABOUT 4 CM WAS ADVANCED. WRITER REMOVED THE MIDLINE, AND THE PATIENT WAS FOUND TO STILL HAVE ABOUT 4 CM OF CATHETER STILL IN THEIR ARM. MD NOTIFIED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02121 | 2023-05-25 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 3275118 |  | IT WAS REPORTED THAT THE TIP OF THE STYLET INSIDE THE POWERPICC CATHETER (3275118) WAS FOUND BIFURCATED IN THE LAST STEP OF STYLET REMOVAL DURING CATHETER PLACEMENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02122 | 2023-05-25 | 2023-05-07 | I, n, j, u, r, y | Manufacturer report | FULL KIT 20G X 8 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F320088PT |  | IT WAS REPORTED VAT RN ATTEMPTED TO PLACE 8 CM POWERGLIDE MIDLINE. REPORTED TO MANAGER THAT MIDLINE MET RESISTANCE AND WHEN VAT RN REMOVED DEVICE NOTICED TIP OF CATHETER NOT INTACT. FOLLOWED UP WITH POLICIES AND PROCEDURES FOR POSSIBLE DEVICE FRACTURE. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02125 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIG JAPAN HANG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405J |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS FLUSHED, THE EXTENSION TUBE OF THE CATHETER CONNECTOR WAS INFLATED LIKE A BALLOON. NO RESISTANCE WAS FELT TO BOTH POSITIVE AND NEGATIVE PRESSURE AND IT WAS ABLE TO USE. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02128 | 2023-05-25 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE CORD IS PULLING OUT OF THE SENSOR/EXPOSED WIRING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF CORD IS PULLING OUT OF THE SENSOR/EXPOSED WIRING WAS CONFIRMED. THE USB CABLE IS DISCOLORED AND HAS A CUT IN THE CABLE INSULATION, EXPOSING THE BRAIDED GROUND WIRING. THESE ISSUES ARE DUE TO DEVICE USE. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02110 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT,4FR POWERMIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P4154108 |  | IT WAS REPORTED THAT WIRE SHEARED BEFORE PLACING IN PATIENT. ADDITIONAL INFORMATION RECEIVED: THE EVENT WAS DISCOVERED BY THE PERSON PERFORMING THE PROCEDURES WITNESSING THEM IN REAL TIME. NO REPORTED NEGATIVE OUTCOMES, JUST HAD TO GET NEW KITS AND START AGAIN. THERE WAS NOT A SERIOUS DELAY IN TREATMENT AS THEY WERE NOT EMERGENT LINE PLACEMENTS. THE ONE WAS FOR DE-ESCALATING PICCS PER OUR PROTOCOL SO THE PATIENT HAD AN ALTERNATE LINE. I CANNOT RECALL THE OTHER PATIENT BUT THE ONLY DELAY WAS GOING TO GET NEW KITS. THIS FILE ADDRESSES THE DELAY IN PROCEDURE FOR MIDLINE PLACEMENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02099 | 2023-05-25 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 3174118J |  | IT WAS REPORTED THAT THE PATIENT HAD A POWERPICC IMPLANTED VIA THE LEFT BASILIC VEIN AND THE CATHETER WAS PLACED IN THE AXILLARY VEIN ON (B)(6) 2023. ON (B)(6) 2023, BLOODING AND EXUDATE WERE OBSERVED FROM THE INSERTION SITE. ON (B)(6) 2023, BLOOD RETURN COULD NOT BE CONFIRMED. THEREFORE, THE DEVICE WAS REMOVED FROM THE PATIENT¿S BODY. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. THIS REPORT ADDRESSES THE EVENT OCCURRING (B)(6) 2023. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING CATHETER IS UNCONFIRMED BECAUSE THE ALLEGED PROBLEM COULD NOT BE REPRODUCED. ONE 4 FR S/L POWERPICC WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED LITTLE USE RESIDUES THROUGHOUT THE RETURNED DEVICE, AND NOTHING REMARKABLE WAS OBSERVED UPON AN INITIAL VISUAL OBSERVATION. A FUNCTIONAL TEST OF ATTEMPTING TO INFUSE AND ASPIRATE WATER INTO THE DEVICE USING A 12 ML SYRINGE REVEALED THE CATHETER TO BE PATENT TO INFUSION AND ASPIRATION WITH NO OBSERVED LEAKS. THE DEVICE WAS THEN PRESSURIZED, AND AN ATTEMPT TO INFUSE WATER INTO THE CATHETER DID NOT REVEAL ANY LEAKS. DUE TO THE INABILITY TO RECREATE THE STATED FAILURE, THE COMPLAINT OF A LEAKING CATHETER CANNOT BE CONFIRMED. POSSIBLE CAUSES OF THE OBSERVED REPORTED EVENT MAY INCLUDE TIP LOCATION AND FIBROUS ENCAPSULATION OF THE CATHETER. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02094 | 2023-05-25 | 2023-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK CV PLUS SLIDING POST, TRICOT HYDROCOLLOID CRESCENT PAD, PIGTAIL PAD | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | CV0220CE |  | IT WAS REPORTED THAT WHEN PERFORMING DAILY MAINTENANCE OF THE PICC CATHETER IN THIS PATIENT, THE TUMOR DEPARTMENT OPENED THE FIRST INDIVIDUALLY-PACKAGED STATLOCK AND FOUND THAT THE POST WAS SEPARATED AND WAS NOT ADHERED TO THE BASE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02098 | 2023-05-25 | 2023-05-07 | I, n, j, u, r, y | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | Z097FO |  | IT WAS REPORTED THAT THE PATIENT WAS INSERTED 40CC IABP INTRAOP TO ASSIST WITH CORONARY PERFUSION. IABP WAS INSERTED WITH TOE GUIDANCE AND TIPS CONFIRMED IN THE CORRECT POSITION. THE PROXIMAL END OF IABP WAS SECURED IN PLACE WITH STITCH. THE DISTAL PART WAS SECURED WITH PROVIDED CLIPS (AS OPPOSED TO SUTURING IN PLACE). CXR POST OP - IABP WAS IN THE CORRECT POSITION AND WORKED WELL. POST OP DAY 2 - OVERNIGHT, THERE WAS AN INCREASED LACTATE AND DECREASED URINE OUTPUT. THE PATIENT EXPERIENCED HYPOGLYCEMIA, DERANGED LIVER, RENAL FUNCTION AND RAISED AMYLASE. THERE WAS A CONCERN ABOUT BOWEL ISCHEMIA. THE PATIENT'S ABDOMINAL CT WAS PERFORMED AND THERE WAS NO OBVIOUS ISCHEMIA, HOWEVER IABP WAS FOUND IN THE ABDOMINAL AORTA. THE PATIENT WAS INSPECTED. THE DISTAL CLIP OF THE IABP WAS DISCONNECTED, LEADING TO THE IABP SLIPPING DOWN. IABP WAS REMOVED AND AFTER THE PATIENT'S HEALTH IMPROVED (URINE OUTPUT INCREASED, LFTS NORMALIZED). | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00365 | 2023-05-25 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 014 | PTA BALLOON DILATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | US14150530 | US14150530 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE SUPERFICIAL FEMORAL ARTERY, THE PTA BALLOON WAS ALLEGEDLY DIFFICULT TO DEFLATE AFTER INFLATION AT FIRST INFLATION ATTEMPT AT NOMINAL PRESSURE. IT WAS FURTHER REPORTED THAT THE BALLOON WAS DEFLATED GRADUALLY. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 09/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02089 | 2023-05-25 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT APPROXIMATELY ONE MONTH POST DIALYSIS CATHETER PLACEMENT, THE EXTENSION LEG WAS ALLEGEDLY BROKEN AND LEAKED. THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 05/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03748 | 2023-05-25 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THERE WAS STRONG RESISTANCE WHEN DRAINING WATER FROM THE FOLEY CATHETER AND THE WATER COULD NOT BE DRAINED. WHEN THE CATHETER WAS LEFT IN PLACE, IT CAME OUT AFTER A WHILE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) SUCCESSFULLY AND WITH ONLY SLIGHT RESISTANCE USING IN-HOUSE LUER LOCK SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. DEFLATED BALLOON PASSIVELY IN 2 MINUTES AND 31 SECONDS. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00366 | 2023-05-25 | 2023-04-28 | I, n, j, u, r, y | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U3513066 | U3513066 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY DETACHED AND FELL INTO THE HEART. IT WAS FURTHER REPORTED THE CATHETER WAS REMOVED FROM THE BODY. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 10/2024). | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
17001623 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDEX® ALL-SILICONE FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | 165816 | 165816 |  | BARD/BARDEX FOLEY CATHETER PACKAGING IDENTIFIES THAT THE BALLOON IS 5CC'S. THE CATHETER IDENTIFIES THE BALLOON AS BEING 10CC. PRODUCTS NOT USED ON A PATIENT, NO KNOWN PATIENT HARM. 4 CATHETERS WITH THIS ISSUE HAVE BEEN ISOLATED. MANUFACTURER RESPONSE FOR CATHETER, RETENTION TYPE, BALLOON, BARDEX® ALL-SILICONE FOLEY CATHETER (PER SITE REPORTER) WILL OBTAIN. MANUFACTURER RESPONSE FOR CATHETER, COUDE, BARDEX® LUBRICATH® FOLEY CATHETER (PER SITE REPORTER) WILL OBTAIN. MANUFACTURER RESPONSE FOR CATHETER, COUDE, BARD® FOLEY CATHETER SILICONE COATED (PER SITE REPORTER) WILL OBTAIN. MANUFACTURER RESPONSE FOR CATHETER, RETENTION TYPE, BALLOON, BARDEX® ALL-SILICONE FOLEY CATHETER (PER SITE REPORTER) WILL OBTAIN. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03752 | 2023-05-25 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. IT WAS STATED THAT THE PATIENT WAS HOSPITALIZED WITH A BAD URINARY TRACT INFECTION WHILE USING THE MAGIC 3 GO FEMALE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE CUSTOMER STATED THAT THEY DID NOT HAVE COMPLAINT BUT WANTED TO SWITCH CATHETERS DUE TO THE PREVIOUS ONES NOT FITTING. THE CUSTOMER WAS ABLE TO SWITCH CATHETERS AND THEY STATED THAT THEY WERE WORKING WELL . | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INDICATIONS FOR USE THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA, REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL DISPOSAL. 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03755 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THERE WAS NO FOLEY CATHETER IN THE FOLEY TRAY. | THE REPORTED EVENT WAS CONFIRMED ¿ UNKNOWN CAUSE. BASED ON THE EVALUATION, OBSERVED DRAINAGE BAG, FORCEPS, GAUZE, COTTON BALLS, SYRINGE WERE RETURNED. OBSERVED PVI, JELLY, CATHETER, SAMPLE PORT WERE NOT RETURNED. OBSERVED DRAINAGE BAG, GAUZE, SYRINGE WERE NOT USED WHILE COTTON BALL HAS ALREADY BEEN USED. EVALUATED THE INLET TUBE AND OBSERVED NO EVIDENCE OF ASSEMBLING SAMPLE PORT IN THE INLET TUBE. HOWEVER, THE EXACT OF CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS POTENTIAL FAILURE MODE COULD BE DUE TO OPERATOR'S HANDLING METHOD. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX INTENDED USE & EFFECT- EFFICACY: THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03753 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE CATHETER BALLOON RUPTURED. THERE WAS A URINE LEAKAGE FOUND DURING THE CATHETER USE ON THE PATIENT AND UPON CHECKING THE CATHETER, THE BALLOON WAS FOUND TO BE RUPTURED. | THE REPORTED EVENT WAS CONFIRMED AND UNKNOWN CAUSE. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE SAMPLE WAS EVALUATED AND VISUAL INSPECTION NOTED SAC BURST ALONG LUMEN WITHOUT MISSING PIECES. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03760 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT CUSTOMER NOTICED WHILE OPENING THE BOX OF INTERMITTENT CATHETER WITH LOT# JUGX1192, CONTENTS APPEARED TO BE PACKED IMPROPERLY AND ALSO REPORTED THAT THE NON-INSERTION END OF THE CATHETERS ARE JAGGED AND DID NOT LUBRICATE LIKE THEY SHOULD AND CAUSED MINOR BLEEDING AND DISCOMFORT. CUSTOMER SERVICE HAD BEEN NOTIFIED BECAUSE THEY WERE UNABLE TO TRANSFER TO THEM THIS MORNING AND THEY HAVE BEEN ASKED TO CONTACT PATIENT TO DO AN EXCHANGE OF THIS PRODUCT. PATIENT HAD USED A FEW OUT OF ONE BOX AND HAS AN ENTIRE UNOPENED BOX OF THE SAME LOT#. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03763 | 2023-05-25 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE VALVE WHEN INJECTING STERILE WATER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03761 | 2023-05-25 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT UPON INJECTING STERILE WATER, THE BASE OF THE INFLATION FUNNEL SWELLED INSTEAD OF THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03762 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 996101 | 996101 |  | IT WAS REPORTED THAT A PHYSICIAN IN AN ONLINE SURVEY STATED THAT THEY WERE NOT ABLE TO SUCCESSFULLY DILATE THE NEPHROSTOMY TRACT. THE PHYSICIAN WAS NOT ABLE TO SUCCESSFULLY DILATE THE NEPHROSTOMY TRACT FOR THE DEVICE 996101 - BALLOON DILATION CATHETER BECAUSE OF PERSONAL EXPECTATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03765 | 2023-05-25 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD BLOOD IN URINE AND DOCTOR INDICATED THEY HAD AN INFECTION WHILE USING THE MAGIC 3 GO MALE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. ". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03764 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT STERILE WATER LEAKED BETWEEN THE SYRINGE AND THE VALVE OF THE FOLEY CATHETER WHEN THE USER ATTEMPTED TO INJECT STERILE WATER DURING THE PRETEST. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. A DHR REVIEW IS NOT REQUIRED FOR THIS COMPLAINT. A RISK REVIEW AND LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02092 | 2023-05-25 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 24CM (SHORT-TERM) (DUAL-LUMEN) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | NA | 5593240 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN USING DIALYSIS CATHETER FOR PATIENTS WITH CARDIOPLASMIC SHOCK SEPSIS, THE GUIDE WIRE WAS BENT IN MANY PLACES AND THE TIP OF THE SHEATH WAS EASILY BROKEN. NO PATIENT HARM REPORTED, NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03768 | 2023-05-25 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT AFTER INSERTION ON 20TH AND ATTEMPTED REMOVAL THE NEXT DAY ON 21ST, 10 ML OF STERILE WATER WAS WITHDRAWN, BUT THE TIP OF THE CATHETER WAS STUCK AND COULD NOT BE REMOVED. IT WAS LUBRICATED WITH GEL AND REMOVED WITH LIGHT PRESSURE. ON CHECKING THE CATHETER, THE LOWER PART WAS FOUND BENT. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON 28APR2023, STATED THAT THE BALLOON WAS MUSHROOMED AT THE BOTTOM. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THE FAILURE COULD BE "INCORRECT CATHETER POSITIONING / RESTRAINT". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03773 | 2023-05-25 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | URETHRAL CATHETER KIT | C.R. BARD INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE ISSUES WITH THE INTERMITTENT CATHETERS. THE URINE WOULD NOT COME OUT AND IT WAS CLOGGED. NO IODINE CAME IN PACKAGE AND ONLY ONE GLOVE CAME IN PACKAGE. PER FOLLOW UP VIA EMAIL ON 02MAY2023, CUSTOMER RECEIVED THE MEDICAL SUPPLIES, BUT WAS STILL MISSING THE CATHETERS. CUSTOMER SAID THEY HAD BEEN RUNNING OUT OF THE SUPPLIES WHILE AWAITING FOR THE REPLACEMENTS. PER FOLLOW VIA PHONE ON 09MAY2023, IT WAS REPORTED THAT CUSTOMER WAS STILL WORKING WITH CUSTOMER SERVICE TO GET REPLACEMENT CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿INCORRECT OPERATION¿. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03776 | 2023-05-25 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION. THE PATIENT WAS UNSURE IF THE MAGIC 3 INTERMITTENT CATHETER CAUSED IT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL. 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. " | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03777 | 2023-05-25 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT STATED THAT THE PATIENT RECEIVED 2 PLUS URINARY TRACT INFECTION THROUGHOUT THE YEAR WHILE USING INTERMITTENT CATHETER. PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03782 | 2023-05-25 | 2023-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PATIENT WAS PLACED WITH AN INDWELLING FOLEY CATHETER CONNECTED TO THE BLADDER FOR CONTINUOUS IRRIGATION ON (B)(6) 2023. AFTER STOPPING THE BLADDER IRRIGATION ON (B)(6) 2023. THE PATIENT'S CATHETER FELL OUT, AND THE BALLOON WAS FOUND TO BE RUPTURED. WITHOUT CATHETERIZATION, THE PATIENT COULD URINATE BY THEMSELVES. AS PER DESCRIPTION OF EVENT CAUSE ANALYSIS, STATED THAT IT MIGHT BE DUE TO THE QUALITY OF THE CATHETER, OR IT MIGHT BE DUE TO THE HIGH PRESSURE OF THE BALLOON. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 19MAY2023, IT WAS STATED THAT THERE NO MISSING PIECES OF BALLOON INSIDE THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02095 | 2023-05-25 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS 3CG, GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "DL PICC KINKING IN PATIENT". ADDITIONAL INFORMATION RECEIVED 05/08/2023: CUSTOMER REPORTED THIS WAS THE FIRST OF TWO KINKED CATHETERS WITH THE SAME LOT NUMBER. THIS PARTICULAR PATIENT WAS AN INPATIENT. VAS WAS CONSULTED X 2 TO ASSESS CATHETER BECAUSE RN WAS HAVING DIFFICULTY INFUSING. CHEST X-RAY RECOMMENDED. KINK NOTED IN ARM. PICC PULLED BACK AND WORKING. PT MEDS OFF SCHEDULE. REQUIRED MULTIPLE ASSESSMENTS, MANIPULATION AND TESTING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A KINKED CATHETER IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. TWO RADIOGRAPHIC IMAGES OF WHAT APPEARS TO BE AN IMPLANTED CATHETER WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE RADIOGRAPHIC IMAGES SHOWED A LINE WHICH APPEARS TO BE AN IMPLANTED CATHETER. THE LINE IN ONE IMAGE APPEARS TO BE EVENLY PLACED. THE LINE IN THE OTHER IMAGE APPEARS TO BEND AT AN ANGLE, SUGGESTING THE CATHETER MAY BE KINKED WITHIN THE PATIENT. WHILE THE EXACT CAUSE OF THE OBSERVED KINKED CATHETER COULD NOT BE DETERMINED FROM THE RETURNED PHOTOGRAPH, POSSIBLE CAUSES OF THE KINKED CATHETER INCLUDE WEAKENED MATERIAL DUE TO FLEXURAL FATIGUE, PLACEMENT TECHNIQUE, SECUREMENT TECHNIQUE, AND PATIENT ANATOMICAL FEATURES. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02093 | 2023-05-25 | 2022-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT ICU OF SURGERY DEPARTMENT PLACED A 3275118 IN THIS PATIENT WITH SEVERE TRAUMA. DURING INFUSION IN THE SAME AFTERNOON, THE CONNECTION BETWEEN ONE PURPLE END AND THE FITTING LEAKED LIQUID. THE FITTING WAS REPLACED TWO DIFFERENT FITTINGS, BUT LIQUID LEAKS CONTINUED. CONSIDERING THE CONTINUOUS USE OF THE NEWLY PLACED CATHETER, THE CLINICIAN CLAMPED THE PURPLE END WITH A THUMBSCREW AND ONLY USED THE OTHER LUMEN TO MAINTAIN INFUSION TO DATE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02096 | 2023-05-25 | 2022-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT ICU OF SURGERY DEPARTMENT PLACED A 3275118 IN THIS SEVERELY ILL PATIENT IN COMA. DURING INFUSION AT THAT NIGHT, THE CONNECTION BETWEEN ONE PURPLE END AND THE FITTING LEAKED LIQUID. THE FITTING WAS REPLACED TWO DIFFERENT FITTINGS, BUT LIQUID LEAKS CONTINUED. CONSIDERING THE CONTINUOUS USE OF THE NEWLY PLACED CATHETER, THE CLINICIAN CLAMPED THE PURPLE END WITH A THUMBSCREW AND ONLY USED THE OTHER END TO MAINTAIN INFUSION TO DATE. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A HUB LEAK IS INCONCLUSIVE DUE TO POOR SAMPLE CONDITION. ONE PHOTO SAMPLE OF A 4 FR POWERPICC CATHETER WAS PROVIDED FOR EVALUATION. THE LUER HUBS ARE BEING SHOWN IN THE IMAGE. BOTH LUER HUBS ARE ATTACHED TO EXTENSION SET CONNECTORS AND THE PURPLE HUB IS CIRCLED. NO APPARENT DAMAGE ON THE LUER HUBS WAS NOTED AND THE CHARACTERISTICS COULD NOT BE INSPECTED FROM THE IMAGE. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE DAMAGED LUER CONNECTOR AND CONNECTION INTERFERENCE. SINCE THE REPORTED EVENT OF A LEAK WAS NOT OBSERVED FROM THE IMAGE PROVIDED, THE COMPLAINT COULD NOT BE CONFIRMED AND REMAINS INCONCLUSIVE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02097 | 2023-05-25 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THIS PATIENT HAD THE PICC CATHETER PLACED IN PICC CLINIC OF THE HOSPITAL ON (B)(6) 2023 DUE TO THE NEED FOR INFUSION. ON APRIL 3, THE PATIENT VISITED THE PICC CLINIC DUE TO PAIN AND ARM SWELLING. WHEN CONDUCTING CATHETER MAINTENANCE, THE NURSE FOUND SEEPAGE AT THE INSERTION SITE AND THICKENED CIRCUMFERENCE OF THE ARM. UPON RETRACTION, THE CATHETER WAS FOUND DAMAGED. THE CATHETER WAS THEN TRIMMED AND SECURED FOR CONTINUOUS USE. ON MAY 3, THE CATHETER WAS REMOVED DUE TO INTOLERABLE PAIN WITH THE AFFECTED ARM AND SEEPAGE AT THE INSERTION SITE. A BREACH WAS FOUND WITH THE CATHETER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER LEAK IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. TWO PHOTO SAMPLES AND A VIDEO OF A 4 FR GROSHONG NXT CATHETER WERE PROVIDED FOR EVALUATION. THE FIRST TWO PHOTOS SHOWED THE CATHETER OUTSIDE OF PATIENT USE. THE CATHETER IS BEING BENT AND A CIRCUMFERENTIAL SPLIT APPEARS TO BE VISIBLE APPROXIMATELY 5 CM FROM THE PROXIMAL END. THE VIDEO SAMPLE APPEARS TO SHOW THE CONNECTOR AND CATHETER TRIMMED NEAR THE DAMAGED REGION; HOWEVER, THE CONDITION COULD NOT BE CLEARLY DETERMINED FROM THE INFORMATION PROVIDED. SINCE A BREAK IN THE CATHETER WAS FOUND TO BE VISIBLE ON THE CATHETER, THE COMPLAINT IS CONFIRMED BUT THE EXACT FACTORS REMAIN UNKNOWN. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03783 | 2023-05-25 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD BEEN DEALING WITH SPERM TUBE INFECTION FOR 2 MONTHS FROM USING THESE MAGIC3 GO MALE COUDE CATHETER AND IT CAUSED THE PATIENT'S TESTICLES TO SWELL AND BECAME VERY SORE AND THE PATIENT SAID THAT THE SPERM HAD BECOME INFECTED. PATIENT SPOKE WITH THE DOCTOR ABOUT THIS AND THE DOCTOR TOLD THAT IT WAS NOT ANYTHING PATIENT DID AND SAID THE PATIENT RECEIVED THAT THEY RECEIVED A RECALL LETTER STATED THAT IT WAS NOT ANYTHING FROM THE PATIENT SIDE BUT IT WAS THE CATHETERS THAT WERE DEFECTIVE. PATIENT RECEIVED THESE CATHETERS FROM HOME HEALTH AGENCY. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02100 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 3174118J |  | IT WAS REPORTED THAT THE PATIENT HAD A POWERPICC IMPLANTED VIA THE LEFT BASILIC VEIN AND THE CATHETER WAS PLACED IN THE AXILLARY VEIN ON (B)(6) 2023. ON (B)(6) 2023, BLOODING AND EXUDATE WERE OBSERVED FROM THE INSERTION SITE. ON (B)(6) 2023, BLOOD RETURN COULD NOT BE CONFIRMED. THEREFORE, THE DEVICE WAS REMOVED FROM THE PATIENT¿S BODY. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. THIS REPORT ADDRESSES THE EVENT OCCURRING (B)(6) 2023. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03772 | 2023-05-25 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT CUSTOMER NOTICED WHILE OPENING THE BOX OF INTERMITTENT CATHETER WITH LOT# JUGX1192, CONTENTS APPEARED TO BE PACKED IMPROPERLY AND ALSO REPORTED THAT THE NON-INSERTION END OF THE CATHETERS ARE JAGGED AND DID NOT LUBRICATE LIKE THEY SHOULD AND CAUSED MINOR BLEEDING AND DISCOMFORT. CUSTOMER SERVICE HAD BEEN NOTIFIED BECAUSE THEY WERE UNABLE TO TRANSFER TO THEM THIS MORNING AND THEY HAVE BEEN ASKED TO CONTACT PATIENT TO DO AN EXCHANGE OF THIS PRODUCT. PATIENT HAD USED A FEW OUT OF ONE BOX AND HAS AN ENTIRE UNOPENED BOX OF THE SAME LOT#. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02124 | 2023-05-25 | 2023-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D4 |  | IT WAS REPORTED THERE IS SALINE SEEPING OUT OF RUBBER STOPPER. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02090 | 2023-05-25 | 2023-03-27 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1618000 | 1618000 |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR NINE MONTHS AND SIX DAYS POST PORT PLACEMENT, THE PORT ALLEGEDLY HAD DIFFICULTY IN FLUSHING. IT WAS FURTHER REPORTED THAT THE PORT ALLEGEDLY HAD SUCTION ISSUE. REPORTEDLY, UPON COMPUTED TOMOGRAPHY WITH CONTRAST SHOWED THE TIP ALLEGEDLY HAD A FIBROUS FLAP. THE CURRENT STATUS OF THE PATIENT IS REPORTED TO BE STABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 03/2022). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02091 | 2023-05-25 | 2023-04-14 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1618000 | 1618000 |  | IT WAS REPORTED THAT APPROXIMATELY TWO WEEKS POST PORT PLACEMENT PROCEDURE, THE ACCESS OF THE PORT WAS ALLEGEDLY ATTEMPTED BUT UNSUCCESSFUL AND THE PORT ALLEGEDLY HAD SUCTION PROBLEM. IT WAS FURTHER REPORTED THAT THE UPON X-RAY THE PORT ALLEGEDLY WAS DISPLACED WITH ITS TIP NOTED ABOUT THE RIGHT LOWER NECK. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 04/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03790 | 2023-05-26 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 365714 | 365714 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS DEFECTIVE. THE CATHETER FELL OUT OF PATIENT. NO APPARENT HARM REACHED THE PATIENT. THEY HAD INCONVENIENCE. STATED THAT THE PATIENT REQUIRED UNEXPECTED OR PROLONGED CARE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03789 | 2023-05-26 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATED THAT URINARY TRACT INFECTION DID DEVELOP. PATIENT WAS GIVEN ANTIBIOTICS 10-DAY REGIMEN. PATIENT STOPPED USING INTERMITTENT CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS. NONE KNOWN. WARNINGS. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS. INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS. URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03786 | 2023-05-26 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT GOT URINARY TRACT INFECTION WHILE USING THESE MAGIC 3 MALE CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. THE PATIENT REPORTED ISSUES WITH LOT# JUGT1000 , LOT# JUGS0046. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03788 | 2023-05-26 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123622A | 123622A |  | IT WAS REPORTED THAT DURING CHEMO MEDICATION IT WAS SEEN POWER PICC CATHETER 5FR DL LUMEN CAP IS BROKEN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03784 | 2023-05-26 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123514A | 123514A |  | IT WAS REPORTED THAT THE FOLEY CATHETER CAUSED PATIENT A URINARY TRACT INFECTION THAT NEEDED HOSPITALIZATION. STATED THAT THEY WERE WAITING ON THE DOCTOR TO DETERMINE THE NEXT STEPS AND PATIENT NOW HAD FOLEY FROM THE HOSPITAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03787 | 2023-05-26 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 365714 | 365714 |  | IT WAS REPORTED THAT THE FOLEY BALLOON WAS DEFECTIVE AND THE CATHETER FELL OUT OF THE PATIENT. THERE WAS AN UNEXPECTED OR PROLONGED CARE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00370 | 2023-05-26 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE TIP OF THE CATHETER ALLEGEDLY SEPARATED. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN D2 AND G4. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 11/2024). THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02134 | 2023-05-26 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0606150J |  | IT WAS REPORTED THAT THE DURING A PORT PLACEMENT PROCEDURE VIA THE RIGHT UPPER ARM, BLEEDING WAS ALLEGEDLY OCCURRED. HOWEVER, POST PORT PLACEMENT NO BLEEDING WAS OBSERVED. IT WAS FURTHER REPORTED THAT BLEEDING AND LEAKAGE WERE ALLEGEDLY OBSERVED FROM THE INSERTION SITE ON THE SAME DAY. REPORTEDLY THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 05/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02130 | 2023-05-26 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | (01)00801741098246 | 7707540J |  | IT WAS REPORTED THAT PRIOR TO A PORT PLACEMENT PROCEDURE, THE HUBER NEEDLE HAD A WHITE CAP ON THE PART WITH THE YELLOW CAP AND NOTHING ON THE PART WITH THE WHITE CAP. THERE WAS NO PATIENT CONTACT. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE PORT INFUSION SET ATTACHED TO A TERUMO SYRINGE WAS RECEIVED FOR EVALUATION. VISUAL EVALUATION WAS PERFORMED. NO CAP WAS RECEIVED WITH THE INFUSION SET AND NO SAFETY MECHANISM WAS NOTED ON THE NEEDLE. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED COMPONENT MISSING AND COMPONENT MISASSEMBLED ISSUES, AS THE ORIGINAL STATE OF THE PACKAGE AT THE TIME OF OPENING IS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3, H6 (METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03814 | 2023-05-26 | 2023-05-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT PATIENT HAD URINARY TRACT INFECTION FOR 2 MONTHS. PATIENT WAS CURRENTLY ON MAINTENANCE ANTIBIOTICS TO TREAT IT AND THE STATED THAT ABOUT 10 WEEKS OF TREATMENT LEFT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS. NONE KNOWN. WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS. URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. ". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03816 | 2023-05-26 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT PATIENT DID NOT SPECIFY THE TIME FRAME OF THE URINARY TRACT INFECTION. THE PATENT STATES THAT THEY WERE IN HOSPITAL FOR THREE DAYS WITH URINARY TRACT INFECTION AND UNSURE IF IT WAS BECAUSE OF THE CATHETERS. THE REPRESENTATIVE APOLOGIZED AND ADVISED THAT THEY WOULD STILL DOCUMENT IT SINCE WE DO HAVE THIS PRODUCT BEING RECALLED. PER FOLLOW-UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE CUSTOMER DEVELOPED URINARY TRACT INFECTION AND SOUGHT MEDICAL INTERVENTION. CUSTOMER WAS PRESCRIBED ANTIBIOTICS AND LIBERATOR REPLACED CATHETERS OF DIFFERENT TYPE WHICH HAD CAUSED NO ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03817 | 2023-05-26 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A BAD BACTERIAL INFECTION THAT WILL NOT GO AWAY, AND THEY WONDER WHETHER IT WAS DUE TO THE RECALL. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03815 | 2023-05-26 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT STATED THAT PATIENT BEEN GETTING URINARY TRACT INFECTIONS FOR A YEAR NOW AND FOR THE LAST THREE MONTHS HAS HAD REALLY BAD URINARY TRACT INFECTIONS. PATIENT STATED THAT PATIENT DID SEEK TREATMENT WITH HER DOCTOR, AND PATIENT DOCTOR DID NOT KNOW WHY THE PATIENT WAS GETTING THE URINARY TRACT INFECTIONS. PATIENT ALSO STATED THAT NOW THAT PATIENT RECEIVED THE RECALL LETTER, PATIENT WILL LET PATIENTS DOCTOR KNOW THAT'S WHAT CAUSED THEM. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03820 | 2023-05-26 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS PLUS CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT CUSTOMER RECENTLY SWITCHED THEIR CATHETERS FROM THE RUBBER TO THE VINYL ONES ON THIS ORDER. THEY STATED THAT A PATIENT WAS BLEEDING WHEN THEY TRIED TO INSERT THE CATHETER. NO MEDICAL INTERVENTION REQUIRED. (MATERIAL# 4A5144, MATERIAL# 4A5142) PER FOLLOW UP VIA MAIL ON (B)(6) 2023, THE INCIDENT THAT CUSTOMER WAS NOTIFIED OF WAS THE ONLY ONE AND THEY DO NOT HAVE ANY OTHER INFORMATION EXCEPT THAT THE PATIENT WAS BLEEDING AND WAS UNCOMFORTABLE WHEN NURSES WERE INSERTING THE CATHETER. CUSTOMER WAS JUST ASKED TO SEE IF THERE WAS AN ALTERNATIVE TO THE VINYL BESIDES THE RUBBER. THEY HAVE NOT RECEIVED ANY MORE FEEDBACK FROM NURSING REGARDING THESE CATHETERS. THEY WERE KEPT A CASE OF THE RUBBER CATHETERS JUST IN CASE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03822 | 2023-05-26 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT 5 OF THE MAGIC 3 HYDROPHILIC MALE INTERMITTENT CATHETERS DID NOT HAVE ENOUGH WATER IN THE BURST PACKET TO FULLY LUBRICATE THE CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02148 | 2023-05-26 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS CATHETER KIT | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | NA | 5613150 |  | IT WAS REPORTED BY THE CUSTOMER THAT DURING INSERTION OF POWER TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS CATHETER, THE CATHETER GOT HUNG UP ON THE GUIDEWIRE DURING THE NORMAL EXCHANGE. INITIALLY, PROVIDER THOUGHT THERE WAS A PROBLEM WITH THE CATHETER, LEFT THE GUIDEWIRE IN PLACE AND REQUESTED A SECOND CATHETER. THE SECOND CATHETER ALSO HUNG UP ON THE GUIDEWIRE IN THE SAME WAY AS THE FIRST CATHETER. UPON FURTHER INSPECTION, THE GUIDEWIRE IS FRAYED AT THE DISTAL END. NO IMMEDIATE HARM TO PATIENT DETECTED. NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED 05/18/2023: THE EVENT OCCURRED DURING A NON-EMERGENT TRIALYSIS CATHETER INSERTION FOR RENAL REPLACEMENT THERAPY. EVENT DISCOVERED THE INSERTING PROVIDER NOTED THE FRAYED GUIDEWIRE WHILE ATTEMPTING TO REMOVE THE GUIDEWIRE AFTER THREADING THE CATHETER INTO PLACE. THE WIRE REMAINED IN 1 PIECE, NO FRAGMENTS NOTED. NO COMPLICATIONS NOTED. NO PATIENT HARM. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03823 | 2023-05-26 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM BETWEEN THE SYRINGE AND THE VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03824 | 2023-05-26 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT WHEN THE SYRINGE WAS ATTACHED TO THE INFLATION VALVE TO INJECT STERILE WATER, THE WATER LEAKED. WHEN INJECTED WITH ANOTHER SYRINGE, IT WAS ABLE TO BE INJECTED. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. PACKAGING LOTS NUMBER MANUFACTURED WITH AFFECTED SYRINGE BATCHES IDENTIFIED AND DOCUMENTED. THE DHR REVIEW IS NOT REQUIRED AND THE RISK REVIEW AND LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. CORRECTIONS: F, H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
17009702 | 2023-05-26 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | CK000746 | CK000746 |  | DURING PICC (PERIPHERALLY INSERTED CENTRAL CATHETER) LINE PLACEMENT, STAFF REPORTED THE FOLLOWING, THE VEIN WAS ACCESSED AND THEN THE WIRE WAS INSERTED THROUGH THE NEEDLE PER PROCEDURE. THE WIRE DID NOT THREAD PROPERLY. DIFFICULTY MET WHEN ATTEMPTING TO REMOVE THE WIRE. STAFF NOTED THE NEEDLE DISLODGED AND THE WIRE REMAINED IN THE TISSUE. WIRE WAS RESISTANT TO REMOVAL AND DURING REMOVAL, IT BEGAN TO UNBRAID. X-RAYS WERE OBTAINED TO CONFIRM THERE WAS NO RETAINED FOREIGN BODY AS A RESULT OF THE WIRE UNRAVELING. MANUFACTURER RESPONSE FOR VASCULAR, POWER PICC (PER SITE REPORTER): SUPPLY CHAIN SPECIALIST TO FOLLOW-UP WITH MANUFACTURER REP. RESPONSE IS UNKNOWN TO THIS WRITER AT THIS TIME. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03792 | 2023-05-26 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 165814 | 165814 |  | IT WAS REPORTED THAT WHEN THE LPN WENT TO REMOVE 10CC OF NORMAL SALINE TO DEFLATE THE BALLOON AT 1300H, IT WAS ABLE TO ONLY ASPIRATE BACK 3CC OF NORMAL SALINE. THE NURSE TRIED MULTIPLE TIMES BUT KEPT FACING RESISTANCE. WHEN THEY TRIED TO PULL THE CATHETER OUT IN CASE THE BALLOON WAS DEFLATED, THE BALLOON DID NOT COME OUT. THEN THEY TRIED TO CUT THE CATHETER WHICH WOULD AUTOMATICALLY ALLOW THE NORMAL SALINE TO COME OUT, THE URINE CONTINUED TO FLOW OUT OF THE CATHETER BUT THE BALLOON DID NOT DEFLATE. TROUBLESHOOTING INCLUDED TRYING TO DEFLATE THE BALLOON BY INSERTING A NEEDLE INTO THE LUMEN THAT GOES TO THE BALLOON BUT COULD NOT ASPIRATE BACK ANY FLUID AFTER MULTIPLE ATTEMPTS WITH DIFFERENT SIZED NEEDLES. ULTRASOUND WAS PERFORMED TO ASSESS THE PATIENT'S BLADDER WHERE THE BALLOON APPEARED TO STILL BE INFLATED. A UROLOGIST WILL HAVE TO SEE THE PATIENT LATER THAT WEEK TO USE A NEEDLE AND ULTRASOUND TO PUNCTURE THE BALLOON. THE CATHETER WAS CLAMPED IN THE MEANTIME AND THE END OF THE REMAINING 14F CATHETER WAS INSERTED INTO A 22F 3 WAY CATHETER WHICH IS CONNECTED TO A FOLEY BAG UNTIL THE PATIENT CONSULTS THE UROLOGIST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03794 | 2023-05-26 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAD UTI (URINARY TRACT INFECTION). PATIENT STATES THAT THEY HAD A UTI (URINARY TRACT INFECTION) BUT COULD NOT DETERMINE IF IT WAS CAUSED BY THE RECALL. PATIENT STATES THAT HAD BEEN TAKING MEDICINE FOR 2 YEARS. FOR LOT#JUGT0289, FOR LOT#JUGU1969, FOR LOT#JUGU0783, FOR LOT#JUGV0665. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS. URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03793 | 2023-05-26 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT PATIENT HAVING 3 SEPARATE BLADDER INFECTIONS IN MALE INTERMITTENT CATHETER ON (B)(6) 2022.. THE REPRESENTATIVE SUGGESTED TO PATIENT IF THEY HAVE AN ISSUE TO REACH OUT TO US SO THEY CAN CORRECT ISSUE. UNKNOWN MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03818 | 2023-05-26 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT STATED THAT ONE OF THESE CATHETERS WAS DEFECTIVE AND HAD A CURVED TIP AND CAUSED SOME BLEEDING. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03821 | 2023-05-26 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53812G | 53812G |  | IT WAS REPORTED THAT THE PATIENT GOT AN INFECTION DURING THE TIME WHEN THE PATIENT WAS USING THESE MAGIC 3 GO MALE CATHETER AND SAID THAT MIGHT HAVE CAUSED THE KIDNEY DISEASE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03809 | 2023-05-26 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER LEAKED FROM THE BALLOON AREA. ACCORDING TO THE INQUIRY SHEET ATTACHED, THERE WAS NO BALLOON RUPTURE OR TEAR. NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER AND THE CATHETER WAS INSERTED BY HAND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03807 | 2023-05-26 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT DURING A PRETEST, AFTER STERILE WATER WAS INJECTED, THE STERILE WATER COULD NOT BE REMOVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02145 | 2023-05-26 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1295108FD4 |  | CUSTOMER REPORTED THAT DURING PICC PLACEMENT THE USER EXPERIENCED DIFFICULT BREAKING THE INTRODUCER SHEATH WHEN FINALIZING THE PLACEMENT OF A PICC. THE INTRODUCER DID NOT CRACK CLEANLY, AND THE TEAM ALMOST HAD TO COMPLETELY RESTART BUT THEY WERE ABLE TO BREAK AWAY ENOUGH OF THE INTRODUCER TO SAFELY REMOVE IT AND FINALIZE PICC PLACEMENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02146 | 2023-05-26 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1174108D |  | CUSTOMER REPORTED THAT AFTER PICC INSERTION, PIECE OF PLASTIC FROM THE INTRODUCER GOT STUCK ON THE LINE AFTER BREAKING AWAY THE INTRODUCER. NO PATIENT HARM REPORTED, NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02144 | 2023-05-26 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD DL POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED, "IN A CALL, NURSE STATES HER PATIENT HAS A POWERPICC. THE CLAMP HAS BROKEN FROM NORMAL USE AND NURSE WANTS TO KNOW WHAT SHE CAN DO. RESPONSE: INFORMED THAT A NEW CLAMP CAN BE PURCHASED FROM ANOTHER COMPANY. ADVISED TO CONSIDER THE TYPE OF CLAMP THEY GET IN REGARDS TO HOW IT WILL NEED TO BE PUT ONTO THE CATHETER." | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02143 | 2023-05-26 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED, THAT DURING PLACEMENT OF THE PICC CATHETER. THE CUSTOMER OPENED THE MST AND INSERTED THE GUIDEWIRE INTO THE BODY OF THIS PATIENT. THE INTRODUCER NEEDLE WAS UNABLE TO SEPARATE FROM THE GUIDE WIRE. AN EXAMINATION REVEALED, A BARB AT THE DISTAL END OF THE GUIDE WIRE, WHICH LED TO THE IMPOSSIBILITY OF THE SEPARATION. THE DISTAL END WAS CUT OFF WITH SCISSORS, AND THE INTRODUCER NEEDLE WAS REMOVED. CAUSING NO IMPACT ON THE PATIENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02139 | 2023-05-26 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR NINE MONTHS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, WHEN THE PORT SYSTEM WAS REMOVED AFTER THE COMPLETION OF THE CHEMOTHERAPY THE CATHETER WAS ALLEGEDLY FOUND TO BE FRACTURED. REPORTEDLY, PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, AN IMAGE AND PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 01/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03804 | 2023-05-26 | 2023-03-06 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT WAS BROUGHT INTO SURGICAL DECISION UNIT AMBULATORY AREA WITH A BLOCKED CATHETER OR UNSUCCESSFUL CATHETERISATION. THE BLADDER SCAN SHOWED 560MLS IN BLADDER. BLEEDING AND CLOTS WERE NOTED TO END OF PENIS. THE CATHETER APPEARED VERY LONG AND IT WAS NOT INSERTED PROPERLY. TEN ML OF STERILE WATER WAS DEFLATED FROM BALLOON. THE CATHETER WAS REPOSITIONED AND INSERTED FURTHER INTO BLADDER. THEN THE CATHETER WAS DRAINING CLEAR URINE. TEN ML OF STERILE WATER WAS REINFLATED INTO BALLOON. IT APPEARED THAT THE CATHETER BALLOON HAD BEEN INFLATED INTO PATIENTS URETHRA AND WAS CAUSING TRAUMA. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02141 | 2023-05-26 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | BARD POWERGLIDE MIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED, "THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE MIDLINE INSERTION SITE IN UPPER ARM." NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02142 | 2023-05-26 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SHERLOCK 3CG (5F) (TRIPLE-LUMEN) (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | CK000803 |  | IT WAS REPORTED THAT BD PICC CATHETER TO LEAK SALINE FROM THE ¿RUBBER HUB¿ BY THE ENTRY PORT FOR THE SHERLOCK WIRE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03801 | 2023-05-26 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT AFTER THE CATHETER INSERTION, URINE COULD NOT BE DRAINED. ACCORDING TO THE INQUIRY SHEET ATTACHED, THERE WAS NO KINK OR BENT ON THE CATHETER AND THERE WAS NO PATIENT'S URINE STONE. THERE WAS NO ABNORMALITY OBSERVED WITH THE PATIENT'S BLADDER CONDITION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03800 | 2023-05-26 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PT STATED KEPT GETTING URINARY TRACT INFECTION WHILE USING CATHETERS HAD TOTAL OF 4. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: REVIEW THE SELF CATHETERIZATION: PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03798 | 2023-05-26 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THEY WERE WARMING PATIENT TO TARGET TEMPERATURE OF 36.8C ON THE ARCTIC SUN DEVICE. ESOPHAGEAL AND TEMPERATURE SENSING FOLEY PROBES IN PLACE FOR THERAPY. PHYSICIAN ORDERED HER TO SWAP OUT TO ESOPHAGEAL PROBE AS PRIMARY SOURCE FOR PATIENT TEMPERATURE INPUT. AFTER THAT THE PATIENT TEMPERATURE DROPPED IN THE LAST HOUR FROM 34.4C TO 34C. PATIENT TEMPERATURE NOW AT 34.2C, TARGET TEMPERATURE WAS 36.8C, WATER TEMPERATURE WAS 34.1C AND WATER FLOW RATE WAS 3.1 L/M. SIZE XS PADS IN PLACE. NURSE VERIFIED REWARM RATE SET TO 0.5C/HR. TREND SHOWS 3 BARS TRENDING UPWARD. MIS EXPLAINED DEVICE APPEARS TO BE WORKING APPROPRIATELY AND RESPONDING TO DROP IN PATIENT TEMPERATURE. THE WATER TEMPERATURE STILL INCREASING. PER SECOND CALL RECEIVED, IT WAS REPORTED THAT THEY HAVE A PATIENT THAT WAS REWARMING ON THE ARCTIC SUN DEVICE. NURSE STATED THAT THEY BEGAN REWARMING AROUND 10 AM, HOWEVER THEY HAVE CHANGED THE SETTINGS A FEW TIMES. THE MOST RECENT CHANGE TO SETTINGS WAS MADE ABOUT AN HOUR AGO. THE CURRENT TARGET TEMPERATURE WAS 36.8C, PATIENT TEMPERATURE 33.2C AND WATER FLOW RATE WAS 3.2 L/MIN. NURSE STATED THAT THE WATER TEMPERATURE WAS IN THE 30'S AND NOW THE WATER TEMPERATURE WAS 25C. ALSO STATED THAT WANTING TO KNOW IF THE WATER SHOULD BE WARMER. MIS CONFIRMED WITH NURSE THEY HAVE THE XX SMALL PADS CONNECTED WITH ADEQUATE SKIN COVERAGE, THE PATIENT WAS 20KG. THEY WERE USING A FOLEY TEMPERATURE PROBE. MIS WALKED NURSE THROUGH ACCESSING EVENT LOG, THE DEVICE HAD AN ALARM 11 (PATIENT TEMPERATURE 1 BELOW LOW PATIENT ALERT) RECENTLY, NO OTHER ALARMS. NURSE CONFIRMED THEIR REWARMING SETTINGS WERE CURRENTLY SET TO REWARM PATIENT FROM 33C AT A RATE OF 0.5C/HOUR TO TARGET TEMPERATURE OF 36.8C. NURSE WAS UNSURE OF THE PATIENT STARTING TEMPERATURE WHEN THEY BEGAN REWARMING, NURSE WAS NOT THE PRIMARY NURSE. NURSE STATES WHEN THEY BEGAN REWARMING, THEY STARTED A NEW PATIENT THERAPY SO THEY COULD NOT SEE THE DATA. MIS WALKED NURSE THROUGH ACCESSING THE SYSTEM DIAGNOSTICS OUTLET MONITOR TEMPERATURE (T1) WAS 32.2C, OUTLET CONTROL TEMPERATURE (T2) WAS 32.3C, INLET TEMPERATURE (T3) WAS 30.1C, CHILLER TEMPERATURE (T4) WAS 4.9C, CIRCULATION PUMP COMMAND WAS 68 PERCENTAGE, MIXING PUMP COMMAND WAS 0 PERCENTAGE, HEATER COMMAND WAS 41 PERCENTAGE, WATER RESERVOIR LIMIT WAS 3, SYSTEM HOURS WERE 373 AND PUMP HOURS WERE 223. NURSE CONFIRMED PATIENT WAS NOT SHOWING ANY SIGNS OF HEAT GENERATION, NO MICRO SHIVERING, AND NO MEDICATIONS GIVEN. NURSE STATED THAT THE WATER TEMPERATURE WAS INCREASING TO 33.8C. NURSE STATED THAT WHEN INITIALLY WALKED INTO THE PATIENT ROOM, THE TREND INDICATOR HAD TWO ARROWS POINTING UP, AND THEN NURSE NOTICED THE WATER TEMPERATURE DROP. NOW THERE WAS ONE ARROW POINTING UP ON THE TREND INDICATOR. MIS EXPLAINED THE DEVICE APPEARS TO BE WORKING APPROPRIATELY. THE TREND INDICATOR ARROWS POINTING UP EXPLAINED THE PATIENT TEMPERATURE WAS EXPECTED TO INCREASE. MIS DISCUSSED THE DEVICE MIGHT HAVE DETECTED HEAT GENERATION OR THE PATIENT REWARMING TOO QUICKLY AND THE WATER TEMPERATURE DECREASED IN RESPONSE. MIS INFORMED NURSE EVEN THOUGH THE PATIENT WAS REWARMING, THE WATER TEMPERATURE MIGHT NOT BE THE WARMEST POSSIBLE, IT WOULD FLUCTUATE AS NEEDED. BY THE END OF THE CALL, THE WATER TEMPERATURE WAS 34.7C. NURSE WOULD MONITOR PATIENT AND CALL BACK IF THEY HAVE AN ISSUE REWARMING. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "SENSOR WIRE TOO WEAK / THERMISTOR WIRE TOO WEAK". THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03799 | 2023-05-26 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L20 |  | IT WAS REPORTED THAT THE BALLOON PORTION OF THE CATHETER WAS RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
17010102 | 2023-05-26 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET REFERENCE # S1275108FD5, LOT # REHN2437 & REGY0842 NOTED LEAKAGE WITH FLUSHING AT RUBBER-LIKE STOPPER. REHN2437 ULTIMATELY SUCCESSFULLY PLACED. STYLETS RETAINED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03795 | 2023-05-26 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE PACKED IMPROPERLY WHEN THEY OPENED THE BOX. THE END TIPS, OPPOSITE OF THE INSERTION END, ARE JAGGED AND NOT CUT PROPERLY. THE PATIENT STATED THAT THEY APPEAR TO NOT BE LUBRICATING CORRECTLY AND HAVE CAUSED MINOR BLEEDING. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: INSTRUCTIONS FOR USE, INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER. H11: A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03803 | 2023-05-26 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 |  | FOL0102 |  | IT WAS REPORTED THAT THE STATLOCK WAS NOT COMPATIBLE WITH THE LINC ALL SILICONE OPEN TIP CATHETER. IT DID NOT SECURE AND THE SIZE WAS NOT ATTACHING TO THE CATHETER PROPERLY. THE PATIENT HAD NOT ORDERED THE STATLOCK SINCE (B)(6) 2022, SO MOST LIKELY THE ISSUE WAS WITH BAGS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INAPPROPRIATE SNAP FIT". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "APPLICATION TECHNIQUE PREP 1. PLACE FOLEY CATHETER INTO RETAINER. DIRECTIONAL ARROW SHOULD POINT TOWARDS CATHETER TIP, AND BALLOON INFLATION ARM SHOULD BE NEXT TO THE HINGE. 2. CLOSE LID, BEING CAREFUL TO AVOID PINCHING THE CATHETER. 3. IDENTIFY SECUREMENT SITE BY LAYING THE DEVICE RETAINER ON THE FRONT OF THE THIGH, LEAVING 2.5 CM OF CATHETER SLACK BETWEEN INSERTION SITE AND THE STATLOCK® DEVICE RETAINER. 4. AFTER PLACING THE STATLOCK® STABILIZATION DEVICE OFF TO THE SIDE, CLEANSE AND DEGREASE THE SECUREMENT SITE WITH ALCOHOL PER HOSPITAL POLICY. LET SKIN DRY. 5. APPLY SKIN PROTECTANT, IN DIRECTION OF HAIR GROWTH, TO AREA LARGER THAN SECUREMENT SITE. ALLOW TO DRY COMPLETELY (10-15 SECONDS). 6. USING PERMANENT MARKER, WRITE INITIALS AND DATE OF APPLICATION ON THE STATLOCK® DEVICE ANCHOR PAD. NOTE: ALWAYS SECURE CATHETER INTO THE STATLOCK® DEVICE RETAINER BEFORE APPLYING ADHESIVE PAD ON SKIN. PLACE AND PEEL 7. ALIGN THE STATLOCK® STABILIZATION DEVICE OVER SECUREMENT SITE LEAVING 2.5 CM OF CATHETER SLACK. MAKE SURE LEG IS FULLY EXTENDED. 8. WHILE HOLDING THE RETAINER TO KEEP THE PAD IN PLACE, PEEL AWAY PAPER BACKING, ONE SIDE AT A TIME AND PLACE TENSION-FREE ON SKIN. REMOVAL TECHNIQUE DISENGAGE 1. OPEN RETAINER BY PRESSING RELEASE BUTTON WITH THUMB, THEN LIFT TO OPEN. 2. REMOVE FOLEY CATHETER FROM THE STATLOCK® DEVICE. DISSOLVE 3. WIPE THE EDGE OF THE PAD USING AT LEAST 5-6 ALCOHOL PADS UNTIL A CORNER LIFTS. THEN CONTINUE TO STROKE UNDERSURFACE OF PAD WITH ALCOHOL TO DISSOLVE ADHESIVE PAD AWAY FROM SKIN. DO NOT PULL OR FORCE PAD TO REMOVE." THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03813 | 2023-05-26 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT STATED THEY HAD THE FIRST URINARY TRACT INFECTION ON (B)(6) 2021 AND HAD 7 SINCE. MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03819 | 2023-05-26 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OFF. ON (B)(6), A FOLEY CATHETER WAS PLACED TO THE PATIENT FOR THE PURPOSE OF URINARY DRAINAGE FOR SURGERY. IMMEDIATELY AFTER SURGERY, 3 HOURS AFTER PLACEMENT, THE FOLEY CATHETER WAS FOUND TO HAVE FALLEN OUT OF THE PATIENT. ANOTHER FOLEY CATHETER WAS PLACED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 15MAY2023, STATED THAT THERE WERE NO OBVIOUS VISIBLE DEFECTS TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02133 | 2023-05-26 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716000J |  | IT WAS REPORTED THAT EIGHT MONTHS POST PORT PLACEMENT, THE PORT SYSTEM ALLEGEDLY COULD NOT BE INFUSED. THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02137 | 2023-05-26 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT DUO M.R.I. IMPLANTABLE PORT, CHRONOFLEX DUAL-LUMEN, 9.5F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1829570 | 1829570 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT WAS ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT THE PORT ALLEGEDLY LEAKED WHEN FLUSHING. IT WAS ALSO REPORTED THAT THE UPON PALPATION AND INSPECTION PRIOR TO REMOVAL, LATERAL LUMEN WAS ALLEGEDLY IRREGULAR AND RAISED. REPORTEDLY THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 12/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02136 | 2023-05-26 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT PRIOR TO A PORT PLACEMENT PROCEDURE, IT WAS NOTICED THAT ONE-THIRD OF THE ENTIRE CIRCUMFERENCE WAS ALLEGEDLY CRACKED AT THE TIP OF THE CATHETER UPON OPENING THE PACKAGE. IT WAS FURTHER REPORTED THAT THE TIP OF CATHETER WAS ALLEGEDLY COMPLETELY BROKEN WHILE THE PHYSICIAN TOUCHED THE CRACKED SITE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02140 | 2023-05-26 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616070 | 1616070 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CATHETER WAS INSERTED INTO THE SHEATH, AND THE SHEATH WAS REMOVED IN THE EXTRACORPOREAL DIRECTION OF THE CATHETER WITHOUT PEELING AWAY THE SHEATH TO COMPLETE THE PLACEMENT OF THE CATHETER. IT WAS FOUND THAT THE BASE OF THE SHEATH WAS ALLEGEDLY DAMAGED AFTER THE PORT SYSTEM PLACEMENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 04/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02154 | 2023-05-27 | 2023-04-27 | I, n, j, u, r, y | Manufacturer report | POWERPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1708000 | 1708000 |  | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT PROCEDURE, THE DEVICE WAS ALLEGEDLY FRACTURED NEAR THE INSERTION OF TUBING. IT WAS FURTHER REPORTED THAT THE RESERVOIR AND PROXIMAL PORTION OF THE CATHETER WERE REMOVED. IT WAS ALSO REPORTED THAT THE DISTAL PORTION COULD NOT SAFELY BE REMOVED AND WAS LEFT IN THE SUBCLAVIAN VEIN AND REQUIRED VASCULAR INTERVENTIONAL RADIOLOGY AT AN OUTSIDE FACILITY FOR REMOVAL. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02152 | 2023-05-27 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK APHERESIS PORT ACCESS SYS |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT ALLEGEDLY WAS UNABLE TO BE ACCESSED DUE TO TISSUE OVERGROWTH. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02156 | 2023-05-28 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT MRI ISP W/SP, ATTACH 9.6 FR. SIL S/L CATHETER | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1809660 | 1809660 |  | IT WAS REPORTED THAT APPROXIMATELY ONE MONTH POST PORT PLACEMENT IN THE RIGHT SIDE OF CHEST, THE PATIENT ALLEGEDLY EXPERIENCED PAIN IN THE NECK AREA. IT WAS FURTHER REPORTED THAT PAIN WAS TREATED WITH TYLENOL. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 12/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02155 | 2023-05-28 | 2023-04-28 | I, n, j, u, r, y | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5883690 | 5883690 |  | IT WAS REPORTED THAT TWO DAYS POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY CRACKED. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY LEAKED. FURTHERMORE, THE PATIENT ALLEGEDLY EXPERIENCED BLEEDING WHICH DID NOT STOP, AND THE TUBING WAS LATER CLAMPED WITH HEMOSTATS WHICH SUCCESSFULLY STOPPED THE BLEEDING. REPORTEDLY, THE CATHETER WAS REPLACED IN THE INTERVENTIONAL RADIOLOGY. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | ADDITIONAL INFORMATION WAS RECEIVED, AND THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE REPORTABLE AS SERIOUS INJURY. MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRY DATE: 05/2024), SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03878 | 2023-05-30 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT PATIENT HAVING 3 SEPARATE BLADDER INFECTIONS IN MALE INTERMITTENT CATHETER ON (B)(6) 2022. THE PATIENT FELT LIKE THEY WERE TRYING TO PUSH THIS PRODUCT. THE REPRESENTATIVE SUGGESTED TO PATIENT IF THEY HAVE AN ISSUE TO REACH OUT TO US SO THEY CAN CORRECT ISSUE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA AND IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02163 | 2023-05-30 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH 15 CM PRE-CURVE KIT | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5554150 |  | IT WAS REPORTED BY THE CUSTOMER THAT ISSUES WITH EITHER THE GUIDEWIRE OR INTRODUCER NEEDLE SUCH THAT THE WIRE COULD NOT BE THREADED THROUGH THE NEEDLE. 05/05/2023 - THE CUSTOMER REPORTED THIS OCCURRED WITH TWO DEVICES HOWEVER ONLY ONE WAS RETURNED FOR EVALUATION. THE RETURNED GUIDEWIRE EXHIBITED A FRACTURED CORE WIRE AND DEFORMED NEEDLE BEVEL. THIS REPORT ADDRESSES THE RETURNED DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF DIFFICULTY INSERTING THE GUIDEWIRE THROUGH THE NEEDLE WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 18GA INTRODUCER NEEDLE AND ONE 0.037¿ J-TIP GUIDEWIRE. THE NEEDLE WAS RECEIVED WITH THE GUIDEWIRE INSERTED THROUGH THE NEEDLE SHAFT. THE J-TIP PROTRUDED PAST THE TIP OF THE NEEDLE. TACTILE INSPECTION OF THE SAMPLE REVEALED THE WIRE TO BE STUCK WITHIN THE NEEDLE SHAFT. THE CORE WIRE WAS FOUND TO BE FRACTURED DURING TACTILE INSPECTION. THE COIL WIRE AND WELD TIP WERE INTACT. MICROSCOPIC INSPECTION OF THE NEEDLE REVEALED DEFORMATION ALONG THE PROXIMAL EDGE OF THE BEVEL. FOLLOWING REMOVAL OF THE WIRE FROM THE NEEDLE, INSPECTION REVEALED BIOLOGICAL RESIDUE ADHERED TO THE COIL WIRE. THE NEEDLE BEVEL DAMAGE WAS CONSISTENT WITH THE WIRE BEING WITHDRAWN THROUGH THE NEEDLE SHAFT. SUCH RETRACTION CAN CAUSE THE WIRE TO BECOME STUCK BOTH BY DAMAGING THE WIRE AND BY DRAWING BIOLOGICAL MATERIAL INTO THE NEEDLE SHAFT, CAUSING INTERFERENCE BETWEEN THE WIRE AND THE NEEDLE. IT APPEARED THAT THE WIRE BECAME STUCK AND SUBSEQUENT PULLING CAUSED A FRACTURE IN THE CORE WIRE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03872 | 2023-05-30 | 2023-04-30 | I, n, j, u, r, y | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON COULD NOT BE FILLED WITH WATER DURING INDWELLING CATHETERIZATION AFTER THE OPERATION. THE PATIENT HAD DELAYED TREATMENT. THE PATIENT WAS GIVEN CATHETER REPLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03871 | 2023-05-30 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE FEW INTERMITTENT SILICON CATHETER DIDN'T HAVE THE TUBES AND IT HAD ONLY HANDLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03857 | 2023-05-30 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT STATED THAT PATIENT BEEN GETTING URINARY TRACT INFECTIONS FOR A YEAR (PR#(B)(4)) NOW AND FOR THE LAST THREE MONTHS (PR#(B)(4)) HAS HAD REALLY BAD URINARY TRACT INFECTIONS. PATIENT STATED THAT PATIENT DID SEEK TREATMENT WITH HER DOCTOR, AND PATIENT DOCTOR DID NOT KNOW WHY THE PATIENT WAS GETTING THE URINARY TRACT INFECTIONS. PATIENT ALSO STATED THAT NOW THAT PATIENT RECEIVED THE RECALL LETTER, PATIENT WILL LET PATIENTS DOCTOR KNOW THAT'S WHAT CAUSED THEM. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03858 | 2023-05-30 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER LEAKED FROM FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03877 | 2023-05-30 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS URETHRAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE PATIENT WAS UNABLE TO INSERT SOME OF THE TOUCHLESS PLUS CATHETERS AND HAD TO THROW THEM AWAY AND TRIED ANOTHER ONE AND SAID THE PATIENT HAD TO DISCARD ABOUT 10 CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OUTSIDE DIAMETER IS TOO LARGE ROUGH OR IRREGULAR SHAPE PROTRUSIONS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: PREPARING FOR CATHETERIZATION: 1. OPEN KIT AND REMOVE CONTENTS. 2. THE GLOVES ARE NOT NECESSARY TO MAINTAIN ASEPTIC CONDITIONS DURING THE PROCEDURE. THE GLOVES ARE FOR THE OPERATOR'S PROTECTION. MALE/FEMALE CATHETERIZATION: 1. REMOVE CAP FROM INSERTION TIP WHILE SQUEEZING TABS. SAVE CAP FOR CLOSING THE BAG. 2. SLIDE THE CATHETER HALFWAY INTO THE INSERTION TIP. 3. MALE - PREPARE THE MALE URETHRAL MEATUS AND THE SURROUNDING AREA WITH POVIDONE-IODINE SWABS PROVIDED. FEMALE - PREPARE THE FEMALE URETHRAL MEATUS BY HOLDING THE OUTER LABIA APART AND PREP THE URETHRAL MEATUS AND SURROUNDING AREA WITH POVIDONE IODINE SWABS PROVIDED. 4. MALE - HOLDING THE PENIS, ADVANCE THE INSERTION TIP INTO THE URETHRA NO FURTHER THAN THE FLANGE BASE. FEMALE - SPREAD THE INNER LABIA, ADVANCE THE INSERTION TIP INTO THE URETHRA NO FURTHER THAN THE FLANGE BASE. RELEASE THE INNER LABIA. 5. PLACE YOUR NONDOMINANT HAND ON THE FINGER CONTROL GUIDE TO STABILIZE CATHETER IN URETHRA. WITH YOUR DOMINANT HAND, GRASP CATHETER THROUGH BAG APPROXIMATELY 1" BELOW THE FINGER CONTROL GUIDE AND PUSH CATHETER INTO URETHRA. THE CATHETER SHOULD BE INTRODUCED BY SHORT, REPETITIVE PUSHING MOTIONS. REPEAT MOTIONS UNTIL CATHETER REACHES BLADDER AND URINE STARTS TO FLOW. 6. ALLOW URINE TO FLOW FREELY, MAKING CERTAIN THE CATHETER GUIDE IS ELEVATED AT LEAST 4" ABOVE LOWER PORTION OF THE BAG. ALLOW FLOW UNTIL BLADDER IS EMPTY OR UNTIL BAG IS FILLED. PRECAUTION: IT IS RECOMMENDED THAT THE COLLECTION BAG BE HELD. THE CATHETER COULD POSSIBLY SEPARATE FROM THE COLLECTION BAG WHEN URINE INCREASES WEIGHT OF THE BAG. 7. WITHDRAW THE CATHETER FROM THE URETHRA. REMOVE THE REMAINING PORTION OF THE CATHETER FROM THE COLLECTION BAG BY PULLING IT THROUGH THE INSERTION TIP. NOTE: THE CATHETER IS DESIGNED TO PASS THROUGH THE INSERTION TIP. 8. THE FILLED COLLECTION BAG MAY BE CLOSED BY REPLACING THE CAP OVER THE INSERTION TIP. MAKE SURE THE CAP SNAPS ON SECURELY. SPECIMEN COLLECTION: TO COLLECT A SPECIMEN, OBTAIN A SAMPLE CUP/TUBE AND ALCOHOL WIPE. REMOVE GUIDE TIP BY PULLING TAB(S) UPWARD FROM BAG. WIPE PORT WITH ALCOHOL. POUR SPECIMEN THROUGH PORT AT TOP OF BAG INTO CUP/TUBE. DRAINING BAG: REMOVE INSERTION TIP BY PULLING TAB(S) UPWARD FROM BAG. DRAIN URINE THROUGH PORT AT TOP OF BAG." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03854 | 2023-05-30 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LUBRI-SIL® I.C. COMPLETE CARE® ALL-SILICONE FOLEY CATHETER TRAY | FOLEY CATHETER TRAY | C.R. BARD INC. (COVINGTON) -1018233 |  | 900114A |  | IT WAS REPORTED THAT SYRINGE OF THE FOLEY TRAY DID NOT CONTAIN THE SPECIFIED VOLUME OF STERILE WATER. AN ALTERNATIVE SYRINGE WAS USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03859 | 2023-05-30 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THERE WAS A HOLE ON THE FOLEY CATHETER. WHEN THE CATHETER WAS USED, THERE WAS A HOLE IN THE INFLATION VALVE AREA, AND LEAKAGE FROM THE HOLE WAS OBSERVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02164 | 2023-05-30 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH 15 CM PRE-CURVE KIT | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5554150 |  | IT WAS REPORTED BY THE CUSTOMER THAT ISSUES WITH EITHER THE GUIDEWIRE OR INTRODUCER NEEDLE SUCH THAT THE WIRE COULD NOT BE THREADED THROUGH THE NEEDLE. 05/05/2023 - THE CUSTOMER REPORTED THIS OCCURRED WITH TWO DEVICES HOWEVER ONLY ONE WAS RETURNED FOR EVALUATION. THE RETURNED GUIDEWIRE EXHIBITED A FRACTURED CORE WIRE AND DEFORMED NEEDLE BEVEL. THIS REPORT ADDRESSES THE DEVICE THAT WAS NOT RETURNED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3: OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03887 | 2023-05-30 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATED THAT PATIENT HAD A URINARY TRACT INFECTION EVERY MONTH WHILE USING THE INTERMITTENT CATHETER. PATIENT STATED THAT PATIENT WAS SEEN BY THE DOCTOR AND WAS TREATED. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03881 | 2023-05-30 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT THE FOLEY CATHETER OF LOT# NGGY5874 WAS INSERTED INTO PATIENT, WHEN PATIENT WENT HOME THE CATHETER HAD FALLEN OUT AS THE BALLOON DEFLATED AND UPON INSPECTION A SMALL PINHOLE WAS DISCOVERED. PATIENT HAD TO RETURN TO THE EMERGENCY ROOM. PER FOLLOW UP VIA PHONE ON (B)(6)2023, IT WAS REPORTED THAT THE CUSTOMER HAD AN ADDITIONAL DEFECTIVE SAMPLE THAT THEY WOULD RETURN FOR EVALUATION. PER FOLLOW UP VIA PHONE ON (B)(6)2023, IT WAS REPORTED THAT THEY SENT BACK SIX CATHETERS TO THE MANUFACTURER. PER SAMPLE FORM RECEIVED ON (B)(6)2023, IT WAS STATED THAT THE PATIENT HAD A FOLEY IN PLACE AND WAS DISCHARGED FROM THE EMERGENCY ROOM. WHILE THE PATIENT WAS AT HOME THE FOLEY FELL OUT. PATIENT RETURNED AND A 2ND CATHETER WAS INSERTED. THIS DEFLATED AS WELL, WHILE THE PATIENT WAS STILL IN THE HOSPITAL AND THIRD FOLEY WAS INSERTED. ALSO STATED THAT THE 2ND FOLEY WAS THROWN AWAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03885 | 2023-05-30 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT SAID ONLY GENERALLY THAT THEY GOT URINARY TRACT INFECTION IN (B)(6)2023. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTIONS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER. WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03886 | 2023-05-30 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATED THAT PATIENT GOT THE RECALL LETTER AND STATES PATIENT HAD REOCCURRING URINARY TRACT INFECTIONS BUT NO LONGER HAS ANY OF THESE PRODUCTS OR LOT #S TO VERIFY IF THEY WERE ON THE LIST ADVISED. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03851 | 2023-05-30 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD TO URINARY TRACT INFECTION WITH THESE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS :THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03884 | 2023-05-30 | 2023-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE SHAFT OF THE FOLEY CATHETER. THE EVENT OCCURRED HOURS AFTER THE PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE SHAFT. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03888 | 2023-05-30 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD UTI ( URINARY TRACT INFECTION). IT WENT SEPTIC AND ALMOST WENT INTO KIDNEY FAILURE. REPRESENTATIVE SENT REPLACEMENT AND ADVISED TO DISPOSE OF AFFECTED CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. (LOT# JUGR0224, LOT# JUGR1388, LOT# JUGV1279). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03889 | 2023-05-30 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD UTI ( URINARY TRACT INFECTION). IT WENT SEPTIC AND ALMOST WENT INTO KIDNEY FAILURE . REPRESENTATIVE SENT REPLACEMENT AND ADVISED TO DISPOSE OF AFFECTED CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. ( LOT# JUGR0224 ,LOT# JUGR1388, LOT# JUGV1279 ). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03890 | 2023-05-30 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119416M | 119416M |  | IT WAS REPORTED THAT THE CUSTOMER HAD MULTIPLE EVENTS WHERE THEY HAD TO CHANGE THE FOLEY CATHETER DURING PROCEDURE BECAUSE OF NO URINE OUTPUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
17022295 | 2023-05-30 | 2023-05-15 | I, n, j, u, r, y | User Facility report | SEEKER | CATHETER, PERCUTANEOUS | BARD |  | SK13535 |  | ON (B)(6) 2023, PATIENT UNDERWENT DUPLEX GUIDED ANGIOGRAM WITH RIGHT FEMORAL ACCESS. PATIENT WAS DISCHARGED POST PROCEDURE. ON (B)(6) 2023, PATIENT PRESENTED TO OFFICE FOR FOLLOW UP VISIT. PATIENT REPORTED PAIN, NUMBNESS, LEG DISCOLORATION. PATIENT SENT TO HOSPITAL FOR URGENT REVASCULARIZATION FOR LIMB ISCHEMIA. SURGEON FOUND RETAINED VASCULAR CATHETER. 8 CM REMOVED, ADDITIONAL VASCULAR CATHETER REMAINED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03879 | 2023-05-30 | 2023-03-25 | I, n, j, u, r, y | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE PATIENT FOUND URINE LEAKAGE WITH USING THE FOLEY CATHETER. THE NURSE IMMEDIATELY REPORTED TO THE DOCTOR AND INJECTED NORMAL SALINE (30ML) INTO THE BALLOON, AND THE URINE LEAKAGE IMPROVED (THE PATIENT HAD LEAKED URINE SEVERAL TIMES) BUT AFTER A WHILE THE LEAKAGE OF URINE OCCURRED AGAIN. THEREFORE, THE CATHETER WAS REMOVED COMPLETED WITHOUT MISSING PIECES AND REPLACED WITH A NEW CATHETER IMMEDIATELY. DUE TO THE FREQUENT RE-CATHETERIZATION, THE PATIENT GOT THE URETHRAL INJURY AND URINARY TRACT INFECTION WHICH REQUIRES ANTIMICROBIAL TREATMENT AND PROLONGS THE HOSPITAL STAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03846 | 2023-05-30 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT GOT URINARY TRACT INFECTION WHILE USING THESE MAGIC 3 MALE CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. THE PATIENT REPORTED ISSUES WITH LOT# JUGT1000, LOT# JUGS0046. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03845 | 2023-05-30 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 996081 | 996081 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO TO THE QUESTION, WERE YOU ABLE TO USE THE REFOLDING TOOL TO REFOLD THE BALLOON DILATION CATHETER FOR REINSERTION. THE PHYSICIAN WAS NOT ABLE TO USE THE REFOLDING TOOL TO REFOLD THE 996081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, W/ INFLATION DEVICE, STERILE BECAUSE DIFFICULT TO MANEUVER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03826 | 2023-05-30 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 996101 |  | IT WAS REPORTED THAT IN AN ONLINE SURVEY THE PHYSICIAN STATED THAT THEY WERE NOT ABLE TO SUCCESSFULLY SUSTAIN PRESSURE, PRESSURE MONITORING, AND DEFLATION OF THE BALLOON FOR THE UNKNOWN/DO NOT KNOW PRODUCT CODE (SKU) DUE TO MECHANICAL FAULT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02162 | 2023-05-30 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5616000 | 5616000 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY EXPERIENCED ITCHING WHEN SHE RECEIVED HER SALINE INJECTIONS THROUGH HER PORT. IT WAS FURTHER REPORTED THAT PATIENT DOES NOT HAVE THE REACTION WHEN THE PORT WAS NOT IN USE. REPORTEDLY, THERE IS NO INFORMATION PROVIDED REGARDING ADDITIONAL INTERVENTION OR MEDICATION PRESCRIBED TO TREAT ITCHING. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 01/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00371 | 2023-05-30 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SEEKER CROSSING SUPPORT | SUPPORT CATHETER | BARD PERIPHERAL VASCULAR, INC. | SK15018 | SK15018 |  | IT WAS REPORTED THAT DURING PREPARATION OF A RECANALIZATION PROCEDURE, THE DEVICE ALLEGEDLY HAD RED SLIM SUBSTANCE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 02/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00373 | 2023-05-30 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER IQ ULTRASONIC | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | XCTO146 | XCTO146 |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE TIP OF THE CATHETER ALLEGEDLY BROKE OFF WHILE INSIDE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 10/2023). THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02157 | 2023-05-30 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5883690 | 5883690 |  | IT WAS REPORTED THAT SOMETIME POST THE DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY LEAKED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00374 | 2023-05-30 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER IQ ULTRASONIC | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | XCTO146 | XCTO146 |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE IN THE RIGHT SUPERFICIAL FEMORAL ARTERY VIA LEFT GROIN ACCESS, THE TIP OF THE CATHETER ALLEGELDY DETACHED FROM THE CATHETER. IT WAS FURTHER REPORTED THAT THE TIP DID NOT FULLY DETACH FROM THE CATHETER, BUT ALLEGEDLY PULLED FROM THE CATHETER ITSELF. FURTHERMORE, THE DEVICE TIP WAS ALLEGEDLY CAUGHT ON THE SUPPORT CATHETER AND WOULD NOT RETRACT. REPORTEDLY, THE PHYSICIAN HAD TO REMOVE EVERYTHING FROM THE PATIENT WHICH LOST THE ACCESS. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 07/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00375 | 2023-05-30 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CRUS6A | CRUS6A |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE IN THE LEFT SUPERFICIAL ARTERY VIA RIGHT GROIN ACCESS, THE CATHETER ALLEGEDLY BECAME STUCK IN THE LESION. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY COMPLETELY BROKEN AND STUCK INSIDE THE SUPPORT CATHETER. FURTHERMORE, THE CATHETER ALLEGEDLY HAD DIFFICULTY IN REMOVING FROM THE PATIENT. REPORTEDLY, THE PHYSICIAN WAS ABLE TO REMOVE THE DEVICE FROM THE PATIENT WITHOUT RIPPING THE VESSEL APART. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 08/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03843 | 2023-05-30 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE USER DISCONTINUED THE USE OF THE FOLEY CATHETER DUE TO DIFFICULTY IN WITHDRAWING WATER DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03828 | 2023-05-30 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED AFTER THE OPERATION. THEN IT WAS FOUND THAT THE FLOW OF DRAINAGE WAS NOT SMOOTH AND THE CATHETER FELL OUT. PATIENT EXPERIENCED PAIN WHILE CATHETER REINSERTION. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6)2023, CLARIFIED THAT THE PAIN WAS DUE TO THE CATHETER USED AND THERE WAS DAMAGE ON THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02161 | 2023-05-30 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT EIGHT MONTHS AND FOUR DAYS POST THE PORT PLACEMENT, THE CATHETER ALLEGEDLY HAD A LEAKAGE. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN. REPORTEDLY, THE PORT AND THE CATHETER WERE REMOVED COMPLETELY. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 09/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03829 | 2023-05-30 | 2023-04-26 | I, n, j, u, r, y | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT ON THE MORNING OF (B)(6) 2023, THE CATHETERIZATION WAS PERFORMED FOR BLADDER IRRIGATION AFTER THE OPERATION. UPON CHECKING, IT WAS FOUND THAT THE BALLOON WAS BURST AND THEN WAS REPLACED WITH A NEW CATHETER. AND THE PATIENT HAS EXPERIENCED MUCOUS MEMBRANE DAMAGE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03831 | 2023-05-30 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DUE TO STERILE WATER LEAKAGE. THE EVENT OCCURRED HOURS AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAK LOCATION WAS UNKNOWN BECAUSE THE EVENT WAS CATHETER FELL OUT AND THEY DID NOT TRY TO REPRODUCE THE EVENT. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. THE LOT WAS UNKNOWN, BUT THE CUSTOMER REQUESTED WHETHER THE OCCURRENCE OF SIMILAR EVENTS HAD INCREASED IN THE MONTHS BEFORE AND AFTER THE REPORT ISSUANCE. PER MAIL RECEIVED FROM IBC ON 16MAY2023, STATED THAT THE BALLOON RUPTURE WAS OBSERVED TO THE RETURNED SAMPLE. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 19MAY2023, STATED THAT NO MISSING PIECES WERE OBSERVED TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03840 | 2023-05-30 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE PATIENT STATED THAT THE PATIENT GOT A URINARY TRACT INFECTION IN JANUARY AND IT TURNED INTO A PROSTATE INFECTION USING INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03836 | 2023-05-30 | 2023-05-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE PATIENT HAD 3 URINARY TRACT INFECTION IN THE LAST 12 MONTHS AND WAS HOSPITALIZED FOR IT. THE CUSTOMER STATED URINARY TRACT INFECTIONS WERE FROM URINE BACKING UP. THE PATIENT WAS NOT SURE IF IT BACKED UP ALL THE WAY INTO THE KIDNEYS BUT WAS TO THE POINT PATIENT COULD NOT CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03834 | 2023-05-30 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT SOMETIMES THE PATIENT FELT LIKE THE PUREWICK FEMALE EXTERNAL CATHETER WAS PULLING TOO HARD AND IT HURTS THE PATIENT. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03842 | 2023-05-30 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED, THAT PATIENT HAVING 3 SEPARATE BLADDER INFECTIONS, IN MALE INTERMITTENT CATHETER ON (B)(6) 2022. THE PATIENT FELT LIKE THEY WERE TRYING TO PUSH THIS PRODUCT. THE REPRESENTATIVE SUGGESTED TO PATIENT, IF THEY HAVE AN ISSUE TO REACH OUT TO US SO THEY CAN CORRECT ISSUE. IT WAS UNKNOWN, WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN, WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN, WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN. THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR, WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE, WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE. AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY, AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY (1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED, TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03849 | 2023-05-30 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK ACCESSORY CANNISTER ARRIVED CRACKED. IT WAS NOTED THAT THE PUREWICK ACCESSORY REPLACEMENT KIT WAS SHIPPED WITH THE ORIGINAL PUREWICK URINE COLLECTION SYSTEM ON (B)(6) 2023. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE BOTTOM OF THE PUREWICK FEMALE EXTERNAL CATHETER WAS PLACED TOO FAR UNDER THEM, AND THEY SAT ON IT, AND IT WAS NOT SUCTIONING CORRECTLY. STATED THAT THEY WERE WALKED THROUGH STEPS TO MAKE SURE IT WORKED CORRECTLY AND TOOK A BREAK FROM MACHINE BECAUSE THEY HAD A URINARY TRACT INFECTION AND WAS ON MEDICATION BUT PLANS TO GO BACK ON IT ONCE THEY WERE BETTER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03844 | 2023-05-30 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® I.C. COMPLETE CARE FOLEY CATHETER | COMPLETE CARE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | A319416AM |  | IT WAS REPORTED THAT THE SYRINGE IN THE KIT TO INFLATE THE BALLOON HAD AIR IN IT AND NOT STERILE WATER. WHILE I WENT TO INSERT THERMISTER FOLEY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03832 | 2023-05-30 | 2023-05-06 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MAGIC3 GO MALE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE CATHETERS ON THE RECALL LIST GAVE THE PATIENT BLADDER ISSUES. THEY SAID THAT THEY STILL HAD A FULL BOX OF 30 AND EXTRA 9 CATHETERS LEFT AND WANTED TO KNOW WHAT TO DO WITH THEM. REPRESENTATIVE ADVISED THEM TO DISCARD. THEY DID NOT SPECIFY THE TIMEFRAME OF THE ISSUE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA AND IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02159 | 2023-05-30 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT FIVE YEARS TWO MONTHS AND TWENTY-SEVEN DAYS POST THE PORT PLACEMENT, THE DEVICE ALLEGEDLY HAD NO BLOOD RETURN. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY FOUND TO BE FRACTURED UPON CHEST X-RAY EXAMINATION. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER IN TWO SEGMENTS WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATION WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED ON THE DISTAL END OF THE ATTACHED CATHETER AND THE PROXIMAL END OF THE DISTAL CATHETER SEGMENT. THE PROXIMAL END OF THE DISTAL CATHETER SEGMENT RETURNED WAS ELLIPTICAL IN SHAPE AND THE EDGES WERE NOTED TO BE UNEVEN. THE EDGES OF THE COMPLETE CIRCUMFERENTIAL BREAK ON THE DISTAL END OF THE ATTACHED CATHETER WERE NOTED TO BE JAGGED. THE SURFACES OF THE BREAKS WERE NOTED TO BE ROUND AND SMOOTH IN ONE REGION AND GRANULAR IN THE OTHER REGION. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND THE IDENTIFIED MATERIAL SEPARATION ISSUES. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED SUCTION ISSUE AS THE EXACT CIRCUMSTANCE AT THE TIME OF THE EVENT REPORTED WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 01/2021), G3, H6 (DEVICE, METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02166 | 2023-05-31 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY HAD A HOLE IN THE ARTERIAL LEG. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 05/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02165 | 2023-05-31 | 2022-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NAVARRE OPTI DRAIN | DRAINAGE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NOD8LPT | NOD8LPT |  | IT WAS REPORTED THAT DURING AN ULTRASOUND-GUIDED PERCUTANEOUS TRANSHEPATIC CHOLANGIOGRAPHY DRAINAGE CATHETER PLACEMENT PROCEDURE, THE CONNECTOR ALLEGEDLY FRACTURED. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02167 | 2023-05-31 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT SOMETIME POST THE PORT PLACEMENT, THE PATIENT ALLEGEDLY COMPLAINED ABOUT THE SIZE OF THE PORT AND TRIED TO WARM IT UP TO MAKE IT EXPAND. IT WAS FURTHER REPORTED THAT THE BRACELET DID NOT FIT AROUND THE PATIENT'S WRIST, WHICH MADE THE PATIENT'S ARM SORE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03893 | 2023-05-31 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT AFTER 4 DAYS OF CATHETER PLACEMENT TO THE PATIENT, THE BALLOON WAS FOUND BURST. AS PER THE FOLLOW UP INFORMATION RECEIVED ON (B)(6) 2023, CLARIFIED THAT THERE WAS NO MISSING PIECES OF BALLOON AND THE PATIENT DID NOT UNDERGO ANY MEDICAL INTERVENTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02168 | 2023-05-31 | 2023-04-27 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT TWO YEARS FIVE MONTHS AND EIGHT DAYS POST PORT PLACEMENT, BLOOD RETURN COULD NOT BE ALLEGEDLY CONFIRMED. IT WAS FURTHER REPORTED THAT CHEST X-RAY SHOWED THAT THE CATHETER WAS ALLEGEDLY RUPTURED AT THE INSERTION SITE AND THE RUPTURED PORTION ALLEGEDLY FELL INTO THE RIGHT VENTRICLE. REPORTEDLY, THE CATHETER AND THE PORT BODY WERE REMOVED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 10/2021). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03897 | 2023-05-31 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51810 | 51810 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATED THAT PATIENT WAS CALLING FOR THE RECALL. PATIENT STATED THAT PATIENT DID HAVE A FEW URINARY TRACT INFECTIONS WHILE USING THE MAGIC 3 GO BUT CANNOT SAY WHAT CAUSED IT, PATIENT WAS TREATED BY THE DOCTOR. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. IT WAS UNKNOWN IF THE DEVICE CAUSED URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03904 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAD FOUR MAJOR URINARY TRACT INFECTIONS. PATIENT DIDN'T SPECIFY THE TIMEFRAME OF THE INFECTION. REPRESENTATIVE ADVISED TO DISCONTINUE USING CATHETERS AND WAS REPLACING. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. BATCH #JUGU0783, BATCH #JUGV0317. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN; WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA; REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03896 | 2023-05-31 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03932 | 2023-05-31 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER WAS UNABLE TO BE INJECTED DUE TO WATER LEAKAGE FROM THE INFLATION VALVE. UPON CHECKING THE CATHETER, THE INFLATION VALVE GOT DISCONNECTED. NO PROBLEM WAS OBSERVED WITH THE CATHETER DURING INFLATION ACCORDING TO THE INQUIRY SHEET ATTACHED. WATER WAS INJECTED GENTLY. USUALLY A BALLOON CAN BE INFLATED WITHIN SECONDS BUT IT TOOK MORE THAN USUAL INFLATING THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03930 | 2023-05-31 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MALE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE URINARY TRACT INFECTION WAS SO SEVERE, WENT INTO BLOOD STREAM STARTING BACK IN DECEMBER AND THE MOST RECENT STARTED MARCH. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿PATIENT SENSITIVITY TO MATERIALS¿. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA AND IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03935 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC 3 GO MALE COUDE CATHETER | MALE COUDE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAD FOUR MAJOR URINARY TRACT INFECTIONS. PATIENT DIDN'T SAYS THAT TIMEFRAME OF THE URINARY TRACT INFECTION. REPRESENTATIVE ADVISED TO DISCONTINUE USING CATHETER'S AND ITS REPLACING. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. FOR BATCH # JUGU0783. FOR BATCH # JUGV0317. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA AND IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03934 | 2023-05-31 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT STERILE WATER COULD NOT BE DRAINED IN THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03937 | 2023-05-31 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | INTP16 | INTP16 |  | IT WAS REPORTED THAT TWO BARD INTERMITTENT CATHETER KITS FROM DIFFERENT LOTS (NGHN0240, NGGY1050) WHEN OPENED, HAD THE CATHETER STICKING OUT OF STERILE WRAP WHEN THE FIRST FLAP WAS UNFOLDED. UPON REVIEW OF THE MEDWATCH FORM PROVIDED BY (B)(6) ON (B)(6) 2023, THIS WAS MERELY A REPORTED EVENT AND NOT A COMPLAINT AGAINST THE SURESTEP INTERMITTENT TRAY. THE CUSTOMER HAD MADE IT CLEAR THAT THERE WERE NO SAMPLES AVAILABLE TO RETURN TO THE MANUFACTURER FOR EVALUATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03925 | 2023-05-31 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS WITHDRAWN, AND WHEN CHECKED, THERE WAS NO STERILE WATER, AND WHEN STERILE WATER WAS INJECTED, A PINHOLE COULD BE SEEN AROUND 5 CM AT THE TIP. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03938 | 2023-05-31 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | INTP16 | INTP16 |  | IT WAS REPORTED THAT TWO BARD INTERMITTENT CATHETER KITS FROM DIFFERENT LOTS (NGHN0240, NGGY1050) WHEN OPENED, HAD THE CATHETER STICKING OUT OF STERILE WRAP WHEN THE FIRST FLAP WAS UNFOLDED. UPON REVIEW OF THE MEDWATCH FORM PROVIDED BY (B)(6) ON (B)(6) 2023, THIS WAS MERELY A REPORTED EVENT AND NOT A COMPLAINT AGAINST THE SURESTEP INTERMITTENT TRAY. THE CUSTOMER HAD MADE IT CLEAR THAT THERE WERE NO SAMPLES AVAILABLE TO RETURN TO THE MANUFACTURER FOR EVALUATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03940 | 2023-05-31 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT WAS DEALING WITH A PERSISTENT URINARY TRACT INFECTION, BUT SWITCHED TO A CATHETER FROM A DIFFERENT MANUFACTURER AND THAT ONE WAS WORKING BETTER FOR THEM. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03943 | 2023-05-31 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 37303 | 37303 |  | IT WAS REPORTED THAT THESE MALE EXTERNAL CATHETERS WERE TOO STICKY AND VERY HARD TO GET OFF. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: 1) VERIFY CORRECT SIZE PRIOR TO USE. 2) TRIM PUBIC HAIR IF NECESSARY. 3) WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. 4) OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. 5) PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. 6) UNROLL THE CATHETER OVER PENIS. 7) GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. 8) CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: 1) ENSURE DRAINAGE BAG IS EMPTY. 2) DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. 3) GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03947 | 2023-05-31 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT AFTER THE CATHETER PLACEMENT TO THE PATIENT, IT WAS FOUND THAT THERE WAS A LEAKAGE AT THE CONNECTIONS IN THE LUMEN. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 21MAY2023, THE DAMAGE ON THE BALLOON WAS UNKNOWN BUT WATER LEAKAGE HAS BEEN CONFIRMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03926 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION AND THEY HAD 3 URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA AND IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03941 | 2023-05-31 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT 30 DAYS AFTER THE FOLEY CATHETER PLACEMENT, STERILE WATER LEAKED DURING USE ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03945 | 2023-05-31 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT AND UPON CHECKING, THE BALLOON WAS FOUND BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03946 | 2023-05-31 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED, THAT THE FOLEY CATHETER BALLOON HAD BURST. AFTER CONDUCTING A PRETEST AND CONFIRMING THAT THERE WERE NO ISSUES WITH THE BALLOON, THE CATHETER WAS PLACED TO THE PATIENT. AFTER THE PROCEDURE, A POPPING SOUND WAS HEARD AND UPON CHECKING THE CATHETER, THERE WAS NO EXTRACTION AT THAT TIME. LATER, IT FELL OFF SPONTANEOUSLY. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023. STATED, THAT THERE WERE NO MISSING PIECES OF THE BALLOON. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03942 | 2023-05-31 | 2023-04-23 | I, n, j, u, r, y | Manufacturer report | TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE ELECTRODE OF THE TEMPORARY PACING ELECTRODE WAS KINKED. THE USER WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03951 | 2023-05-31 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED FOR CONTINUOUS BLADDER IRRIGATION WITH SALINE. THEN THE CATHETER WAS FOUND TO HAVE FALLEN OUT, WHICH CAUSED THE PATIENT TO SUFFER FINANCIAL LOSS AND PSYCHOLOGICAL DAMAGE. A NEW 18F CATHETER WAS REPLACED. ACCORDING TO THE DEVICE FAILURE PERFORMANCE, THE CATHETER WAS BURST AND PROLAPSED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03898 | 2023-05-31 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATED THAT THE RECALLED CATHETERS CAUSED A URINARY TRACT INFECTION, AND PATIENT DID NOT KNOW BECAUSE PATIENT GOT THE LETTER AFTER USE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03954 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MAGIC3 GO MALE COUDE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD SEVERAL URINARY TRACT INFECTIONS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE. THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. ". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03955 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MAGIC3 GO MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. THE PATIENT STATED THAT THEY HAD FREQUENT URINARY TRACT INFECTION. IT WAS UNKNOWN IF THE INTERMITTENT CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. PER FOLLOW UP VIA PHONE ON 26MAY2023, IT WAS REPORTED THAT THE CUSTOMER DEVELOPED URINARY TRACT INFECTION AND SOUGHT MEDICAL INTERVENTION. ALSO STATED THAT THEY DEVELOPED CLOSTRIDIUM DIFFICILE (CDIFF) AND WAS PRESCRIBED ANTIBIOTICS. THEY STATED THAT THEY HAD PROSTATE SURGERY SO NO LONGER NEEDED TO USE THE CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03957 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT PATIENT SAID THAT INTERMITTENT CATHETER POSSIBLY CAUSED URINARY TRACK INFECTION WHILE USING INTERMITTENT AND PATIENT WAS ON ANTIBIOTICS. ITS UNKNOWN IF DEVICE CAUSED URINARY TRACK INFECTION PER FOLLOW UP VIA PHONE ON 26MAY2023, IT WAS REPORTED THAT THE CUSTOMER DEVELOPED URINARY TRACT INFECTION AND SOUGHT MEDICAL INTERVENTION. CUSTOMER WAS PRESCRIBED ANTIBIOTICS AND LIBERATOR REPLACED CATHETERS WHICH HAVE CAUSED NO ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02176 | 2023-05-31 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS 3CG, GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY CUSTOMER THAT THEY HAD A KINKED 4F DL PICC. CATHETER WAS KINKED BY THE CLAVICLE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-02177 | 2023-05-31 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT WHEN CONDUCTING PUNCTURE IN THE AFTERNOON, (B)(6)2023, THE NURSE OPENED THE PACKAGE OF THE PICC PRODUCT AND INSPECTED THE CATHETER FOR INTEGRITY, FINDING THAT THE STYLET HAD BEEN BENT. WHEN PRE-FLUSHING THE CATHETER, THE NURSE FOUND THAT THE BENT GUIDE WIRE HAD SCRATCHED THE CATHETER, MAKING THE PRODUCT UNUSABLE. A NEW CATHETER WAS THEN USED AND NO INJURY WAS CAUSED TO THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03944 | 2023-05-31 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MAGIC 3 GO MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50812G | 50812G |  | IT WAS REPORTED, THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION (UTI), WHILE USING THE MAGIC 3 MALE COUDE CATHETER. REPRESENTATIVE ADVISED TO THROW AWAY AND REPLACED. IT IS UNKNOWN, WHAT MEDICAL INTERVENTION IS PROVIDED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN, WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN, WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED, AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR, WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE, WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE, AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY (1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED, TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03928 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT CURRENTLY HAD A URINARY TRACT INFECTION. ALSO STATED THEY WAS ON CEPHALEXIN REGIMEN ANTIBIOTICS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03953 | 2023-05-31 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD UTI ( URINARY TRACT INFECTION). IT WENT SEPTIC AND ALMOST WENT INTO KIDNEY FAILURE . REPRESENTATIVE SENT REPLACEMENT AND ADVISED TO DISPOSE OF AFFECTED CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. FOR LOT# JUGR0224, FOR LOT# JUGV1279. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03924 | 2023-05-31 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT WAS CATHING 4 TIMES A DAY AND SOMETIMES THE CATHETER WOULD CLOG AND HAD WHITE STUFF IN IT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03902 | 2023-05-31 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE PACK OF THREE POVIDONE-IODINE SWAB STICKS THAT ARE ALWAYS INCLUDED AND USUALLY SATURATED WITH POVIDONE IODINE, THAT WERE COMPLETELY DRY IN THIS STRAIGHT CATHETER KIT. THE STRAIGHT CATHETER PACKAGE WAS NOT OPENED, AND THE PACKAGE FOR THE IODINE WAS ALSO NOT OPENED AT ALL UNTIL USE. THE PATIENT THAT THE CATHETER WAS TO BE USED FOR WAS NOT AFFECTED, BUT A NEW KIT HAD TO BE USED AS THEY HAD TO BREAK STERILITY TO GET A NEW KIT WITH IODINE TO CLEAN THE PATIENT. THEY COULD ALSO HAVE JUST GOTTEN IODINE OR BETADINE TO CLEAN, BUT THEY DID NOT HAVE THOSE ROUTINELY STOCKED IN ROOMS, SO IN ORDER TO MINIMIZE THE TIME SPENT GATHERING SUPPLIES, THEY USED ANOTHER KIT THAT WAS ALREADY PRESENT. OTHERWISE, THE REST OF THE KIT HAD ALL THE EXPECTED COMPONENTS. PER ADDITIONAL INFORMATION RECEIVED ON 20MAY2023, IT WAS STATED THAT THE PATIENT WAS DISCHARGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03899 | 2023-05-31 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATES THAT EVERY TIME SHE REORDERS THE AGENT ASKS PATIENT HOW MANY URINARY TRACT INFECTIONS PATIENT HAD IN A 12MONTH CYCLE. PATIENT SAYS PATIENT KNOWS THAT SHE KEEPS DEVELOPING THE URINARY TRACT INFECTIONS FROM THE USE OF THE M3GO CATHETERS. PER FOLLOW UP VIA PHONE ON 25MAY2023, IT WAS REPORTED THAT THE CUSTOMER DEVELOPED UTI AND SOUGHT MEDICAL INTERVENTION. CUSTOMER WAS PRESCRIBED ANTIBIOTICS AND LIBERATOR REPLACED CATHETERS OF DIFFERENT TYPE WHICH HAD CAUSED NO ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03903 | 2023-05-31 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED IN THE OPERATING ROOM AND THE PATIENT WAS TRANSFERRED TO THE WARD AFTER SURGERY. ON THE NEXT DAY, WHEN THE PATIENT LEANED OVER FOR WIPING, THE BALLOON FELL OUT. WHEN THE CUSTOMER PUT THE STERILIZED WATER IN THE BALLOON AND CHECKED IT, IT SWELLED FOR THE TIME BEING, BUT THERE WAS A POSSIBILITY OF A PINHOLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03901 | 2023-05-31 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT DURING A SURGERY. UPON INJECTING WATER AND CHECKING THE CATHETER, THERE WAS WATER LEAKAGE FROM AROUND THE BALLOON AREA. THE EVENT OCCURRED ONE HOUR AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BALLOON AND THE SHAFT. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. AS PER MAIL RECEIVED FROM IBC REPRESENTATIVE ON (B)(6) 2023, STATED THAT THE LEAK LOCATION OF THIS COMPLAINT WAS ONE. SINCE THE EXACT LEAK LOCATION WAS UNKNOWN, THE SALES REPRESENTATIVE MARKED BALLOON AND SHAFT IN THE INQUIRY SHEET. REPORTED COMPLAINT WAS THAT THE WATER LEAKAGE FROM THE BALLOON DURING A WATER INJECTION. STATED THERE WERE NO OBVIOUS VISIBLE DEFECTS TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03907 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION AND PATIENT GOT URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03908 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED, THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. AND THEY HAD 2 URINARY TRACT INFECTION DURING USE OF INTERMITTENT CATHETER. IT WAS UNKNOWN, WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN, WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN, WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN. THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR, WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE, WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE. AND ADVANCE THE CATHETER TIP INTO THE URETHRA SLOWLY. AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY (1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED, TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
17030119 | 2023-05-31 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP¿ FOLEY TRAY SYSTEM BARD® LUBRI-SIL® FOLEY CATHETER TRAY | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | A942216 | A942216 |  | INDWELLING CATHETER FELL OUT AFTER PATIENT WAS IN PRONE POSITION. UPON INSPECTION OF CATHETER, TIP OF CATHETER RIPPED OFF CAUSING THE BALLOON TO DEFLATE AND CATHETER TO FALL OUT. BALLOON STILL INTACT, APPROXIMATELY 1CM OF THE TIP OF THE CATHETER IS MISSING. NO BLOOD NOTED ON CATHETER TO INDICATE ANY TRAUMA. CATHETER WAS NOT TUGGED ON OR PULLED DURING POSITIONING. TIP NOT FOUND ON INITIAL ASSESSMENT OF AREA. DOCTORS WERE MADE AWARE. PLAN TO INSPECT THE MEATUS AFTER THE CASE WHEN PATIENT IS SUPINE PER DOCTOR, NOTIFIED UROLOGY IS UNABLE TO FOUND BROKEN TIP. NURSE MANAGER NOTIFIED OF EVENT AND UPDATED ON PLAN OF CARE REGARDING THE MISSING CATHETER TIP. CATHETER LOT INFORMATION SAVED, AND PRODUCT FAILURE FORM COMPLETED. NEURO CSL AWARE AS WELL. UROLOGY CONSULT PAGED AND INTRAOPERATIVE CYSTO COMPLETED TO RULE OUT FOREIGN BODY. CYSTO WAS NEGATIVE FOR FOREIGN BODY. |  | N | Y |  |  | 2023-05
1018233-2023-03906 | 2023-05-31 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE BALLOON BURST ON THE FIRST DAY AFTER THE FOLEY CATHETER PLACEMENT TO THE PATIENT AND FELL SPONTANEOUSLY OFF. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03909 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACK INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTIONS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03911 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT WAS UPSET AND REQUESTED TO SPEAK WITH BARD. THEY RECEIVED A LETTER STATING THAT THE CATHETERS THEY WERE USING WERE NOT GOOD. THEY HAD ALREADY USED ALL OF THE CATHETERS THEY HAD. AS A RESULT, THEY WERE HOSPITALIZED FOR 5 DAYS (ADMITTED AND RELEASED ON (B)(6) 2023). THEY ALSO STATED THEY WERE IN THE HOSPITAL WITH SEPSIS AND A REALLY BAD INFECTION, HEART PROBLEMS, AND KIDNEY AND BLADDER ISSUES. THEY ADVISED THEIR FAMILY THAT THEY DID NOT KNOW IF THEY WOULD PULL OUT OF IT. THEY WERE INSISTING ON SPEAKING WITH BARD TO EXPLAIN THEIR SITUATION, AND THEY SAID THEIR DOCTORS AND EVERYONE ELSE WERE REALLY UPSET BECAUSE THEY COULD NOT DETERMINE WHAT WAS GOING ON WITH THEM. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03910 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAVE URINARY TRACK INFECTION. PER FOLLOW UP VIA PHONE ON (B)(6) 2023,CUSTOMER DID DEVELOP A URINARY TRACT INFECTION AND WAS TREATED WITH ANTIBIOTICS. THEY STILL USING THE CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03914 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT PATIENT HAD A URINARY TRACT INFECTION DIAGNOSED AND CURRENTLY ON ANTIBIOTICS WHILE USING INTERMITTENT CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS. URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03929 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT STATED THEY HAD THE FIRST URINARY TRACT INFECTION ON (B)(6) 2021 AND HAD 7 SINCE. MEDICAL INTERVENTION WAS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03923 | 2023-05-31 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51612 | 51612 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION. MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTERMITTENT CATHETER HYDROPHILIC COATED SILICONE. INSTRUCTIONS FOR USE. INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. THE CATHETER BECOMES SLIPPERY WHEN WETTED WITH WATER, ELIMINATING THE NEED FOR A SEPARATE LUBRICANT. FOR YOUR ADDED CONVENIENCE, THIS CATHETER IS PACKAGED WITH ITS OWN STERILE WATER. SIMPLY RELEASE THE WATER FROM ITS FOIL PACKET AND THEN TIP THE UN-OPENED CATHETER PACKAGE END-TO-END. THE CATHETER ACTS LIKE A MAGNET TO ATTRACT THE WATER AND ACTIVATE ITS SLIPPERY COATING. FOLLOW THESE STEPS FOR BEST RESULTS. 1. RELEASE THE WATER. PRIOR TO OPENING THE SEALED CATHETER POUCH: A. APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. B. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. 2. WET THE CATHETER. A. HOLD PACKAGE WITH PRINTED SIDE UP. B. TIP PACKAGE END-TO-END THREE TO SIX TIMES TO WET CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. 3. USE THE CATHETER. A. PEEL BACK PACKAGE TO EXPOSE FUNNEL END OF CATHETER. B. IF DESIRED, USE THE SELF-ADHESIVE TAPE ON THE PACKAGE TO TEMPORARILY ATTACH THE POUCH TO ANY DRY VERTICAL SURFACE. C. REMOVE CATHETER AND USE ACCORDING TO PHYSICIAN¿S INSTRUCTIONS. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE, DO NOT RESTERILIZE. MADE OF SILICONE ELASTOMER.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03912 | 2023-05-31 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT PATIENT STATED THEY WERE HOSPITALIZED WITH MAJOR URINARY TRACT INFECTION AND WAS PEEING BLOOD. PATIENT WAS ADAMANT THAT THE INTERMITTENT CATHETERS CAUSED THE URINARY TRACT INFECTION. MEDICAL INTERVENTION WAS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03917 | 2023-05-31 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MALE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE URINARY TRACT INFECTION WAS SO SEVERE, WENT INTO BLOOD STREAM STARTING BACK IN DECEMBER AND THE MOST RECENT STARTED MARCH. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03919 | 2023-05-31 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT STERILE WATER WAS INJECTED BUT THE FOLEY CATHETER BALLOON DID NOT INFLATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03915 | 2023-05-31 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC 3 GO MALE COUDE CATHETER | MALE COUDE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAD URINARY TRACT INFECTION ,9 BOXES WERE FROM JUGX9006, 4 BOXES WERE FROM JUGT0289 AND 1 BOX WAS FROM JUGV0690 AND REQUESTED FOR COMPENSATION AND DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. REPRESENTATIVE ADVISED AT THIS TIME WE WILL REPLACE AFFECTED LOTS AND DISCARD AFFECTED ONES. MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02169 | 2023-05-31 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT DURING THE PREPARATION OF A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY HAD A HOLE IN THE ARTERIAL LEG. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 05/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03913 | 2023-05-31 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE URINARY TRACT INFECTION AND WAS HOSPITALIZED TWICE. PATIENT WAS TREATED WITH MEDICINE FROM DOCTOR. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03916 | 2023-05-31 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. IT WAS UNKNOWN IF THE INTERMITTENT CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-02199 | 2023-06-01 | 2023-03-24 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER W/ SHERLOCK 3CG (TPS) STYLET FULL KIT (4F) (SL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | CUSTOMER REPORTED THAT A PICC LINE HAS LED TO AN EXTRAVASATION. THE LINE WAS INSERTED TO 44CM AND HAD A RUPTURE 1.5 CM ABOVE. PATIENT ATTENDED FOR PUMP REMOVAL. PICC BLED BACK FINE, MINIMAL RESISTANCE ON FLUSHING. ADDITIONAL INFORMATION RECEIVED 5/24/2023: CUSTOMER REPORTED THE OXALIPLATIN AND CALCIUM FOLINATE LEAKED INTO THE TISSUE. PATIENT EXPERIENCED PAIN AROUND THE PICC INSERTION AREA AND INFORMED STAFF. ON INSPECTION THERE WAS OBVIOUS SWELLING TO THE AREA OF FLUID IN THE SUBCUTANEOUS TISSUE. PATIENT RECEIVED THE 1ST LINE TREATMENT WHERE INFUSION WAS STOPPED IMMEDIATELY AND TREATMENT DISCONNECTED. 40 ML OF BLOOD ASPIRATED FROM PICC LINE (NOT FLUSHED). ARM ELEVATED. HEAT APPLIED DUE TO THE ADVERSE EFFECTS OF COLD WITH OXALIPLATIN AND AS PER POLICY. INJECTED HYDRAULINASE AROUND EXTRAVASATION AREA. DISCUSSION HELD WITH PLASTICS TEAM AT (B)(6) AND NO INTERVENTION WANTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02198 | 2023-06-01 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 6F RADIOLOGY BASIC TRAY (DUAL-LUMEN) (135 CM GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3276135 |  | IT WAS REPORTED THAT, PICC LINE FAULTY. LINE LEAKAGE. (B)(6) 2023: PICC WAS CHANGED OUT AFTER THE LEAKAGE WAS OBSERVED. NO PATIENT HARM REPORTED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02204 | 2023-06-01 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | BARD DL POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED "THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. THIS EVENT INVOLVED A DL PICC IN THE UPPER ARM." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02202 | 2023-06-01 | 2023-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER (4F) (SINGLE-LUMEN) (20CM) | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000756 |  | IT WAS REPORTED BY THE CUSTOMER THAT THERE ARE LEAKING FROM INSERTION SITE ON A MIDLINE PLACED ON 4/29 AND REMOVED 5/2. LINE PLACED ON LEFT CEPHALIC. A NEW CATHETER WAS INSERTED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02203 | 2023-06-01 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD DL POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED "THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. THIS EVENT INVOLVED A DL PICC IN THE UPPER ARM." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02205 | 2023-06-01 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F DUAL-LUMEN CATHETER BASIC TRAY (NITINOL GUIDEWIRE 135CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8375335 |  | IT WAS REPORTED THAT THE END OF TIP AT THE CATHETER IS BLOCKED. NO OTHER INFORMATION WAS PROVIDED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02201 | 2023-06-01 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED, EXPERIENCED ACCUCATH USER UTILIZED A 20 GA ACCUCATH. WENT TO SAFETY THE NEEDLE OF THE ACCUCATH IV AND THE NEEDLE FAILED TO SAFETY ADDITIONAL INFO RECEIVED 05/18/2023 "THE RETRACTING MECHANISM WAS FAULTY. SHE ATTEMPTED TO PUSH THE RETRACTION BUTTON AND IT BECAME STUCK IN THE DOWNWARD POSITION. ANOTHER VASCULAR RN HAPPEN TO STOP IN THE ROOM AND ATTEMPTED TO TROUBLE SHOOT ALSO BUT WAS UNSUCCESSFUL. THE NEEDLE, CATH AND GUIDE WIRE WERE ALL WITHDRAWN FROM THE PATIENT. ONCE OUT OF THE PATIENT THE RN ATTEMPTED TO RETRACT THE NEEDLE BY PRESSING IT ON THE COUNTER. IT DID RETRACT BUT POPPED RIGHT BACK OUT AS IF WHATEVER KEEPS THE NEEDED IN THE CHAMBER WAS NOT WORKING." ADDITIONAL INFORMATION RECEIVED 05/22/2023: IT APPEARS THE ISSUE WAS BOTH A NEEDLE RETRACTION AND A SAFETY SHIELD ACTIVATION FAILURE. NO PATIENT HARM WAS REPORTED, BUT PATIENT WAS STUCK AND EVERYTHING HAD TO BE REMOVED. PATIENT WAS A FEMALE IN HER 80'S. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02206 | 2023-06-01 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK CV PLUS SLIDING POST, TRICOT HYDROCOLLOID CRESCENT PAD, PIGTAIL PAD | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | CV0220CE |  | IT WAS REPORTED AT 8:10 A.M. (B)(6)2023, IT WAS FOUND THAT THE CATHETER FIXATION DEVICE OF THE PATIENT'S THORACIC CLOSED DRAINAGE TUBE FELL OFF DURING THE MORNING SHIFT. AFTER OPENING THE CENTRAL VENOUS CATHETER PACKAGE OF PERIPHERAL INTUBATION, IT WAS FOUND THAT THE CATHETER FIXATION DEVICE WAS SEPARATED FROM THE APPLICATION, WHICH WAS IMMEDIATELY REPLACED AND REPORTED TO THE MEDICAL EQUIPMENT DEPARTMENT. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: PATIENT HAD NO SYMPTOMS RELATED TO THE DRAIN DISLODGEMENT AND A NEW DRAIN WAS NOT PLACED. CUSTOMER STATES "THE STATLOCK COULD NOT BE FIXED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03958 | 2023-06-01 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MAGIC3 GO MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAD URINARY TRACT INFECTION ,9 BOXES WERE FROM JUGX9006, 4 BOXES WERE FROM JUGT0289 AND 1 BOX WAS FROM JUGV0690 AND REQUESTED FOR COMPENSATION AND DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. REPRESENTATIVE ADVISED AT THIS TIME WE WILL REPLACE AFFECTED LOTS AND DISCARD AFFECTED ONES. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "ROUGH OR IRREGULAR SHAPE PROTRUSIONS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02178 | 2023-06-01 | 2023-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5733690 | 5733690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER SHEATH ALLEGEDLY DEFORMED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORD IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2024) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02180 | 2023-06-01 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PRISTINE LONG-TERM HEMODIALYSIS CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5403190 | 5403190 |  | IT WAS REPORTED THAT SOME TIME POST A DIALYSIS CATHETER PLACEMENT, THE TUBING NEAR THE HUB ALLEGEDLY KINKED. IT WAS FURTHER REPORTED THAT DEVICE ALLEGEDLY HAD POOR FLOW. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02179 | 2023-06-01 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393190 | 5393190 |  | IT WAS REPORTED THAT SOME TIME POST A DIALYSIS CATHETER PLACEMENT, THE TUBING NEAR THE HUB ALLEGEDLY KINKED. IT WAS FURTHER REPORTED THAT DEVICE ALLEGEDLY HAD POOR FLOW. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03959 | 2023-06-01 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0167SI24 | 0167SI24 |  | IT WAS REPORTED THAT CUSTOMER HAD GOTTEN COMPLAINTS FROM A COUPLE OF THE UROLOGY DOCTORS ABOUT FOLEY CATHETER. CUSTOMER HAD FOLEY CATHETER WITH ISSUES OF THEIR BALLOONS LEAKING. ONE SURGEON FILLED THEIR BALLOON WITH 30 CC¿S, PATIENT WAS TRANSFERRED TO THE FLOOR, AND THIS AM, IT WAS FOUND THAT THE CATHETER HAD MIGRATED TO THE PATIENT¿S PROSTATE. ANOTHER LAST NIGHT FILLED THEIR BALLOON IN THE OPERATION ROOM AND BY THE TIME THEY GOT TO PACU, IT CAME OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03961 | 2023-06-01 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0167SI24 | 0167SI24 |  | IT WAS REPORTED THAT CUSTOMER HAD GOTTEN COMPLAINTS FROM A COUPLE OF THE UROLOGY DOCTORS ABOUT FOLEY CATHETER. CUSTOMER HAD FOLEY CATHETER WITH ISSUES OF THEIR BALLOONS LEAKING. ONE SURGEON FILLED THEIR BALLOON WITH 30 CC¿S, PATIENT WAS TRANSFERRED TO THE FLOOR, AND THIS AM, IT WAS FOUND THAT THE CATHETER HAD MIGRATED TO THE PATIENT¿S PROSTATE. ANOTHER LAST NIGHT FILLED THEIR BALLOON IN THE OPERATION ROOM AND BY THE TIME THEY GOT TO PACU, IT CAME OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02200 | 2023-06-01 | 2023-05-03 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER W/ SHERLOCK 3CG (TPS) (4F) (SL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED, PATIENT ATTENDED FOR C5 FOLFOX, NO PROBLEMS DOCUMENTED ON ACCESS. PATIENT HAD RECEIVED 225ML OF TREATMENT AND EXPRESSED CONCERN WITH PAIN AROUND PICC SITE AND SURROUNDING TISSUE. STOPPED IMMEDIATELY AND LINE ASPIRATED 20ML OF BLOOD EASILY. COMMENCED TREATMENT FOR EXTRAVASATION. LINE REMOVED AND FLUSHED AT CHAIR SIDE. RUPTURE IN LINE 1.5 AM ABOVE INSERTION SITE, LINE WAS INSERTED TO 44CM. ADDITIONAL INFORMATION 5/18/23: NEW LINE WAS NOT PLACED ON THE EXTRAVASATION DAY BUT DID ATTENDED FOR A NEW LINE YESTERDAY BUT IT WAS NOT PLACED AS CONCERN OF CELLULITIS THROMBUS DELAYED TISSUE DAMAGE (PLASTICS TEAM DID NOT FEEL THIS WAS DELAYED TISSUE DAMAGE). REGULAR CLINICAL REVIEWS WERE PERFORMED AND AFFECTED AREA FULLY RESOLVED BUT HAS NOW DEVELOP ERYTHEMA AND SWELLING (AS ABOVE) SO ACTUAL HARM NOT YET KNOWN. THE TREATMENT ON THE DAY OF EXTRAVASATION WAS ABANDONED AS HE NEEDED TO GO HOME WITH A PUMP AND THE CONSULT DECIDED TO OMIT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02183 | 2023-06-01 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5883690 | 5883690 |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY LEAKED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02207 | 2023-06-01 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGLIDE PRO, REINFORCED TIP 18G 10 CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED THE NEEDLE WAS NOT ABLE TO BE SAFETY ENGAGED. NO OTHER INFORMATION WAS PROVIDED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02197 | 2023-06-01 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED WHEN THE PUNCTURE NEEDLE IS REMOVED, THE PROTRUDING PART OF THE END OF THE GUIDE WIRE FRICTS THE PUNCTURE NEEDLE. WHEN THE INTUBATION SHEATH IS SENT, THE END OF THE GUIDE WIRE CANNOT ENTER THE INTUBATION SHEATH. ONLY AFTER THE PROTRUDING PART OF THE END OF THE GUIDE WIRE IS CUT WITH SCISSORS, CAN THE END OF THE GUIDE WIRE ENTER THE INTUBATION SHEATH. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GUIDEWIRE IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE PHOTOGRAPH AND ONE VIDEO OF A GUIDEWIRE WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE RETURNED PHOTOGRAPH AND VIDEO SHOWED A GUIDEWIRE WITH A SLIGHT BEND NEAR THE MIDDLE OF ITS LENGTH. ONE END OF THE GUIDEWIRE APPEARS TO BE DAMAGED; HOWEVER, NO DETAIL OF THAT APPARENT DAMAGE COULD BE DISCERNED. THERE WERE NO DISTINGUISHING FEATURES IN THE RETURNED PHOTOGRAPH OR VIDEO OF THE DEVICE WHICH COULD AID IN IDENTIFICATION OF A SPECIFIC ROOT CAUSE. THEREFORE, THIS COMPLAINT WAS CONFIRMED AS CAUSE UNKNOWN. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02186 | 2023-06-01 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED, "THE SPECIALIST NURSE PLACED CATHETERIZATION ON THE PATIENT AND FOUND THAT THERE WAS A FLOC FOREIGN BODY IN THE CATHETER PACKAGE WHEN TEARING THE CENTRAL VENOUS CATHETER PACKAGE FOR PERIPHERAL INTUBATION, AND IMMEDIATELY REPLACED A NEW CENTRAL VENOUS CATHETER PACKAGE FOR PERIPHERAL INTUBATION. NO ABNORMALITY WAS FOUND AND IT WAS USED NORMALLY." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02194 | 2023-06-01 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 20CM (SHORT-TERM) (DUAL-LUMEN) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5593200 |  | IT WAS REPORTED THE PATIENT WAS ADMITTED TO HOSPITAL DUE TO DRUG POISONING. DURING TEMPORARY RENAL DIALYSIS CATHETER PLACEMENT IN FEMORAL VEIN ON (B)(6) 2023.20, DIALYSIS CATHETER WAS USED, BUT THE PUNCTURE NEEDLE COULD NOT BE PULLED OUT NORMALLY AFTER THE GUIDE WIRE WAS SENT INTO THE PUNCTURE NEEDLE, AND THE PUNCTURE NEEDLE AND GUIDE WIRE WERE IMMEDIATELY PULLED OUT OF THE PATIENT. AFTER THE GUIDE WIRE WAS PULLED OUT FORCIBLY, THE GUIDE WIRE WAS DEFORMED AND COULD NOT BE USED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03963 | 2023-06-01 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166V12S |  | IT WAS REPORTED THAT THE PATIENT WAS INSERTED A FOLEY CATHETER AFTER THE EPIDURAL PROCEDURE TO CONTINUOUSLY DRAIN URINE. THE PATIENT'S BLADDER WAS DISCOVERED TO BE FULL PRIOR TO THEIR IMPENDING CAESAREAN SECTION. ANOTHER FOLEY CATHETER FROM THE SAME BATCH WAS REINSERTED. IT WAS NOT DRAINING WELL AND FOUND THE BLADDER FULL DURING THE PATIENT'S CAESAREAN SECTION. THAT WAS THEIR SECOND CASE. THE AFFECTED CATHETERS WERE CUT INTO 2 SECTIONS BY THE DOCTOR. AS PER THE ADDITIONAL INFORMATION RECEIVED VIA IBC ON 25MAY2023, CLARIFIED THAT IT WAS DIFFICULT TO REMOVE URINE AND ACCORDING TO THE END USERS, IT SEEMED TO BE A BLOCKAGE FOR THE URINE NOT TO DRAIN OUT AND A LARGER SIZED FOLEY WAS SET TO REMOVE URINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03967 | 2023-06-01 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT PATIENT USING MALE INTERMITTENT CATHETER AND THEY RETURNED HOSPITAL . DUE TO A URINARY TRACK INFECTION AND STATED NONE OF THE RECALLED CATHETERS LEFT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTIONS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-03974 | 2023-06-01 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE PATIENT'S WIFE DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT'S WIFE STATED THAT HER HUSBAND HAS A URINARY TRACT INFECTION, AND THEY THINK IT COULD BE FROM THE MAGIC 3. PATIENT'S WIFE STATED THAT THE CATHETERS DO NOT HAVE ENOUGH LUBRICANT WHICH MAKES IT DIFFICULT TO USE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. PER FOLLOW UP VIA PHONE ON 25MAY2023, IT WAS REPORTED THAT THE CUSTOMER DID DEVELOP A UTI FROM CATHETERS THAT WERE NOT LUBRICATED. HE WAS PRESCRIBED ANTIBIOTICS BUT OPTED TO KEEP USING THE CATHETERS WITH A NEW LOT #. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-03975 | 2023-06-01 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD A BAD URINARY TRACT INFECTION BUT NOT SURE IF IT WAS FROM USING THESE MAGIC 3 MALE GO INTERMITTENT CATHETERS OR NOT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN; WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-03977 | 2023-06-01 | 2023-05-13 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAD A URINARY TRACT INFECTION BACK IN FEBRUARY. REPRESENTATIVE ADVISED TO STOP USAGE AND THROW AWAY. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-03973 | 2023-06-01 | 2023-05-13 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50812G | 50812G |  | IT WAS REPORTED THAT PATIENT EXPERIENCED AN URINARY TRACT INFECTION AND WAS ON ANTIBIOTIC. THEY THOUGHT IT MIGHT HAVE BEEN FROM RECALL. | CONCOMITANT MEDICAL PRODUCTS: W1DR01 IPG, IMPLANTED (B)(6) 2022 MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-03976 | 2023-06-01 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MAGIC3 GO MALE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD SEVERAL URINARY TRACT INFECTIONS FROM USING THESE MAGIC 3 MALE GO INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS, MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
17038087 | 2023-06-01 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET LOT # REHN2437 RUBBER-LIKE STOPPER LEAKING WITH FLUSHING. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03978 | 2023-06-01 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MAGIC3 GO MALE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD A LOT OF URINARY TRACT INFECTION. IT WAS UNKNOWN IF THE INTERMITTENT CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02185 | 2023-06-01 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED, PATIENT ATTENDED FOR C2 FOLFOX, PICC DID NOT BLEED BACK. UROKINASE USED AND THEN BLED AND FLUSHED FINE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02184 | 2023-06-01 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED, PATIENT ATTENDED FOR PICC CARE AND BLOODS. LINE NOT BLEEDING BACK, UROKINASE USED. BLED BACK BUT STILL SOME RESISTANCE ON FLUSHING. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03980 | 2023-06-01 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT FOLEY CATHETER FELL OUT FROM THE PATIENT DUE TO WATER LEAK IMMEDIATELY AFTER THE PLACEMENT. ALSO STATED THAT WATER LEAK OBSERVED FROM THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03965 | 2023-06-01 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 | FOL0102 | FOL0102 |  | IT WAS REPORTED THAT THE HITCH OF THE CLAMP WAS SEPARATING FROM THE ADHESIVE WHEN REMOVING FROM LEG. PER SAMPLE RECEIVED ON (B)(6) 2023, THE CUSTOMER RETURNED TWO USED DEVICES. PER ADDITIONAL INFORMATION RECEIVED VIA SAMPLE FORM ON (B)(6) 2023, CUSTOMER STATED THAT THIS TYPE UNIT HAD FAILED SEVERAL TIMES CAUSING THE CATHETER TO SEPARATE ITSELF FROM THE BAG AT THE ATTACH POINT. DEVICE WAS REPLACED. CUSTOMER STATED THAT UPON INSPECTION OF EACH FAILED DEVICE THEY ALL HAVE FAILED FOR THE SAME REASON, IT DID NOT APPEAR THAT THE GLUE TYPE OR APPLICATION PROCEDURE WAS WORKING OUT IN ATTACHING THE STATLOCK SWIVEL UNIT TO THE CLOTH ADHESIVE BASE. CUSTOMER SHOWER EACH MORNING WITH UNIT ATTACHED, USING SOAP AND WATER ON THE BAG AND EXPOSED PORTION OF THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02187 | 2023-06-01 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | PPICC, IR, 135 CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275335J |  | IT WAS REPORTED THAT DURING A CATHETER PLACEMENT PROCEDURE, BECAUSE THE MOVEMENT OF THE GUIDEWIRE WORSENED, THE TIP OF THE GUIDEWIRE WAS BROKEN WHEN THE GUIDEWIRE WAS WITHDRAWN WITH FORCE. THE GUIDEWIRE SEGMENT WAS REMOVED FROM THE PATIENT¿S BODY. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02190 | 2023-06-01 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PPICC PROVENA 4F DL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000951 |  | IT WAS REPORTED VIA MEDWATCH, "PATIENT ORDERED TO RECEIVE A DOUBLE LUMEN PICC LINE FOR TPN. THE IV TEAM NURSE WAS SETTING UP THE TRY FOR INSERTION AND PRE-FLUSHING THE CATHETER WAS DONE PRIOR TO STARTING THE PROCEDURE. THE T-PORT THAT WAS ATTACHED TO THE PICC LINE WAS LEAKING FROM THE STOPPER WHEN FLUSHED IN THE LOCATION OF WHERE THE WIRE WAS INSERTED. THE PRODUCT WAS NOT USED FOR INSERTION AND WAS DISPOSED OF. A TRIPLE LUMEN PICC WAS INSERTED INSTEAD WITH DIFFERENT TYPE OF T-PORT ATTACHMENT. NO COMPLICATIONS WERE OBSERVED. NO HARM TO THE PATIENT." | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION, AS THE DEVICE WAS DISCARDED AFTER THE EVENT OCCURRED. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02189 | 2023-06-01 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED, PATIENT ATTENDED FOR C4 FOLFOX, PUMP PRESSURE CAUSING ALARMS. PICC REDRESSED AND REPOSITIONED. NO FURTHER PROBLEMS. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02188 | 2023-06-01 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER (4F) (SINGLE-LUMEN) (20CM) | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000756 |  | IT WAS REPORTED BY THE CUSTOMER THAT THERE ARE LEAKING FROM INSERTION SITE ON A MIDLINE THAT WAS PLACED ON 5/2 AND REMOVED ON 5/4/2023. A NEW MIDLINE WAS PLACED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03985 | 2023-06-01 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THESE INTERMITTENT CATHETERS CAUSED THE PATIENT TO DEVELOP A URINARY TRACT INFECTION MID-FEBRUARY OR CLOSE TO IT AND WAS STILL ON MEDICATION AND HAS AN APPOINTED WITH UROLOGY ON (B)(6) 2023. THE PATIENT DISCARDED PRODUCT URINARY TRACT INFECTION NOTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02193 | 2023-06-01 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | A710962 | A710962 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, A BUTTONHOLE HAS ALLEGEDLY DEVELOPED IN THE AREA OF ACCESS. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03989 | 2023-06-01 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER WAS UNABLE TO BE INJECTED DUE TO WATER LEAKAGE FROM THE INFLATION VALVE. UPON CHECKING THE CATHETER, THE INFLATION VALVE GOT DISCONNECTED. NO PROBLEM WAS OBSERVED WITH THE CATHETER DURING INFLATION ACCORDING TO THE INQUIRY SHEET ATTACHED. WATER WAS INJECTED GENTLY. USUALLY A BALLOON CAN BE INFLATED WITHIN SECONDS BUT IT TOOK MORE THAN USUAL INFLATING THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00387 | 2023-06-01 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQF75104 | CQF75104 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE PERIPHERAL DIALYSIS FISTULA VIA LEFT ELBOW, THE PTA BALLOON ALLEGEDLY RUPTURED AT 15 ATM. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE CONQUEST 40 PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, NO KINKS WERE NOTED TO THE CATHETER SHAFT AND NO ANOMALIES TO THE LUERS/Y-BODY. DURING FUNCTIONAL EVALUATION, AN IN-HOUSE PRESTO INFLATION DEVICE WAS USED TO INFLATE THE BALLOON AND WATER WAS NOTED LEAKING FROM THE BALLOON. THE BALLOON FIBERS WERE STRIPPED AND UNDER MICROSCOPIC EXAMINATION, A COMPOUND RUPTURE WAS NOTED TO THE BALLOON. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED BALLOON RUPTURE AS A COMPOUND RUPTURE WAS NOTED TO THE BALLOON UNDER MICROSCOPIC EXAMINATION. A DEFINITIVE ROOT CAUSE FOR THE REPORTED COMPOUND BALLOON RUPTURE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. D4 (EXPIRATION DATE: 09/2025). THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02196 | 2023-06-01 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F DUAL-LUMEN CATHETER BASIC TRAY (NITINOL GUIDEWIRE 135CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8275335 |  | IT WAS REPORTED THAT THE END OF TIP AT THE CATHETER IS BLOCKED. NO OTHER INFORMATION WAS PROVIDED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02195 | 2023-06-01 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT MRI ISP W/O SUTURE PLUGS, 9.6 FR SILICONE SL CATHETER | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1809661 | 1809661 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, WHEN TRIED TO REMOVE THE INNER SHEATH PART OF THE LUER-LOCK PIECE THAT IS ATTACHED TO THE OUTER GREY SHEATH ALLEGEDLY BROKE OFF AND BLOOD ALLEGEDLY STARTED TO SHOOT OUT. IT WAS FURTHER REPORTED THAT THE GUIDE WIRE WAS ALLEGEDLY DID NOT ENTER THROUGH THE OUTER SHEATH AND IT ALLEGEDLY OCCLUDED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 07/2024). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02182 | 2023-06-01 | 2018-04-24 | I, n, j, u, r, y | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1808001 |  | IT WAS REPORTED THAT TWO YEARS SIX MONTHS AND ELEVEN DAYS PORT DEVICE IMPLANT PROCEDURE, THE PATIENT COMPLAINED FOR ACUTE ONSET OF RIGHT SIDED CHEST PAIN. FURTHER, THE PATIENT UNDERWENT EMERGENT SURGERY TO REMOVE THE FAILED DEVICE AND REPLACE IT WITH ANOTHER DEVICE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02181 | 2023-06-01 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1678300 | 1678300 |  | IT WAS REPORTED THAT THREE DAYS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY FELT LUMP IN THROAT. IT WAS FURTHER REPORTED THAT THE PATIENT WAS ABLE TO FEEL THE PORT ON THE SIDE WHERE IT WAS PLACED WHILE SWALLOWING. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02224 | 2023-06-02 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG 4F SINGLE-LUMEN PICC PERIPHERALLY INSERTED CENTRAL VENOUS CATHETER | CATHETER, PERCUTANEOUS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7715405 |  | IT WAS REPORTED BY THE CUSTOMER THAT DURING THE FIRST SESSION OF CHEMOTHERAPY FOLLOWING CATHETER PLACEMENT ON (B)(6) 2023 AND THE SECOND SESSION OF TREATMENT, THE PATIENT REPORTED PAIN NEAR THE INSERTION SITE. PAIN WAS STILL FELT IN SITU DURING THE THIRD SESSION OF CHEMOTHERAPY. PART OF THE CATHETER WAS REMOVED, AND LIQUID WAS FOUND LEAKING FROM THE SITE 1.5 CM AWAY FROM THE INSERTION SITE. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: THE LEAKING FLUID WAS SALINE. THERE WERE NO COMPLICATIONS, AND NO TREATMENT WAS NEEDED. THE EVENT DID NOT CAUSE A CLINICALLY SIGNIFICANT DELAY IN TREATMENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02214 | 2023-06-02 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 5F MICROEZ UNIVERSAL PTFE MICROINTRODUCER SAFETY KIT | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0678950 |  | IT WAS REPORTED BY THE CUSTOMER ¿I WAS DOING A PICC LINE TODAY AND IN THE MICROEZ MICROINTRODUCER KIT, THE END OF THE GUIDEWIRE WAS BENT AND GOT CAUGHT ON THE END OF THE NEEDLE WHEN TRYING TO EXTRACT IT. THANKFULLY I WAS TRYING TO DO SO AFTER THE WIRE AND NEEDLE HAD BEEN PULLED OUT OF THE PATIENT¿S ARM.¿ NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02217 | 2023-06-02 | 2023-04-28 | I, n, j, u, r, y | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (20G)(8CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120080T |  | IT WAS REPORTED VIA MEDWATCH ¿THIS RN WAS ASKED TO RE-VISIT A POSSIBLE PICC LINE FOR CONTINUED ELECTROLYTE SUPPLEMENTATION AND FREQUENT LAB DRAWS. PATIENT DECLINED "I HAVE A HEART ISSUE, I DON'T LIKE THE IDEA OF THAT". PATIENT EDUCATED THAT PICC LINE DOES NOT GO INTO THE HEART BUT DECLINED. DISCUSSED THE OPTION OF A MIDLINE. EXPLAINED THAT IT IS 8CM AND WOULD NOT GO PAST SHOULDER. PATIENT AGREED. ARM PREPPED WITH A MINUTE. CHG SCRUB, AND POWER MIDLINE ATTEMPTED. ULTRASOUND USED. BLOOD RETURN NOTED AND ADVANCEMENT. GUIDEWIRE WITHOUT DIFFICULTY. RESISTANCE MET WHILE ADVANCING CATHETER. ADVANCEMENT STOPPED. IMMEDIATELY AND REMOVAL OF ENTIRE POWERGLIDE SYSTEM. CATHETER NOTED TO NOT BE INTACT. APPROXIMATELY. HALF OF THE CATHETER NOTED IN THE POWERGLIDE SYSTEM. TORNIQUET APPLIED AND DIRECT PRESSURE. PHLEBOTOMIST IN ROOM AND ASSISTED WITH PRESSURE. PROVIDERS NOTIFIED. PATIENT WITHOUT SHORTNESS OF BREATH. NO COMPLAINTS.¿ | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02216 | 2023-06-02 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY CUSTOMER THAT PICC LINE NOTED TO BE LEAKING WHEN FLUSHED. FLUID NOTED TO BE LEAKING FROM CATHETER AND CONNECTOR. NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02215 | 2023-06-02 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 |  | 7617405J |  | IT WAS REPORTED THAT LEAKAGE WAS FOUND NEAR THE CATHETER CONNECTOR. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKAGE NEAR THE CATHETER CONNECTOR WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR NXT GROSHONG CATHETER WITH THE TWO-PIECE CONNECTOR. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH CATHETER DAMAGE CAUSED BY CONTACT WITH A SHARP EDGED INSTRUMENT SUCH AS A SCALPEL. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED ON THE DISTAL END OF THE EXTENSION TUBING OF THE PROXIMAL SEGMENT OF THE TWO-PIECE CONNECTOR. MICROSCOPIC EXAMINATION OF THE CATHETER SPLIT CONFIRMED IT WAS TYPICAL OF CONTACT WITH A SHARP BLADED INSTRUMENT, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: THE FRACTURE SURFACE WAS REFLECTIVE IN NATURE AND CONTAINED A STRIATED PATTERN. THIS CAN OCCUR DUE TO PATTERN TRANSFER OF THE SHARPENED EDGE TYPICALLY FOUND ON A SCALPEL-TYPE INSTRUMENT. SHARPLY FORMED FRACTURE EDGES. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02219 | 2023-06-02 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2174108 |  | IT WAS REPORTED BY THE CUSTOMER "THE CATHETER IS DISSOLVED ON THE OUTSIDE, AT THE WINGS." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02220 | 2023-06-02 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED VIA MS&S ¿PATIENT HAS A POWERPICC PROVENA INSERTED JUST ABOVE THE ELBOW. IT WILL PUMP INFUSE TPN WITH NO DIFFICULTY ACCEPT FOR THE OCCASIONAL ALARM. IT WILL FLUSH AND GET BLOOD RETURN, BUT WITH SOME RESISTANCE. IT WILL NOT INFUSE NORMAL SALINE VIA GRAVITY AND THE PATIENT WOULD LIKE TO KNOW WHAT TO DO. THE HCP WILL NOT ORDER A CATHFLOW TREATMENT BECAUSE THEY ARE ABLE TO INFUSE AT THE INFUSION CLINIC. THE TROUBLES ARE ALL AT HOME.¿ MS&S RESPONSE TO PATIENT ¿CONFIRMED THAT THE BAG IS RAISED AS HIGH AS IT WILL GO. ENCOURAGED THE PATIENT TO TRY REPOSITIONING THE ARM, ABOVE HEAD, ACROSS THE CHEST, OR TO THE SIDE BUT ELEVATED ON A PILLOW OR TABLE. OTHERWISE THERE IS NOT A LOT THAT CAN BE DONE AND A CONVERSATION NEEDS TO HAPPEN WITH THE HCP. PATIENT REPORTED THAT THEY WILL TRY REPOSITIONING AND HAVE A TALK WITH THE PROVIDER IN A FEW DAYS WHEN THEY GO INTO THE CLINIC FOR ANOTHER INFUSION.¿ ADDITIONAL INFORMATION RECEIVED MAY 23 2023: PATIENT DID NOT REPORT ANY VISUAL DEFECTS TO THE CATHETER. THE EVENT WAS NOT URGENT OR LIFE THREATENING BUT DID CAUSE A DELAY IN ADMINISTRATION OF MEDICATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION, AS THE DEVICE REMAINS IN THE PATIENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02229 | 2023-06-02 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | CUSTOMER REPORTS INSTANCES OF PICCS LEAKING. NO OTHER DETAILS PROVIDED FOR EVENTS. THE EXACT QUANTITY OF LEAKING PICCS WAS NOT PROVIDED TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02211 | 2023-06-02 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD DUAL LUMEN POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. DUAL LUMEN PLACED IN THE UPPER ARM. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02231 | 2023-06-02 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. TRIPLE LUMEN PLACED IN UPPER ARM. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02228 | 2023-06-02 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 2194108 |  | CUSTOMER REPORTED THAT THE PICC LINE CATHETER FOR CENTRAL VENOUS USE WAS USED TO DELIVER CYTOTOXIC SUBSTANCES. PATIENT GIVEN PICCLINE ON (B)(6) 2023, LAID ON 4 CM, PRIOR TO CHEMOTHERAPY TREATMENT- CISPLATIN AND WITH PREHYDRATION WHICH IS FINE TO GIVE. PATIENT IS WITH HIS DOCTOR IN THE WARD (B)(6) 2023 WHO SEES AFTER REARRANGEMENT OF PICCLINE. WHEN FLUSHING THAT IT LEAKS. SODIUM CHLORIDE IN THE NEWLY APPLIED DRESSING. THE PATIENT IS ASKED TO CONTACT US. PATIENT CALLS AND ARRIVES. NURSE REMOVES DRESSING. PUTS COMPRESS UNDER INJECTION SITE AND SECURA CATH. SEES LEAKAGE. PICCLINE NURSE IS CALLED. THE LID OF THE SECURACATH IS LIFTED OFF. THE PICCLINE NURSE PULLS OUT THE PICCLINE A LITTLE AT A TIME. AFTER 1.5 CM A HOLE CAN BE SEEN ON THE CATHETER, I.E. AT 5.5 CM. A NEW CATHETER NEEDS TO BE INSERTED. NEW ACCESS NEEDED, EXTENDED WAITING TIME. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02222 | 2023-06-02 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 2275108 |  | CUSTOMER REPORTED A LEAKING STOPPER. ADDITIONAL INFORMATION RECEIVED 5/31/2023: IT WAS REPORTED BY THE CUSTOMER "THE PRODUCT WAS USED ON A PATIENT HOWEVER WE HAD TO TROUBLESHOOT THE INSERTIONS IN ORDER TO ENSURE A SAFE INSERTION. THE DEVICE WAS DISCARDED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02230 | 2023-06-02 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | IT WAS REPORTED BY THE CUSTOMER THAT PRIOR TO INSERTION IN THIS PATIENT, THE NURSE OPENED THE PACKAGE OF THE PICC CATHETER, PULLED OUT SOME OF THE GUIDE WIRE AND FOUND WAX-LIKE SUBSTANCES ON ITS SURFACE. THEREFORE, THE NURSE DARED NOT TO CONTINUE USING IT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF MATERIAL ON THE STYLET WAS CONFIRMED; HOWEVER, THE EXACT ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS 4FR S/L POWERPICC CATHETER WITH INLAID STYLET. THE STYLET HAD BEEN PARTIALLY WITHDRAWN THROUGH THE T-LOCK ASSEMBLY. SEVERAL BENDS WERE OBSERVED IN THE STYLET. MATERIAL APPEARED TO BE FLAKING OFF OF THE WIRE. MICROSCOPIC INSPECTION OF THE SAMPLE CONFIRMED THE PRESENCE OF WHITE MATERIAL ON THE STYLET. THE MATERIAL APPEARED BUNCHED AND PEELING. THE MATERIAL APPEARED TO PARTIALLY COAT THE STYLET WIRE. IT APPEARED THAT THE HYDROPHILIC COATING APPLIED TO THE STYLET WIRE DURING PRODUCT MANUFACTURE WAS BUNCHING AND PEELING. IT COULD NOT BE DETERMINED WHAT CAUSED THE DELAMINATION OF THE COATING. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE EXPOSURE TO INCOMPATIBLE CHEMICALS AND UNIDENTIFIED ENVIRONMENTAL FACTORS AFFECTING THE PROPERTIES OF THE COATING. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02227 | 2023-06-02 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER BASIC KIT (4F) (SINGLE-LUMEN) (20CM) | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | P6154118 |  | CUSTOMER REPORTED MIDLINE WITH LEAKAGE ON THE T-JUNCTION OF THE PURPLE TUBING, SOAKED DRESSING. IMPACTS ON THE PATIENT OR USER: DIFFICULTY OF ADMINISTRATION OF THERAPY, INFECTIOUS RISK, NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02218 | 2023-06-02 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | CUSTOMER REPORTED THAT ONE MONTH AFTER THE IMPLANT WHICH TOOK PLACE ON (B)(6) 2023, ON (B)(6) 2023 REPORTED PAIN DURING THE ADMINISTRATION OF THE DRUGS, ON PALPATION THE NURSE NOTICED A SWELLING FROM THE EXIT SITE UP TO 4 CM TOWARDS THE AXILLARY AREA. THE INFUSION, WHICH FIRST PROVIDED FOR THE IRRIGATION OF THE CATHETER WITH PHYSIOLOGICAL SOLUTION, WAS INTERRUPTED TO CHECK THE CASE. ONCE THE RUPTURE IN THE INTRAVASCULAR TRACT OF THE CATHETER WAS ASCERTAINED BY THE VASCULAR ACCESS TEST, WHICH WAS NOT THE SAME ONE WHO PRACTICED CHEMOTHERAPY, THE DEVICE WAS REMOVED AND REPLACED. NO PIECES REMAINED INSIDE THE PATIENT, BUT THE CATHETER WAS DAMAGED ABOUT 4 CM FROM THE EXIT SITE. THE PATIENT HAD ALREADY INFUSED FOLFIRINOX IN THE PREVIOUS WEEKS AND ON THE DAY ON WHICH THE RUPTURE WAS NOTICED SHE HAD ONLY INFUSED PHYSIOLOGICAL SOLUTION, THEREFORE THERE WAS NO HARM TO THE PATIENT. THE DEVICE WAS STABILIZED WITH SUTURLESS SYSTEMS WITH SKIN ADHESIVENESS AND TRANSPARENT SEMI-PERMEABLE MEMBRANE (IV ADVANCED DRESSING 3M). ON FOLFRINOX THERAPY, THE CATHERER WAS DAMAGED AT 5 CM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A FLUID LEAK WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR S/L POWERPICC SOLO. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED JUST DISTAL OF THE 4 CM DEPTH MARKER. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED ¿ 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER ¿ OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING) AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02212 | 2023-06-02 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | GROSHONG MID LINE MINI TRAY 4F (EXPORT ONLY) | CATHETER, PERCUTANEOUS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7715435CE |  | IT WAS REPORTED BY THE CUSTOMER ¿TWO DAYS AFTER DM IMPLANTATION, A CHECK OF THE DRESSING APPLIED TO THE SKIN REVEALED LEAKAGE OF INFUSION FLUID AND THE ABSENCE OF A CATHETER. CONSEQUENCE WAS PROLONGED HOSPITALIZATION AND POSTPONEMENT OF DISCHARGE TO HOSPITAL. PATIENT UNDERWENT A CHEST CT SCAN TO HIGHLIGHT THE PLACEMENT OF THE CATHETER WHICH WAS FOUND TO BE IN THE ATRIUM WITH CONSEQUENT TRANSFER TO CARDIOLOGY FOR REMOVAL AND TRANSPORT TO WARD IN MEDICINE. PATIENT RECEIVED CHEST X-RAY, ECHOCARDIOGRAPHY, AND CHEST CT SCAN.¿ NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02221 | 2023-06-02 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2275108 |  | IT WAS REPORTED BY THE CUSTOMER "I USED A PICC TODAY AND UNFORTUNATELY HAD ONE DROP OF LEAKAGE. I INSERTED THE PICC LINE SUCCESSFULLY THOUGH." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED 5/30/2023: IT WAS REPORTED THE EVENT DID NOT INVOLVE ANY LIFE-THREATENING MEDICAL SITUATION. THE PICC LINE WAS LEAKING FROM THE YELLOW STOPPER AND NO MEDICATION WAS INFUSING. THERE WAS NO PATIENT HARM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02209 | 2023-06-02 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 20G X 10 CM WITH PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F220108PT |  | IT WAS REPORTED VIA MEDWATCH THAT "MIDLINE CATHETER FROM BARD BEING INSERTED, GUIDEWIRE HAD BECOME DISLODGED." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02210 | 2023-06-02 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER 4F BASIC TRAY (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6194118 |  | IT WAS REPORTED BY THE CUSTOMER THAT "PATIENT WAS RECEIVING IV ANTIBIOTICS THROUGH HER PICC WHEN SHE NOTICED SWELLING IN HER UPPER ARM AND LOWER FOREARM. IT APPEARED THE PICC WAS EXTRAVASATING IN HER ARM. THE PICC WAS REMOVED BY THE DISTRICT NURSE. ANOTHER PICC INSERTED THE FOLLOWING DAY. PATIENT HAD GOOD RECOVERY WITH NO SIGNIFICANT DAMAGE. PICC WAS ALSO NOTED TO BE LEAKING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING CATHETER IS UNCONFIRMED BECAUSE THE LEAK COULD NOT BE IDENTIFIED. ONE 4 FR POWERPICC SOLO WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED BLOOD USE RESIDUES THROUGHOUT THE RETURNED DEVICE, AND SOME INDENTATIONS FROM THE SECURACATH WERE VISIBLE ALONG THE CATHETER. A FUNCTIONAL TEST OF INFUSING THE DEVICE WITH WATER USING A 12 ML SYRINGE REVEALED THE CATHETER WAS PATENT TO INFUSION. A LEAK WAS NOT OBSERVED DURING INFUSION OF WATER. THE CATHETER WAS SUBSEQUENTLY PRESSURIZED WITH WATER, AND A LEAK COULD NOT BE IDENTIFIED ANYWHERE ALONG THE CATHETER. A MICROSCOPIC OBSERVATION REVEALED SOME INDENTATIONS AT THE 3 CM DEPTH MARKER AND AT THE 6 CM DEPTH MARKER FROM THE SECURACATH, BUT NO SPLITS WERE FOUND DURING A MICROSCOPIC EVALUATION. THE COMPLAINT OF A LEAKING POWERPICC SOLO IS UNCONFIRMED BECAUSE THE FAILURE COULD NOT BE REPRODUCED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03990 | 2023-06-02 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MAGIC3 GO MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT PATIENT HAD URINARY TRACT INFECTION ,9 BOXES WERE FROM JUGX9006, 4 BOXES WERE FROM JUGT0289 AND 1 BOX WAS FROM JUGV0690 AND REQUESTED FOR COMPENSATION AND DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-03991 | 2023-06-02 | 2023-05-13 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT HAD AN URINARY TRACT INFECTION IN LAST MONTH. THEIR DOCTOR DETERMINED THERE WAS ECOLI IN URINE SO NOW THAT SHE GOT RECALL FORM. SHE WAS CONCERNED IT MIGHT HAD BEEN AS RESULT OF THAT. REPRESENTATIVE ADVISED TO DISCARD THE PRODUCTS AND OFFERED TO REPLACE. BUT THE PATIENT STATES THAT THEY STOPPED CATHING AND JUST WANTED AN INSURANCE REFUND. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-03992 | 2023-06-02 | 2023-04-12 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THESE PUREWICK FEMALE EXTERNAL CATHETERS WERE IRRITATING THE LABIA. REPRESENTATIVE RECOMMENDED PATIENT TO DISCONTINUE THE USE. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. NO MEDICAL INTERVENTION WAS REPORTED. AS PER FOLLOW UP VIA PHONE (B)(6) 2023, CUSTOMER STATED THEY HAD AN INFECTION BEFORE THEY EVER USED THE PURE WICK AND THE INFECTION WORSENED WITH USE AND ALSO STATED THEY WERE TAKING UNKNOWN MEDICATION TO CLEAR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-03993 | 2023-06-02 | 2023-05-11 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER COUDE TIP | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0168L16 | 0168L16 |  | IT WAS REPORTED THAT PATIENT SEEN IN SCC FOR TREATMENT BLADDER INSTALLATION OF GEMZAR CHEMO. FOLEY CATHETER INSERTED BY REGISTER NURSE WITH NO ISSUES. TREATMENT PROCEEDED WITH NO ISSUES. UPON COMPLETION OF TREATMENT, REGISTER NURSE ATTEMPTED TO DEFLATE BALLOON AND REMOVE THE CATHETER. BALLOON DID NOT DEFLATE WITH NO SALINE WITHDRAWN FROM BALLOON UPON ATTEMPT. REGISTER NURSE TROUBLESHOOTED, FOLEY BALLOON WOULD STILL NOT DEFLATE. ON CALL UROLOGIST CONSULTED, REGISTER NURSE CUT CATHETER BY VALVE OF BALLOON AS INSTRUCTED. NO FLUID RETURN AND BALLOON DID NOT INFLATE. REGISTER NURSE DISCUSSED FURTHER WITH DOCTOR AND PLAN IN PLACE FOR PATIENT TO PROCEED TO DOVER OFFICE OF MANCHESTER FOR FURTHER EVALUATION. FOLEY BAG CHANGED PER CHEMO PRECAUTIONS TO NEW EMPTY BAG, AS IT HAD BEEN DRAINING WELL. PATIENT LEFT AMBULATORY AND STABLE, FAMILY DRIVING TO UROLOGIST OFFICE. UNEXPECTED MEDICAL INTERVENTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02208 | 2023-06-02 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE PHYSICIAN COULDN¿T BE ABLE TO PLACE THE GUIDEWIRE INTO VESSEL TO SUPERIOR VENA CAVA AND FOUND THAT THE WIRE WAS IN INTERNAL JUGULAR UP TO THE NECK. IT WAS FURTHER REPORTED THAT THE GUIDEWIRE ALLEGEDLY GOT STUCK AND KINKED WHEN TRIED TO PULL THE WIRE OUT OF THE VESSEL. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 08/2024). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-03998 | 2023-06-02 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OFF. IMMEDIATELY AFTER THE CATHETER PLACEMENT, THE BALLOON GOT DEFLATED AND SPONTANEOUSLY FELL OUT OF THE PATIENT. THERE WAS A PINHOLE ON THE BALLOON AND THE DISTILLED WATER HAD LEAKED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04000 | 2023-06-02 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT DURING A SURGERY. UPON INJECTING WATER AND CHECKING THE CATHETER, THERE WAS WATER LEAKAGE FROM AROUND THE BALLOON AREA. THE EVENT OCCURRED ONE HOUR AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BALLOON AND THE SHAFT AS PER INQUIRY SHEET. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. AS PER MAIL RECEIVED FROM IBC REPRESENTATIVE ON (B)(6) 2023, STATED THAT THE LEAK LOCATION OF THIS COMPLAINT WAS ONE. SINCE THE EXACT LEAK LOCATION WAS UNKNOWN, SO THE SALES REPRESENTATIVE MARKED BALLOON AND SHAFT IN THE INQUIRY SHEET. REPORTED COMPLAINT WAS THAT THE WATER LEAKAGE FROM THE BALLOON DURING A WATER INJECTION. STATED THERE WERE NO OBVIOUS VISIBLE DEFECTS TO THE RETURNED SAMPLE. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. SAMPLE WAS INFLATED WITH WATER AND OBSERVED INTERLUMEN LEAKAGE AS WATER LEAKING THROUGH DRAINAGE EYE. THE CATHETER WAS DISSECTED AND OBSERVED TEAR ON RUBBERIZE LAYER OF INFLATION LUMEN. TEAR WAS EXAMINED UNDER MICROSCOPE AND NOTED NOT TO BE ROUND AND SMOOTH. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON UNIFORM RUBBERIZED THICKNESS ETC CAUSING TORN RUBBERIZE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04001 | 2023-06-02 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MALE COUDE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD SEVERAL URINARY TRACT INFECTIONS. MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04008 | 2023-06-02 | 2023-05-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION AND STATED THAT THE PATIENT WAS HOSPITALIZED FOR HEART, AND THEY DEVELOPED THE URINARY TRACT INFECTION IN THE HOSPITAL AND RAN A HIGH FEVER USING THE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04009 | 2023-06-02 | 2023-05-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51816 | 51816 |  | IT WAS REPORTED THAT PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACK INFECTION. THE PATIENT HAS SEVERAL URINARY TRACK INFECTION. THE RMA DONE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | OTHER |  | 2023-06
1018233-2023-03994 | 2023-06-02 | 2023-05-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MAGIC3 GO MALE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION AND BLEEDING FROM USING THESE MAGIC 3 MALE GO INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS, MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04041 | 2023-06-05 | 2023-05-16 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MAGIC3 GO MALE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT PATIENT HAD A SEVERE URINARY TRACT INFECTION AND OTHER HEALTH ISSUE. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AND OTHER HEALTH ISSUES AT THIS TIME AND MEDICAL INTERVENTION WAS UNKNOWN | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04037 | 2023-06-05 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT MALE INTERMITTENT CATHETER WAS NO LUBRICANT IN THE PACKAGE. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04039 | 2023-06-05 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | ALL-SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0170SI12 | 0170SI12 |  | IT WAS REPORTED THAT THE TUBING WAS SPLIT JUST ABOVE THE COUDE TIP CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04040 | 2023-06-05 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51614 | 51614 |  | IT WAS REPORTED THAT THE PATIENT DID NOT UNDERSTAND HOW TO USE THE HYDROPHILIC INTERMITTENT CATHETERS. REPRESENTATIVE WAS ABLE TO EXPLAIN HOW TO USE THE HYDRO CATHETERS BUT THEY STILL SEEMED CONFUSED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT/ MISSING TRANSLATION; MISSING INSTRUCTIONS; VENDOR/PRINTER ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. THE CATHETER BECOMES SLIPPERY WHEN WETTED WITH WATER, ELIMINATING THE NEED FOR A SEPARATE LUBRICANT. FOR YOUR ADDED CONVENIENCE, THIS CATHETER IS PACKAGED WITH ITS OWN STERILE WATER. SIMPLY RELEASE THE WATER FROM ITS FOIL PACKET AND THEN TIP THE UN-OPENED CATHETER PACKAGE END-TO-END. THE CATHETER ACTS LIKE A MAGNET TO ATTRACT THE WATER AND ACTIVATE ITS SLIPPERY COATING. FOLLOW THESE STEPS FOR BEST RESULTS. RELEASE THE WATER PRIOR TO OPENING THE SEALED CATHETER POUCH: APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. WET THE CATHETER. HOLD PACKAGE WITH PRINTED SIDE UP. TIP PACKAGE END-TO-END THREE TO SIX TIMES TO WET CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. USE THE CATHETER. PEEL BACK PACKAGE TO EXPOSE FUNNEL END OF CATHETER. IF DESIRED, USE THE SELF-ADHESIVE TAPE ON THE PACKAGE TO TEMPORARILY ATTACH THE POUCH TO ANY DRY VERTICAL SURFACE. REMOVE CATHETER AND USE ACCORDING TO PHYSICIAN¿S INSTRUCTIONS. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE, DO NOT RESTERILIZE. MADE OF SILICONE ELASTOMER." CORRECTION: F,H. H11: SECTION A THROUGH F THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02261 | 2023-06-05 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRO 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN THE PATIENT IS UNDERGOING CHEMOTHERAPY, WHEN THE PICC IS INSERTED, THE PERIPHERALLY INSERTED CENTRAL VENOUS CATHETER HAS RESISTANCE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02265 | 2023-06-05 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED, "I MAKE 3 ATTEMPTS TO PASS A PICC CATHETER, 2 IN LEFT SUPERIOR MEMBER AND 1 IN RIGHT SUPERIOR MEMBER. I PERFORM INSERTION OF A 46 CM, BUT WITHOUT SUCCESS. IN MSD I MADE A 36 CM INSERTION, WHEN I REMOVED THE GUIDE WIRE FROM THE CATHETER, IT DIDN¿T COME OUT, I MADE NUMEROUS ATTEMPTS TO REMOVE THE GUIDEWIRE, WITHOUT SUCCESS. PICC CATHETER WAS STRETCHED AND THE GUIDEWIRE WAS DAMAGED, AND IT WAS POSSIBLE TO REMOVE IT AFTER REMOVAL OF THE PATIENT¿S CATHETER. PATIENT GUIDANCE, PHARMACEUTICAL COMMUNICATION, MEDICAL COMMUNICATION FOR A NEW PICC REQUEST." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04044 | 2023-06-05 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT ONLY ABOUT 7-8 CC OF STERILE WATER COULD BE DRAINED FROM THE FOLEY CATHETER INSTEAD OF 10CC. | THE REPORTED EVENT WAS NOT CONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS NOT CONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02262 | 2023-06-05 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 15CM PRE-CURVED SHORT-TERM DL | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5594150 |  | IT WAS REPORTED THAT, "ON MONDAY, (B)(6), THE PLACEMENT OF A 15CM, 13.5 FR NIAGARA CATHETER IS SCHEDULED FOR INSERTION. AT THE TIME OF PUNCTURE, IT IS OBSERVED THAT THERE IS NO HERMETIC SEAL BETWEEN THE NEEDLE AND THE SYRINGE BECAUSE ADEQUATE SUCTION IS NOT OBSERVED AFTER THE PUNCTURE, WHICH IT IS NECESSARY TO OBTAIN VENOUS RETURN FROM THE VESSEL TO WHICH ONE WISHES TO ACCESS, WHEN ASPIRATING THE SOLUTION AND PRESSING THE PLUNGER, A LEAK IS OBSERVED BETWEEN IT, IT IS DECIDED TO USE ANOTHER NEW 20 CM CATHETER, 13.5 FR, HOWEVER, THE SAME SITUATION OCCURS, DIFFERENT SYRINGES ARE USED BOTH 5CC AND 10CC STERILE WITHOUT OBTAINING ADEQUATE FUNCTIONALITY, SO THE PROCEDURE IS SUSPENDED." ADDITIONAL INFORMATION RECEIVED ON MAY 18, 2023: THE PATIENT COULD NOT RECEIVE HIS TREATMENT, DUE TO THE PROBLEM PRESENTED. THIS REPORT ADDRESSES THE FIRST INSERTION ATTEMPT WITH A15 CM CATHETER. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02238 | 2023-06-05 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE UNIT IS GIVING INACCURATE READING. LOLLIPOP MOVES AROUND THE SCREEN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF GIVING INACCURATE READING. LOLLIPOP MOVES AROUND THE SCREEN WAS CONFIRMED. THE LOLLIPOP MOVING AROUND THE SCREEN IS INTERMITTENT. AFTER THE SENSOR IS CALIBRATED, THE LOLLIPOP IS IN THE CORRECT ORIENTATION. THE ROOT CAUSE OF THE REPORTED FAILURE WAS IDENTIFIED AS AN INTERNAL FAILURE OF THE SHERLOCK SENSOR MAGNET TRACKING PCB. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04016 | 2023-06-05 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THERE WAS NO OUTFLOW OF URINE AFTER THE FOLEY CATHETER WAS INSERTED INTO THE PATIENT. | THE REPORTED EVENT WAS NOT CONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS NOT CONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04042 | 2023-06-05 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT WATER LEAKED FROM THE FOLEY CATHETER SHAFT. AFTER PLACEMENT, WATER LEAKED FROM THE CATHETER SHAFT AND THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. THE EXACT LEAK LOCATION WAS NOT KNOWN, BUT IT WAS SUSPECTED THAT THE CATHETER HAD A DAMAGE SOMEWHERE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04027 | 2023-06-05 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE SYRINGE AND THE FOLEY CATHETER CONNECTION. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. THIS INVESTIGATION DOES NOT REQUIRE A DHR REVIEW DUE TO A CLOSURE LETTER NOT REQUIRED FOR THIS COMPLAINT. A LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04017 | 2023-06-05 | 2023-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER FELL OUT OF THE PATIENT AFTER INSERTION. UPON CHECKING, THE BALLOON WAS FOUND TO BE TORN. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. VISUAL INSPECTION NOTED IRREGULAR BURST WITHOUT MISSING PIECES. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04026 | 2023-06-05 | 2023-05-12 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT UPON ATTEMPTING TO REPLACE THE FOLEY CATHETER, THE STERILE WATER COULD NOT BE DRAINED SO THE USER ADDED WATER IN CONTRAST. THE ADDITIONAL INJECTION OF WATER WAS POSSIBLE BUT COULD NOT BE DRAINED AND THE BALLOON HAD BURST INSIDE THE PATIENT'S BLADDER. THE REMNANTS OF THE BALLOON REMAINED INSIDE THE PATIENT'S BODY. ACCORDING TO THE INQUIRY SHEET PROVIDED, THE EVENT HAS OCCURRED AFTER 7 DAYS OF CATHETER PLACEMENT TO THE PATIENT AND THERE WAS NO PATIENT'S URINE STONE. NO MATERIALS HAVE COME IN CONTACT WITH THE BALLOON AND THE CATHETER WAS NOT PULLED BY HAND. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. AS PER THE NOTIFICATION RECEIVED ON 22MAY2023, CLARIFIED THAT THE FRAGMENTS WERE STILL INSIDE THE PATIENT'S BODY AND THE ATTENDING PHYSICIAN DECIDED TO KEEP AN EYE ON IT FOR A WHILE. THEY WILL NOTIFY ONCE THE FRAGMENT IS REMOVED. | THE REPORTED EVENT IS INCONCLUSIVE DUE TO POOR SAMPLE CONDITION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE THIN RUBBERIZE WALL (EXAMPLE: CAUSING COLLAPSE/ PINCHED INFLATION LUMEN UNDER THE BALLOON). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02232 | 2023-06-05 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN, WHITE, 4.2F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 0601610 | 0601610 |  | IT WAS REPORTED THAT DURING PREPARATION OF A CATHETER REPAIR PROCEDURE, THE GLUE HAD THE INTACT FOIL SEAL OVER THE END. IT WAS FURTHER REPORTED THAT THE GLUE WAS ROCK HARD WHEN PUNCTURED THE SEAL WITH THE CAP. THERE WAS NO PATIENT CONTACT. | H10: A VOLUNTARY RECALL HAS BEEN INITIATED FOR CATHETER REPAIR KITS WHICH ARE PRODUCT CATALOG/LOT NUMBER SPECIFIC. REPORTEDLY THE CATHETER REPAIR KITS HAVE HARDENED/COAGULATED ADHESIVE. HARDENED OR COAGULATED ADHESIVE HAS THE POTENTIAL TO CAUSE DELAYS WHILST AN ALTERNATIVE REPAIR KIT, ADHESIVE OR REPLACEMENT CATHETER IS OBTAINED; THEREFORE, POTENTIALLY PROLONGING SURGERY OR NECESSITATING EXCHANGE. TO DATE THERE HAVE BEEN NO ADVERSE EVENTS WORLDWIDE REPORTED RELATED TO THIS ISSUE. AS RESULT OF THE FIELD ACTION, THIS EVENT IS BEING REPORTED AS A MALFUNCTION REPORTABLE EVENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10:D4 (EXPIRY DATE: 09/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04014 | 2023-06-05 | 2023-05-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE URINARY TRACT INFECTIONS FROM USING THE MAGIC 3 GO INTERMITTENT CATHETER (FOR LOT#JUGT3350), FOR (LOT#JUGT3297). IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INDICATIONS FOR USE. THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS. NONE KNOWN. WARNINGS. THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS. INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS. URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04019 | 2023-06-05 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | ALL-SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN STERILE WATER WAS INJECTED TO INFLATE THE FOLEY CATHETER, THE STERILE WATER COULD BE SEEN MOVING THROUGH THE LUMEN OF THE CATHETER, BUT THE BALLOON WAS NOT INFLATED AS IT WAS IN CLOSE CONTACT TO THE CATHETER. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED A 3-WAY TEMP SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. ATTEMPTED TO INFLATE BALLOON WITH RETURNED SYRINGE AND IN-HOUSE LUER LOCK SYRINGE AND SOLUTION COULD NOT GO INTO BALLOON. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED POTENTIAL ROOT CAUSES COULD BE: LUMEN COLLAPSES UNDER THE PRESSURE OF THE BALLOON, PINCHED LUMEN; KINKED LUMEN. THE DHR REVIEW DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER." THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04015 | 2023-06-05 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | ALL-SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRETEST, THE USER FELT MORE RESISTANCE THAN CONVENTIONAL PRODUCTS WHEN INJECTING STERILE WATER AND WHEN TRIED TO DRAIN, THE STERILE WATER COULD NOT BE DRAINED OUT. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON 24MAY2023, STATED THAT THAT RESISTANCE WAS FELT DURING INFLATION. AS PER THE NOTIFICATION RECEIVED ON 24JUN2023, THE RETURNED SYRINGE DID NOT DEFLATE THE BALLOON ACCORDING TO PRODUCT SPECIFICATIONS. | THE REPORTED EVENT WAS UNCONFIRMED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION CATHETER KINKING, DAMAGE, RUPTURE. DIFFICULTY OR FAILURE TO REMOVE THE DEVICE. OCCLUSION OF CATHETER INNER LUMENS. ENCRUSTATION. ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE . DEVICE DAMAGE DUE TO INAPPROPRIATE USE. FAILURE TO MEASURE TEMPERATURE. IMPROPER TEMPERATURE INDICATION. 2) ADVERSE EVENTS URINARY-TRACT INFECTION; HEMORRHAGE, HEMATURIA; ALLERGY REACTION TO THE DEVICE; CALCULUS FORMATION; EDEMA; PAIN; DISCOMFORT; INJURY OF BLADDER OR URETHRAL; URETHRITIS, URINARY INCONTINENCE; RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04018 | 2023-06-05 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | ALL-SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE BALLOON DID NOT DEFLATE AFTER BEING INFLATED IN THE PRE-TEST OF THE FOLEY CATHETER. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. A POTENTIAL FAILURE MODES ARE COLLAPSED LUMEN AND BLOCKED LUMEN. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00391 | 2023-06-05 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQ5064 | CQ5064 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE VIA LEFT FOREARM, ULTRASOUND REVEALED THAT THE PTA BALLOON WAS ALLEGEDLY UNABLE TO BE DEFLATED COMPLETELY. IT WAS FURTHER REPORTED THAT THE BALLOON WAS FOUND TO BE BURST AND WAS LEAKING WATER. REPORTEDLY, THE BALLOON WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE CONQUEST BALLOON CATHETER WAS RETURNED FOR EVALUATION. PEELED PEBAX, UNRAVELED FIBER DISTURBANCE, AND FRAYED FIBERS WERE NOTED AT THE PROXIMAL END OF THE BALLOON; NO OTHER SPECIFIC ANOMALIES WERE NOTED DURING THE VISUAL EVALUATION. DURING THE FUNCTIONAL TESTING, THE BALLOON WAS INFLATED WITH AN IN-HOUSE PRESTO INFLATION DEVICE, AND DURING THE INFLATION, WATER WAS NOTED TO BE STREAMING FROM THE BALLOON. FURTHER, THE BALLOON WAS STRIPPED, AND UNDER MICROSCOPIC OBSERVATION, A LONGITUDINAL RUPTURE WAS NOTED ON THE BALLOON. ONE VIDEO WAS REVIEWED. THE VIDEO BALLOON HAD BEEN PRESSURIZED WITH AN UNKNOWN INFLATION DEVICE, AND UPON INFLATING THE BALLOON, WATER WAS NOTED TO BE STREAMING OUT OF THE BALLOON. NO OTHER SPECIFIC ANOMALIES ARE NOTED. THE SUBMITTED VIDEO SHOWS THE LEAKAGE ON THE BALLOON, AND DURING THE SAMPLE ANALYSIS, IT WAS CONFIRMED THAT THE LEAKAGE ON THE BALLOON WAS CAUSED BY THE LONGITUDINAL RUPTURE THAT WAS OBSERVED UNDER MICROSCOPIC OBSERVATION. HOWEVER, NO OBJECTIVE EVIDENCE FOR THE DEFLATION ISSUE COULD BE NOTED DURING THE VIDEO REVIEW AND SAMPLE ANALYSIS. PEELED PEBAX, UNRAVELED FIBER WAS ALSO NOTED ON THE RETURNED BALLOON DURING THE VISUAL EVALUATION. THEREFORE, THE INVESTIGATION WAS CONFIRMED FOR THE REPORTED BALLOON RUPTURE AND IDENTIFIED PEELED PEBAX, UNRAVELED FIBER AND THE INVESTIGATION REMAINS INCONCLUSIVE FOR THE REPORTED DEFLATION ISSUE. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BALLOON RUPTURE, DEFLATION ISSUE AND IDENTIFIED PEELED PEBAX AND UNRAVELED FIBERS COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. B5, D4 (EXPIRATION DATE: 01/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00390 | 2023-06-05 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | AT75164 | AT75164 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY FOUND TO BE BURST WHEN THE PRESSURE WAS RELEASED. IT WAS FURTHER REPORTED THAT THERE WAS DIFFICULTY IN RETRACTING THE BALLOON THROUGH THE INTRODUCER SHEATH AND THE INTRODUCER SHEATH WAS REMOVED AWAY TOGETHER. REPORTEDLY, THE BALLOON WAS REMOVED OUT OF THE PATIENT AND THOUGH THE BALLOON ENLARGED THE SKIN WOUND, THE SURGEON DID COMPRESSION OVER THE WOUND TO STOP BLEEDING. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. ONE PHOTO WAS REVIEWED. THE PHOTO SHOWS THE BALLOON CATHETER NOTED TO BE LOADED WITHIN AN UNKNOWN INTRODUCER SHEATH. THE BALLOON WAS NOTED TO HAVE A BREAK NEAR THE PROXIMAL END OF THE CATHETER WHICH LED TO THE EXPOSURE OF THE INNER GUIDE WIRE LUMEN. THE BALLOON WAS NOTED TO BE BUNCHED ALL TOGETHER NEAR THE DISTAL TIP OF THE CATHETER. THE INTRODUCER SHEATH IS ALSO NOTED TO BE DEFORMED ALL ALONG THE LENGTH. NO OBJECTIVE EVIDENCE OF BALLOON RUPTURE COULD BE NOTED FOR THE DEVICE'S CONDITION. NO OTHER ANOMALIES NOTED. FROM THE PHOTO REVIEW, THE BALLOON IS NOTED TO BE STUCK WITHIN THE UNKNOWN SHEATH AND THE BALLOON WAS ALSO NOTED TO HAVE BROKEN FROM PROXIMAL AND BUNCHED TO THE DISTAL END, EXPOSING THE INNER GUIDE WIRE LUMEN. HENCE THE INVESTIGATION WAS CONFIRMED THE REPORTED SHEATH REMOVAL DIFFICULTIES AND IDENTIFIED BALLOON BREAK. NO EVIDENCE OF BALLOON RUPTURE WAS NOTED, HENCE THE INVESTIGATION REMAINS INCONCLUSIVE FOR THE REPORTED BALLOON RUPTURE. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BALLOON RUPTURE, SHEATH REMOVAL DIFFICULTY AND IDENTIFIED BREAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 12/2025), G3, H6 (DEVICE, METHOD) H11: H6 (RESULT, CONCLUSION) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02237 | 2023-06-05 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE TIP OF THE SHEATH INTRODUCER WAS ALLEGEDLY DAMAGED LIKE A HANGNAIL. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10:D4 (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02235 | 2023-06-05 | 2023-04-19 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT TWO YEARS FIVE MONTHS AND FOURTEEN DAYS POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY PERFORATED. IT WAS FURTHER REPORTED THAT THE CONTRAST MEDIUM ALLEGEDLY LEAKED. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 01/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02234 | 2023-06-05 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE GUIDEWIRE WAS ALLEGEDLY BENT. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE J-TIP GUIDEWIRE LOADED TO A GUIDEWIRE HOOP WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC AND DIMENSIONAL EVALUATIONS WERE PERFORMED. A BEND WAS NOTED IN THE GUIDEWIRE. THE FLAT CORE WIRE WAS NOTED TO HAVE A COMPLETE BREAK AND PROTRUDING FROM A GUIDEWIRE PORTION. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED DEFORMATION DUE TO COMPRESSIVE STRESS ISSUE AND THE IDENTIFIED MATERIAL PROTRUSION AND MATERIAL SEPARATION ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. EXPIRY DATE: 04/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02260 | 2023-06-05 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER, "DURING PLACEMENT OF A PICC I NOTICED THAT WHEN I ASPIRATED TO CHECK FOR BLOOD RETURN THAT AIR WAS COMING INTO MY SYRINGE. I REMOVED THE SYRINGE FROM THE INJECTION PORT AND EXPELLED THE AIR THEN REATTACHED AND FLUSHED THE LINE. I NOTICED THAT THE SALINE WAS LEAKING OUT FROM AROUND THE WIRE COMING OUT OF THE RUBBER (UNSURE OF MATERIAL HERE) STOPPER. AFTER ENSURING THAT ALL MY CONNECTIONS WERE TIGHTLY SECURED, I ATTEMPTED TO ASPIRATE AGAIN AND PULLED AIR INTO MY SYRINGE. I REMOVED THE SYRINGE, EXPELLED THE AIR, REATTACHED AND ATTEMPTED TO FLUSH A SECOND TIME NOTING FLUID AGAIN LEAKING FROM THE SAME AREA WHEN FLUSHING. SYRINGE REMOVED AND PLACED ON THE OTHER LUMEN (PURPLE) AND WAS ABLE TO ASPIRATE BLOOD AND FLUSH WITHOUT ISSUE. I WAS ABLE TO LOOK CLOSER AND NOTICED THE LEAKING. PATIENT WAS SENT HOME WITH THE PICC THE NEXT DAY FOR 8 WEEKS OF IV ANTIBIOTICS. THE EVENT WAS NOT URGENT OR LIFE THREATENING. THERE WAS NO HARM TO THE PATIENT. PATIENT HAS NO RECORDED INFECTIOUS DISEASES." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04030 | 2023-06-05 | 2023-05-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTIONS FROM USING THE MAGIC 3 GO INTERMITTENT CATHETER FOR THE LAST 1.5 MONTH. PATIENT HAD BEEN ON 2 MEDICATIONS AND THE PATIENT WAS GOING TO CALL THE DOCTOR AND LET THE PATIENT KNEW THIS WAS THE CAUSE AND STATED THAT THE PATIENT ALSO REPORTED AN ORDER THEY GOT CRUSHED BOXES AND THE PATIENT SAID THEY WOULD KEEP AS THE CATHETERS WERE OK. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL : ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02264 | 2023-06-05 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/SHERLOCK 3CG TPS STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED, "GUIDEWIRE IN PICC TRAY CAME APART WHILE IN THE PATIENT AND DEVELOPED A KNOT ON THE END OF IT, CAUSING A PROBLEM GETTING IT OUT OF THE PATIENT." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04023 | 2023-06-05 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165PL10 | 0165PL10 |  | IT WAS REPORTED THAT THERE WAS A CRACK AT 17 CM FROM THE TIP OF THE FOLEY CATHETER. THERE WAS NO STERILE WATER AND THE BALLOON CAME OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02263 | 2023-06-05 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 5F FULL TRAY DL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2275108 |  | IT WAS REPORTED, "CUSTOMER EXPERIENCED INSERTION CHALLENGE REGARDING AN AZYGOUS INSERTION. I TALKED HER THROUGH IT FROM A CLINICAL PERSPECTIVE." NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02267 | 2023-06-05 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 20CM PRE-CURVED SHORT-TERM DL | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5594200 |  | IT WAS REPORTED THAT, "ON MONDAY, APRIL 10, THE PLACEMENT OF A 15CM, 13.5 FR NIAGARA CATHETER IS SCHEDULED FOR INSERTION. AT THE TIME OF PUNCTURE, IT IS OBSERVED THAT THERE IS NO HERMETIC SEAL BETWEEN THE NEEDLE AND THE SYRINGE BECAUSE ADEQUATE SUCTION IS NOT OBSERVED AFTER THE PUNCTURE, WHICH IT IS NECESSARY TO OBTAIN VENOUS RETURN FROM THE VESSEL TO WHICH ONE WISHES TO ACCESS, WHEN ASPIRATING THE SOLUTION AND PRESSING THE PLUNGER, A LEAK IS OBSERVED BETWEEN IT, IT IS DECIDED TO USE ANOTHER NEW 20 CM CATHETER, 13.5 FR, HOWEVER, THE SAME SITUATION OCCURS, DIFFERENT SYRINGES ARE USED BOTH 5CC AND 10CC STERILE WITHOUT OBTAINING ADEQUATE FUNCTIONALITY, SO THE PROCEDURE IS SUSPENDED. THE PATIENT COULD NOT RECEIVE HIS TREATMENT, DUE TO THE PROBLEM PRESENTED." THIS REPORTED ADDRESSES THE SECOND INSERTION ATTEMPT WITH A 20CM CATHETER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF UNABLE TO ASPIRATE IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE VIDEO SAMPLE OF AN 18 GA INTRODUCER NEEDLE ATTACHED TO A SYRINGE WAS PROVIDED FOR EVALUATION. THE USER IS ATTEMPTING TO ASPIRATE FLUID WITHIN A BOWL; HOWEVER, NO LIQUID IS FOUND TO ASPIRATE. WHEN FLUID IS INFUSED THROUGH THE NEEDLE, THE FLUID IS FOUND TO LEAK AT THE HUB. THE CHARACTERISTICS OF THE DAMAGE COULD NOT BE INSPECTED FROM THE VIDEO PROVIDED. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE DAMAGE DURING HANDLING AND OVER-TIGHTENING. SINCE A HUB LEAK WAS OBSERVED, THE COMPLAINT IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02272 | 2023-06-05 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED A 10 CM HAIR IN THE KIT WHEN IT WAS OPENED. KIT WAS NOT USED ON THE PATIENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | N | Y | OTHER |  | 2023-06
3006260740-2023-02270 | 2023-06-05 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "AFTER PLACEMENT OF THE PICC, WHILE USING THE SHERLOCK SYSTEM OR THE ECG TRACE, WE NOTICE THAT THE CATHETER TIP DOES NOT APPEAR (ONLY ECG BASE TRACE). ONCE THE GUIDE WIRE REMOVED FROM THE CATHETER, IT WAS NOTICED THAT THE WIRE WAS ¿CHOCKED¿ ALMOST BROKEN. IT BECAME NECESSARY TO PERFORM THORACIC EXAMINATION TO VERIFY CORRECT PICC POSITIONING. THE PATIENT HAS BEEN SUBJECTED TO AN ADDITIONAL CHECK X-RAY TO CHECK THE CORRECT POSITIONING." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF BROKEN STYLET WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 3CG STYLET WITH A T-LOCK ASSEMBLY. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND CONFIRMED TO BE BROKEN. THE FOLLOWING OBSERVATIONS WERE MADE WHICH WERE CONSISTENT WITH THIS FAILURE TYPE: ¿ MICROSCOPIC EXAMINATION REVEALED DEFORMATION OF THE STYLET TUBING AT MAGNET EDGE(S). ¿ MICROSCOPIC EXAMINATION REVEALED LOCALIZED DELAMINATION OF THE STYLET TUBING AT THE DAMAGE SITE(S), INDICATIVE OF TUBING KINKING/FOLDING. MEASUREMENTS OF THE STYLET TUBING, STYLET CORE WIRE, AND MAGNET WERE TAKEN AND CONFIRMED TO BE WITHIN SPECIFICATIONS. ALTHOUGH THE BREAK IN THE WIRE COULD BE CONFIRMED, THE ROOT CAUSE COULD NOT BE DETERMINED THROUGH EXAMINATION OF THE RETURNED SAMPLE. THE OBSERVED FEATURES WERE INDICATIVE OF CIRCUMSTANCES WHICH INCLUDED STYLET KINKING, LIKELY DURING THE INSERTION PROCESS. IT APPEARED THAT ADDITIONAL UNIDENTIFIED FACTORS MAY HAVE ALSO CONTRIBUTED. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02241 | 2023-06-05 | 2023-01-10 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT ONE DAY POST THE PORT PLACEMENT VIA RIGHT INTERNAL JUGULAR VEIN, THE PATIENT ALLEGEDLY DEVELOPED FEVER, REDNESS, SWELLING AND INFECTION AT THE SITE OF POCKET. IT WAS FURTHER REPORTED THAT THE SYMPTOMS WERE NOT CONTROLLED EFFECTIVELY AFTER ONE WEEK OF ANTI-INFECTIVE THERAPY. FURTHERMORE ANTI-INFLAMMATORY MEDICATION, RED LIGHT PHYSIOTHERAPY AND ENHANCED DRESSING CHANGE WERE PROVIDED FOR TREATMENT. REPORTEDLY, THE PORT WAS REMOVED AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION, ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS THE POWER PORT ATTACHED TO THE CATHETER IN A PACKAGE. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 11/2023), G3, H6 (PATIENT, METHOD). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02266 | 2023-06-05 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 5F FULL TRAY DL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2275108 |  | IT WAS REPORTED BY CUSTOMER, "LEAKAGE OF PICCS." NO OTHER INFORMATION WAS PROVIDED. A SPECIFIC NUMBER OF AFFECTED DEVICES WAS NOT PROVIDED BY THE COMPLAINANT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. INITIAL MEDWATCH REPORT WAS SUBMITTED, UPON FURTHER REVIEW IT WAS FOUND THAT THIS MEDWATCH REPORT 3006260740-2023-02266 IS A DUPLICATE FILE AND HAS BEEN VOIDED. THE ORIGINAL EVENTS WERE SUBMITTED ON MEDWATCH REPORTS 3006260740-2023-02108, 3006260740-2023-02221, AND 3006260740-2023-02222. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02240 | 2023-06-05 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR (RADEL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770352 |  | IT WAS REPORTED BY THE CUSTOMER "UNIT IS GIVING INACCURATE READINGS. OFF BY 3CM". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE REPORTED ISSUE THAT THE SENSOR IS GIVING INACCURATE READINGS AND IS OFF BY 3CM IS UNCONFIRMED. THE SENSOR WAS TESTED AND PASSED THE MAGNET TRACKING AND ECG FUNCTIONALITIES. THERE IS NO ROOT CAUSE DUE TO THE REPORTED ISSUE BEING UNCONFIRMED. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04029 | 2023-06-05 | 2023-05-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT PATIENT HAD URINARY TRACT INFECTION. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER: WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02271 | 2023-06-05 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "WHEN PLACING PICC THEY STUCK THE VEIN AND GOT BLOOD. INSITE THE GUIDEWIRE AND THE INTRODUCER OVER THE WIRE. PULLED OUT THE GUIDEWIRE. WHEN INSERT THE CATHETER IT GOT STUCK AT AXILARY LEVEL. THEY TRIED TO REVERSE THE WIRE IN THE CATHETER AND INSERT THE CATHETER WITH NO WIRE IN THE TIP. DIDN'T GO. THEY USE FLOROSCOPY TO SEE WHAT HAPPENED AND THE CATHETER WAS NOT IN THE VEIN. WHEN THEY PULLED THE CATHETER OUT THE SAW THAT 3 CM OF THE 3CG WIRE TIP WAS LOST IN PATIENT." ADDITIONAL INFORMATION RECEIVED 05/23/2023: "I HAVE TALKED TO THE CUSTOMER AND THE PATIENT HAS NOT SUFFER ANY HARM. THE WIRETIP IS STILL IN PATIENT'S ARM AND THE SURGEON SAID IT COULD STAY THERE BECAUSE IT IS IN SUBCUTANEOUS TISSUE." ADDITIONAL INFORMATION RECEIVED: IT WAS STATED THE PRODUCT HAS BEEN DISCARDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02295 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "THE STYLET IN THE PICC HAS CHANGED AND IT LEAKS WHEN THE PICC IS FLUSHED AND AIR ENTERS THE SYRINGE WITH THE REFLUX." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 10 DEVICES. THIS REPORT ADDRESSES THE SIXTH DEVICE. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02305 | 2023-06-06 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED, 3 OF CLINICIAN'S PICCS TODAY HAVE HAD THE SAME ISSUE. AGAIN ITS LEAKING AROUND THE WIRE THAT COMES OUT OF THE STOPPER. I ALSO HAD THIS HAPPEN WHEN I WAS WITH CLINICIAN THE OTHER DAY. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE FOURTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02298 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "THE STYLET IN THE PICC HAS CHANGED AND IT LEAKS WHEN THE PICC IS FLUSHED AND AIR ENTERS THE SYRINGE WITH THE REFLUX." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 10 DEVICES. THIS REPORT ADDRESSES THE NINTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02300 | 2023-06-06 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT, "A FEW MONTHS AGO, SHE COULD NOT GET BLOOD RETURN DURING A PLACEMENT CHECK FOR A PICC LINE. SHE DID NOT NOTICE THE LEAKING AT THE T-LOCK. THE PROCEDURE WAS ABORTED BECAUSE SHE DID NOT WANT TO INTRODUCE AIR INTO THE LINE AND POTENTIALLY CAUSE AN AIR EMBOLISM.". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02299 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "THE STYLET IN THE PICC HAS CHANGED AND IT LEAKS WHEN THE PICC IS FLUSHED AND AIR ENTERS THE SYRINGE WITH THE REFLUX." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 10 DEVICES. THIS REPORT ADDRESSES THE TENTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
17071119 | 2023-06-06 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP¿ FOLEY TRAY SYSTEM LUBRI-SIL® IC COMPLETE CARE® | CATHETER, UROLOGICAL (ANTIMICROBIAL) AND ACCESSORIES | C. R. BARD, INC. | A319516AM | A319516AM |  | OPERATING ROOM NURSE INSERTED FOLEY WITHOUT ISSUES ONCE PATIENT WAS PUT UNDER ANESTHESIA. DURING THE SURGICAL CASE PERFUSION AND ANESTHESIA NOTED THERE WAS NO URINE OUTPUT. ONCE OPERATING ROOM NURSE HAD ACCESS TO CHECK UNDER THE DRAPES, THE FOLEY WAS FOUND LYING ON THE BED, NOT INSIDE THE PATIENT. FOLEY WAS TAKEN OFF THE BED BY OPERATING ROOM NURSE AND INSPECTED. THE BALLOON WAS REINFLATED WITH 10 CC OF SALINE AND OPERATING ROOM NURSE OBSERVED THE FOLEY BALLOON SLOWLY DEFLATE OVER SERVAL MINUTES. |  | N | Y |  |  | 2023-06
1018233-2023-04065 | 2023-06-06 | 2023-05-16 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO PEDIATRIC CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 52810G | 52810G |  | IT WAS REPORTED THAT THESE INTERMITTENT CATHETERS CAUSED PATIENT AN INFECTION AND WAS HOSPITALIZED AND CAUGHT TETANUS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR INFECTION AND TETANUS. PER FOLLOW UP VIA PHONE ON 22MAY2023, IT WAS REPORTED THAT THE CUSTOMER RECEIVED A VOLUNTARY EMERGENCY NOTICE FROM BD REGARDING A RECALL FOR A NONCONFORMITY WITH THE MAGIC GO 14 FR CATHETERS. THEY WERE REPORTED IN A COMPANY-WIDE NOTICE AS HAVING NONSTERILE COUDE AND THEY HAPPENED TO BE ONE OF THOSE IMPACTED BY IT. AT THE TIME, THEY WERE SENT OUT FROM 180 MEDICAL. THIS LED TO NUMEROUS BLADDER INFECTIONS AND SEPSIS, AND THIS MADE THINGS DIFFICULT FOR THEM AS THEY COULD NOT WALK. THEY STILL HAD SAMPLES AVAILABLE AND PROVIDED THE LOT# JUGW1445. PER SAMPLE RETURNED ON 14JUN2023, CUSTOMER HAD RETURNED WITH MATERIAL #50814G AND BACTH#JUFT0536. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02302 | 2023-06-06 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED, 3 OF CLINICIAN'S PICCS TODAY HAVE HAD THE SAME ISSUE. AGAIN ITS LEAKING AROUND THE WIRE THAT COMES OUT OF THE STOPPER. I ALSO HAD THIS HAPPEN WHEN I WAS WITH CLINICIAN THE OTHER DAY. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02303 | 2023-06-06 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED, 3 OF CLINICIAN'S PICCS TODAY HAVE HAD THE SAME ISSUE. AGAIN ITS LEAKING AROUND THE WIRE THAT COMES OUT OF THE STOPPER. I ALSO HAD THIS HAPPEN WHEN I WAS WITH CLINICIAN THE OTHER DAY. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02304 | 2023-06-06 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED, 3 OF CLINICIAN'S PICCS TODAY HAVE HAD THE SAME ISSUE. AGAIN ITS LEAKING AROUND THE WIRE THAT COMES OUT OF THE STOPPER. I ALSO HAD THIS HAPPEN WHEN I WAS WITH CLINICIAN THE OTHER DAY. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02306 | 2023-06-06 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH (TPS) STYLET 5F FULL TRAY (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2295108 |  | IT WAS REPORTED "AIR ENTERS WHEN PURGING THE LINE. THE CUSTOMER TELLS ME THAT DESPITE BEING THE SAME REFERENCE AS ALWAYS ((B)(4)) WE HAVE MODIFIED THE PARTS OF IT. THERE IS A RUBBER CONNECTOR THAT HAS A DIFFERENT SHAPE AND WHEN PURGING THE LINE, THEY SAY THAT AIR ENTERS THROUGH THAT PIECE, AND TO PURGE IT PROPERLY, THEY HAVE TO COVER THAT PIECE WITH THEIR FINGER TO PREVENT AIR FROM ENTERING. THIS DID NOT HAPPEN BEFORE." NO OTHER INFORMATION PROVIDED. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAK THROUGH THE T-LOCK SEPTUM WAS CONFIRMED. FOUR VIDEOS WERE PROVIDED FOR EVALUATION OF THIS COMPLAINT. THE VIDEOS WERE OF IMPLANTED POWERPICC SOLO CATHETERS. THE SOLO FUNNEL CONNECTORS AND T-LOCK ASSEMBLIES WERE ATTACHED TO THE LUER ADAPTORS ON THE CATHETERS. TWO OF THE VIDEOS WERE OF A NON-COMPLAINANT S/L POWERPICC CATHETER WITH THE PREVIOUS T-LOCK SEPTUM. WHEN AN ATTEMPT WAS MADE TO ASPIRATE THE CATHETER THROUGH THE T-LOCK ASSEMBLY, BLOOD WAS SEEN ASPIRATING THROUGH THE DEVICE AND INTO THE SYRINGE BARREL, PER DESIGN. THE OTHER TWO VIDEOS WERE OF THE IMPLICATED D/L POWERPICC SOLO CATHETER WITH THE NEWER T-LOCK SEPTUM. WHEN AN ATTEMPT WAS MADE TO ASPIRATE THE CATHETER THROUGH THE T-LOCK ASSEMBLY, AIR WAS SEEN ENTERING INTO THE SYRINGE BARREL. THE INTRODUCTION OF AIR INTO THE SYRINGE BARREL INDICATED A LEAK IN THE SYSTEM. THE YELLOW SEPTUM WAS THEN COVERED BY THE USER¿S GLOVED HAND AND AN ADDITIONAL ATTEMPT WAS MADE TO ASPIRATE FLUID THROUGH THE CATHETER. WHEN THE SEPTUM SURFACE WAS OCCLUDED, BLOOD WAS ASPIRATED THROUGH THE CATHETER AND INTO THE SYRINGE BARREL. THIS FUNCTIONAL TEST INDICATED THAT THE LEAK OCCURRED AT THE SEPTUM AND WAS LIKELY LOCATED WHERE THE STYLET TRAVERSED THROUGH THE T-LOCK SEPTUM. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02310 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | CUSTOMER REPORTED 2 OCCURRENCES OF A LEAKING T-LOCK IN THE LAST WEEK. NO OTHER INFORMATION PROVIDED. ADDITIONAL INFORMATION ON 30 MAY 2023 LEAKING OF FLUSH THROUGH STOPPER ON T CONNECTOR AT THE WIRE EXIT AREA. LEAKING WHEN FLUSHED, AND PULLING AIR INTO SYRINGE WHEN ASPIRATING TO CHECK BLOOD RETURN. NO PATIENT INJURY OR HARM, HOWEVER SIGNIFICANT RISK TO PATIENT (INFECTION, AIR EMBOLUS) AND INSERTER (BLOOD EXPOSURE). THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02308 | 2023-06-06 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH (TPS) STYLET 5F FULL TRAY (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2295108 |  | IT WAS REPORTED "AIR ENTERS WHEN PURGING THE LINE. THE CUSTOMER TELLS ME THAT DESPITE BEING THE SAME REFERENCE AS ALWAYS (2295108) WE HAVE MODIFIED THE PARTS OF IT. THERE IS A RUBBER CONNECTOR THAT HAS A DIFFERENT SHAPE AND WHEN PURGING THE LINE, THEY SAY THAT AIR ENTERS THROUGH THAT PIECE, AND TO PURGE IT PROPERLY, THEY HAVE TO COVER THAT PIECE WITH THEIR FINGER TO PREVENT AIR FROM ENTERING. THIS DID NOT HAPPEN BEFORE." NO OTHER INFORMATION PROVIDED. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02309 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | CUSTOMER REPORTED 2 OCCURRENCES OF A LEAKING T-LOCK IN THE LAST WEEK. NO OTHER INFORMATION PROVIDED. ADDITIONAL INFORMATION ON 30 MAY 2023: LEAKING OF FLUSH THROUGH STOPPER ON T CONNECTOR AT THE WIRE EXIT AREA. LEAKING WHEN FLUSHED, AND PULLING AIR INTO SYRINGE WHEN ASPIRATING TO CHECK BLOOD RETURN. NO PATIENT INJURY OR HARM, HOWEVER SIGNIFICANT RISK TO PATIENT (INFECTION, AIR EMBOLUS) AND INSERTER (BLOOD EXPOSURE). THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02294 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "THE STYLET IN THE PICC HAS CHANGED AND IT LEAKS WHEN THE PICC IS FLUSHED AND AIR ENTERS THE SYRINGE WITH THE REFLUX." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 10 DEVICES. THIS REPORT ADDRESSES THE FIFTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02312 | 2023-06-06 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL KIT W/BIOPATCH, PROBE COVER, & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE LEAKS BRISKLY AROUND STYLETTE WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. NO PATIENT OR CLINICIAN HARM REPORTED. IT WAS REPORTED THIS OCCURRED WITH EIGHT DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02313 | 2023-06-06 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL KIT W/BIOPATCH, PROBE COVER, & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE LEAKS BRISKLY AROUND STYLETTE WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. NO PATIENT OR CLINICIAN HARM REPORTED. IT WAS REPORTED THIS OCCURRED WITH EIGHT DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02314 | 2023-06-06 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL KIT W/BIOPATCH, PROBE COVER, & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE LEAKS BRISKLY AROUND STYLETTE WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. NO PATIENT OR CLINICIAN HARM REPORTED. IT WAS REPORTED THIS OCCURRED WITH (B)(4) DEVICES. ONLY (B)(4) SAMPLES WERE RETURNED FOR EVALUATION. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02319 | 2023-06-06 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL KIT W/BIOPATCH, PROBE COVER, & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE LEAKS BRISKLY AROUND STYLETTE WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. NO PATIENT OR CLINICIAN HARM REPORTED. IT WAS REPORTED THIS OCCURRED WITH EIGHT DEVICES. THIS REPORT ADDRESSES THE FIFTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02315 | 2023-06-06 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL KIT W/BIOPATCH, PROBE COVER, & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE LEAKS BRISKLY AROUND STYLETTE WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. NO PATIENT OR CLINICIAN HARM REPORTED. IT WAS REPORTED THIS OCCURRED WITH 8 DEVICES. ONLY 7 SAMPLES WERE RETURNED FOR EVALUATION. THIS REPORT ADDRESSES THE FOURTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02318 | 2023-06-06 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL KIT W/BIOPATCH, PROBE COVER, & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE LEAKS BRISKLY AROUND STYLETTE WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. NO PATIENT OR CLINICIAN HARM REPORTED. IT WAS REPORTED THIS OCCURRED WITH EIGHT DEVICES. THIS REPORT ADDRESSES THE SIXTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02316 | 2023-06-06 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL KIT W/BIOPATCH, PROBE COVER, & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE LEAKS BRISKLY AROUND STYLETTE WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. NO PATIENT OR CLINICIAN HARM REPORTED. IT WAS REPORTED THIS OCCURRED WITH EIGHT DEVICES. THIS REPORT ADDRESSES THE EIGHTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02317 | 2023-06-06 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL KIT W/BIOPATCH, PROBE COVER, & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE LEAKS BRISKLY AROUND STYLETTE WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. NO PATIENT OR CLINICIAN HARM REPORTED. IT WAS REPORTED THIS OCCURRED WITH 8 DEVICES. ONLY 7 SAMPLES WERE RETURNED FOR EVALUATION. THIS REPORT ADDRESSES THE SEVENTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02307 | 2023-06-06 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH (TPS) STYLET 5F FULL TRAY (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2295108 |  | IT WAS REPORTED "AIR ENTERS WHEN PURGING THE LINE. THE CUSTOMER TELLS ME THAT DESPITE BEING THE SAME REFERENCE AS ALWAYS ((B)(4)) WE HAVE MODIFIED THE PARTS OF IT. THERE IS A RUBBER CONNECTOR THAT HAS A DIFFERENT SHAPE AND WHEN PURGING THE LINE, THEY SAY THAT AIR ENTERS THROUGH THAT PIECE, AND TO PURGE IT PROPERLY, THEY HAVE TO COVER THAT PIECE WITH THEIR FINGER TO PREVENT AIR FROM ENTERING. THIS DID NOT HAPPEN BEFORE." NO OTHER INFORMATION PROVIDED. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02293 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "THE STYLET IN THE PICC HAS CHANGED AND IT LEAKS WHEN THE PICC IS FLUSHED AND AIR ENTERS THE SYRINGE WITH THE REFLUX." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 10 DEVICES. THIS REPORT ADDRESSES THE FOURTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02282 | 2023-06-06 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG* TCS SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770131 |  | FOUND DURING EVALUATION AT BD SERVICE FACILITY: THE SENSORS HOUSINGS ARE CHIPPED AND CRACKED. THE USB CABLE IS PULLING OUT OF THE SENSOR HOUSING, EXPOSING WIRES. THESE ISSUES ARE DUE TO USE OF THE DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: FOUND DURING EVALUATION AT BD SERVICE FACILITY: THE SENSORS HOUSINGS ARE CHIPPED AND CRACKED. THE USB CABLE IS PULLING OUT OF THE SENSOR HOUSING, EXPOSING WIRES. THESE ISSUES ARE DUE TO USE OF THE DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02291 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "THE STYLET IN THE PICC HAS CHANGED AND IT LEAKS WHEN THE PICC IS FLUSHED AND AIR ENTERS THE SYRINGE WITH THE REFLUX." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 10 DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02273 | 2023-06-06 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DECATHLON DF LONG-TERM HEMODIALYSIS CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | DF23SH28K | DF23SH28K |  | IT WAS REPORTED THAT DURING THE PREPARATION OF A DIALYSIS CATHETER PLACEMENT PROCEDURE, A SPLIT WAS ALLEGEDLY NOTICED IN THE TIP OF THE CATHETER UPON OPENING THE PACKAGE. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02274 | 2023-06-06 | 2023-03-11 | I, n, j, u, r, y | Manufacturer report | HICKMAN PORT SINGLE LUMEN | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0603880C |  | IT WAS REPORTED THAT FIVE DAYS POST THE PORT PLACEMENT VIA RIGHT INTERNAL JUGULAR VEIN, THE PORT WAS ALLEGEDLY UNABLE TO BE INFUSED. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY FOUND TO BE DETACHED UNDER DIGITAL SUBTRACTION ANGIOGRAPHY. FURTHERMORE, AFTER THE CATHETER DISCONNECTION, THE CATHETER HAD ALLEGEDLY FALLEN INTO THE HEART AND A CATCHER WAS USED TO CAPTURE THE CATHETER. REPORTEDLY THE TISSUE FIBERS WERE STRIPPED FROM THE PORT BODY WITH SURGICAL INSTRUMENTS AND SNARE WAS USED TO TAKE OUT THE PORT BODY. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 04/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02275 | 2023-06-06 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT SOMETIMES POST A PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED SKIN INFECTION. IT WAS FURTHER REPORTED THAT ANTIBIOTICS WERE GIVEN TO TREAT INFECTION. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02276 | 2023-06-06 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393190 | UNK GLIDEPATH |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT VIA RIGHT INTERNAL JUGULAR VEIN, THE CATHETER ALLEGEDLY KINKED IN THE PLASTIC TUBING THAT HOOKS UP TO THE LINES. IT WAS FURTHER REPORTED THAT DUE TO THE CATHETER TUBE KINK IT WAS ALLEGEDLY UNABLE TO PULL BLOOD DURING THE DIALYSIS TREATMENT. REPORTEDLY, THE CATHETER WERE REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE GLIDEPATH D/L 23CM CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, FUNCTIONAL AND TACTILE EVALUATIONS WERE PERFORMED. THE BLUE LUER CLAMP WAS NOTED TO BE FRACTURED. DISCOLORATION WAS NOTED THROUGHOUT THE SAMPLE. THE BLUE LUMEN WAS NOTED TO BE KINKED DISTAL TO THE CLAMP. BOTH THE EXTENSION LEGS WERE NOTED TO BE OFF-WHITE IN COLOR. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED DEFORMATION DUE TO COMPRESSIVE STRESS AND THE IDENTIFIED FRACTURE, DISCOLORATION, MATERIAL OPACIFICATION AND MATERIAL SEPARATION ISSUES. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED SUCTION ISSUES AS THERE WERE NO ASPIRATION DIFFICULTY WAS IDENTIFIED DURING EVALUATION AND THE SAMPLE EVALUATION RESULTS INDICATING NO DIFFICULTY UNDER LABORATORY CONDITIONS ARE NOT BY THEMSELVES SUFFICIENT TO CONFIRM THAT THIS EVENT DID NOT OCCUR UNDER CLINICAL CONDITIONS. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02277 | 2023-06-06 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393190 | 5393190 |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT VIA RIGHT INTERNAL JUGULAR VEIN, THE CATHETER ALLEGEDLY KINKED IN THE PLASTIC TUBING THAT HOOKS UP TO THE LINES. IT WAS FURTHER REPORTED THAT DUE TO THE CATHETER TUBE KINK IT WAS ALLEGEDLY UNABLE TO PULL BLOOD DURING THE DIALYSIS TREATMENT. REPORTEDLY, THE CATHETER WERE REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, G3, H6 (DEVICE) H11: E1, H6 (METHOD, RESULT, CONCLUSION) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02278 | 2023-06-06 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT SOMETIMES POST A PORT PLACEMENT, THE DEVICE ALLEGEDLY COULD NOT GET A BLOOD RETURN. IT WAS FURTHER REPORTED THAT THE DEVICE ALLEGEDLY HAD DIFFICULT TO INFUSE THE MEDICATIONS. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02279 | 2023-06-06 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT SOMETIMES POST A PORT PLACEMENT, THE DEVICE ALLEGEDLY COULD NOT GET A BLOOD RETURN. IT WAS FURTHER REPORTED THAT THE DEVICE ALLEGEDLY HAD DIFFICULT TO INFUSE THE MEDICATIONS. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00397 | 2023-06-06 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | AT120144 | AT120144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE RIGHT ARM ATRIOVENTRICULAR FISTULA, THE PTA BALLOON CATHETER ALLEGEDLY GOT STUCK WHILE REMOVING IT FROM THE SHEATH. IT WAS FURTHER REPORTED THAT AFTER SEVERAL ATTEMPTS OF REMOVAL, THE BALLOON ALLEGEDLY SEPARATED FROM THE SHAFT WHILE IN THE SHEATH. REPORTEDLY, THE DETACHED BALLOON WAS REMOVED FROM THE PATIENT COMPLETELY. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 01/2026). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04050 | 2023-06-06 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414 |  | IT WAS REPORTED THAT THE WATER WAS UNABLE TO REMOVED FROM THE FOLEY CATHETER DURING A PRETEST. ABOUT 0 ML OF WATER WAS ASPIRATED BY SYRINGE. THE USER TRIED TO DEFLATE BALLOON BY SPONTANEOUS OR GENTLE DEFLATION BY SYRINGE AND ASPIRATION BY SYRINGE MANUALLY. USUALLY, BALLOON CAN BE DEFLATED WITHIN 20-40 SECONDS. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEM OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDO-ELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS. 1) MALFUNCTION. - CATHETER KINKING, DAMAGE, RUPTURE. - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE. - OCCLUSION OF CATHETER INNER LUMENS. - ENCRUSTATION. - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE. - FAILURE TO MEASURE TEMPERATURE. - IMPROPER TEMPERATURE INDICATION. 2) ADVERSE EVENTS. - URINARY-TRACT INFECTION. - HEMORRHAGE, HEMATURIA. - ALLERGY REACTION TO THE DEVICE. - CALCULUS FORMATION. - EDEMA. - PAIN. - DISCOMFORT. - INJURY OF BLADDER OR URETHRAL. - URETHRITIS, URINARY INCONTINENCE .- RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]. 1. STORAGE. STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04046 | 2023-06-06 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED, THAT THERE WAS A PROTRUSIONS ON FOLEY CATHETER BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04047 | 2023-06-06 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UROFORCE® BALLOON DILATION CATHETER | DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 887410 | 887410 |  | IT WAS REPORTED THAT A PHYSICIAN STATED IN AN ONLINE SURVEY THAT THEY WERE NOT ABLE TO SUCCESSFULLY SUSTAIN AND/OR MONITOR PRESSURE WITH THE EAGLE INFLATION DEVICE. STATED THAT THE UROFORCE BALLOON DILATION CATHETER WITH EAGLE INFLATION DEVICE WAS NOT SUCCESSFUL IN SUSTAINING AND/OR MONITORING PRESSURE BECAUSE OF LACK OF NURSING STAFF EXPERIENCE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04051 | 2023-06-06 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DUE TO LEAK OF THE FIXATION WATER. THE CUSTOMER HAD REQUESTED AN INVESTIGATION FOR THE CAUSE OF THE LEAK. | THE REPORTED EVENT WAS NOT CONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS NOT CONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04054 | 2023-06-06 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT URINE WAS UNABLE TO DRAINED FROM THE FOLEY CATHETER. AFTER 4 WEEKS OF DEVICE PLACEMENT, URINE DID NOT FLOW AND THE PATIENT'S ABDOMEN WAS DISTENDED, SO THE CATHETER WAS REPLACED AND ABOUT 1400 ML OF URINE FLOWED OUT. WHILE THE CATHETER WAS REMOVED, ABOUT 10 ML OF AIR CAME OUT ALONG WITH 6 ML OF FIXATION WATER. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. BASED ON EVALUATION FINDING, OBSERVED SALT ACCUMULATION THAT CAUSING DRAINAGE LUMEN OCCLUSION WHICH CAN CAUSE FLOWRATE ISSUE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE)/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE/SALT ACCUMULATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04056 | 2023-06-06 | 2023-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED, THAT THE PATIENT COMPLAINED OF ABDOMINAL DISTENSION THREE HOURS AFTER THE OPERATION. WITH THE CATHETER NOT IN PLACE SMOOTHLY AND ONLY A SMALL AMOUNT OF BLOODY FLUID WAS DRAINED. AS THE DRAINAGE SPEED WAS NOT EQUAL TO THE IRRIGATION RATE. UPON CHECKING THE CATHETER, IT WAS FOUND BURST AND THEN WAS REPLACED WITH A NEW CATHETER. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, CLARIFIED THAT THERE WAS RESTRICTED FLOW OF FLUID. AS IT WAS MENTIONED AS ONLY A FEW AMOUNT OF BLOODY FLUID WAS DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02283 | 2023-06-06 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR (RADEL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770352 |  | IT WAS FOUND DURING EVALUATION AT BD SERVICE FACILITY: THE SENSORS USB CABLE IS CUT, EXPOSING WIRES DUE TO USE OF THE DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: FOUND DURING EVALUATION AT BD SERVICE FACILITY: THE SENSORS USB CABLE IS CUT, EXPOSING WIRES DUE TO USE OF THE DEVICE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04061 | 2023-06-06 | 2023-05-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT PATIENT HAD URINARY TRACT INFECTION, BLEEDING, PAIN DURING MAGIC 3 GO CATHETER USE. THEY SET ASIDE FULL 340 QTY PER DOCTOR INSTRUCTION. MEDICAL INTERVENTION WAS UNKNOWN. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND TO BE A DUPLICATE OF AN EVENT PREVIOUSLY REPORTED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04059 | 2023-06-06 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DUE TO THE BALLOON BURST AFTER PLACEMENT. THE CATHETER WAS DISCONTINUED TO USE. THERE WERE NO BLADDER STONES. | THE REPORTED EVENT WAS CONFIRMED ¿ UNKNOWN CAUSE. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04060 | 2023-06-06 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DUE TO THE BALLOON BURST AFTER PLACEMENT. THE CATHETER WAS DISCONTINUED TO USE. AS PER MAIL RECEIVED ON 22MAY2023, STATED THAT THE BALLOON RUPTURE WAS NOT OBSERVED TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04063 | 2023-06-06 | 2023-05-16 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT PATIENT HAD MANY SERIOUS URINARY TRACT INFECTION DURING THE TIME FRAME SPECIFIED IN THE RECALL NOTICE. PATIENT WAS ADVISED TO DISCARD THE REMAINING CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. PER ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023, PATIENT STATED THAT PATIENT BROKE OUT IN A HORRIBLE RASH AND HAD A BAD URINARY TRACT INFECTION WITH USE OF INTERMITTENT CATHETERS. PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATED THAT PATIENT WANTED TO BE COMPENSATED FOR PATIENTS' DOCTORS' VISITS. REPRESENTATIVE ADVISED THAT WITHOUT LOT NUMBER NO REPLACEMENT COULD BE DONE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS, URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02290 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "THE STYLET IN THE PICC HAS CHANGED AND IT LEAKS WHEN THE PICC IS FLUSHED AND AIR ENTERS THE SYRINGE WITH THE REFLUX." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 10 DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02292 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "THE STYLET IN THE PICC HAS CHANGED AND IT LEAKS WHEN THE PICC IS FLUSHED AND AIR ENTERS THE SYRINGE WITH THE REFLUX." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 10 DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02296 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "THE STYLET IN THE PICC HAS CHANGED AND IT LEAKS WHEN THE PICC IS FLUSHED AND AIR ENTERS THE SYRINGE WITH THE REFLUX." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 10 DEVICES. THIS REPORT ADDRESSES THE SEVENTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02288 | 2023-06-06 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVE, NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D5 |  | CUSTOMER REPORTS INDWELLING SINGLE LUMEN PICC LINE WAS DAMAGED WHEN THE PATIENT PULLED ON THE EXTENSION LEGS, CAUSING THE EXTENSION LEGS TO BREAK AWAY FROM THE DEVICE. PICC LINE WAS REMOVED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR S/L POWERPICC CATHETER. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. THE EXTENSION TUBING WAS COMPLETELY DETACHED FROM THE MOLDED JOINT AND WAS NOT RETURNED FOR EVALUATION. MICROSCOPIC INSPECTION OF THE PROXIMAL END OF THE MOLDED JOINT REVEALED WHAT APPEARED TO BE A FRAGMENT OF EXTENSION TUBING REMAINING WITHIN. THE PRESENCE OF A TUBING FRAGMENT WAS CONFIRMED WHEN THE SAMPLE WAS DISSECTED. THE FRAGMENT EXHIBITED A GRANULAR FRACTURE SURFACE WITH A FLAT, EVEN SURFACE. THE TUBING BREAK FEATURES WERE TYPICAL OF A TEARING FAILURE OF THE MATERIAL. THESE OBSERVATIONS WERE CONSISTENT WITH THE EVENT DESCRIPTION, WHICH INDICATED THAT THE CATHETER WAS EXPOSED TO PULLING FORCE WHEN IN PLACE. IT APPEARED THAT PULLING RESULTED IN TENSILE FAILURE OF THE EXTENSION TUBING. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02289 | 2023-06-06 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX KIT W/BIOPATCH, PROBE COVER, NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CLINICIAN REPORTS NORMAL SALINE LEAKS BRISKLY AROUND STYLETTE WIRE THROUGH THE T-LOCK STOPPER WHILE PREPARING THE CATHETER FOR INSERTION. NO PATIENT OR CLINICIAN HARM REPORTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02280 | 2023-06-06 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT ALLEGEDLY ERODED OVER THE PATIENT'S THIN SKIN. REPORTEDLY, THE PORT WAS REMOVED. THE PATIENT IS REPORTED TO BE STABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02297 | 2023-06-06 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "THE STYLET IN THE PICC HAS CHANGED AND IT LEAKS WHEN THE PICC IS FLUSHED AND AIR ENTERS THE SYRINGE WITH THE REFLUX." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH 10 DEVICES. THIS REPORT ADDRESSES THE EIGHTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04066 | 2023-06-07 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED AND THE FOLEY CATHETER BALLOON DID NOT INFLATE DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04090 | 2023-06-07 | 2023-05-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED, THAT THE PATIENT HAD MULTIPLE URINARY TRACT INFECTIONS. FROM USING THE MAGIC 3 GO INTERMITTENT CATHETER (LOT#: JUGT3350), (LOT#: JUGT3297). IT WAS UNKNOWN, WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE, BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN, WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN, WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED, AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT, INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION. COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE. WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS: URINARY TRACT INFECTION BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE, WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE. AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED, TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
2020394-2023-00401 | 2023-06-07 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8074 | VA8074 |  | IT WAS REPORTED THAT DURING THE PREPARATION OF AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY BROKE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE VACCESS PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, A COMPLETE CIRCUMFERENTIAL BREAK TO THE CATHETER AT THE GLUE WELD WAS NOTED LEAVING THE INNER GUIDEWIRE LUMEN EXPOSED. NO ANOMALIES WERE NOTED TO THE LUERS/Y-BODY. NO FUNCTIONAL TESTING WAS PERFORMED DUE TO THE CONDITION OF THE SAMPLE. ALSO, THREE PHOTOS WERE PROVIDED AND REVIEWED. THE FIRST PHOTO SHOWS THE INNER PACKAGING OF THE VACCESS DEVICE. THE LABEL MATCHES WITH THE INFORMATION PROVIDED IN TRACKWISE. THE SECOND AND THIRD PHOTOS SHOW THE BALLOON PORTION OF THE VACCESS CATHETER. A COMPLETE CIRCUMFERENTIAL BREAK IS SEEN PROXIMAL TO THE BALLOON WITH THE INNER GUIDEWIRE LUMEN EXPOSED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED BREAK AS A COMPLETE CIRCUMFERENTIAL BREAK OF THE CATHETER WITH THE INNER GUIDEWIRE LUMEN EXPOSED WAS SEEN BOTH FROM THE PROVIDED PHOTOS AND THE RETURNED SAMPLE. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BREAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 10/2025), G3. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02331 | 2023-06-07 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT SOMETIME POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY EXPELLED OUT OF THE CUFF. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY APPEARED PAPERY WITH NO FIBROSIS. THE CATHETER ALLEGEDLY MIGRATED AND CAME OUT FROM THE PATIENT BODY. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 05/2024) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04080 | 2023-06-07 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 106173P | 106173P |  | IT WAS REPORTED THAT CUSTOMER HAD A PRODUCT FAILURE WITH THE TEMPORARY PACING ELECTRODE CATHETER IN THE INTENSIVE CARE UNIT. THE BALLOON DID NOT INFLATE WHEN PHYSICIAN ATTEMPTED. ALSO STATED THAT THEY HAD THE FAILED COMPONENT IN THEIR OFFICE FOR INVESTIGATION, WHICH WAS A PART OF A KIT THAT THEY DID NOT CARRY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04083 | 2023-06-07 | 2023-05-16 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION WHILE USING THE MAGIC3 GO MALE INTERMITTENT CATHETER AND THE ANTIBIOTICS WERE PRESCRIBED BY THE DOCTOR, AND IT WORKED. THE PATIENT REPORTED ISSUES WITH LOT#: JUGT1691, LOT#: JUGT2738. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, CUSTOMER WAS PRESCRIBED A FIVE DAY COURSE OF BACTRIM FOR THE INFECTION AND IT CLEARED UP WITHIN 10 DAYS WITH NO ADDITIONAL MEDICAL ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04082 | 2023-06-07 | 2023-05-16 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS HAD CAUSED A LOT OF ISSUES WITH THEIR HEALTH. PATIENT HAD BEEN HOSPITALIZED, AND THEY HAD HOME THERAPY, AND HAD BACTERIA AND BLADDER AND KIDNEY ISSUES FROM USING THE CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04081 | 2023-06-07 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT THE FIXATION WATER LEAKED FROM THE LUMEN OF THE FOLEY CATHETER WHEN ATTEMPTING TO PLACE THE CATHETER TO THE PATIENT. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS INFLATED WITH WATER AND OBSERVED WATER LEAKING THROUGH DRAINAGE EYE. DISSECTED THE CATHETER AND OBSERVED TEAR ON RUBBERIZE LAYER OF INFLATION LUMEN. TEAR WAS EXAMINED UNDER MICROSCOPE AND NOTED TO BE ROUNDED AND NOT SMOOTH AND IDENTIFIED THIS IS RELATED TO MANUFACTURING. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO OPERATOR ERROR-"POOR STRIPPING OR ROUGH STRIPPING". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARDEX® SILVER LUBRICATH® FOLEY CATHETER SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04075 | 2023-06-07 | 2023-05-16 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50812G | 50812G |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION ON (B)(6) 2023. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04074 | 2023-06-07 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT 4 OF THE MAGIC 3 GO MALE INTERMITTENT CATHETER HAD NO LUBE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
2020394-2023-00399 | 2023-06-07 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQ12064 | CQ12064 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY STUCK IN THE SHEATH UPON REMOVAL. IT WAS FURTHER REPORTED THAT BALLOON WAS ALLEGEDLY UNABLE TO BE REMOVED THROUGH THE SHEATH. REPORTEDLY SHEATH AND BALLOON WERE REMOVED INTACT AND AN ADDITIONAL ACCESS SITE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | ADDITIONAL INFORMATION WAS RECEIVED, AND THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE REPORTABLE AS MALFUNCTION. MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE CONQUEST PTA DILATATION CATHETER WAS RETURNED FOR EVALUATION. IT WAS NOTED THAT THE CATHETER WAS NOTED TO BE LOADED WITH AN UNKNOWN SHEATH AND GUIDE WIRE. THE SHEATH WAS NOTED TO BE BUNCHED AND THE DISTAL END OF THE BALLOON WAS NOTED TO BE KINKED. NO OTHER SPECIFIC ANOMALIES WERE NOTED DURING THE VISUAL EVALUATION. DURING THE FUNCTIONAL TESTING, A NEGATIVE PRESSURE WAS APPLIED TO REMOVE THE REMAINING AIR FROM THE BALLOON. A HUGE FORCE WAS APPLIED TO REMOVE THE LOADED SHEATH FROM THE CATHETER BUT IT WAS UNSUCCESSFUL. DURING THE TESTING, THE CATHETER WAS STRETCHED INADVERTENTLY. THE INSERTED GUIDE WIRE WAS REMOVED FROM THE CATHETER WITHOUT ANY PROBLEMS. NO FURTHER TESTING WAS PERFORMED. AS THE RETURNED CATHETER WAS NOTED TO BE LOADED WITHIN AN UNKNOWN SHEATH AND UNABLE TO BE REMOVED DURING THE VISUAL EVALUATION, THE INVESTIGATION FOR THE REPORTED SHEATH REMOVAL DIFFICULTY WAS CONFIRMED. A DEFINITIVE ROOT CAUSE FOR THE REPORTED SHEATH REMOVAL DIFFICULTY COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. D4 (EXPIRY DATE: 10/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04076 | 2023-06-07 | 2023-05-16 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION WHILE USING THE MAGIC3 GO MALE INTERMITTENT CATHETER AND THE ANTIBIOTICS WERE PRESCRIBED BY THE DOCTOR, AND IT WORKED. THE PATIENT REPORTED ISSUES WITH LOT#: JUGT1691, LOT#: JUGT2738. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, CUSTOMER WAS PRESCRIBED A FIVE DAY COURSE OF BACTRIM FOR THE INFECTION AND IT CLEARED UP WITHIN 10 DAYS WITH NO ADDITIONAL MEDICAL ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04077 | 2023-06-07 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123620A |  | IT WAS REPORTED THAT THE PATIENT GOT CATHETERIZATION FOR 6 DAYS AND WHEN THE PATIENT WAS IN THE TOILET, THE CATHETER FELL OUT. AFTER CHECKING, IT WAS FOUND THAT THE BALLOON HAD BURST. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO THIN SAC CAUSING BY SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04071 | 2023-06-07 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT CUSTOMER NOTICED WHILE OPENING THE BOX OF INTERMITTENT CATHETER WITH LOT# JUGX1192, CONTENTS APPEARED TO BE PACKED IMPROPERLY AND ALSO REPORTED THAT THE NON-INSERTION END OF THE CATHETERS ARE JAGGED AND DID NOT LUBRICATE LIKE THEY SHOULD AND CAUSED MINOR BLEEDING AND DISCOMFORT. CUSTOMER SERVICE HAD BEEN NOTIFIED BECAUSE THEY WERE UNABLE TO TRANSFER TO THEM THIS MORNING, AND THEY HAVE BEEN ASKED TO CONTACT PATIENT TO DO AN EXCHANGE OF THIS PRODUCT. PATIENT HAD USED A FEW OUT OF ONE BOX AND HAS AN ENTIRE UNOPENED BOX OF THE SAME LOT#. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04070 | 2023-06-07 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER WAS SPONTANEOUSLY REMOVED FROM THE FOLEY CATHETER BALLOON A DAY AFTER THE CATHETER PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04069 | 2023-06-07 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED HOLD CM 6.7THAT 15 OF THE MAGIC 3 HYDROPHILIC INTERMITTENT COUDE CATHETERS HAD BEEN OPENED AND THE RING THAT PULLS THE CATHETER OPEN WAS ALREADY OPEN TO THE POINT THE SURE GRIP SLEEVE WAS SHOWING AND THE HYDROPHILIC POUCH WAS NOT IN THERE OR HAD BEEN SMASHED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04067 | 2023-06-07 | 2023-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER FELL OUT OF THE PATIENT AFTER INSERTION. UPON CHECKING, THE BALLOON WAS FOUND TO BE TORN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02327 | 2023-06-07 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602833 | 0602833 |  | IT WAS REPORTED THAT ONE YEAR TWO MONTHS AND SEVENTEEN DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, THE PATIENT ALLEGEDLY EXPERIENCED REDNESS OF THE NECK DURING CHEMOTHERAPY AND FEVER WAS NOTED. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 12/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02326 | 2023-06-07 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602833 | 0602833 |  | IT WAS REPORTED THAT ONE YEAR THREE MONTHS AND TWENTY EIGHT DAYS POST PORT PLACEMENT, UPON REMOVAL A CRACK WAS ALLEGEDLY OBSERVED AT A SITE COINCIDING WITH THE REDDENED NECK PORTION OF THE REMOVED CATHETER. REPORTEDLY, THE PORT WAS REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 12/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02325 | 2023-06-07 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN PORT SINGLE LUMEN | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0603880C |  | IT WAS REPORTED THAT SOMETIME POST THE PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, DURING ASPIRATION AND INFUSION, THERE WAS NO BLOOD RETURN, AND INFUSION WAS ALLEGEDLY DIFFICULT TO GIVE USING THE PORT. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY FOUND TO BE FRACTURED AT THE CONNECTION OF THE LOCK, AND IT WAS CONFIRMED THROUGH FILMING THAT THE CATHETER HAD ALLEGEDLY FALLEN OFF. REPORTEDLY, THE CATHETER AND PORT BODY WERE IMMEDIATELY RETRIEVED WITH A RETRIEVAL DEVICE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 04/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04091 | 2023-06-07 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE MISSHAPEN. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02335 | 2023-06-07 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO HF CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1395108QD |  | IT WAS REPORTED, TRIPLE LUMEN PICC LINE WAS PLACED TODAY AND WHEN REMOVING THE STYLE IT WAS CURLED. ADDITIONAL INFORMATION 5/22/2023: PATIENT WAS NOT COMPROMISED, AND NO INJURY OCCURRED. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A CURLED STYLET WAS CONFIRMED AND THE CAUSE APPEARED TO BE USE-RELATED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE ECG TLS STYLET. THE WIRE WAS INLAID THROUGH THE T-LOCK ASSEMBLY. A REGION OF THE STYLET WIRE BETWEEN THE PLASTIC-COATED REGION AND THE T-LOCK ASSEMBLY WAS TIGHTLY COILED. MICROSCOPIC INSPECTION OF THE STYLET CONFIRMED THAT A PORTION OF THE WIRE WAS TIGHTLY COILED. THE PLASTIC-COATED REGION WAS INTACT. THE OBSERVED PERMANENT STYLET WIRE SHAPE MEMORY WAS REPLICATED USING A NON-COMPLAINANT DEVICE BY WITHDRAWING THE STYLET WHEN THE CATHETER WAS CONSTRAINED AND WAS NOT FLUSHED PRIOR TO STYLET WITHDRAWAL. THE OBSERVED STYLET DEFORMATION APPEARED CONSISTENT WITH DAMAGE CAUSED BY STYLET REMOVAL AGAINST RESISTANCE, LIKELY DURING CATHETER PLACEMENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04088 | 2023-06-07 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER LEAKED FROM THE BALLOON AREA IN FOLEY CATHETER. | THE REPORTED EVENT WAS CONFIRMED-MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS INFLATED WITH WATER AND OBSERVED WATER LEAKING THRU DRAINAGE EYE. DISSECTED THE CATHETER AND OBSERVED TEAR ON RUBBERIZE LAYER OF INFLATION LUMEN. TEAR WAS EXAMINED UNDER MICROSCOPE AND NOTED TO BE ROUNDED AND SMOOTH AND IDENTIFIED THIS IS RELATED TO MANUFACTURING. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE CONTRIBUTED BY TORN RUBBERIZE LAYER THAT MAY LEAD TO LUMEN-TO-LUMEN LEAK. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "SINGLE USE ONLY [WARNINGS] 1.METHOD FOR USE (1)DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2)DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2.APPLICABLE PATIENTS (1)PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1)DO NOT REUSE. (2)DO NOT RESTERILIZE. (3)BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4)DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1)DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA®BIOCATH®FOLEYCATHETER IS A BALLOON CATHETER.THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS.SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR 1. BALLOONCATHETER INFLATION VALVE IRRIGATION FUNNEL DRAINAGE FUNNEL INFLATION FUNNEL BALLOON <CROSS SECTION OF CATHETER> DRAINAGE LUMEN IRRIGATION LUMEN INFLATION LUMEN CAP 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT-EFFICACY] THIS DEVICE ISDESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1)CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2)LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3)INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4)PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5)TO DEFLATE BALLOON AND REMOVE CATHETER, INSERTA LUERTIP(NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATERTO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE.[THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATEURINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNELAND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1.PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1)EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1)WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2)WHEN ANY PART OFTHEBALLOON AND/OR THE CATHETERIS MISSING,CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3)WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON,TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAYREMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1)ATTACH LUER TIP SYRINGE TO THE INFLATIONVALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2)IF SITUATIONWOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) 3)IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) 4)IF SITUATION WOULDN'T BE IMPROVED WITH3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) 5)IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) B. BALLOON-RUPTURE METHOD 1)INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) MINERAL OIL 2)IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A)UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER.(FIG.6) B)IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUMUNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) C)IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) 3.MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION -CATHETER KINKING, DAMAGE, RUPTURE -DIFFICULTY OR FAILURE TO REMOVE THE DEVICE -OCCLUSION OF CATHETER INNER LUMENS -ENCRUSTATION -ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE -DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS -URINARY-TRACT INFECTION -HEMORRHAGE, HEMATURIA -ALLERGY REACTION TO THE DEVICE -CALCULUS FORMATION -EDEMA -PAIN -DISCOMFORT -INJURY OF BLADDER OR URETHRAL -URETHRITIS, URINARY INCONTINENCE -RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02337 | 2023-06-07 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD TRIALYSIS HD CATHETER 3F | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED, "THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. EVENT INVOLVES A 3F CATHETER INSERTED IN THE NECK." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION AND SPECIFIC PATIENT DETAILS WERE NOT DISCLOSED AS PART OF THE SURVEY. AS SUCH, ADDITIONAL INFORMATION AND THE SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02328 | 2023-06-07 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT THREE MONTHS AND TEN DAYS POST THE PORT PLACEMENT VIA THE LEFT INTERNAL JUGULAR VEIN, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN INSIDE THE PATIENT, AND THE END OF THE CATHETER WAS LOCATED AT THE JUNCTION OF THE SUPERIOR VENA CAVA AND RIGHT ATRIUM. IT WAS FURTHER REPORTED THAT THE DEVICE ALLEGEDLY HAD SUBCUTANEOUS DRUG LEAKAGE. FURTHERMORE, THE PATIENT EXPERIENCED SWELLING IN THE NECK AND UPPER PART OF THE CHEST, WHICH WAS TREATED WITH HOT COMPRESSION AND OVER-THE-COUNTER CREAM. REPORTEDLY, THE PORT WAS REMOVED AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORD IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A VIDEO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 11/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02324 | 2023-06-07 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT ONE YEAR AND EIGHT MONTHS POST THE PORT PLACEMENT, WHEN THE PORT SYSTEM WAS REMOVED AFTER THE COMPLETION OF THE CHEMOTHERAPY, A CRACK WAS ALLEGEDLY FOUND AT APPROXIMATELY FOURTEEN CENTIMETERS FROM THE TIP OF THE REMOVED CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT TO A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. IN ADDITION TO THE RETURNED SAMPLE ANALYSIS, TWO ELECTRONIC IMAGES WERE PROVIDED FOR REVIEW. THE PORT POSITIONING APPEARS NORMAL ON THE CHEST WALL WITH THE CATHETER ASCENDING TO THE INTERNAL JUGULAR VEIN INSERTION AND THEN DESCENDING INTO THE SUPERIOR VENA CAVA. THE PORT IS IN THE CORRECT PLACE WITH INSERTION SITE AT THE INTERNAL JUGULAR VEIN. FUNCTIONAL, GROSS VISUAL, TACTILE, MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. A SPLIT WAS NOTED APPROXIMATELY 9.7CM FROM THE DISTAL END OF THE CATH-LOCK. UPON INFUSION OF THE PORT BODY, A LEAK FROM THE SPLIT ON THE CATHETER WAS OBSERVED WHILE THROMBUS EXITED THE DISTAL END. ASPIRATION OF WATER INTO THE SYRINGE WAS ATTEMPTED AND WAS UNSUCCESSFUL. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 08/2023), G3, H6 (METHOD). H11: H6 (RESULT, CONCLUSION) . H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02338 | 2023-06-07 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY THE CUSTOMER THAT "DURING A TRAINING A MOST OF THE ACCUCATHS WERE VERY DIFFICULT TO SAFETY AND SOME DID NOT SAFETY AT ALL. YOU WOULD PRESS THE BUTTON AND THE NEEDLE JUST WOULD NOT RETRACT." ADDITIONAL INFO RECEIVED ON 24MAY2023: IT WAS REPORTED "THE SAFETY BUTTON WAS ABLE TO BE PRESSED BUT BECAME STUCK DOWN ONCE IT WAS PRESSED AND THE NEEDLE DID NOT RETRACT. IF THE NEEDLE ITSELF WAS PHYSICALLY WIGGLED, IT WOULD THEN RETRACT. THERE WAS NO VIABLE BLOCKAGE OF THE NEEDLE." THE EXACT QUANTITY OF INVOLVED ACCUCATHS WAS NOT PROVIDED TO BD VASCULAR DEVICES FIELD ASSURANCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02321 | 2023-06-07 | 2023-04-26 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1606052 | 1606052 |  | IT WAS REPORTED THAT SOMETIMES PORT PLACEMENT PROCEDURE, THE PATIENT EXPERIENCED SKIN INFECTION. IT WAS FURTHER REPORTED ANTIBIOTICS WERE GIVEN TO TREAT INFECTIONS AND THEN PORT WAS REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 03/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02322 | 2023-06-07 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6CF INT W SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806060 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY KINKED AT THE PUNCTURE SITE. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY HAD NO BACKFLOW AND FORWARD FLOW. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL SAMPLE WAS NOT RETURNED FOR EVALUATION. ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS THE IMPLANTED CATHETER WHICH WAS NOTED TO HAVE A KINK AT THE INSERTION SITE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED DEFORMATION DUE TO COMPRESSIVE STRESS ISSUE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FLUSHING PROBLEM AND SUCTION ISSUES AS THE EXACT CIRCUMSTANCE AT THE TIME OF THE EVENT REPORTED WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 07/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02320 | 2023-06-07 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR AND TEN MONTHS POST THE PORT PLACEMENT PROCEDURE, A CRACK WAS ALLEGEDLY FOUND ON THE REMOVED CATHETER WHEN THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, THREE CHEST X-RAYS WERE PROVIDED FOR REVIEW. THE IMAGE SHOWS THE DEVICE INSERTED VIA THE RIGHT JUGULAR VEIN AND NO DISRUPTION OF THE CATHETER IS EVIDENT. IN VISUAL EVALUATION, A COMPOUND BREAK WAS NOTED APPROXIMATELY 10.0CM FROM THE DISTAL END OF THE CATH-LOCK. THE EDGES OF THE COMPOUND BREAK ON THE ATTACHED CATHETER WERE NOTED TO BE UNEVEN AND THE SURFACE WAS NOTED TO BE ROUND AND SMOOTH IN BOTH REGIONS. UPON INFUSION, A LEAK FROM THE COMPOUND BREAK WAS OBSERVED WHILE WATER WITH THROMBUS EXITED THE DISTAL END OF THE CATHETER. ASPIRATION WAS ATTEMPTED AND WAS UNSUCCESSFUL. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02323 | 2023-06-07 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT ONE YEAR AND TWENTY SEVEN DAYS POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN DURING A ROUTINE PORT FLUSHING. IT WAS FURTHER REPORTED THAT THE DEVICE ALLEGEDLY LEAKED. REPORTEDLY, THE DISTAL CATHETER SEGMENT AND THE PORT BODY WERE REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION AND TWO IMAGES WERE PROVIDED FOR REVIEW. VISUAL, MICROSCOPIC FUNCTIONAL AND TACTILE EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED TO THE DISTAL END OF THE ATTACHED CATHETER AND THE PROXIMAL END OF THE DISTAL CATHETER SEGMENT. THE EDGES OF THE COMPLETE CIRCUMFERENTIAL BREAKS WERE NOTED TO BE UNEVEN. THE SURFACES OF BOTH THE BREAKS WERE NOTED TO BE GRANULAR IN ONE REGION, AND ROUND/SMOOTH IN THE OTHER REGION. THE IMAGE SHOWS THE PORT WITH THE CENTRAL CATHETER REMOVED AND THE REMAINING PORTION ATTACHED TO THE PORT. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION ISSUES. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED LEAK ISSUE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 07/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02339 | 2023-06-07 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 2FR SL W/O GW BULK NON | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | SA5167436B |  | IT WAS REPORTED BY THE CUSTOMER TECHNICAL AREA, "PICC RUPTURED SPONTANEOUSLY AFTER 8 DAYS OF USE IN THE NB, WHILE INFUSING INTRAVENOUS HYDRATION. IT RUPTURED ON THE OUTSIDE, WAS REMOVED IMMEDIATELY, NEONATOLOGIST INFORMED." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED 05/29/2023: THE BREAK OCCURRED ON THE OUTSIDE. THERE WAS NO HARM TO THE PATIENT, BUT IT WAS NECESSARY TO REMOVE THIS PICC AND PLACE ANOTHER ONE. THEREFORE, THERE WAS A NEED FOR A NEW PROCEDURE. THERE WAS NO NEED FOR MEDICAL/SURGICAL INTERVENTION, ONLY PICC REPLACEMENT. INTRAVENOUS HYDROCHLORIDE WAS BEING INFUSED. THERE WAS LEAKAGE OF THIS INTRAVENOUS HYDRATION. ADDITIONAL INFORMATION RECEIVED 26 JUNE 2023: THE PROFESSIONALS WHO HANDLE THIS MATERIAL HAVE A SPECIFIC COURSE CERTIFICATE FOR THIS PURPOSE. ONLY NURSES WITH THIS CERTIFICATION HANDLE THE PICC. THE PRODUCT WAS USED IN ACCORDANCE WITH THE CONDITIONS OF USE. THERE WAS NO TRACTION MOVEMENT DURING CATHETER INSERTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02334 | 2023-06-07 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS 3CG, GUARDIVA, PROBE COVER, BELDAR NG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER, OCCURRENCES OF A LEAKING T-LOCK IN THE LAST WEEK. ADDITIONAL INFORMATION RECEIVED (B)(6)2023: "LEAKING OF FLUSH THROUGH STOPPER ON T CONNECTOR AT THE WIRE EXIT AREA. LEAKING WHEN FLUSHED, AND PULLING AIR INTO SYRINGE WHEN ASPIRATING TO CHECK BLOOD RETURN. NO PATIENT INJURY OR HARM, HOWEVER SIGNIFICANT RISK TO PATIENT (INFECTION, AIR EMBOLUS) AND INSERTER (BLOOD EXPOSURE). EVENT DID NOT INTERRUPT THE ADMINISTRATION OF ANY MEDICATIONS OR CAUSE A CLINICALLY SIGNIFICANT DELAY IN MEDICATION." IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02333 | 2023-06-07 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS 3CG, GUARDIVA, PROBE COVER, BELDAR NG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER, OCCURRENCES OF A LEAKING T-LOCK IN THE LAST WEEK. ADDITIONAL INFORMATION RECEIVED MAY 31, 2023: "LEAKING OF FLUSH THROUGH STOPPER ON T CONNECTOR AT THE WIRE EXIT AREA. LEAKING WHEN FLUSHED, AND PULLING AIR INTO SYRINGE WHEN ASPIRATING TO CHECK BLOOD RETURN. NO PATIENT INJURY OR HARM, HOWEVER SIGNIFICANT RISK TO PATIENT (INFECTION, AIR EMBOLUS) AND INSERTER (BLOOD EXPOSURE). EVENT DID NOT INTERRUPT THE ADMINISTRATION OF ANY MEDICATIONS OR CAUSE A CLINICALLY SIGNIFICANT DELAY IN MEDICATION." IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02336 | 2023-06-07 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO HF CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1395108QD |  | IT WAS REPORTED BY THE CUSTOMER THAT, 2 TRIPLE LUMEN PICC LINES WERE PLACED TODAY AND WHEN REMOVING THE STYLET IN BOTH THE LINES WE PLACED THE STYLET CURLED. ADDITIONAL INFORMATION PROVIDED 22 MAY 2023 IT WAS REPORTED BY THE CUSTOMER "PATIENTS WERE NOT COMPROMISED, AND NO INJURY OCCURRED." THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. INITIALLY THE CUSTOMER REPORTED TWO EVENTS HOWEVER ADDITIONAL INFORMATION WAS PROVIDED BY THE CUSTOMER WHICH REQUIRED A SEPARATE REPORT BE GENERATED. MEDWATCH REPORT 3006260740-2023-02336 WILL BE CANCELLED AS A DUPLICATE AND THE EVENT WILL BE SUBMITTED UNDER MEDWATCH REPORT 3006260740-2023-02905. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02341 | 2023-06-07 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM, STYLET BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 4275118 |  | CUSTOMER REPORTED ONCOLOGY PATIENT WITH POWER PICC SINCE FEBRUARY 24, TODAY ATTENDS ROUTINE MAINTENANCE, RETURN DUE TO INSERTION POINT IS EVIDENCED, DEVICE IS REMOVED, EVIDENCE OF FRACTURE. NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED ON 26.MAY.2023: IT WAS REPORTED BY THE CUSTOMER ¿NO EVIDENCE OF COMPLICATION ARE CONFIRMED, APART FROM REMOVAL AND REINSERTION DUE TO LEAKS IN THE CATHETER, WHICH DOES NOT GENERATE SERIOUS DAMAGE OR ADDITIONAL TREATMENTS TO THE PATIENT. USE FOR NEURO ONCOLOGICALS AND ANTIBIOTICOHERAPY WITH APPARENT WASHING BEFORE AND AFTER USE. IT COULD NOT BE ESTABLISHED WITH WRITTEN EVIDENCE WHETHER THEY USED THE CALMING SEQUENCE, FORCE AND TECHNIQUE FOR MAINTENANCE, OR THE TYPE OF SYRINGE (10ML DIAMETER). | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02357 | 2023-06-08 | 2022-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER (20G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202252 |  | IT WAS REPORTED VIA MEDWATCH, "ICU RN CALLED BY SUPERVISOR TO PLACE ACCUCATH FOR PATIENT IN ROOM 286-1 WITH DIFFICULT IV ACCESS. RN ARRIVED TO ROOM TO PLACE ACCUCATH. PATIENT'S RIGHT UPPER EXTREMITY WAS UTILIZED, AS NO LIMB ALERT PRESENT. ACCUCATH PLACED UNDER ULTRASOUND GUIDANCE, BRACHIAL VEIN ACCESSED, AT SUPERFICIAL LOCATION NEAR ANTECUBITAL FOSSA. GUIDEWIRE THREADED WITHOUT RESISTANCE, CATHETER ADVANCED OFF GUIDEWIRE. VENOUS BLOOD RETURN WAS PRESENT AT THIS TIME. GUIDEWIRE AND NEEDLE WERE SUBSEQUENTLY RETRACTED. RESISTANCE WAS MET TO GUIDEWIRE WITHDRAWAL. ATTEMPT TO GENTLY REMOVE CATHETER REVEALED ONCE AGAIN RESISTANCE AND THE DEVICE WAS STUCK IN THE PATIENT'S ARM. PATIENT DISCOMFORT NOTED WITH THIS ATTEMPT. THE DEVICE WAS SECURED, AND THE MD/SUPERVISOR WERE IMMEDIATELY NOTIFIED. DEVICE REMOVED BY MD WITHOUT IMMEDIATE COMPLICATION. UPON REMOVAL THE DEVICE, CATHETER, AND GUIDEWIRE WERE INTACT, HOWEVER THE GUIDEWIRE WAS CAUGHT AT THE DISTAL END OF THE CATHETER AND UPON RETRACTION CURLED THE CATHETER BACK ONTO ITSELF, KINKED AT MULTIPLE LOCATIONS, WHICH OCCURRED INSIDE THE PATIENT¿S VESSEL. GUIDEWIRE WAS REMOVED FROM CATHETER FOR CLOSER INSPECTION AND REMAINED INTACT. PRODUCT CATALOG NUMBER: AC1202252. 20GX2.25 IN ACCUCATH ACE INTRAVASCULAR CATHETER WITH 0.008 IN. NITINOL GUIDEWIRE TIP. GUIDEWIRE FOR PLACEMENT OF ACCUCATH BECAME BENT AND DIFFICULT TO REMOVE FROM PATIENT¿S VEIN PATIENT COMPLAINED OF SOME DISCOMFORT WHEN TENSION WAS MET UPON ATTEMPTED REMOVAL OF GUIDEWIRE. ICU RN NOTIFIED MHO AND PHYSICIAN IMMEDIATELY AFTER SECURING THE LINE. PHYSICIAN WAS ABLE TO REMOVE GUIDEWIRE AND THERE WAS NO PATIENT HARM." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02348 | 2023-06-08 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY 5F DUAL-LUMEN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4275118 |  | IT WAS REPORTED, "NEUROLOGICAL-ONCOLOGICAL PATIENT WITH POWER PICC SINCE 03-03-23, WHO GOES FOR EVALUATION, RECEIVES A CALL FROM THE ONCOLOGY STAFF OF THE OUTPATIENT CLINIC DUE TO EVIDENCE OF CATHETER LEAKAGE, AT EVALUATION AND LAVAGE THE LEAK IS CONFIRMED. WHEN THE DEVICE IS REMOVED WHEN CHECKING THE CATHETER, THERE IS EVIDENCE OF A LEAK IN THE BUTTERFLY BASE AND AN ADDITIONAL ONE." THERE WAS NO PATIENT HARM REPORTED. NO OTHER INFORMATION WAS PROVIDED. __ ADDITIONAL INFORMATION RECEIVED ON 26.MAY.2023: IT WAS REPORTED BY THE CUSTOMER ¿NO EVIDENCE OF COMPLICATION ARE CONFIRMED, APART FROM REMOVAL AND REINSERTION DUE TO LEAKS IN THE CATHETER, WHICH DOES NOT GENERATE SERIOUS DAMAGE OR ADDITIONAL TREATMENTS TO THE PATIENT. USE FOR NEURO ONCOLOGICALS AND ANTIBIOCOTHERAPY WITH APPARENT WASHING BEFORE AND AFTER USE. IT COULD NOT BE ESTABLISHED WITH WRITTEN EVIDENCE WHETHER THEY USED THE CALMING SEQUENCE, FORCE AND TECHNIQUE FOR MAINTENANCE, OR THE TYPE OF SYRINGE (10ML DIAMETER)." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02350 | 2023-06-08 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | CUSTOMER REPORTED A LEAKAGE DETECTED AFTER FLUSHING WITH SODIUM CHLORIDE FROM THE INSERTION POINT. TWO HOLES ON EITHER SIDE OF THE 15CM LINE, DETECTING THE HOLES AFTER THE DRAWING OF THE PICCLINE DURING USE. ADDITIONAL INFORMATION RECEIVED 24-MAY-2023: IT WAS REPORTED BY THE CUSTOMER "NO HARM OR IMPACT, THE PATIENT GOT A NEW PICC-LINE. A NEW CATHETER WAS PLACED. SITERITE8 3CG." ADDITIONAL INFORMATION RECEIVED 5/29/2023: IT WAS REPORTED BY THE CUSTOMER "LENGTH OF INSERTED CATHETER IS 42 CM." NO OTHER INFORMATION WAS PROVIDED | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02351 | 2023-06-08 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 |  | 7617405J |  | IT WAS REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, AFTER BLOOD SAMPLING, THE CATHETER WAS FLUSHED WITH 10CC OF SALINE SOLUTION. THE EXTENSION TUBE OF THE CATHETER CONNECTOR WAS DAMAGED WHILE FLUSHING THE SECOND 10CC OF SALINE SOLUTION. THE CATHETER CONNECTOR WAS REPLACED TO A NEW ONE. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DAMAGE TO EXTENSION TUBING WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE SEGMENT OF A 4 FR S/L GROSHONG CATHETER AND ITS NXT CONNECTOR. THE DAMAGE OBSERVED IN THE RETURNED SAMPLE WAS CHARACTERISTIC OF OVER-PRESSURIZATION (BURST) DAMAGE. THIS CAN OCCUR THROUGH THE USE OF SYRINGES SMALLER THAN 10ML, BY FLUSHING AGAINST AN OCCLUSION, OR DUE TO EXCESSIVE FORCE APPLIED DURING INFUSION PROCEDURES. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED IN THE EXTENSION TUBING JUST PROXIMAL OF THE MOLDED SUTURE WING. THIS CATHETER DAMAGE WAS TYPICAL OF A BURST, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ 'S' SHAPED SPLIT. ¿ TENSILE WEAKNESS AT THE FRACTURE SITE (DUE TO MATERIAL BALLOONING PRIOR TO BURST). ¿ GRANULAR FRACTURE SURFACE TEXTURE (TYPICAL OF TEARING FAILURE MODES). AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02352 | 2023-06-08 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 20G X 10 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F320108PT |  | IT WAS REPORTED BY THE CUSTOMER "I ADVANCED THE GUIDE WIRE WITH NO ISSUES. ONCE I ADVANCED THE WINGS THERE WAS A POP SENSATION AND THE WINGS ADVANCED. UPON REMOVAL OF POWERGLIDE BASE, I NOTED NO NEEDLE SAFETY SHIELD ON THE END OF NEEDLE. ONLY GUIDE WIRE AND NEEDLE EXPOSED. UPON OPENING BASE TO ENSURE PARTS WERE LEFT IN PATIENT, NOTED SAFETY AT BASE OF NEEDLE." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED 6/12/2023: IT WAS REPORTED BY THE CUSTOMER ¿THE PATIENT IS OVER AGE 50. THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION, PATIENT HARM, INJURY, COMPLICATION, OR SIGNIFICANT DELAY IN TREATMENT.¿ | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A SAFETY MECHANISM NOT ACTIVATING IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE PHOTOGRAPH AND ONE VIDEO OF A POWERGLIDE PRO WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH AND VIDEO SHOWED PIECES OF THE POWERGLIDE PRO SEPARATED. THE NEEDLE AND SAFETY MECHANISM WERE DETACHED FROM THE HOUSING, AND THE SAFETY MECHANISM WAS AT THE PROXIMAL END OF THE NEEDLE AND THE NEEDLE TIP WAS EXPOSED. THE TOP HOUSING PORTION OF THE DEVICE SHOWED THE GUIDEWIRE CONNECTED TO THE COUPLER AND THE GUIDEWIRE WAS POINTING UPWARD. THE CLEAR PORTION OF THE HOUSING, THE CATHETER, AND THE HANDLE/WINGS COULD NOT BE SEEN IN THE RETURNED PHOTOGRAPH AND VIDEO. SOME USE RESIDUE WAS OBSERVED ON THE SAMPLE IN THE RETURNED PHOTOGRAPH. NO OBVIOUS DAMAGE WAS OBSERVED. WHILE THE EXACT CAUSE OF THE OBSERVED SAFETY MECHANISM NOT ACTIVATING COULD NOT BE DETERMINED FROM THE RETURNED PHOTOGRAPH, POSSIBLE CAUSES INCLUDE DETACHMENT OF SAFETY MECHANISM FROM CARRIER DURING USE/MANIPULATION, DAMAGED COMPONENT, AND MISASSEMBLED DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04127 | 2023-06-08 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS DIFFICULT TO REMOVE AFTER 28 DAYS OF PLACEMENT. WHEN ASPIRATING STERILE WATER AND PULLING THE CATHETER OUT OF THE PATIENT TO REMOVE THE CATHETER, THE USER FELT RESISTANCE. WHEN THE CATHETER WAS FORCEFULLY REMOVED, A KNOT WAS FOUND AT THE TIP OF THE CATHETER. PER FOLLOW-UP INFORMATION RECEIVED FROM (B)(6) ON 23MAY2023, STATED THAT WHEN REMOVING THE CATHETER, SOME RESISTANCE WAS FELT BUT NOT STRONG, SO THE CATHETER WAS REMOVED WITH FORCE. IMMEDIATELY AFTER REMOVAL, THE KNOT WAS OBSERVED. THE PATIENT, A FEMALE, HAD HEMATURIA IMMEDIATELY AFTER THE CATHETER REMOVAL, BUT LATER RETURNED TO NORMAL URINE. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02363 | 2023-06-08 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D3 |  | IT WAS REPORTED BY CUSTOMER THAT STERILE GOWN IN KIT HAS STAIN ON IT. ADDITIONAL INFO RECEIVED ON 24MAY2023: IT WAS REPORTED BY THE CUSTOMER "THE STAIN WAS NOTICED UPON UNWRAPPING STERILE PACKAGING AND DONNING GOWN. THE PACKAGE WAS INTACT PRIOR TO OPENING. STERILE WRAPPING WAS INTACT. WE ARE UNABLE TO IDENTIFY THE STAIN- DOES HAVE A "SNAG" NEAR THE STAIN." IT WAS REPORTED THIS OCCURRED WITH TWO GOWNS. THIS REPORT ADDRESSES THE FIRST GOWN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A STAIN ON THE GOWN WAS CONFIRMED BUT THE EXACT CAUSE IS UNKNOWN. IMAGES OF A SURGICAL GOWN WERE PROVIDED FOR EVALUATION. THE BLUE SECTION OF THE GOWN AND BOTH WRIST CUFFS WERE SHOWN IN THE IMAGE. A BROWNISH DISCOLORATION WAS NOTED ON ONE OF THE WHITE CUFF SECTIONS. THE DISCOLORATION WAS SPREAD OVER THE TOP VISIBLE AREA SHOWN IN THE IMAGE. A FIBER, CONSISTENT WITH HAIR, APPEARED TO BE VISIBLE ON THE CUFF. THE CONDITION OF THE KIT COMPONENTS AND GOWN PRIOR TO UNBOXING AND HANDLING ARE UNKNOWN. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE EXPOSURE TO CONTAMINANTS PRIOR TO OR AFTER KIT HANDLING. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02364 | 2023-06-08 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D3 |  | IT WAS REPORTED BY CUSTOMER THAT STERILE GOWN IN KIT HAS STAIN ON IT. ADDITIONAL INFO RECEIVED ON 24MAY2023: IT WAS REPORTED BY THE CUSTOMER "THE STAIN WAS NOTICED UPON UNWRAPPING STERILE PACKAGING AND DONNING GOWN. THE PACKAGE WAS INTACT PRIOR TO OPENING. STERILE WRAPPING WAS INTACT. WE ARE UNABLE TO IDENTIFY THE STAIN- DOES HAVE A "SNAG" NEAR THE STAIN." IT WAS REPORTED THIS OCCURRED WITH TWO GOWNS. THIS REPORT ADDRESSES THE SECOND GOWN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02362 | 2023-06-08 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERLOC SAFETY INFUSION SET 20G X 0.75IN WITH Y-SITE | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0672034 |  | IT WAS REPORTED BY CUSTOMER CRACKED AND LEAKED AT THE Y WHEN FLUSHED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04126 | 2023-06-08 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE FOLEY CATHETER. BEFORE PLACING THE CATHETER, THE PHYSICIAN CONFIRMED THAT THE BALLOON WAS ABLE TO INFLATE. WHEN THE CATHETER WAS PLACED IN THE PATIENT AND WAS INJECTED WITH FIXATION WATER, IT WAS FOUND THAT THE CATHETER WAS DAMAGED AND CATHETER WAS LEAKING THE FIXATION WATER. THE CATHETER WAS IMMEDIATELY REMOVED AND A NEW ONE WAS REPLACED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02359 | 2023-06-08 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER W/SHERLOCK 3CG TIP POSITIONING SYSTEM STYLET FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | CK000375 |  | IT WAS REPORTED, CLINICIAN NOTICED PICC WAS SPLIT WHEN FLUSHING WITH SALINE. NO HARM TO PATIENT. PATIENT IMPACT IS THAT A NEW PICC LINE NEEDED TO BE INSERTED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED CATHETER WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE PHOTOGRAPH WHICH DEPICTED A 4FR S/L POWERPICC CATHETER. THE PHOTOGRAPH DEPICTED A PORTION OF CATHETER SHAFT WHICH EXHIBITED A PARTIALLY CIRCUMFERENTIAL SPLIT. THE CATHETER WAS PULLED TAUGHT AGAINST A FINGER AND WHITENING WAS VISIBLE AT THE CORNERS OF THE SPLIT. CATHETER DAMAGE WAS EVIDENT IN THE SUBMITTED PHOTOGRAPH; HOWEVER, INSPECTION OF THE PHOTOGRAPH WAS INSUFFICIENT TO IDENTIFY THE CAUSE OF THE DAMAGE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04119 | 2023-06-08 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS DIFFICULT TO REMOVE AFTER 28 DAYS OF PLACEMENT. WHEN ASPIRATING STERILE WATER AND PULLING THE CATHETER OUT OF THE PATIENT TO REMOVE THE CATHETER, THE USER FELT RESISTANCE. WHEN THE CATHETER WAS FORCEFULLY REMOVED, A KNOT WAS FOUND AT THE TIP OF THE CATHETER. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 23MAY2023, STATED THAT WHEN REMOVING THE CATHETER, SOME RESISTANCE WAS FELT BUT NOT STRONG, SO THE CATHETER WAS REMOVED WITH FORCE. IMMEDIATELY AFTER REMOVAL, THE KNOT WAS OBSERVED. THE PATIENT, A FEMALE, HAD HEMATURIA IMMEDIATELY AFTER THE CATHETER REMOVAL, BUT LATER RETURNED TO NORMAL URINE. NO MEDICAL INTERVENTION WAS REPORTED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 25MAY2023, IT WAS REPORTED THAT WHEN THE CATHETER WAS REMOVED OUTSIDE THE PATIENT, THERE HAD ALREADY BEEN THE KNOT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02360 | 2023-06-08 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT W/GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT THERE IS LEAKAGE FROM T-LOCK CONNECTION OF PICC WHEN FLUSHING. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02356 | 2023-06-08 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY CUSTOMER THAT THEY NOTICED IN THE INSIDE OF THE NEEDLE, SAFETY HYPODERMIC 25G. LOOKED LIKE A LONG DRY TISSUE LIKE OR OLD BLOODY STRING LIKE WRAPPED AROUND THE TOP PART OF THE NEEDLE AND ON THE WALLS OF THE SAFETY CAP. ADDITIONAL INFORMATION RECEIVED (B)(6)2023: IT WAS REPORTED BY THE CUSTOMER NO PATIENT INVOLVEMENT; MATERIAL WAS FOUND BY THE PICC NURSE PRIOR TO INSERTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02358 | 2023-06-08 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT PRIOR TO CATHETER PLACEMENT, THE CUSTOMER UNPACKED AND PRE-FLUSHED THE PICC CATHETER KIT AND FOUND THAT THE CATHETER LEAKED LIQUID. ANOTHER KIT WAS UNPACKAGED FOR PLACEMENT IN THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. ONE OPENED 4 FR S/L GROSHONG NXT CLEARVUE CATHETER BASIC KIT WAS RETURNED FOR EVALUATION. A MICROSCOPIC OBSERVATION REVEALED A ¿U¿ SHAPED SPLIT BETWEEN THE 21 CM AND 22 CM DEPTH MARKERS. THE SPLIT EXHIBITED SHARP FRACTURE EDGES, AND THE FRACTURE SURFACE OF THE SPLIT WAS OBSERVED TO HAVE STRIATIONS. SHARP ANGLES WERE OBSERVED AT THE SPLIT SITE. THE SPLIT WAS OBSERVED TO HAVE OCCURRED FROM THE OUTSIDE OF THE DEVICE DUE TO SMALL SECTIONS OF THE SPLIT BEING CONNECTED BY LITTLE PIECES OF CATHETER WITH MORE OF THE SPLIT FORMING FROM THE OUTSIDE OF THE DEVICE. THE CHARACTERISTICS OF THE OBSERVED SPLIT APPEAR TO BE FROM SHARP INSTRUMENT CONTACT DURING HANDLING OF THE GROSHONG CATHETER. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02361 | 2023-06-08 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | CK000883 |  | IT WAS REPORTED THAT DRESSINGS SEEM TO BE NOT ADHERING AND THE GEL SEEMS TO BE MELTING AWAY AND DISCOLORING. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02340 | 2023-06-08 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER THEY HAD A PICC THAT WAS RESISTIVE TO FLUSH AND UNABLE TO ASPIRATE. THEY ATTEMPTED CATHFLO, AFTER 2 HOURS OF BEING UNSUCCESSFUL THE MD GAVE AN ORDER TO REMOVE THE PICC AS THE PATIENT WAS ALMOST DONE WITH HER IV MEDS. CUSTOMER REMOVE THE LINE AND NOTED IT WAS CRACKED AT THE 4CM MARKING WHEN IT STARTED DRAINING FLUID AND BLOOD ON HER ARM. THE LINE DID REMOVE INTACT. PICC WAS PLACED (B)(6) 2023. ADDITIONAL INFORMATION RECEIVED 5/24/2023: IT WAS REPORTED BY THE CUSTOMER ¿EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. PICC DID NOT FUNCTION PROPERLY THEREFORE ALTERNATIVE ACCESS WAS ESTABLISHED IN THE COMMUNITY WITH A SALINE LOCK. THE CRACK WAS AT THE 4CM MARKING SO POTENTIALLY IN THE VEIN NOT SUBCUTANEOUS OR EXTERNAL, LINE HAD BEEN SECURED WITH STATLOCK AND PICC DRESSING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04102 | 2023-06-08 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50612 | 50612 |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION, NOT SURE WHAT CAUSED. USING MULTIPLE INTERMITTENT CATHETERS AFFECTED KIDNEYS AND CAUSED BACTERIAL INFECTIONS OF STOMACH AND INTESTINES. THIS CAUSED HOSPITALIZED. REPRESENTATIVE SENT REPLACEMENTS. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. . THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04117 | 2023-06-08 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER WAS UNABLE TO INJECT. | THE REPORTED EVENT WAS NOT CONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS NOT CONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04107 | 2023-06-08 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION FROM USING THE MAGIC 3 GO MALE COUDE CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04116 | 2023-06-08 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT STERILIZED WATER LEAKED DUE TO THE BALLOON RUPTURE. AFTER THE CATHETER USE ON THE PATIENT, STERILIZED WATER LEAKED FROM THE BALLOON AREA. THE BALLOON RUPTURE WAS CONFIRMED. | THE REPORTED EVENT WAS CONFIRMED ¿ UNKNOWN CAUSE. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE SAMPLE WAS EVALUATED AND VISUAL INSPECTION NOTED IRREGULAR BURST WITHOUT MISSING PIECES. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE LOW LATEX VISCOSITY, SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS: (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS: (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX. (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER IS A BALLOON CATHETER WITH TEMPERATURE- SENSING. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14 TO 18 EVERY 2FR. 1. BALLOON CATHETER WITH TEMPERATURE-SENSING. 2. [INTENDED USE & EFFECT- EFFICACY]. 3. THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. THE SURFACE OF CATHETER IS COATED FIRST WITH A TRACE AMOUNT OF METALLIC SILVER, AND THEN WITH POLYURETHANE ONTO THAT, TO INHIBIT PROLIFERATION OF BACTERIA THAT MAY ADHERE TO THE CATHETER. 4. [DIRECTIONS FOR USE] 1. METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. 3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 5) CONNECT LEAD WIRE TO MONITOR THROUGH EXTENSION CABLE. 6) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 7) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 8) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. 9) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 10) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 11) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 12) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 13) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 14) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 15) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 16) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 17) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS). 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD: WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD: 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS: (1) MALFUNCTION: - CATHETER KINKING, DAMAGE, RUPTURE, - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE, - OCCLUSION OF CATHETER INNER LUMENS, - ENCRUSTATION, - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE, - DEVICE DAMAGE DUE TO INAPPROPRIATE USE, - FAILURE TO MEASURE TEMPERATURE, AND - IMPROPER TEMPERATURE INDICATION. (2) ADVERSE EVENTS: - URINARY-TRACT INFECTION, - HEMORRHAGE, HEMATURIA, - ALLERGY REACTION TO THE DEVICE, - CALCULUS FORMATION, - EDEMA, - PAIN, - DISCOMFORT, - INJURY OF BLADDER OR URETHRAL, - URETHRITIS, URINARY INCONTINENCE, AND - RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04101 | 2023-06-08 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT WHEN STERILE WATER WAS INJECTED INTO THE FOLEY CATHETER THE BALLOON INFLATED BUT THE STERILE WATER LEAKED OUT AND THE BALLOON DEFLATED. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. BASED ON THE EVALUATION, A TEAR WAS FOUND ON THE RUBBERIZED LAYER CAUSING INTERLUMEN LEAKAGE. THE REPORTED EVENT IS CONFIRMED AS MANUFACTURING RELATED ISSUE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARDEX® SILVER LUBRICATH® FOLEY CATHETER SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02346 | 2023-06-08 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716000J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE GUIDEWIRE DID NOT ADVANCE WHEN INSERTED AFTER PUNCTURE. IT WAS FURTHER REPORTED THAT THE GUIDEWIRE WAS ALLEGEDLY STUCK INTO THE INTRODUCER NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02342 | 2023-06-08 | 2023-04-18 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT ONE MONTH AND TWENTY THREE DAYS POST A PORT PLACEMENT, THE PORT BODY ALLEGEDLY SLIGHTLY EXPOSED OUTSIDE THE SKIN. IT WAS FURTHER REPORTED THAT THE SURROUNDING TISSUES WERE RED, SWOLLEN, AND SCABBED. REPORTEDLY, THE PATIENT SKIN WAS DISINFECTED AND THE PATIENT HEALED AND HAD NO SYMPTOMS. THE PATIENT IS REPORTED TO BE STABLE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04095 | 2023-06-08 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE INFLATION VALVE AND SYRINGE CONNECTION OF THE FOLEY CATHETER DURING PRETEST AND THERE WAS NO BALLOON RUPTURE OR TEAR. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND SOLUTION IMMEDIATELY LEAKED OUT A 0.1565" TEAR FOUND UNDER INFLATION CAP. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INCORRECT MOLDING OF EITHER THE INFLATION VALVE OR RETENTION CAP. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04093 | 2023-06-08 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE FLOATING BALLOON CANNOT BE INFLATED. AS PER FOLLOW-UP INFORMATION RECEIVED VIA (B)(6) ON 10FEB2023, STATED THAT THE DEVICE WAS USED ON THE PATIENT AND THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. AS PER NOTIFICATION RECEIVED FROM INVESTIGATOR ON 17MAY2023, IT WAS STATED THAT A TEAR WAS NOTED IN THE TIP OF THE CATHETER WHERE THE BALLOON WAS ATTACHED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04094 | 2023-06-08 | 2023-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON BURST AND CATHETER FELL OUT. ON (B)(6) CATHETERIZATION WAS PERFORMED FOR THE PATIENT AND THE PROCEDURE WAS A SUCCESS. 30ML SALINE WAS INJECTED INTO THE BALLOON AND DOUBLE FIXATION WAS MADE, WITH SMOOTH DRAINAGE AND NO URINE LEAKAGE. ON (B)(6) THE CATHETER WAS FOUND FELL OUT WHEN THE PATIENT TURNED OVER. UPON CHECKING, THE BALLOON HAD A 2 CM BREAK WHILE THE DOUBLE FIXATION WAS STILL AT POSITION. THEN REPLACED A NEW CATHETER FOR THE PATIENT AND THE PROCEDURE WAS SUCCESSFUL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04100 | 2023-06-08 | 2023-05-16 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT CAUSED URINARY TRACT INFECTION FROM THE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND TO BE A DUPLICATE OF AN EVENT PREVIOUSLY REPORTED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04103 | 2023-06-08 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LUBRI-SIL® I.C. COMPLETE CARE® ALL-SILICONE FOLEY CATHETER TRAY | FOLEY CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 |  | 900114A |  | IT WAS REPORTED THAT DURING A PRETEST, WATER LEAKED FROM THE BALLOON AREA OF THE FOLEY CATHETER AND WAS DISCONTINUED USE. ACCORDING TO THE INQUIRY SHEET ATTACHED, THERE WAS NO BALLOON RUPTURE OR TEAR. THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. PACKAGING LOTS NUMBER MANUFACTURED WITH AFFECTED SYRINGE BATCHES IDENTIFIED AND DOCUMENTED. THE DHR REVIEW IS NOT REQUIRED AND THE RISK REVIEW AND LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04109 | 2023-06-08 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | ALL-SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DURING SURGERY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04105 | 2023-06-08 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L22 |  | IT WAS REPORTED THAT THERE WAS A HOLE IN THE FOLEY CATHETER BALLOON AND THE STERILE WATER DRAINED OUT. | THE REPORTED EVENT WAS CONFIRMED-MANUFACTURING RELATED. SAMPLE HAS BEEN EVALUATED AND OBSERVED PINHOLE ON SHAFT OF CATHETER. THE ROOT CAUSE OF THIS FAILURE MODE IS COULD BE MANUFACTURING RELATED DUE TO BUBBLE DURING DIPPING PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE(S) THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX SHAPE, CONFIGURATION AND PRINCIPLES: BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR. 1. BALLOON CATHETER. 2. ACCESSORIES. SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) INTENDED USE & EFFECT- EFFICACY: THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. DIRECTIONS FOR USE: 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. CORRECTION- G H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04106 | 2023-06-08 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 33618 | 33618 |  | IT WAS REPORTED THAT THE CUSTOMER REALIZED THAT THERE WAS A LEAK IN THE FOLEY CATHETER BALLOON AT THE MOMENT THEY WERE GOING TO USE THE PROBE AND RETURNED TO THE PHARMACY TO HAVE THEIR CHANGE MADE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "WATER LOST VIA PERMEATION". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04115 | 2023-06-08 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THERE WAS A LEAK FROM FOLEY CATHETER. WATER LEAKED WHEN STERILE WATER WAS INJECTED. ALSO CHECKED THE LEAK BY INJECTING FIXED WATER AND DETAILS WERE UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04112 | 2023-06-08 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED BY THE NURSE FOUND THAT FOLEY CATHETER FELL OUT FROM THE PATIENT AT 16:10 ON (B)(6) 2023 AFTER INSPECTION, THE CATHETER BALLOON WAS RUPTURED. THEN THE NURSE COMFORTED THE PATIENT AND REPLACED THE DISPOSABLE STERILE CATHETER. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC N (B)(6) 2023 CONFIRMED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04113 | 2023-06-08 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THERE WAS DIFFICULTY IN DRAINING THE CATHETER. ALSO STATES THAT ABOUT 10ML OF WATER WAS ASPIRATED BY SYRINGE AND THE BALLOON WAS DEFLATED COMPLETELY. THE USER TRIED TO DEFLATE USING SYRINGE FOR AROUND 15 SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
MW5118288 | 2023-06-08 | 2023-05-25 | I, n, j, u, r, y | Voluntary report | FOLEY CATHETER | CATHETER, UROLOGICAL | C.R. BARD, INC. | 0167L20 | 0167L20 |  | THE PATIENT HAD A THREE-WAY INDWELLING CATHETER (20 F) THAT WAS BEING MANUALLY IRRIGATED EVERY 1 TO 2 HOURS. THE ORDERS INDICATED TO IRRIGATE WITH 30-60 ML EVERY 1 TO 2 HOURS. AT 1300 THE NURSE WAS GOING TO IRRIGATE THE CATHETER WHEN SHE NOTICED THE CATHETER WAS ALMOST TOTALLY OUT OF THE URETHRA. THE NURSE CONTACTED HOSPITALIST WHO FURTHER REVIEWED THE CATHETER, WHICH WAS NOT TOTALLY INTACT, WITH A MISSING PORTION OF THE CATHETER IN AN AREA OF 3 TO 6 CM FROM THE TIP OF THE CATHETER. THE PORTION OF THE CATHETER MISSING APPEARS TO BE THE BALLOON PORTION OF CATHETER. THERE IS QUESTION OF WHETHER ANY OF THE MANUAL IRRIGATION WAS UNDERTAKEN IN THE PORT WHERE THE BALLOON IS FILLED TO 10 ML TO HELP SECURE THE CATHETER PLACEMENT. THE CATHETER WAS REPLACED WITH ANOTHER 20 F, 3-WAY CATHETER. PATIENT REQUIRED CYSTOSCOPY COMPLETED 5/26 - RETAINED FOLEY CATHETER BALLOON WAS VISUALIZED AND GRASPED USING FORCEPS AND REMOVED IN IT'S ENTIRETY. PATIENT TOLERATED THE PROCEDURE WELL. SCOPE WAS REMOVED. NO IMMEDIATE COMPLICATIONS IDENTIFIED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02372 | 2023-06-09 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS PASSED THROUGH THE OVERSLEEVE, THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE. THE OVERSLEEVE AND THE CATHETER CONNECTOR WERE CONNECTED WITH A CLICK, BUT THE CONNECTION WAS LOOSE. THERE WAS NO REPORTED PATIENT INJURY. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02369 | 2023-06-09 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2174108 |  | IT WAS REPORTED BROKEN CATHETER 16 CM FROM EXIT-SITE. DURING INFUSION, PATIENT COMPLAINS OF AXILLARY DISCOMFORT. UNUSUAL NOISE DURING FLUSHING MANEUVERS. PICC REMOVED WITH PRESENCE OF FISSURE AT 16 CM. PICC IMPLANTED ON (B)(6), FIXED WITH SECUR-A-CATH AND GLUE ON EXIT-SITE. ADDITIONAL INFORMATION RECEIVED 6/18/2023: "THE PATIENT HAD NO COMPLICATIONS, ONLY DISCOMFORT. THE PICC USED IT FOR CHEMOTHERAPY". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A SPLIT IN THE CATHETER WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. A SINGLE VIDEO WAS SUPPLIED OF THE IMPLICATED SINGLE-LUMEN POWERPICC CATHETER. THE CATHETER HAD BEEN REMOVED FROM THE PATIENT PRIOR TO TAKING THE VIDEO. WHEN THE USER ATTEMPTED TO FLUSH THE CATHETER WITH A CLEAR LIQUID, A SPRAYING LEAK WAS OBSERVED EMANATING FROM THE CATHETER SHAFT BETWEEN THE 16CM AND 17CM DEPTH MARKINGS. THERE WERE NO DISTINGUISHING FEATURES VISIBLE IN THE RETURNED VIDEO WHICH COULD AID IN IDENTIFICATION OF A SPECIFIC ROOT CAUSE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04134 | 2023-06-09 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51818 | 51818 |  | IT WAS REPORTED THAT THE PATIENT HAD A BAD URINARY TRACT INFECTION FROM THE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INDICATIONS FOR USE THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA, REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04135 | 2023-06-09 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTIONS LAST YEAR BUT WAS UNABLE TO VERIFY IF DUE TO INTERMITTENT CATHETER. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AND DISCOMFORT AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04137 | 2023-06-09 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. THE USER CONFIRMED THE CATHETER HAD NO PROBLEM BY PERFORMING A PRETEST. HOWEVER, AFTER A SURGERY WITH THE CATHETER PLACEMENT, IT FELL OFF. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02373 | 2023-06-09 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS PASSED THROUGH THE OVERSLEEVE, THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE. THE OVERSLEEVE AND THE CATHETER CONNECTOR WERE CONNECTED WITH A CLICK, BUT THE CONNECTION WAS LOOSE. THERE WAS NO REPORTED PATIENT INJURY. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04141 | 2023-06-09 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT WAS HAVING INSERTION ISSUES WHILE USING THE INTERMITTENT CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OUTSIDE DIAMETER IS TOO LARGE". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER. CAUTION: FEDERAL (USA) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02370 | 2023-06-09 | 2023-05-24 | I, n, j, u, r, y | Manufacturer report | 18G POWERGLIDE PRO RT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE MIDLINE INSERTION SITE. THE EVENT INVOLVED AN 18G POWERGLIDE PRO RT INSERTED IN THE RIGHT UPPER ARM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02368 | 2023-06-09 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PARTIAL CATHETER RUPTURE DURING USE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04130 | 2023-06-09 | 2023-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT STATED THERE WAS NO LUBRICANT ON SOME OF THE MAGIC 3 INTERMITTENT CATHETER AND THEN THERE HAS BEEN CATHETERS WITH TOO MUCH LUBRICANT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEM OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02367 | 2023-06-09 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602830 | 0602830 |  | IT WAS REPORTED THAT APPROXIMATELY FOUR YEARS POST PORT PLACEMENT, THE RUPTURE OF THE CATHETER WAS ALLEGEDLY FOUND AT ABOUT FOUR CENTIMETERS FROM THE RESERVOIR. REPORTEDLY, THE DEVICE WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE BARDPORT MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER IN TWO SEGMENTS WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED ON THE ATTACHED CATHETER APPROXIMATELY 1.1CM FROM THE DISTAL END OF THE CATH-LOCK AND A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED AT THE DISTAL END OF THE ATTACHED CATHETER. ALSO A PARTIAL CIRCUMFERENTIAL BREAK WAS ON NOTED APPROXIMATELY 0.7CM FROM THE PROXIMAL END OF THE DISTAL CATHETER SEGMENT. THE SURFACE WAS NOTED TO BE ROUGH AND SMOOTH ON BOTH REGIONS. THE EDGES OF THE COMPLETE CIRCUMFERENTIAL BREAK ON THE DISTAL END OF THE ATTACHED CATHETER WERE NOTED TO BE UNEVEN. THE SURFACE WAS NOTED TO BE GRANULAR IN ONE REGION AND ROUND AND SMOOTH ON THE OTHER REGION. A SMALL LONGITUDINAL SPLIT WAS NOTED ON THE BORDER OF BOTH REGIONS. UPON INFUSION, A LEAK WAS OBSERVED EXITING FROM THE PARTIAL CIRCUMFERENTIAL BREAK ON THE ATTACHED CATHETER WHILE WATER EXITED THE DISTAL END. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, ONE ELECTRONIC PHOTO AND THREE ELECTRONIC IMAGES WERE PROVIDED FOR REVIEW. THE PHOTO SHOWS ONE PORT ATTACHED TO A GROSHONG CATHETER. THE IMAGES DEMONSTRATE THE PORT CATHETER IN THE RIGHT ATRIUM AND SUPERIOR VENA CAVA AND THE CATHETER FRACTURED AT THE POINT THAT IT ENTERS THE INTERNAL JUGULAR VEIN. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND IDENTIFIED MATERIAL SEPARATION, WEAR AND DEFORMATION ISSUES. THE IDENTIFIED BREAK IS TYPICAL OF FLEXURAL FATIGUE, WHICH IS DUE TO THE REPETITIVE, CYCLIC KINKING OF THE CATHETER. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: B5, D4 (EXPIRATION DATE: 04/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02376 | 2023-06-09 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 4F SL CATHETER BASIC TRAY (W/ 70CM NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6194355 |  | IT WAS REPORTED PICC INSERTED INTO LEFT BASILIC VEIN ON MONDAY (B)(6) 2023. CHEMO THERAPY COMMENCED. PICC HAS BEEN RAPIDLY BLEEDING SINCE INSERTION, FROM THE INSERTION SITE. DRESSING CHANGED X4 ON MONDAY, PATIENT NOT ON ANY MEDICATION THAT MAY CAUSE INCREASED BLEEDING. CHEMO ATTACHED VIA INFUSOR FOR 46/24 INFUSION AT HOME. PATIENT RETURNED WEDNESDAY FOR DISCONNECT & REPLACED DRESSING AS HEAMOSEROUS OOZE, PLACED STERILE GAUZE AROUND SITE AND AFTER 2 DAYS, WAS SOAKED THROUGH. REDRESSED SITE AND SENT PATIENT HOME WITH NO ADDITIONAL CONNECTION OTHER THAN END CAP/VALVE. PATIENT CONTACTED WARD YESTERDAY DUE TO "DAMP ARM, NOT BLOOD' CAME BACK LAST NIGHT FOR REVIEW, NURSES NOTED A WHITE FLUID, COMING FROM INSERTION SITE, REDRESSED, AND CHECKED ALL CONNECTIONS AND RE-FLUSHED, . REPORTED THAT WHEN FLUSHING IT APPEARED TO "SPURT BACK FROM THE INSERTION SITE" FLUSHING AGAIN WITH GAUZE AT SITE WHICH BECAME WET UPON FLUSHING. ONCOLOGIST CONTACTED AND PICC REMOVED AS PER MEDICAL INSTRUCTION. FURTHER INVESTIGATION BY NURSING STAFF ONCE REMOVED BY FLUSHING. APPEARED NO LEAKS IN TUBING. PATIENT UNABLE TO RECEIVE TREATMENT THROUGH DEVICE, AS REMOVED. A SECONDARY DEVICE(PORT) WILL NEED TO BE INSERTED AS AN ALTERNATIVE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04136 | 2023-06-09 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MAGIC3 GO MALE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE PATIENT HAD A BLADDER INFECTION AND URINARY TRACT INFECTION. PATIENT WAS HOSPITALIZED AND ON ANTIBIOTICS (LOT#JUGU1957 AND LOT#JUGT1013). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04131 | 2023-06-09 | 2023-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT STATED THERE WAS NO LUBRICANT ON SOME OF THE MAGIC 3 INTERMITTENT CATHETER AND THEN THERE HAS BEEN CATHETERS WITH TOO MUCH LUBRICANT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04145 | 2023-06-09 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123514A | 123514A |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAD BURST. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON BURST. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE PATIENT USE ONLY. DO NOT REUSE AND RESTERILIZE. FOR UROLOGICAL USE ONLY. USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PROTOCOL, THE MAY BE CUT OFF. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." CORRECTION: E. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04166 | 2023-06-09 | 2023-05-18 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTIONS. THE PATIENT STATED THAT THEY HAD 3 URINARY TRACT INFECTIONS AND THEY SAW UROLOGIST. IT WAS UNKNOWN IF THE INTERMITTENT CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. PER ADDITIONAL INFORMATION RECEIVED ON 31MAY2023, IT WAS REPORTED THAT CUSTOMER DID NOT RECEIVE CONFIRMATION THAT CATHETERS CAUSED URINARY TRACT INFECTIONS. THE CUSTOMER STATED THEY ONLY USED 8 WICKS THEN WAS ADVISED TO STOP USE BY THEIR PROVIDER. IT WAS UNKNOWN IF THE INTERMITTENT CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04179 | 2023-06-09 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 129416M | 129416M |  | IT WAS REPORTED, THAT UPON ATTEMPTING TO FLUSH FOLEY CATHETER. BLUE HUB BROKE AND DETACHED. NO PATIENT HARM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE, A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04181 | 2023-06-09 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED, THAT THE FOLEY CATHETER CAME OUT OF THE PATIENT FULLY INTACT. THEY WERE UNABLE TO FIND AN ISSUE WITH THE FOLEY. IT JUST CAME OUT. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED, THAT THE FOLEY WAS REPLACED IN THE DOCTOR'S OFFICE. THE PATIENT HAD NO INJURY. AND WAS UNSURE IF THE DOCTOR'S OFFICE HELD ON TO THE FOLEY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE, A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04183 | 2023-06-09 | 2023-05-19 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD 2 URINARY TRACT INFECTIONS IN THE PAST MONTH. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04170 | 2023-06-09 | 2023-05-18 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED, THAT THE INTERMITTENT CATHETER (LOT#: JUGT1013, LOT#: JUGU1957) WAS CAUSING BLADDER INFECTION. AND URINARY TRACT INFECTION. MEDICAL INTERVENTION WAS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE, A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04167 | 2023-06-09 | 2023-05-19 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50812G | 50812G |  | IT WAS REPORTED THAT CAREGIVER DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION AND THEY THOUGHT INTERMITTENT CATHETER CAUSED URINARY TRACT INFECTION FOR PATIENT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04159 | 2023-06-09 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | ALL-SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER WAS UNABLE TO INJECT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04161 | 2023-06-09 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT ON INJECTING STERILE WATER INTO THE FOLEY CATHETER DURING A PRETEST, THE STERILE WATER LEAKED OUT OF THE TRIFURCATION. ANOTHER FOLEY CATHETER WAS REPLACED AND INSERTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04154 | 2023-06-09 | 2023-05-19 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT PATIENT HAD DRASTIC INCREASE IN URINARY TRACT INFECTIONS AND BELIEVES IT WAS RELATED TO THE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTIONS. PER FOLLOW UP VIA EMAIL ON 24MAY2023, IT WAS REPORTED THAT THE SAMPLE IS AVAILABLE. A RETURN KIT WILL BE SENT TO COLLECT IT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04156 | 2023-06-09 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PLACED TO THE PATIENT WITHOUT A PRETEST, AND STERILE WATER WAS DIFFICULT TO REMOVE DURING THE CATHETER REMOVAL. THE DETAILS HAS TO BE CONFIRMED AT THE FACILITY SITE AT A LATER DATE. THERE WAS NO IMPACT TO THE PATIENT. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) USING IN-HOUSE LUER LOCK SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 3 MINUTES WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. A LABELING REVIEW WAS NOT REQUIRED AS THE REPORTED EVENT IS UNCONFIRMED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04155 | 2023-06-09 | 2023-05-19 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50812G | 50812G |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION FROM USING THE MAGIC 3 GO MALE COUDE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1 TO 1.5 OR 2.5 TO 3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04152 | 2023-06-09 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MAGIC3 GO MALE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT CUSTOMER STATED THAT THEY HAD A BAD URINARY TRACT INFECTION AND THEY DID NOT WANT TO USE THEM. THIS WAS THE FIRST TIME THE PATIENT HAD USED THIS PRODUCT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04160 | 2023-06-09 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | ALL-SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER WAS UNABLE TO REMOVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04146 | 2023-06-09 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175814 | 175814 |  | IT WAS REPORTED THAT FOLEY DISLODGED AFTER INSERTION IN PATIENT. CATHETER TIP SEEMED TO BE DEFECTIVE. WHEN TRIED TO INFLATE THE BALLOON, THE SALINE WAS COMING OUT FROM THE TIP OF THE CATHETER AND THE BALLOON DOES NOT INFLATE MUCH. PER FOLLOW UP VIA EMAIL RECEIVED ON (B)(6) 2023, IT WAS REPORTED THE UROLOGIST CUT THE TIP OF THE CATHETER FOR USE AND THAT THE PATIENT WAS STABLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02379 | 2023-06-09 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED "THEY HAVE OPEN A NEW CATHETER JUST FOR A TEST WITH THE SAME LOT AND THE TIP WENT OFF. THIS IS NOT USED WITH PATIENT, JUST TESTED IN THEIR OFFICE." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BROKEN STYLET WAS CONFIRMED. THE PRODUCTS RETURNED FOR EVALUATION WERE ONE 3CG STYLET WITHIN THE RED LUMEN OF A 4 FR SINGLE LUMEN POWERPICC SOLO CATHETER, ONE REMOTE CONTROL HOLDER, TWO BLUE ELASTIC BANDS, TWO PIECES OF GAUZE, AND ONE SEALED 4 FR SINGLE LUMEN POWERPICC SOLO CATHETER WITH 3CG KIT IN A SEALED BOX. THE RETURNED PRODUCT SAMPLES WERE EVALUATED AND ONE OF THE STYLETS WAS CONFIRMED TO BE BROKEN. THE FOLLOWING OBSERVATIONS WERE MADE WHICH WERE CONSISTENT WITH THIS FAILURE TYPE: ¿ MICROSCOPIC EXAMINATION REVEALED DEFORMATION OF THE STYLET TUBING AT MAGNET EDGE(S). ¿ MICROSCOPIC EXAMINATION REVEALED LOCALIZED DELAMINATION OF THE STYLET TUBING AT THE DAMAGE SITE(S), INDICATIVE OF TUBING KINKING/FOLDING. ¿ SOME KINKING OF THE SURROUNDING STYLET TUBING, LOCATED AT MAGNET INTERFACE(S). MEASUREMENTS OF THE STYLET TUBING OF EACH SAMPLE WERE TAKEN AND BOTH WERE CONFIRMED TO BE WITHIN SPECIFICATIONS. ALTHOUGH THE BREAK IN THE WIRE COULD BE CONFIRMED, THE ROOT CAUSE COULD NOT BE DETERMINED THROUGH EXAMINATION OF THE RETURNED SAMPLES. THE OBSERVED FEATURES WERE INDICATIVE OF CIRCUMSTANCES WHICH INCLUDED STYLET KINKING, LIKELY DURING THE INSERTION PROCESS; HOWEVER, IT APPEARED THAT ADDITIONAL UNIDENTIFIED FACTORS MAY HAVE ALSO CONTRIBUTED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02365 | 2023-06-09 | 2023-05-09 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT THREE YEARS TWO MONTHS AND TWENTY-TWO DAYS POST PORT PLACEMENT VIA THE RIGHT SUBCLAVIAN, CONTRAST MEDIUM ALLEGEDLY LEAKED NEAR THE INSERTION SITE. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY BROKEN. REPORTEDLY, THE DISTAL CATHETER SEGMENT MIGRATED TO THE PULMONARY ARTERY. THE DISTAL CATHETER SEGMENT AND THE PORT BODY WERE REMOVED PERCUTANEOUSLY. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 10/2022). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02377 | 2023-06-09 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 15CM (SHORT-TERM) (DUAL-LUMEN) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5593150 |  | IT WAS REPORTED THE CANNULA BROKE OFF WITHOUT MUCH FORCE. 02 JUNE 2023 ADDITIONAL INFORMATION FROM THE CUSTOMER: EVENT HAPPENED ON 22.5. 23 DURING USE. PATIENT HAD TO BE PUNCTURED SEVERAL TIMES AND CAUSED A DELAY IN THERAPY. 1 NEEDLE BROKEN DURING PUNCTURE, 1 NEEDLE BROKEN DURING PREPARATION. THIS FILE ADDRESSES THE ISSUE WITH THE NEEDLE BROKEN DURING PREPARATION. | THE INITIAL COMPLAINT APPEARED TO BE A REPORTABLE OCCURRENCE. ONCE THE SAMPLE WAS RETURNED AND EVALUATED IT WAS DETERMINED THAT A REPORTABLE EVENT HAD NOT OCCURRED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04142 | 2023-06-09 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRETEST, IT WAS DIFFICULT TO INJECT WATER INTO THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02375 | 2023-06-09 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7274118 |  | CUSTOMER REPORTED THAT A MINOR MALE PATIENT, HOSPITALIZED IN THE PEDIATRIC INTENSIVE CARE UNIT, IS A CARRIER OF A 4FR BILUMEN CATHETER INSERTED ON (B)(6) 2023 IN THE LEFT ARM. THE NURSE IN CHARGE OF THE PATIENT REPORTED DYSFUNCTION OF THE DEVICE (LEAKAGE OF LIQUID AT THE POINT OF INSERTION). THE CATHETER WAS REMOVED PER RECOMMENDATIONS FROM THE PICC GROUP STAFF. WHEN EVALUATING THE CATHETER, THE DEVICE IS OBSERVED CUT BY 19 CM WHEN IRRIGATING THROUGH BOTH LUMENS, FLUID LEAKAGE (RUPTURE) IS OBSERVED AT A LEVEL 1 CM AFTER ZERO. DESCRIPTION OF THE IMPACT ON THE PATIENT FAILURE TO COMPLETE TREATMENT, REQUIRES PLACING ANOTHER CATHETER TO CONTINUE TREATMENT. NO SURGICAL INTERVENTION WAS REQUIRED. ADDITIONAL INFORMATION RECEIVED ON 31.MAY.2023: CUSTOMER CONFIRMED A NEW CATHETER WAS REQUIRED AND REPORTS THE PATIENT IS IN "ADEQUATE" CONDITION. IT WAS STATED THERE WAS NO SURGICAL INTERVENTION REQUIRED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02371 | 2023-06-09 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS PASSED THROUGH THE OVERSLEEVE, THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE. THE OVERSLEEVE AND THE CATHETER CONNECTOR WERE CONNECTED WITH A CLICK, BUT THE CONNECTION WAS LOOSE. THERE WAS NO REPORTED PATIENT INJURY. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02378 | 2023-06-09 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 15CM (SHORT-TERM) (DUAL-LUMEN) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5593150 |  | IT WAS REPORTED THE CANNULA BROKE OFF WITHOUT MUCH FORCE. 02 JUNE 2023 ADDITIONAL INFORMATION FROM THE CUSTOMER: EVENT HAPPENED ON (B)(6) 2023 DURING USE. PATIENT HAD TO BE PUNCTURED SEVERAL TIMES AND CAUSED A DELAY IN THERAPY. 1 NEEDLE BROKEN DURING PUNCTURE, 1 NEEDLE BROKEN DURING PREPARATION. THIS FILE ADDRESSES THE ISSUE WITH THE NEEDLE BROKEN DURING PREPARATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04150 | 2023-06-09 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0170SI12 | 0170SI12 |  | IT WAS REPORTED THAT THE CUSTOMER HAD TWO PRODUCTS THAT WERE NOT WORKING, AND THE PRODUCTS WERE BOTH FROM THE SAME LOT. THEY BOTH HAD PARTS OF THE TUBE THAT SEEMED TO BE PINCHED OR BROKEN WHEN THEY WERE REMOVED FROM THE PACKAGE, MAKING THEM UNABLE TO USE THEM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04151 | 2023-06-09 | 2023-05-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51808 | 51808 |  | IT WAS REPORTED THAT CUSTOMER EXPERIENCED URINARY TRACT INFECTION FROM MAGIC 3 GO. INTERMITTENT SILICON CATHETER DATED (B)(6)2023, WHICH CAUSED KLEBSIELLA AMMONIA. DISTRIBUTER HAVE GIVEN NEW SUPPLIES OF CATHETERS BUT CAUSED HOSPITAL STAY OF 6 DAYS AND NOW CUSTOMER WAS ON OXYGEN AND HEALTH HAS FAILED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04140 | 2023-06-09 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0167SI24 | 0167SI24 |  | IT WAS REPORTED THAT CUSTOMER HAD GOTTEN COMPLAINTS FROM A COUPLE OF THE UROLOGY DOCTORS ABOUT FOLEY CATHETER. CUSTOMER HAD FOLEY CATHETER WITH ISSUES OF THEIR BALLOONS LEAKING. ONE SURGEON FILLED THEIR BALLOON WITH 30 CC¿S, PATIENT WAS TRANSFERRED TO THE FLOOR AND IT WAS FOUND THAT THE CATHETER HAD MIGRATED TO THE PATIENT¿S PROSTATE . ANOTHER LAST NIGHT FILLED THEIR BALLOON IN THE OPERATION ROOM AND BY THE TIME THEY GOT TO PACU, IT CAME OUT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "WATER LOST VIA PERMEATION". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE BARDEX® I.C. LATEX FOLEY CATHETER IS INDICATED FOR USE IN THE DRAINAGE AND/OR COLLECTION AND/OR MEASUREMENT OF URINE. GENERALLY, DRAINAGE IS ACCOMPLISHED BY INSERTING THE CATHETER THROUGH THE URETHRA AND INTO THE BLADDER. HOWEVER, DRAINAGE IS SOMETIMES ACCOMPLISHED BY SUPRAPUBIC OR OTHER PLACEMENT OF THE CATHETER, SUCH AS NEPHROSTOMY TRACT. DIRECTIONS FOR USE: 1. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. 2. USING PROPER ASEPTIC METHODS, REMOVE CATHETER FROM PACKAGE. 3. PREPARE PATIENT PER HOSPITAL/NURSING RECOMMENDED PROCEDURE. 4. PROCEED WITH CATHETERIZATION USING STANDARD TECHNIQUES. 5. INFLATE THE BALLOON WITH STERILE WATER TO THE VOLUME STATED ON THE INFLATION LUMEN OF THE CATHETER BY USING A WATER-FILLED LUER TIP SYRINGE. 6. CONNECT CATHETER TO COLLECTION CONTAINER. 7. TO DEFLATE THE BALLOON, GENTLY REINSERT THE LUER TIP SYRINGE INTO THE VALVE AND ASPIRATE. CAUTION: TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE IN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. USE NEEDLE. DO NOT EXCEED RECOMMENDED CAPACITIES AS STATED ON THE INFLATION LUMEN OF THE CATHETER. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. CAUTION: DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON TO BURST. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. DO NOT USE IF PACKAGE IS DAMAGED. RECOMMENDED INDWELLING TIME NOT TO EXCEED 28 DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04149 | 2023-06-09 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT CUSTOMER STATED PATIENT RECEIVED 4 INTERMITTENT CATHETERS WITH NO LUBE/STERILE WATER PACKET. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02384 | 2023-06-10 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC INTRODUCER KIT | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | UNK |  | CUSTOMER REPORTED THAT THE PATIENT IMPLANTED THE CATHETER SMOOTHLY, BUT IT WAS FOUND THAT THE CATHETER HAD GREAT RESISTANCE AND COULD NOT BE USED NORMALLY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02383 | 2023-06-10 | 2023-05-22 | I, n, j, u, r, y | Manufacturer report | BARD POWERGLIDE PRO 18G | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE MIDLINE INSERTION SITE. MIDLINE INSERTED IN UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02387 | 2023-06-10 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1274108D4 |  | IT WAS REPORTED BY THE CUSTOMER THAT T-PORT FROM A PROVENA PICC LINE HAD A LEAK WHEN ASPIRATING FLUID THROUGH THE T-PORT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02385 | 2023-06-10 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT DURING PUNCTURE FOR THE GROSHONG CATHETER, THE NURSE FELT RESISTANCE AT THE PROXIMAL END OF THE GUIDE WIRE WHEN REMOVING THE GUIDE WIRE, MAKING THE REMOVAL IMPOSSIBLE. THEREFORE, THE GUIDE WIRE WAS REMOVED TOGETHER WITH THE MICRO-INTRODUCER. A CAREFUL EXAMINATION OF THE GUIDE WIRE REVEALED THAT ITS PROXIMAL END WAS COARSE AND PROTRUDING, WHICH LED TO THE IMPOSSIBILITY OF THE REMOVAL. A NEW MST WAS THEREFORE USED TO CONDUCT RE-PUNCTURE. NO PATIENT INJURY REPORTED, NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED PROXIMAL END OF A GUIDEWIRE IS INCONCLUSIVE DUE TO THE STATE OF THE RETURNED SAMPLE. ONE GUIDEWIRE WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLE. THE PROXIMAL WELD TIP OF THE GUIDEWIRE WAS OBSERVED TO BE MISSING AND WAS NOT RETURNED FOR EVALUATION. A MICROSCOPIC OBSERVATION REVEALED THE CORE WIRE OF THE GUIDEWIRE TO HAVE BEEN CUT AT THE PROXIMAL END OF THE RETURNED GUIDEWIRE. BECAUSE THE PROXIMAL END OF THE GUIDEWIRE WAS CUT OFF AND NOT RETURNED FOR EVALUATION, THIS COMPLAINT IS INCONCLUSIVE AT THIS TIME. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02390 | 2023-06-10 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405J |  | CUSTOMER REPORTED THAT AFTER THE CATHETER WAS PASSED THROUGH THE OVERSLEEVE, THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02391 | 2023-06-10 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405J |  | CUSTOMER REPORTED THAT AFTER THE CATHETER WAS PASSED THROUGH THE OVERSLEEVE, THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02392 | 2023-06-10 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | CUSTOMER REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, BECAUSE INFUSION WAS DIFFICULT, WHEN THE DRESSING WAS REMOVED AND THE DEVICE WAS CHECKED, THE CATHETER PART WHICH WAS OUTSIDE OF THE PATIENT¿S BODY WAS FOUND TO BE BROKEN. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02386 | 2023-06-10 | 2023-05-14 | I, n, j, u, r, y | Manufacturer report | DOT PROV 4F DL MAX TLS | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S9274108D |  | THE CUSTOMER REPORTED A GUIDEWIRE TIP BREAK OFF DURING A PICC INSERTION WITH ABOUT 1.5 CM OF IT REMAINING IN THE PATIENT. THE CLINICIAN WHO PLACED THE LINE STATED THAT IT WAS A SMOOTH INSERTION, THEY MET RESISTANCE ONCE PULLED BACK AND WERE ABLE TO THREAD THE PICC WITHOUT ANY FURTHER RESISTANCE. THEY DID NOT DISCOVER THE WIRE TIP HAD BROKEN OFF UNTIL DAYS LATER AND THE WIRE WAS ALREADY DISCHARGED. WIRE MADE ITS WAY TO PATIENT'S LUNG. THEY WENT IN YESTERDAY TO RETRIEVE THE WIRE AND WERE UNSUCCESSFUL. DURING THE RETRIEVAL PROCESS THE PATIENT CODED. AT THIS POINT THEY ARE STILL DETERMINING THE BEST COURSE OF ACTION FOR THE PATIENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04187 | 2023-06-11 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | URETHRAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE INTERMITTENT KITS WERE NOT LUBRICATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04186 | 2023-06-11 | 2023-01-31 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT A HOMECARE PATIENT WITH SUPRAPUBIC CATHETER WHEN COMPLETED FLUSHING, THE CATHETER BROKE AND THE BALLOON DID NOT DEFLATE. THE PATIENT HAD TO VISIT ER TO HAVE CENTRAL LINE WIRE TO BREAK BALLOON THROUGH PORT AND REMOVE NUB AND THEN WAS CHANGED A NEW CATHETER. THEY HAVE ALSO HAD THE EXPERIENCE WITH THE CATHETER CRACKING AGAIN IN (B)(6) 2023 WHEN OTHER CATHETER CHANGES COMPLETED AS WELL. FIRST DEVICE WAS A 16F AND THE SECOND DEVICE WAS 18F. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. EVALUATION WAS OBSERVED THE CATHETER WAS CUT ALONG THE DRAINAGE LUMEN AND THERE WAS SALT ACCUMULATION INSIDE THE LUMEN. IT WAS ALSO PRESENCE OF SMOOTH CUT AT THE MIDDLE OF CATHETER SHAFT. THE SURFACE WAS LOOKS LIKE HAS BEEN CUT WITH SHARP OBJECT. THE EXACT CAUSE ON HOW AND WHEN PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO LOWER LATEX/OVER CHLORINATION/USER-RELATED (PULLING BY A USER/ EXPOSE TO SHARP OBJECT). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02381 | 2023-06-12 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | CUSTOMER REPORTED THAT THE INSERTION OF THE PICC CATHETER INTO THIS PATIENT WAS NOT SMOOTH. THE CATHETER WAS EVENTUALLY REMOVED. IT WAS FOUND THAT THE CATHETER LEAKED LIQUID AND THE GUIDE WIRE INSIDE WAS BENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04180 | 2023-06-12 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 52608G | 52608G |  | IT WAS REPORTED THAT THE PATIENT WAS HAVING TROUBLE WITH INSERTION WITH THESE MAGIC 3 INTERMITTENT CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "CATHETER IS NOT STIFF ENOUGH OR CATHETER TOO STIFF FOR INSERTION, ROUGH OR IRREGULAR SHAPE PROTRUSIONS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON'T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION. SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04182 | 2023-06-12 | 2023-05-18 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION. IT IS UNKNOWN IF THE DEVICE CAUSED URINARY TRACT INFECTION AT THIS TIME AND THE MEDICAL INTERVENTION IS UNKNOWN PER CUSTOMER VIA PHONE ON (B)(6)2023 IT WAS REPORTED THAT THE CUSTOMER DEVELOPED UTI AND SOUGHT MEDICAL INTERVENTION. CUSTOMER WAS PRESCRIBED ANTIBIOTICS AND LIBERATOR REPLACED CATHETERS WHICH HAVE CAUSED NO ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04250 | 2023-06-12 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LUBRI-SIL FOLEY CATHETERS | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PORT OF THE FOLEY CATHETER JUNCTION SNAPPED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00409 | 2023-06-12 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CQ7594J |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE FIBER OF THE PTA BALLOON ALLEGEDLY GOT STUCK WITH THE STENT UPON STENT DEPLOYMENT. IT WAS FURTHER REPORTED THAT THE PTA BALLOON WAS DEFLATED TO REMOVE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CONQUEST PTA DILATATION CATHETER PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE CONQUEST PTA DILATATION CATHETER PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02395 | 2023-06-12 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT ONE YEAR THREE MONTHS AND THIRTEEN DAYS POST PORT PLACEMENT, THE CATHETER ALLEGEDLY HAD A SMALL BREAK UPON REMOVAL. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 02/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04173 | 2023-06-12 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT AFTER INSERTING THE FOLEY CATHETER INTO THE PATIENT WITHOUT ANY PROBLEMS DURING THE PRETEST, THERE WAS LITTLE URINE FLOW, SO THE BALLOON WAS CHECKED AND FOUND TO BE LEAKING STERILE WATER, AND THEN STERILE WATER WAS INJECTED INTO THE BALLOON AND THERE WAS A PINHOLE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, IT WAS STATED THAT THE URINE FLOW RATE WAS PROBABLY LOW BECAUSE THE BALLOON HAD DEFLATED AND THE CATHETER WAS BEING DISLODGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04244 | 2023-06-12 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0167SI24 | 0167SI24 |  | IT WAS REPORTED THAT THE FOLEY CATHETER STOPPED DRAINING. THE NURSE WAS GOING TO REPLACE, SO THEY ASPIRATED ON THE BALLOON PORT AND DID NOT GETTING ANY VOLUME. THE BALLOON PORT WAS NOTED TO BE DAMAGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00410 | 2023-06-12 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER IQ ULTRASONIC | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | XCTO146 | XCTO146 |  | IT WAS REPORTED THAT PRIOR TO A RECANALIZATION PROCEDURE, THE DEVICE ALLEGEDLY HAD A CONNECTION PROBLEM. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE CROSSER IQ CATHETER WAS RETURNED FOR EVALUATION. SAMPLE APPEARED TO HAVE RESIDUE THROUGHOUT. UPON FUNCTIONAL EVALUATION, THE POWER CABLE WAS THEN CONNECTED TO THE BD IQ RECANALIZATION SYSTEM AND SCREEN READ CANNOT IDENTIFY DEVICE. THE DEVICE WAS NOTED TO HAVE A NICK IN THE RIBBON. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED CONNECTION ISSUE AND THE IDENTIFIED MATERIAL CUT ISSUE. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED CONNECTION ISSUE AND THE IDENTIFIED MATERIAL CUT ISSUE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 07/2023), G3. H11: B5, H6 (PATIENT). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04188 | 2023-06-12 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER COULD NOT BE INJECTED. ACCORDING TO THE INQUIRY SHEET ATTACHED, THE PRETEST WAS NOT DONE AND THERE WAS NO PROBLEM OBSERVED IN THE CATHETER DURING INFLATION. THERE WAS NO ABNORMALITY OBSERVED ON THE SHAFT OF THE CATHETER AND THE WATER WAS INJECTED GENTLY. USUALLY, BALLOON CAN BE INFLATED WITHIN AROUND 10 SECONDS. AS PER THE ADDITIONAL INFORMATION RECEIVED ON 26 MAY 2023, STATED THAT THE ISSUE WAS THE SLOW RETURN OF STERILE WATER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04222 | 2023-06-12 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L20 |  | IT WAS REPORTED THAT IMMEDIATELY AFTER INSERTION TO THE PATIENT IN THE OPERATING ROOM, URINE BEGAN LEAKING FROM THE THREE-WAY BIFURCATION OF THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. BASED ON THE EVALUATION, IT WAS OBSERVED INTERLUMEN LEAKAGE. DISSECTED THE CATHETER AND OBSERVED TEAR ON RUBBERIZE LAYER. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE TO TORN RUBBERIZE ON RUBBERIZE LAYER THAT MAY LEAD TO LUMEN-TO-LUMEN LEAK. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04228 | 2023-06-12 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER LEAKED FROM THE INFLATION VALVE AREA OF THE FOLEY CATHETER. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE VALVE MENTIONED AS BACKFLOW. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED A 3-WAY TEMP SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. ATTEMPTED TO INFLATE BALLOON WITH 10 ML OF SOLUTION AND IT IMMEDIATELY LEAKED OUT OF 0.030" PINHOLE FOUND ON THE INFLATION ARM AND JUST UNDER THE INFLATION CAP. ALSO RECEIVED 5 PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWS OVERVIEW OF CATHETER WITH USED SYRINGE. SECOND AND THIRD PHOTO SAMPLES SHOW THE LOCATION OF THE PINHOLE PRESENT WITH THE PINHOLE BEING VISIBLE IN THE THIRD PHOTO SAMPLE. FOURTH PHOTO SAMPLE SHOWS TIP OF SYRINGE. FIFTH PHOTO SAMPLE SHOWS END OF SYRINGE INDICATING THE NUMBER 30. BASED ON PHYSICAL SAMPLE RECEIVED PRODUCT DOES NOT MEET SPECIFICATIONS WHICH STATES "CAP AND VALVE MUST BE PRESENT AND ASSEMBLED THE TIP OF THE BLACK MATERIAL OF THE VALVE IS VISIBLE ON THE SURFACE OF THE CAP [2] WITHOUT CUTS, HOLES OR TEARS ON THE INFLATION ARM." ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE VALVE DOES NOT FULLY CLOSE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2). FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04219 | 2023-06-12 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE USER ATTEMPTED TO REPLACE FOLEY CATHETER DUE TO BALLOON ABNORMALITY AND URINE LEAKAGE, BUT THE CATHETER WAS CUT OFF BECAUSE WATER COULD NOT BE DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04217 | 2023-06-12 | 2023-05-22 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD SEVERAL URINARY TRACT INFECTIONS THAT HAD PUT THEM IN THE HOSPITAL FOR THREE DAYS. BUT THE PATIENT WAS CURRENTLY ON MEDICATION FOR A URINARY TRACT INFECTIONS. IT WAS STATED THAT THE PATIENT WAS GETTING THEM SO OFTEN THAT WAS PUT ON A PREVENTATIVE. IT WAS UNKNOWN IF THE DEVICE CAUSED THE INFECTION. THE PATIENT REPORTED ISSUES WITH LOT# JUGU0783, LOT# JUGV0317. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04227 | 2023-06-12 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE DRAINAGE FUNNEL LUMEN OF THE FOLEY CATHETER WAS JAGGED (OR CRACKED). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04216 | 2023-06-12 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT 3DAYS AFTER CATHETER PLACEMENT TO THE PATIENT, THE CATHETER FELL OUT OF THE PATIENT AND THE BALLOON WAS FOUND BURST. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, IT WAS STATED THAT THE THERE WERE NO MISSING PIECES OF THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04221 | 2023-06-12 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0167SI24 | 0167SI24 |  | IT WAS REPORTED THAT THE FOLEY CATHETER STOPPED DRAINING. THE NURSE WAS GOING TO REPLACE, SO THEY ASPIRATED ON THE BALLOON PORT AND DID NOT GETTING ANY VOLUME. THE BALLOON PORT WAS NOTED TO BE DAMAGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04220 | 2023-06-12 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0167SI24 | 0167SI24 |  | IT WAS REPORTED THAT THE CUSTOMER WERE UNABLE TO DEFLATE THE FOLEY BALLOON TO REMOVE DUE TO INABILITY IRRIGATE/DRAIN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04214 | 2023-06-12 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI12 |  | IT WAS REPORTED THAT WHEN CHECKING THE FOLEY CATHETER ONE DAY AFTER PLACEMENT, THE BALLOON BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02397 | 2023-06-12 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | CUSTOMER REPORTED THAT THIS PATIENT WAS ADMITTED TO HOSPITAL ON (B)(6)2023. THE NURSE PERFORMED THE PICC CATHETER MAINTENANCE PLACED IN THE LEFT ARM. WHEN FLUSHING THE CATHETER, LIQUID OOZED FROM THE INSERTION SITE WITH NO BLOOD. X-RAY TEST WAS CONDUCTED ON THE DOCTOR ORDER, SHOWING THAT THE CATHETER WAS POSITIONED PROPERLY. THE HEAD NURSE OF THE ACEOLOGY GROUP AT THE HOSPITAL WAS CONSULTED AND THE CATHETER WAS EXAMINED AT BEDSIDE. THE RUPTURE OF THE CATHETER WAS SUSPECTED. AFTER COMMUNICATION WITH THE DOCTOR, THE CATHETER WAS REMOVED. AN INSPECTION SHOWED A BREACH PARALLEL TO THE CATHETER ABOUT 1.8 CM FROM THE INSERTION SITE. NO PATIENT HARM REPORTED, NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04236 | 2023-06-12 | 2023-05-21 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION (UTI) BUT IS UNSURE IF IT WAS DUE TO THE MAGIC 3 GO COUDE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04242 | 2023-06-12 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE PATIENT STATED THAT ABOVE 10 OUT OF 360 BURST PACKET WAS EMPTY OR HALF FULL. IT WAS STATED THAT THE PATIENT HAD TO USE THEY OWN LUBRICANT IN ORDER TO USED THEM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04245 | 2023-06-12 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0167SI24 | 0167SI24 |  | IT WAS REPORTED THAT THE CUSTOMER WERE UNABLE TO DEFLATE THE FOLEY BALLOON TO REMOVE DUE TO INABILITY IRRIGATE/DRAIN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION". THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DIRECTIONS FOR USE: VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. USING PROPER ASEPTIC METHODS, REMOVE CATHETER FROM PACKAGE. INFLATE THE BALLOON WITH STERILE WATER TO THE VOLUME STATED ON THE INFLATION LUMEN OF THE CATHETER BY USING A WATER-FILLED LUER TIP SYRINGE. CONNECT CATHETER TO COLLECTION CONTAINER. TO DEFLATE THE BALLOON, GENTLY REINSERT THE LUER TIP SYRINGE INTO THE VALVE AND ASPIRATE. CAUTION: TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE IN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04243 | 2023-06-12 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. WHILE INSERTING THE CATHETER INTO THE PATIENT AND CHECKING ITS FIXATION AT THE BLADDER NECK, IT CAME OUT. A PRETEST WAS NOT PERFORMED. AS PER THE ADDITIONAL INFORMATION RECEIVED VIA IBC ON (B)(6)2023, THERE WAS BALLOON BURST OBSERVED IN THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04255 | 2023-06-12 | 2023-05-22 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123614A |  | IT WAS REPORTED, THAT THE PATIENT WAS GIVEN A DISPOSABLE LATEX FOLEY CATHETER TO PREVENT URINARY RETENTION ON (B)(6) 2023. ON (B)(6) 2023, AT 17: 45, IT WAS FOUND THAT THERE WERE ERYTHEMA AND ITCHING ON BOTH BUTTOCKS OF PATIENT, BUT NO SYMPTOMS SUCH AS PALPITATIONS, CHEST TIGHTNESS, OR DYSPNEA. UPON CHECKING, THE VITAL SIGNS WERE NORMAL. AND A 6.0 X 0.7 CM LINE OF RED SPOTS COULD BE SEEN IN THE SACRUM, WHICH WAS CONSIDERED AS CONTACT DERMATITIS. THEN THE USER STOPPED, THE CATHETERIZATION. AND DESONIDE CREAM WAS USED FOR ANTI-ALLERGY SYMPTOMATIC TREATMENT. THE PATIENT WAS KEPT FOR OBSERVATION TILL 22:45 ON (B)(6) 2023. AND THE ERYTHEMA DISAPPEARED. AS PER THE INJURY MANIFESTATION MENTIONED, THAT THE PATIENT HAD RED SKIN AND PAIN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04254 | 2023-06-12 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE VALVE OF THE FOLEY CATHETER WAS GLUED AT AN ANGLE. THE CUSTOMER WANTED TO CHECK IF IT WAS WITHIN STANDARD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04253 | 2023-06-12 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAS BURST. THE CATHETER FELL OUT WHILE THE PATIENT WAS BATHING AND WHEN THE TIP OF THE CATHETER WAS CHECKED, THE BALLOON WAS FOUND TO BE DAMAGED WITH A CRACK. THE EVENT HAS OCCURRED AFTER 7 DAYS OF CATHETER PLACEMENT TO THE PATIENT AND THERE WAS A BALLOON RUPTURE OR TEAR. THE WATER HAS LEAKED FROM THE BALLOON AND NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. THERE WAS NO PATIENT'S URINE STONE AND THE CATHETER WAS NOT PULLED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04251 | 2023-06-12 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER LEAKED FROM THE INFLATION VALVE AREA OF THE FOLEY CATHETER. IT WAS NOTED THAT THERE WAS NO BALLOON RUPTURE OR TEAR FOUND. NO MATERIALS WERE POSSIBLY THERE IN CONTACT WITH THE CATHETER AND THE CATHETER WAS INSERTED BY HAND. THE WATER LEAKED FROM THE VALVE MENTIONED AS BACKFLOW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04233 | 2023-06-12 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE PATIENT HAD A BLADDER INFECTION AND URINARY TRACT INFECTION. PATIENT WAS HOSPITALIZED AND ON ANTIBIOTICS (LOT#JUGU1957 AND LOT#JUGT1013). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
2020394-2023-00415 | 2023-06-12 | 2023-05-16 | I, n, j, u, r, y | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U357594 | U357594 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT THREE CENTIMETER OF THE PTA BALLOON WAS ALLEGEDLY REMAINED IN THE VESSEL AND WAS UNABLE TO REMOVE THE PIECE. REPORTEDLY, THE PTA BALLOON PIECE WAS ATTACHED TO THE VESSEL WALL WITH A STENT. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2025). | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04224 | 2023-06-12 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE USER ATTEMPTED TO REPLACE FOLEY CATHETER DUE TO BALLOON ABNORMALITY AND URINE LEAKAGE, BUT THE CATHETER WAS CUT OFF BECAUSE WATER COULD NOT BE DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02393 | 2023-06-12 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT MRI ISP, 8 FR GROSHONG, INT WO SP, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808561 |  | IT WAS REPORTED THAT TWO YEARS AND NINE MONTHS POST PORT PLACEMENT IN THE SUBCLAVIAN VEIN, THE CATHETER TUBE WAS ALLEGEDLY FOUND TO BE DAMAGED WHEN IT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04190 | 2023-06-12 | 2023-05-22 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MALE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION AROUND THE TIME WHEN USING THE MAGIC 3 GO MALE INTERMITTENT CATHETER BUT WAS NOT SURE IF IT WAS CAUSED. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04189 | 2023-06-12 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER DID NOT DRAIN OUT OF THE FOLEY CATHETER QUICKLY DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04191 | 2023-06-12 | 2023-05-22 | I, n, j, u, r, y | Manufacturer report | MAGIC 3 GO MALE COUDE CATHETER WITH SUREGRIP | MALE COUDE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD SEVERAL URINARY TRACT INFECTIONS THAT HAD PUT THEM IN THE HOSPITAL FOR THREE DAYS. BUT THE PATIENT WAS CURRENTLY ON MEDICATION FOR A URINARY TRACT INFECTIONS. IT WAS STATED THAT THE PATIENT WAS GETTING THEM SO OFTEN THAT WAS PUT ON A PREVENTATIVE. IT WAS UNKNOWN IF THE DEVICE CAUSED THE INFECTION. THE PATIENT REPORTED ISSUES WITH LOT# JUGU0783, LOT# JUGV0317. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
MW5118348 | 2023-06-12 |  | O, t, h, e, r | Voluntary report | POWERGLIDE PRO MIDLINE CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | F220088PT | F220088PT |  | GUIDEWIRE IS MISSING FROM POWERGLIDE WAND, RENDERING IT UNUSABLE. |  | N | N | HEALTH PROFESSIONAL |  | 2023-06
MW5118346 | 2023-06-12 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | ALL SILICONE FOLEY CATHETER 16 | CATHETER, UROLOGICAL (ANTIMICROBIAL) AND ACCESSORIES | C. R. BARD, INC. | A319516AM | A319516AM |  | INFLATED THE BALLOON AND SAW THAT WHEN IT WAS INFLATED WITH THE SALINE, THE FLUID LEAKED OUT OF THE BALLOON AND OUT OF THE CATHETER AND IT DEFLATED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04193 | 2023-06-12 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE PATIENT HAD EIGHT MAGIC 3 MALE INTERMITTENT CATHETERS THAT WERE DEFECTIVE. IT WAS STATED THEY HAD NO FLUID IN THEM AND THE PACKAGE WAS GLUED SO TIGHT, IT WAS HARD TO GET THE PACKAGE OPEN. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
17105654 | 2023-06-12 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARD 18 FR FOLEY CATHETER | CATHETER, UROLOGICAL | C.R. BARD, INC. |  |  |  | LONG CURLY BLACK HAIR FOUND IN STERILE 18 FR FOLEY UPON OPENING PACKAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04194 | 2023-06-12 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE FEW MAGIC 3 INTERMITTENT CATHETERS FROM THIS ORDER WERE TWISTED. REPRESENTATIVE APOLOGIZED AND ADVISED IN FUTURE LET THEM KNOW RIGHT AWAY SO THEY COULD DOCUMENT LOT AND REPLACE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04226 | 2023-06-12 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER COULD NOT BE DRAINED FROM THE FOLEY CATHETER DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04197 | 2023-06-12 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS WITHDRAWN WITHIN A FEW HOURS AFTER PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04204 | 2023-06-12 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE USER ATTEMPTED TO REPLACE FOLEY CATHETER DUE TO BALLOON ABNORMALITY AND URINE LEAKAGE, BUT THE CATHETER WAS CUT OFF BECAUSE WATER COULD NOT BE DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04207 | 2023-06-12 | 2023-05-22 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. STATED THAT THE PATIENT HAD 70 OF THE CATHETERS THAT WERE RECALLED AND HAD BEEN GETTING INFECTIONS AND WAS NOT SURE IF IT WAS DUE TO THE LUBRICANT ON CATHETERS OR OLD ONES THEY HAD FROM PREVIOUS COMPANY. REPRESENTATIVE ADVISED THEY DISCARD THE 70 WITH AFFECTED LOT NUMBERS AND WOULD SEND SAMPLES OF OTHER NON-LUBRICATED CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION. THE AFFECTED LOTS WERE (LOT# JUGT2732 AND LOT# JUGT3397). PER FOLLOW UP VIA PHONE ON 04JUN2023, IT WAS REPORTED THAT THE CUSTOMER DEVELOPED URINARY TRACT INFECTION AND SOUGHT MEDICAL INTERVENTION. CUSTOMER WAS PRESCRIBED ANTIBIOTICS AND LIBERATOR REPLACED CATHETERS OF DIFFERENT TYPE WHICH HAVE CAUSED NO ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04212 | 2023-06-12 | 2023-05-23 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION (UTI) WHILE USING THE MAGIC 3 GO INTERMITTENT CATHETER. PATIENT HAD SEVERAL URINARY TRACT INFECTION (UTI) AND HAD 30 CATHETERS ON HAND. REPRESENTATIVE ADVISED TO THROW AWAY AND THE PATIENT DID NOT TAKE LOT NUMBER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04215 | 2023-06-12 | 2023-05-22 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. STATED THAT THE PATIENT HAD 70 OF THE CATHETERS THAT WERE RECALLED AND HAD BEEN GETTING INFECTIONS AND WAS NOT SURE IF IT WAS DUE TO THE LUBRICANT ON CATHETERS OR OLD ONES THEY HAD FROM PREVIOUS COMPANY. REPRESENTATIVE ADVISED THEY DISCARD THE 70 WITH AFFECTED LOT NUMBERS AND WOULD SEND SAMPLES OF OTHER NON-LUBRICATED CATHETERS. THE AFFECTED LOTS WERE (LOT# JUGT2732) AND (LOT# JUGT3397). PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE CUSTOMER DEVELOPED URINARY TRACT INFECTION AND SOUGHT MEDICAL INTERVENTION. CUSTOMER WAS PRESCRIBED ANTIBIOTICS AND LIBERATOR REPLACED CATHETERS OF DIFFERENT TYPE WHICH HAVE CAUSED NO ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04223 | 2023-06-12 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON RUPTURED THREE TIMES IN A ROW IN THE SAME PATIENT. A CALCULUS WAS SUSPECTED, BUT THERE SEEMS TO BE NO HISTORY. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 30MAY2023, IT WAS STATED THAT DOCTOR DOUBTED IF THERE WERE CALCULUS STONES, BUT THERE WERE NO CALCULUS SUSPECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04225 | 2023-06-12 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT IMMEDIATELY AFTER CATHETER PLACEMENT TO THE PATIENT IN THE OPERATING ROOM, THE CATHETER SPONTANEOUSLY FELL OUT. IT'S SUSPECTED DUE TO A PINHOLE OR A LUMEN TO LUMEN LEAK. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE COULD BE LUMEN COMPROMISED BY A PUNCTURE OR CUT. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3-WAY TEMPERATURE SENSING CATHETER. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE SOLUTION IMMEDIATELY WENT INTO THE DRAINAGE LUMEN AT BIFURCATION AND CREATED LUMEN TO LUMEN LEAK. THIS IS OUT OF SPECIFICATION, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02396 | 2023-06-12 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT ONE YEAR FIVE MONTHS AND TWENTY-NINE DAYS POST PORT PLACEMENT, THE CATHETER ALLEGEDLY OCCLUDED BY LEAVING THE CLAMP OPEN FOR SEVERAL TENS OF MINUTES WHILE PREPARING FOR THE ADMINISTRATION OF THE MEDICINE. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT AND A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. GROSS, MICROSCOPIC VISUAL, TACTILE AND FUNCTIONAL TESTING WERE PERFORMED. UPON INFUSION, THROMBUS WITH WATER WAS OBSERVED EXITING THE PORT STEM AND DISTAL CATHETER SEGMENT. ASPIRATION WAS ATTEMPTED AND WAS SUCCESSFUL AS WATER RETURNED TO THE IN-HOUSE SYRINGE ATTACHED. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED OCCLUSION ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: H6 (METHOD, RESULT, CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02394 | 2023-06-12 | 2022-11-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT MRI ISP, 8FR GROSHONG, INT WO SP, ATTSL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808561 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE LEAKAGE WAS ALLEGEDLY FOUND AT THE GAP BETWEEN THE PUNCTURED PART OF PORT AND THE HARD BASE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 09/2023). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04199 | 2023-06-12 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY STATLOCK DID NOT WORK, AND IT DID NOT CLOSE IN PLACE. REPRESENTATIVE ADVISED PATIENT WAS USING THE STATLOCK INCORRECTLY AND ADVISED ON PROPER USE. AS PER FOLLOW UP VIA PHONE ON (B)(6)2023, IT WAS REPORTED THAT PATIENT WAS USING LATEX FOLEY, BUT THEY WERE SWITCHED TO A SILICONE FOLEY AND THE LATEX STATLOCK DID NOT FIT THE SILICONE FOLEY . AS PER FOLLOW UP VIA PHONE ON (B)(6)023, IT WAS REPORTED THAT THE PATIENT'S UROLOGIST STATED THAT THEIR SUPRAPUBIC TUBE WAS CLOGGED. IT WAS SUPPOSED TO HOLD FOR AT LEAST A MONTH, BUT IT CONTINUALLY FAILED. THEY SWITCHED FROM RUBBER TO PLASTIC, AND THE DRAIN BAG WAS NOT HOLDING ANY URINE AS IT WAS EMPTYING OVER THEIR CLOTHES. THEY HAS TO LEAN BACK IN THEIR WHEELCHAIR TO RELIEVE PRESSURE ON GROIN AND BUTTOCKS. THEY WERE A QUADRIPLEGIC, A RASH AND URINARY TRACT INFECTION DEVELOPED BECAUSE OF THE FAILING STATLOCK. ALSO NOTED THAT THE ISSUE HAD SINCE BEEN RESOLVED AS THEY WERE BACK ON THE OLD STATLOCKS THEY USED YEARS PRIOR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04248 | 2023-06-12 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE BALLOON BURST AND THE FOLEY CATHETER FELL OUT. THEN A NEW CATHETER WAS USED. PER ADDITIONAL INFORMATION VIA EMAIL FROM (B)(6) ON 26MAY2023, IT WAS STATED THAT THE THERE WAS NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02404 | 2023-06-13 | 2023-01-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1606052 | 1606052 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, RESISTANCE WAS ALLEGEDLY FELT AS THE DILATOR WAS INSERTED. IT WAS FURTHER REPORTED THAT THE DILATOR ALLEGEDLY COULD NOT BE INSERTED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE FULLY PEELED 6.5FR PEEL-APART SHEATH AND TWO VESSEL DILATORS, ONE GUIDEWIRE STRAIGHTENER AND ONE J-TIP GUIDEWIRE WERE RETURNED. MANUFACTURING SITE EVALUATION OF THE SAMPLE FOLDS THROUGH THE GRAY SHEATH, TWO VESSEL DILATORS WERE OBSERVED, IN ONE OF THEM IT WAS OBSERVED THAT THE SHEATH VALVE WAS TRAPPED, THIS DILATOR HAD A PRONOUNCED DOUBLE TOWARDS THE DISTANT TIP, NO FRACTURES WERE OBSERVED AT THE TIP OF THE VESSEL DILATOR, SLIGHT RESIDUES OF USE WERE OBSERVED THROUGH THE SAMPLE. THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED DEFORMATION AND MATERIAL SEPARATION ISSUE AS THE SHEATH SHAFTS AND THE VESSEL DILATORS WERE NOTED TO BE DEFORMED AND THE VALVE CAP WAS NOT NOTED ON BOTH T-SEGMENTS. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE TO ADVANCE ISSUE AS THE EXACT CIRCUMSTANCE AT THE TIME OF THE REPORTED EVENT CANNOT BE REPLICATED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 11/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10 | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04280 | 2023-06-13 | 2023-05-23 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53812G | 53812G |  | IT WAS REPORTED THAT PATIENT GOT A URINARY TRACT INFECTION, 5 WEEKS AGO AND IT JUST CLEARED UP THIS WEEK (JUGT2705), (JUGT1012) AND (JUGR2031). IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04273 | 2023-06-13 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED, THAT FOLEY BALLOON BURST INSIDE THE PATIENT'S BLADDER DURING INFLATION WITH 10ML NS. NO IMPACT WAS NOTED TO PATIENT. FOLEY BALLOON WAS IMMEDIATELY REMOVED WITH BURST WAS HEARD. THE BALLOON WAS INSPECTED. AND A TEAR WAS NOTED, ON THE OUTER LAYER OF THE FOLEY TIP, BELOW THE DRAINAGE HOLE. A NEW FOLEY WAS PLACED. PER FOLLOW UP VIA EMAIL ON (B)(6) 2023. IT WAS REPORTED, THAT STAFF INFLATED THE BALLOON, AFTER INSERTION INTO THE BLADDER. THEY HEARD THE BALLOON BURST. STATED, THAT THE FOLEY WAS PULLED OUT OF THE URETHRA AND WAS NOTED, TO BE INTACT WITH NO PIECES OR PORTION OF THE BALLOON MISSING. THE OUTER LINING WAS ONLY SPLIT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04274 | 2023-06-13 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE PATIENT HAD AN INDWELLING FOLEY CATHETER AT 11:40 ON (B)(6) 2023, AND THE URINARY CATHETER RUPTURED AT 11:30 ON (B)(6) 2023. THE USER IMMEDIATELY INFORMED THE DOCTOR, CHECKED WHETHER THE PATIENT'S URETHRAL OPENING WAS NORMAL, MEASURED THE VITAL SIGNS TO BE STABLE AND CONTACTED THE EQUIPMENT DEPARTMENT TO DEAL WITH IT. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 05JUN2023, STATED THAT THERE WAS NO PATIENT INJURY AND THERE WERE NO MISSING PIECES OF BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04269 | 2023-06-13 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED, THAT THE GREEN CLIPS DID NOT HOLD DRAIN IN PLACE AND FELL OUT IN DRAIN BAG. ALSO THE CLAMP DID NOT CLAMP DOWN TIGHT ENOUGH, AND IT STARTED TO LEAK ON FLOOR. CUSTOMER ADDED THAT THE FOLEY CATHETER BALLOON DID NOT INFLATE. AND THEY END UP INTO BREAK THE CONNECTION AND REPLACE THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE, A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04271 | 2023-06-13 | 2023-05-23 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53812G | 53812G |  | IT WAS REPORTED THAT PATIENT GOT A URINARY TRACT INFECTION, 5 WEEKS AGO AND IT JUST CLEARED UP THIS WEEK. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04268 | 2023-06-13 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE URINE LEAKED FROM THE SIDE OF THE CATHETER. THERE HAVE BEEN MULTIPLE EVENTS OF URINE LEAKING FROM THE SIDE OF THE CATHETER WITH THE CATHETER BEING IN PLACE OCCURRED IN ONE DEPARTMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04256 | 2023-06-13 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165V18S |  | IT WAS REPORTED THAT A HOMECARE PATIENT WITH SUPRAPUBIC CATHETER WHEN COMPLETED FLUSHING, THE CATHETER BROKE AND THE BALLOON DID NOT DEFLATE. THE PATIENT HAD TO VISIT ER TO HAVE CENTRAL LINE WIRE TO BREAK BALLOON THROUGH PORT AND REMOVE NUB AND THEN WAS CHANGED A NEW CATHETER. THEY HAVE ALSO HAD THE EXPERIENCE WITH THE CATHETER CRACKING AGAIN IN (B)(6) 2023 WHEN OTHER CATHETER CHANGES COMPLETED AS WELL. FIRST DEVICE WAS A 16F AND THE SECOND DEVICE WAS 18F. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. EVALUATION OBSERVED THE CATHETER WAS CUT ALONG THE DRAINAGE LUMEN AND THERE WAS SALT ACCUMULATION INSIDE THE LUMEN. IT WAS ALSO PRESENCE OF CRACK AT THE MIDDLE OF CATHETER SHAFT. THE SURFACE LOOKS LIKE HAS BEEN CUT WITH APPLIED SOME FORCE. THE EXACT ROOT CAUSE COULD NOT BE IDENTIFIED. ALTHOUGH A ROOT CAUSE COULD NOT BE DEFINITIVELY IDENTIFIED, BASED ON THE RISK DOCUMENTATION REVIEW, AND THE OBSERVATIONS OF THE SAMPLE EVALUATION, THE POTENTIAL ROOT CAUSES FOR THIS TYPE OF FAILURE INCLUDE: OPERATOR ERROR OR MACHINE MALFUNCTION, OVER EXPOSURE TO CHLORINATION DURING LATEX PROCESSING AND OVER EXPOSURE TO OZONE RESULTING IN PRODUCT DEGRADATION, EXPOSURE TO PETROLATUM BASED PRODUCTS ETC. AND USER VIOLATION OF STANDARD PRACTICE. THE EXACT ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION, AND THE EVENT COULD NOT BE CONFIRMED AS USER ERROR OR MANUFACTURING-RELATED EVENT. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04267 | 2023-06-13 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE URINE LEAKED FROM THE SIDE OF THE CATHETER. THERE HAVE BEEN MULTIPLE EVENTS OF URINE LEAKING FROM THE SIDE OF THE CATHETER WITH THE CATHETER BEING IN PLACE OCCURRED IN ONE DEPARTMENT . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04263 | 2023-06-13 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OFF DUE TO THE BALLOON BURST. UPON CHECKING THE CATHETER SINCE THERE WAS NO URINE FLOW OBSERVED, IT WAS FOUND THAT THE FOLEY CATHETER HAS SPONTANEOUSLY FELL OUT OF THE PATIENT. THE SALES REPRESENTATIVE VISUALLY CONFIRMED THAT THE BALLOON HAS BURST. THE EVENT HAS OCCURRED AFTER 1 DAY OF PLACEMENT AND THE WATER LEAKED FROM THE BALLOON SINCE THERE WAS A BALLOON RUPTURE OR TEAR. NO MATERIALS CAME INTO CONTACT WITH THE CATHETER AND WAS NOT PULLED. AND ALSO THERE WAS NO PATIENT'S URINE STONE. AS PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON 30MAY2023, STATED THAT THERE WAS NO BALLOON RUPTURE OBSERVED IN THE RETURNED SAMPLE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D,F,H. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04266 | 2023-06-13 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE URINE LEAKED FROM THE SIDE OF THE CATHETER. THERE HAVE BEEN MULTIPLE EVENTS OF URINE LEAKING FROM THE SIDE OF THE CATHETER WITH THE CATHETER BEING IN PLACE OCCURRED IN ONE DEPARTMENT. | THE REPORTED EVENT WAS NOT CONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. VISUAL AND FUNCTIONAL INSPECTION OBSERVED THE EXTERIOR AND FUNCTIONALITY OF THE SAMPLE AND DID NOT FIND ANY EVIDENCE OF A FAILURE THAT WOULD SUPPORT THE REPORTED EVENT. THE CATHETER SHAFT MEASUREMENT MET SPECIFICATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE LOW LATEX VISCOSITY. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04260 | 2023-06-13 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LUBRI-SIL® I.C. COMPLETE CARE® ALL-SILICONE FOLEY CATHETER TRAY | FOLEY CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 |  | 900114A |  | IT WAS REPORTED THAT DURING THE PRETEST, WATER LEAKED FROM THE CONNECTION BETWEEN THE INFLATION VALVE AND SYRINGE. THEN THE USE WAS DISCONTINUED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04257 | 2023-06-13 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE TEN LUBE PACKETS WERE EMPTY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
2020394-2023-00422 | 2023-06-13 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U35130630 | U35130630 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY GOT STUCK IN THE SHEATH WHEN IT WAS DEFLATED TO PULL OUT OF THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 11/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02402 | 2023-06-13 | 2023-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DECATHLON DF LONG-TERM HEMODIALYSIS CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | DF19SH24K | DF19SH24K |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY BROKEN AWAY FROM TUNNELER. IT WAS FURTHER REPORTED THAT BROKEN CATHETER WAS SEPARATED AND WOULD NOT STAY ON. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE 19CM DECATHLON DF D/L CATHETER WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC AND FUNCTIONAL EVALUATIONS WERE PERFORMED. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DETACHMENT ISSUE AS THE TUNNELER WAS NOT RETURNED FOR EVALUATION. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 01/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02400 | 2023-06-13 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE IN THE RIGHT ATRIUM VIA THE JUGULAR VEIN, AFTER THE PUNCTURE, THE GUIDEWIRE ALLEGEDLY MET RESISTANCE WHILE IMPLANTING AND WAS DIFFICULT TO ADJUST. IT WAS FURTHER REPORTED THAT WHEN THE GUIDEWIRE WAS FOUND DIFFICULT TO ADJUST THROUGH THE PUNCTURE NEEDLE, THE WHOLE DEVICE WAS REMOVED TO ADJUST, AND IT WAS FOUND THAT THE THIN WIRE IN THE GUIDEWIRE WAS ALLEGEDLY FRACTURED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 19CM HEMOSTAR D/L CATHETER, ONE 15.0FR PEEL-APART SHEATH AND VESSEL DILATOR, ONE 8FR & 12FR DILATOR, ONE INTRODUCER NEEDLE, ONE J-TIP GUIDEWIRE, ONE GUIDEWIRE HOOP, ONE TUNNELER WITH A LOADED PROTECTIVE SHEATH AND ONE END CAP WERE RETURNED FOR EVALUATION. GROSS VISUAL, TACTILE, DIMENSIONAL AND FUNCTIONAL EVALUATIONS WERE PERFORMED. THE J-TIP GUIDEWIRE WAS NOTED TO BE UNCOILED IN THE DISTAL PORTION. BOTH CORE WIRES WERE PROTRUDING THE COILS OF THE GUIDEWIRE AND NOTED TO HAVE COMPLETE BREAKS ON THE ENDS. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS A GUIDEWIRE WAS STRETCHED AT THE DISTAL END AND BOTH THE ROUND AND FLAT CORE WIRES WERE NOTED TO BE BROKEN AND PROTRUDING. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION, STRETCHED, AND UNRAVELED ISSUES. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO ADVANCE AND ENTRAPMENT ISSUES AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. BASED ON THE MEASUREMENTS RECORDED DURING SAMPLE EVALUATION AND APPLICABLE DRAWINGS, NO MEASUREMENTS RECORDED COULD BE CONFIRMED TO BE OUT OF TOLERANCE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 03/2024), G3. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02409 | 2023-06-13 | 2023-02-20 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA THE RIGHT INTERNAL JUGULAR VEIN, RESISTANCE WAS ALLEGEDLY FELT WHEN THE GUIDE WIRE WAS BEING REMOVED. IT WAS FURTHER REPORTED THAT THE TIP OF THE GUIDE WIRE WAS FOUND TO BE DRAWN AFTER REMOVING THE GUIDEWIRE OUT OF THE VESSEL. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY H10: D4 (EXPIRY DATE: 02/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02408 | 2023-06-13 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808000 | 1808000 |  | IT WAS REPORTED THAT APPROXIMATELY EIGHT YEARS POST PORT PLACEMENT, THE DEVICE ALLEGEDLY NOT ABLE TO GET A BLOOD RETURN. IT WAS FURTHER REPORTED THAT THE DEVICE COULD NOT BE FLUSHED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 07/2017). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02401 | 2023-06-13 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1716000 | 1716000 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04278 | 2023-06-13 | 2023-05-22 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAD SEVERAL URINARY TRACT INFECTION FROM THE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02405 | 2023-06-13 | 2023-01-04 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 5608062 |  | IT WAS REPORTED THAT ONE YEAR TWO MONTHS AND THREE WEEKS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED HEMATOMA, ARM PAIN AND A HIGH FEVER. IT WAS FURTHER REPORTED THAT AN X-RAY WAS PERFORMED WHICH SHOWED A POTENTIAL BLOODSTREAM INFECTION. REPORTEDLY, PATIENT PRESENTED TO HOSPITAL TO UNDERGO A SURGICAL PROCEDURE TO REMOVE THE INFECTED CATHETER. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. MEDICAL RECORDS WERE PROVIDED AND RECEIVED. APPROXIMATELY ONE YEAR AND THREE MONTHS POST PORT PLACEMENT, THE PATIENT DEVELOPED CATHETER RELATED BLOOD CLOT. AROUND SIX MONTHS AND TWENTY-FOUR DAYS LATER, THE CATHETER WAS REMOVED. THE PATIENT LEFT IN A STABLE CONDITION. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED ADVERSE EVENTS AS NO OBJECTIVE EVIDENCE HAS BEEN PROVIDED TO CONFIRM ANY ALLEGED DEFICIENCY WITH THE PORT. FURTHERMORE, THE CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 08/2022). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04279 | 2023-06-13 | 2023-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT TRANSURETHRAL RIGHT URETEROSCOPY AND LASER LITHOTRIPSY UNDER GENERAL ANESTHESIA, AND CATHETERIZATION WAS PERFORMED. HOWEVER, THE BALLOON OF THE CATHETER BURST. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04281 | 2023-06-13 | 2023-05-23 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53812G | 53812G |  | IT WAS REPORTED THAT PATIENT GOT A URINARY TRACT INFECTION, 5 WEEKS AGO AND IT JUST CLEARED UP THIS WEEK (JUGT2705), (JUGT1012) AND (JUGR2031). IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02421 | 2023-06-13 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D3 |  | IT WAS REPORTED, "WE WERE PLACING A PICC AND WHEN WE WENT TO PUSH THE WIRE BACK THROUGH THE PICC AFTER CUTTING THE PICC TO MEASURED LENGTH, IT WAS NOTICED THAT THE PICC WIRE (2 INCHES FROM THE END) WAS BROKEN AND HANGING ON BY A THREAD. THE WIRE EVENTUALLY FULLY SNAPPED AND BROKE OFF WITH (B)(6) TOUCHING IT". ADDITIONAL INFORMATION RECEIVED (B)(6)2023: "THE INTERNAL PICC WIRE WAS BROKEN AND HINGED PRIOR TO INSERTING PICC LINE SO NO PIECES WERE RETAINED IN THE PATIENT OR NEEDING RETRIEVAL." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02419 | 2023-06-13 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED "THE ORIGINAL PICC WAS LEFT IN INDENTATION ON APRIL 26. WHEN INFUSION WAS GIVEN TO THE PATIENT, IT WAS FOUND THAT THE PLASTIC AT THE TIP OF THE ORIGINAL HEPARIN CAP OF THE PICC WAS WARPED, AND THE PLASTIC DIRECTLY BROKE AND FELL IN THE PROCESS OF IODOPOR DISINFECTION." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02416 | 2023-06-13 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER KIT 13F 15CM TL | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5603150 |  | IT WAS REPORTED BY CUSTOMER THAT THE GUIDEWIRE IN THE TEMP HD CATH VERY DIFFICULT/NOT ABLE TO THREAD THROUGH THE LINE. CLINICIAN HAD TO OPEN A SECOND KIT. ADDITIONAL INFORMATION RECEIVED MAY 31, 2023. EVENT DIRECTLY INVOLVED A PATIENT FOR AN URGENT MEDICAL SITUATION. NO HARM OR INJURY. DELAY IN INSERTION CAUSED DELAY IN DIALYSIS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02417 | 2023-06-13 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) SINGLE-LUMEN MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED THAT, "ON (B)(6), 2022, THE PATIENT WAS IMPLANTED WITH A PERIPHERALLY CANNULATED CENTRAL VENOUS CATHETER KIT MANUFACTURED BY BARD VASCULAR ACCESS SYSTEMS. THE PATIENT WAS ADMITTED ON (B)(6) 2023, AND THE NURSE PERFORMED PICC CATHETER MAINTENANCE, WITH FLUID OOZING FROM THE PUNCTURE POINT WITHOUT REBLEEDING DURING FLUSHING, NO LOCAL OOZING SWELLING, AND NORMAL CATHETER ORIENTATION ON EXAMINATION. THE PATIENT WAS SUSPECTED OF HAVING A RUPTURED CATHETER AND WAS PULLED OUT. THE CATHETER WAS FOUND TO HAVE A RUPTURE PARALLEL TO THE CATHETER AT ABOUT 1.8 CM FROM THE PUNCTURE POINT. THE CATHETER WAS IMMEDIATELY STOPPED AND SEALED AND DELIVERED TO THE MEDICAL EQUIPMENT DEPARTMENT.". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02415 | 2023-06-13 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED BY CUSTOMER THAT, "A 4 FR PROVENA KINK THAT IS STILL IN THE PATIENT THAT WE ARE MANIPULATING EXTERNALLY AS SHE SHOULD FINISH THERAPY SOON." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED ON 5/30/2023: CUSTOMER REPORTED THAT ON (B)(6) AT 18:15 BLOOD RETURN SLUGGISH AND DECLOTTING OF BOTH LUMENS WITH ALTEPLASE WAS UNSUCCESSFUL SO AN X-RAY WAS REQUESTED WHICH SHOWED A KINK. GENTLE PRESSURE ON THE INSERTION SITE IMPROVED FUNCTION. ON (B)(6) AT 02:00 THE LINE WAS NOTED TO BE HARD AND SLUGGISH TO FLUSH IN BOTH LUMENS. ON (B)(6) 2023 THE LINE WAS REMOVED. THERAPY WAS COMPLETED WITH NO ISSUES MENTIONED BY NURSES. CATHETER HAD BEEN INSERTED 12CM ABOVE THE ANTECUBITAL FOSSA WITH 2CM LEFT EXTERNAL TO THE PATIENT. THE EVENT DID NOT INVOLVE AN URGENT OR CRITICAL MEDICAL SITUATION AND DID NOT CAUSE A DELAY IN TREATMENT. THERE WAS NO PATIENT HARM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A KINKED CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR D/L POWERPICC PROVENA CATHETER. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. THREE KINK IMPRESSIONS WERE OBSERVED BETWEEN THE 3CM AND 6CM DEPTH MARKINGS. THE DEPTH MARKINGS WERE WORN IN THE VICINITY OF THE KINK IMPRESSIONS. MICROSCOPIC INSPECTION OF THE SAMPLE REVEALED MATERIAL WEAR AND BUCKLING AT THE KINK SITES. THE SAMPLE KINKED EASILY AND PREFERENTIALLY AT THE KINK IMPRESSION SITES DURING MANUAL MANIPULATION. THE IMPRESSIONS APPEARED TO BE CAUSED BY SHARP KINKING OF THE CATHETER SHAFT. THE USE RESIDUES AND DEPTH MARKING WEAR SUGGESTED THAT KINKING OCCURRED DURING DEVICE USE. THE SHORT PERIOD BETWEEN DEVICE PLACEMENT AND REMOVAL SUGGESTED THAT KINKING MAY HAVE OCCURRED DURING DEVICE PLACEMENT OR SECUREMENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02412 | 2023-06-13 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC 4F S/L 3CG BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 22184118 |  | IT WAS REPORTED THAT "PICC WAS INSERTED BY VAT AND HAD ONLY BEEN IN A FEW DAYS. THE NURSE HAD JUST CHANGED THE DRESSING, THE PATIENT WAS MOVED FROM ONE BED AREA TO A SINGLE ROOM AND THE PICC CATHETER SNAPPED DURING THIS TRANSFER. THE DRESSING WAS NOT DISLODGED OR RIPPED OFF IN ANY WAY, IT WAS FULLY INTACT, THE PICC WAS STILL IN THE STATLOCK AND IT WAS LACERATED IN THE BOX CLAMP." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02414 | 2023-06-13 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH & PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D4 |  | IT WAS REPORTED, "CALLER STATES HE HAS A PATIENT WITH A POWERPICC LOT#REGZ1033 WITH A BROKEN CLAMP. CALLER ASKING IF THEY SHOULD REPLACE THIS PICC. RESPONSE: INFORMED CALLER THE CLAMP IS A SAFETY FEATURE FOR THE OPEN ENDED PICCS. WITHOUT THE PICC CLAMP, AIR CAN ENTER THE LINE OR THE PATIENT CAN BLEED OUT IF THE CONNECTOR BECOMES DISCONNECTED. SUGGESTED CALLER CONSULT THE PROVIDER TO ASSESS THE LINE AND DETERMINE IF IT SHOULD BE REPLACED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04297 | 2023-06-13 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165V16S |  | IT WAS REPORTED THAT THE FLUID IN THE BALLOON COULD NOT BE DRAINED, AND THE FOLEY CATHETER COULD NOT BE PULLED OUT. THE FOLEY CATHETER WAS USED ACCORDING TO THE DOCTOR'S ADVICE. AFTER THE PATIENT UNDERWENT URETEROSCOPIC HOLMIUM LASER LITHOTRIPSY, THE INDWELLING CATHETERIZATION WAS GIVEN ACCORDING TO THE DOCTOR'S ADVICE. WHEN THE CATHETER WAS PULLED OUT, THE BALLOON FLUID COULD NOT BE DISCHARGED, AND THE CATHETER COULD NOT BE PULLED OUT. A SYRINGE WAS USED TO PUNCTURE THE BALLOON AND DRAIN THE LIQUID BEFORE PULLING OUT THE CATHETER FORCEFULLY, WHICH WOULD AFFECT THE PATIENT. IT WAS MORE DIFFICULT TO PULL OUT THE CATHETER, WHICH INCREASED THE PAIN OF THE PATIENT. THIS INCREASED THE TIME FOR THE PATIENT TO TAKE TREATMENT MEASURES. STATED THAT ON (B)(6)2023, THEY INITIATED THERAPEUTIC MEASURES, USED SYRINGE TO PUNCTURE BALLOON, DRAIN FLUID, AND FORCEFULLY PULL OUT THE URINARY CATHETER. STATED THAT THE ADVERSE EVENT IMPROVEMENT TIME WAS ON (B)(6)2023. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02410 | 2023-06-13 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 5F FULL TRAY DL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2275108 |  | IT WAS REPORTED BY CUSTOMER THAT THEY ARE HAVING ISSUES LEAKING PICC'S AND THEY HAD TO TROUBLESHOOT EACH INSERTION TO ENSURE IT REMAINED A CLOSED SYSTEM. ADDITIONAL INFORMATION RECEIVED MAY 15, 2023: IT INVOLVED MULTIPLE PATIENTS WITH DIFFERENT DEMOGRAPHICS. THE EVENT DIRECTLY INVOLVED A PATIENT AND USER. THE EVENT DID NOT INVOLVE AN URGENT MEDICAL SITUATION. PATIENT HARM OR INJURY DID NOT OCCUR AS WE WERE ABLE TO INTERVENE BEFORE INSERTING THE PICC LINE. THE EVENT DID NOT INTERRUPT THE ADMINISTRATION OF MEDICATION. WE HAD TO ORDER CHEST XRAYS TO CONFIRM PLACEMENT. SL WIRES WERE USED IN DL PICCS. A SPECIFIC NUMBER OF AFFECTED PATIENTS OR DEVICES WAS NOT PROVIDED BY THE COMPLAINANT. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04294 | 2023-06-13 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PRE-TESTED AND CONFIRMED THAT EVERYTHING WAS FINE, BUT AFTER INSERTION, THE CATHETER FELL OUT OF THE PATIENT AFTER SURGERY. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6)2023, THE DETAILS RELATED TO DAMAGE ON THE BALLOON WEREN'T PROVIDED FROM CUSTOMER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04292 | 2023-06-13 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER CREEVY MODEL | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 01267520 | 01267520 |  | IT WAS REPORTED THAT AFTER INSERTING THE FOLEY CATHETER INTO THE URETHRA, THERE WAS NO RESISTANCE WHILE INJECTING THE STERILE WATER THROUGH THE INFLATION LUMEN BUT THEN A RESISTANCE WAS FELT WHILE INJECTING THE CLEANING SOLUTION, SO THE CATHETER WAS REMOVED AND THE CLEANING SOLUTION WAS INJECTED AGAIN INTO THE IRRIGATION LUMEN AND TESTED, RESULTING IN BALLOON INFLATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04290 | 2023-06-13 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM NEAR THE VALVE AND COULD NOT BE INJECTED FROM THE SYRINGE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04289 | 2023-06-13 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON COULD NOT BE INFLATED PROPERLY. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 30MAY2023, IT WAS STATED THAT THE SYRINGE MALFUNCTION WAS NOT REPORTED AND NO PATIENT INVOLVEMENT WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04291 | 2023-06-13 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THERE WAS A PINHOLE ON THE BALLOON AND THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 30MAY2023, IT WAS STATED THAT THE CATHETER FELL OUT DUE TO BALLOON BURST, NOT A PINHOLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04287 | 2023-06-13 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON DID NOT INFLATE WHEN THE USER TRIED TO INJECT STERILIZED DISTILLED WATER DURING THE PRETEST. THE REPRESENTATIVE STATED THAT THE CUSTOMER USED WORDING OF THE CATHETER FACTOR, BUT THERE WAS ALSO A PROBLEM WITH THE SYRINGE AND WANTED TO INVESTIGATE THAT AS WELL. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, CLARIFIED THAT THE ISSUE WAS REPORTED AGAINST THE INFLATION PROBLEM OF THE CATHETER. THEN, THE SALES REPRESENTATIVE SUSPECTED THE INFLATION PROBLEM CAUSED BY THE SYRINGE DEFECT. THIS WAS JUST ONLY OPINION OF THE SALES REPRESENTATIVE AND THE SYRINGE ISSUE WAS NOT REPORTED BY THE CUSTOMER. HOWEVER, THE INVESTIGATION WAS NEEDED WHETHER THE INFLATION PROBLEM WAS CAUSED BY THE SYRINGE DEFECT OR NOT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04286 | 2023-06-13 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 129416M | 129416M |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS NOT DRAINING, WAS CHECKING TO SEE IF THE LINE WAS KINKED AND NOTICED PORT WAS MISSING AND FOUND IN THE BED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04296 | 2023-06-13 | 2023-05-24 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION THAT LED TO SEVERE SEPSIS AND HAS 8 FLAGS IN URINE CULTURE IN (B)(6) 2022 AND WAS HOSPITALIZED FOR 4 DAYS AND 2 MONTHS TO RECOVER. PER FOLLOW UP VIA PHONE ON 01JUN2023, IT WAS REPORTED THAT THE CUSTOMER STATED THAT THEY DEVELOPED SEPSIS AND WAS HOSPITALIZED FOR 4 DAYS AND DID NOT HAVE A LOT # BUT STATED THAT THEY WERE TOLD TO TRASH THE CATHETERS IMMEDIATELY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04304 | 2023-06-14 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PLACED IN THE OPERATING ROOM, AND THE CATHETER FELL OUT IN THE WARD. WHEN THE USER LOOKED AT THE BALLOON, THE BALLOON WAS FOUND TO BE RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04305 | 2023-06-14 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON AREA WAS DAMAGED. THE DETAILED INFORMATION WAS NOT OBTAINED. AS PER MAIL RECEIVED FROM IBC ON 30MAY2023, STATED THAT THE BALLOON RUPTURE WAS NOT OBSERVED TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04306 | 2023-06-14 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PLACED AT THE UROLOGY OUTPATIENT DEPARTMENT ON THE PREVIOUS DAY, AND THERE WAS A SMALL CRACK AT THE BASE OF THE 2-WAY PORTION, AND THE CATHETER WAS WITHDRAWN SPONTANEOUSLY. STATED THAT THE PATIENT WOULD COME TO THE OUTPATIENT CLINIC AGAIN ON THE NEXT DAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04307 | 2023-06-14 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILIZED WATER LEAKED FROM THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04308 | 2023-06-14 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS SPONTANEOUSLY WITHDRAWN FROM A PATIENT IN THE HOSPITAL WARD. PER PHOTO SAMPLE FROM IBC ON (B)(6) 2023, IT WAS FOUND THAT THE INFLATION FUNNEL WAS FOUND TO BE BROKEN. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, STATED THAT THE INFLATION FUNNEL WAS RUPTURED. THE DETAILS WEREN'T PROVIDED FROM CUSTOMER. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE INFLATION FUNNEL NEVER ENTERED INSIDE THE BODY OF THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02425 | 2023-06-14 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F DL FT KIT W/GUARDIVA & PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1295108FD2 |  | IT WAS REPORTED BY CUSTOMER ON (B)(6) 2023 "LEAKING AROUND WIRE AND RUBBER PORT". NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02443 | 2023-06-14 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D |  | IT WAS REPORTED BY CUSTOMER THAT 4 FRENCH SINGLE LUMEN PICC WAS OPENED AND THE WIRE FOR THE SHEATH HAD A BIT OF A BURR ON THE END. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02442 | 2023-06-14 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5 F POLYURETHANE CATHETER NURSING BASIC TRAY W/MICROINTRODUCER 55CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3175118 |  | IT WAS REPORTED, "THE CATHETER HUB APPEARS TO HAVE CRACKED WHERE THERE IS LEAKAGE OF ADMINISTERED SOLUTIONS. THE CATHETER HAD TO BE CHANGED." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04328 | 2023-06-14 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER VALVE DAMAGE WAS OBSERVED DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02435 | 2023-06-14 | 2023-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER (4F) (SINGLE-LUMEN) (55CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000796 |  | IT WAS REPORTED BY THE CUSTOMER THAT PULLED A BARD PICC CATHETER CUSTOMER PLACED WITH A TEAR AT ABOUT THE 5CM MARK AFTER 1 WEEK OF USE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04327 | 2023-06-14 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN THE FOLEY CATHETER BALLOON WAS CHECKED, STERILE WATER WAS DIFFICULT TO INJECT AND DIFFICULT TO DRAIN WATER FORM THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04330 | 2023-06-14 | 2023-04-25 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED, THAT THE PATIENT UNDERWENT RESECTION OF THE PROSTATE AT 14:20 ON (B)(6)2023. AFTER THE OPERATION, A DISPOSABLE STERILE FOLEY CATHETER WAS USED FOR INDWELLING CATHETERIZATION. AT 22:06 ON THE SAME DAY, THE CATHETER WAS FOUND TO HAVE SLIPPED OUT BY ITSELF. AND THE CATHETER WAS CHECKED. IT WAS FOUND THAT THE BALLOON WAS RUPTURED, BUT THERE WERE NO MISSING PIECES. AND REPORTED TO THE DOCTOR. THE USER IMMEDIATELY REPLACED THE CATHETER WITH A NEW CATHETER AGAIN, AND THE URINE COLOR WAS CLEAR. STATED, THAT THE CATHETER REMOVAL AND RE-CANNULATION CAUSED SERIOUS INJURY TO THE PATIENT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023 STATED, THAT THERE WAS NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE, BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN, WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN, WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT, EXPOSURE TO PETROLATUM BASED PRODUCTS, MECHANICAL FAILURE, OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN. THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "A REVIEW OF THE DEVICE LABELLING HAVE BEEN CONDUCTED FOR INVESTIGATION PURPOSED. STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS, SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE. AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE. AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02441 | 2023-06-14 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT HE PATIENT NEEDS CHEMOTHERAPY AFTER PICC CATHETERIZATION, AND THE END OF THE GUIDE WIRE IS FOUND TO BE ROUGH DURING THE CATHETERIZATION. IF THE GUIDE WIRE IS USED, IT MAY CAUSE DAMAGE TO THE PICC CATHETER AND AFFECT THE INDWELLING TIME OF THE CATHETER. IF IT IS SERIOUS, THE PICC CATHETER CAN BE BROKEN AND CANNOT BE USED NORMALLY, AND THEN REPUNCTURE AND CATHETERIZATION ARE REQUIRED, WITH EXTREMELY BAD CONSEQUENCES. CONSIDERING THE SERIOUS CONSEQUENCES, THE GUIDE WIRE SHOULD BE IMMEDIATELY ABANDONED AND REPLACED WITH A NEW GUIDE WIRE TO COMPLETE THE FOLLOW-UP OPERATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DEFORMATION ON THE END OF THE GUIDEWIRE WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. TWO PHOTOGRAPHS OF THE IMPLICATED GUIDEWIRE WERE RETURNED FOR EVALUATION. THE GUIDEWIRE CONTAINED A REGION OF DEFORMATION NEAR THE END OF THE WIRE. THIS END OF THE GUIDEWIRE APPEARED ROUGH AND IRREGULAR; WITH WHAT APPEARED TO BE A SLIGHTLY LARGER OUTER DIAMETER COMPARED TO THE REST OF THE WIRE. THE PHOTO QUALITY DID NOT ALLOW FOR A CLEAR VIEW OF THE DAMAGED WIRE SECTION; HOWEVER, THE IRREGULAR PROFILE OF THE END OF THE WIRE WAS LIKELY CAUSED BY THE OUTER WIRE BEING DISPLACED. ALTHOUGH THE DAMAGE TO THE WIRE COULD BE CONFIRMED, THE EXACT ROOT CAUSE OF THE DAMAGE COULD NOT BE DETERMINED FROM THE RETURNED PHOTOGRAPHS. A REVIEW OF THE MANUFACTURE QUALITY RECORDS FOR THE SPECIFIED LOT FOUND THAT NO POTENTIALLY RELATED ISSUES WERE ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. POTENTIAL CONTRIBUTING FACTORS COULD INCLUDE ATTEMPTED INSERTION OF THE WIRE AGAINST RESISTANCE OR MANIPULATION OF THE WIRE CAUSING DISPLACEMENT OF THE OUTER WIRE. THE REPORTED EVENT HAS BEEN DOCUMENTED AND INCLUDED IN COMPLAINT TRENDING AND WILL CONTINUE TO BE MONITORED AS PART OF ONGOING QUALITY EFFORTS. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04324 | 2023-06-14 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 995081 | 995081 |  | IT WAS REPORTED THAT IN AN ONLINE SURVEY A PHYSICIAN STATED NO TO THE QUESTION WHEN THEY WERE ASKED IF THEY WERE ABLE TO USE THE REFOLDING TOOL TO REFOLD THE BALLOON DILATION CATHETER FOR REINSERTION. THE PHYSICIAN WAS NOT ABLE TO USE THE REFOLDING TOOL TO REFOLD THE 995081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, STERILE BECAUSE PHYSICIAN STATED THEY WERE NOT SURE. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE IMPROPER MATERIAL SELECTION. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿DESCRIPTION: THE BALLOON DILATION CATHETER COMES PACKAGED WITH A REFOLDING TOOL AND A WORKING SHEATH. IT IS AVAILABLE WITH OR WITHOUT AN INFLATION DEVICE. IT COMES STERILE AND IS FOR SINGLE USE ONLY. USING THE REFOLDING TOOL: THE BALLOON CATHETER IS SUPPLIED WITH A REFOLDING TOOL. THIS AIDS IN PREPARING THE BALLOON FOR ANOTHER INSERTION DURING THE PROCEDURE. TO USE: MANUALLY COMPRESS THE BALLOON. POSITION THE REFOLDING TOOL AT ONE END OF THE BALLOON. TWIST THE REFOLDING TOOL COUNTERCLOCKWISE AND PUSH DOWN ON THE BALLOON UNTIL THE REFOLDING TOOL TRAVERSES THE ENTIRE LENGTH OF THE BALLOON. 4. ONCE THE BALLOON IS FOLDED, REMOVE THE REFOLDING TOOL AND PROCEED WITH THE PROCEDURE. DISPOSE OF PROPERLY.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04325 | 2023-06-14 | 2023-05-24 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD 5 URINARY TRACT INFECTION (UTI) WHILE USING THESE MAGIC 3 GO INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. PER FOLLOW UP VIA PHONE ON (B)(6)2023, IT WAS REPORTED THAT THE PATIENT WAS STILL USING THE CATHETERS BUT DID REGULAR VISITS FOR THE URINARY TRACT INFECTIONS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04323 | 2023-06-14 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS INFLATED, BUT THE STERILE WATER COULD NOT BE REMOVED DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04321 | 2023-06-14 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 996081 | 996081 |  | IT WAS REPORTED THAT IN AN ONLINE SURVEY A PHYSICIAN STATED NO WHEN THEY WERE ASKED IF THEY WERE ABLE TO SUCCESSFULLY DILATE THE NEPHROSTOMY TRACT. THE PHYSICIAN WAS NOT ABLE TO SUCCESSFULLY DILATE THE NEPHROSTOMY TRACT FOR THE DEVICE 996081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, W/ INFLATION DEVICE, STERILE BECAUSE PHYSICIAN STATED NONE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04320 | 2023-06-14 | 2023-05-26 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD A VERY BAD URINARY TRACT INFECTION WHILE USING THE MAGIC 3 GO MALE COUDE INTERMITTENT CATHETER AND WAS HOSPITALIZED FOR 6 DAYS AROUND (B)(6). IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. AS PER FOLLOW UP VIA PHONE ON 07JUN2023, IT WAS REPORTED THAT THE CUSTOMER DID DEVELOP A URINARY TRACT INFECTION BUT DID NOT RECEIVE CONFIRMATION THAT THE CATHETERS WERE THE CAUSE. THEY WERE PRESCRIBED ANTIBIOTICS AND SPENT 6 DAYS IN THE HOSPITAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04319 | 2023-06-14 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 | FOL0100 | FOL0100 |  | IT WAS REPORTED THAT THE FOLEY STATLOCK BROKE AFTER USING IT FOR 2 DAYS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02434 | 2023-06-14 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 3.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668935 |  | IT WAS REPORTED THAT UPON OPENING OF THE MICRO-INTRODUCER (0668935), THE GUIDE WIRE SCATTERED, UNABLE TO PASS THROUGH THE INTRODUCER NEEDLE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04329 | 2023-06-14 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER WAS DIFFICULT TO REMOVE FROM THE FOLEY CATHETER BALLOON DURING A PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04314 | 2023-06-14 | 2023-05-24 | I, n, j, u, r, y | Manufacturer report | BARD® URETHRAL CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 120614 | 120614 |  | IT WAS REPORTED THAT THE PATIENT EXPERIENCED BLEEDING AND INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02431 | 2023-06-14 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT TI L/P 6 CF INT WOSP ATT SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8716001 |  | IT WAS REPORTED THAT PRIOR TO A PORT PLACEMENT PROCEDURE VIA THE RIGHT INTERNAL JUGULAR VEIN, THE CATHETER ALLEGEDLY FAILED TO ASPIRATE AND DIFFICULT TO BE FLUSHED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT MRI IMPLANTABLE PORT KIT WAS RECEIVED FOR EVALUATION. VISUAL, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. UPON FUNCTIONAL EVALUATION, AN ATTEMPT TO INFUSE THE CATHETER WITH THE RETURNED FLUSHING CONNECTOR WAS PERFORMED AND THE CATHETER WAS PATENT TO BOTH INFUSION AND ASPIRATION WITHOUT ISSUE. ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS THE POWERPORT KIT IN A PACKAGE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO FLUSH AND SUCTION ISSUES AS THE EXACT CIRCUMSTANCE AT THE TIME OF THE EVENT REPORTED WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 03/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02433 | 2023-06-14 | 2021-12-01 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY FOUR YEARS AND THREE MONTHS POST PORT PLACEMENT PROCEDURE, CATHETER ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT PORT BODY AND CATHETER SEGMENTS WERE REMOVED PERCUTANEOUSLY. REPORTEDLY THE CATHETER IS ADHERING TO THE PATIENT'S BODY. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2018). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02426 | 2023-06-14 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F DL FT KIT W/GUARDIVA & PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1295108FD2 |  | IT WAS REPORTED BY CUSTOMER ON (B)(6) 2023 "LEAKING AROUND WIRE AND RUBBER PORT". NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02428 | 2023-06-14 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT MRI ISP W/O SUTURE PLUGS, 9.6 FR SILICONE SL CATHETER | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1809661 | 1809661 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE PORT ALLEGEDLY BROKE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02427 | 2023-06-14 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THE CUSTOMER REVEALED THAT AFTER INSERTION OF CATHETER ON MAY 26, 2023, ATTACH THE LUER FITTING ON ONE LUMEN OF THE CATHETERS WITH A NEEDLE-FREE CONNECTOR, BUT WATER LEAKED. WATER LEAKAGE STILL OCCURRED AFTER MULTIPLE REPLACEMENTS OF THE CONNECTORS. THE EXACT NUMBER OF REPLACEMENTS WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02429 | 2023-06-14 | 2023-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED UPON RECEIPT OF THE FEEDBACK FROM THE CUSTOMER FOR NEEDLESS CONNECTOR LEAK, WE GO TO THE OPERATING SITE TO OBSERVE THE OPERATION OF THE CATHETER PLACEMENT OPERATOR. THERE IS NO ISSUE RELATED TO IMPROPER PLACEMENT OF SHARP DEVICES. PROVIDE THEM WITH TWO MZ CONNECTORS FROM OUR COMPANY FOR REPLACEMENT, BUT WATER LEAKAGE STILL OCCURRED. NO ADVERSE EVENTS REPORTED. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02430 | 2023-06-14 | 2023-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED UPON RECEIPT OF THE FEEDBACK FROM THE CUSTOMER FOR NEEDLESS CONNECTOR LEAK, WE GO TO THE OPERATING SITE TO OBSERVE THE OPERATION OF THE CATHETER PLACEMENT OPERATOR. THERE IS NO ISSUE RELATED TO IMPROPER PLACEMENT OF SHARP DEVICES. PROVIDE THEM WITH TWO MZ CONNECTORS FROM OUR COMPANY FOR REPLACEMENT, BUT WATER LEAKAGE STILL OCCURRED. NO ADVERSE EVENTS REPORTED. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04309 | 2023-06-14 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 39301 | 39301 |  | IT WAS REPORTED, THAT THESE MALE EXTERNAL CATHETERS. CAUSED SORES ON TIP AND SHAFT OF PENIS. THE PATIENT STATED, THAT THE ADHESIVE REMOVED THE SKIN OFF. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02432 | 2023-06-14 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5613200J |  | IT WAS REPORTED THAT THE PART OF THE BLUE LURE CONNECTOR WAS DETACHED. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04310 | 2023-06-14 | 2023-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PATIENT COMPLAINED OF ABDOMINAL DISTENSION THREE HOURS AFTER THE OPERATION WITH THE CATHETER NOT IN PLACE SMOOTHLY AND ONLY A SMALL AMOUNT OF BLOODY FLUID WAS DRAINED AS THE DRAINAGE SPEED WAS NOT EQUAL TO THE IRRIGATION RATE. UPON CHECKING THE CATHETER, IT WAS FOUND BURST AND THEN WAS REPLACED WITH A NEW CATHETER. NO MEDICAL INTERVENTION WAS REPORTED. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, CLARIFIED THAT THERE WAS RESTRICTED FLOW OF FLUID AS IT WAS MENTIONED AS ONLY A FEW AMOUNT OF BLOODY FLUID WAS DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04315 | 2023-06-14 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT EVERY ONCE IN A WHILE, PATIENT GOT SOME THAT MAGIC 3 COUDE INTERMITTENT CATHETER DID NOT WORK. IT DID NOT HAPPEN A LOT. THE PATIENT HAD ISSUES WITH INSERTING THEM INTO THE PENIS THROUGH THE URETHRA BUT NOT INTO THE BLADDER AND HAD AN ISSUE WITH 6 CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED.H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04313 | 2023-06-14 | 2023-05-24 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53812G | 53812G |  | IT WAS REPORTED THAT THE PATIENT ENDED UP GETTING URINARY TRACT INFECTION. THE CAREGIVER DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02424 | 2023-06-14 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2174108 |  | IT WAS REPORTED BROKEN CATHETER 20 CM FROM EXIT-SITE. WET DRESSING DETECTED DURING INFUSION. REMOVED CATHETER FOUND FISSURE 20 CM FROM EXIT-SITE. CATHETER SECURED WITH SECUR-A-CATH AND GLUE ON EXIT-SITE. INCIDENT OCCURRED DURING USE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04311 | 2023-06-14 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® URETHRAL CATHETER TRAY, CLEAR PLASTIC CATHETER | URETHRAL CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | 802214 | 802214 |  | IT WAS REPORTED THAT THE CUSTOMER HAD MISSED LOOKING AT THE NEW CATHETER KITS COMING IN, THE GLOVES INCLUDED IN THE KIT WERE NOT SET UP FOR STERILE USE AT ALL. ALSO STATED THAT THESE WOULD NOT SUFFICE FOR THE NEEDS OF STRAIGHT CATHETER WITHOUT ADDITIONAL SUPPLIES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02455 | 2023-06-15 | 2023-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F DL FT MAX BARRIER KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1295108FD2 |  | IT WAS REPORTED BY CUSTOMER THAT INTRODUCER WOULD NOT STAY TIGHT, HAD TO DROP IN MICROINTRODUCER. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02457 | 2023-06-15 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THE CUSTOMER REVEALED THAT AFTER INSERTION OF CATHETER ON (B)(6) 2023, ATTACH THE LEUR FITTING ON ONE LUMEN OF THE CATHETERS WITH A NEEDLE-FREE CONNECTOR, BUT WATER LEAKED. WATER LEAKAGE STILL OCCURRED AFTER MULTIPLE REPLACEMENTS OF THE CONNECTORS. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02456 | 2023-06-15 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D4 |  | IT WAS REPORTED BY THE CUSTOMER "POST PICC PROCEDURE WIRE WAS INSPECTED TO HAVE A SIGNIFICANT KINK TOWARD THE LOWER THIRD JUST ABOVE THE TECHNOLOGY PORTION OF THE WIRE. NO DAMAGE WAS NOTICED PRE PROCEDURE AND NOTHING WAS NOTED DURING PLACEMENT TO HAVE CAUSED THE KINK OR FEEL THE WIRE HAD A PROBLEM. TECHNOLOGY WORKED AS INTENDED TO CONFIRM VIA ECG." ADDITIONAL INFORMATION RECEIVED 5 JUNE 2023: IT WAS REPORTED BY THE CUSTOMER "NO HARM, INJURY, COMPLICATION, OR NEGATIVE OUTCOME OCCURRED. AFTER FOUND OUT THE BEND PICC WIRE, X-RAY WAS REQUESTED TO MAKE SURE THE TIP WAS IN THE APPROPRIATE LOCATION." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04379 | 2023-06-15 | 2023-05-25 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MAGIC3 GO MALE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAS HAD URINARY TRACT INFECTION WHILE USING THE MAGIC 3 GO MALE INTERMITTENT CATHETER AND WAS UNSURE WHY. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02452 | 2023-06-15 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 3174118 |  | IT WAS REPORTED THAT THIS 13-YEAR-OLD PATIENT NEEDS CHEMOTHERAPY DUE TO GERM-CELL TUMOR AND HAD A POWER PICC CATHETER PLACED FROM THE LEFT CEPHALIC VEIN ON FEBRUARY 23. DURING CHEMOTHERAPY ON MAY 29, LIQUID LEAKAGE OCCURRED. CT SHOWED THAT THE CATHETER TIP WAS POSITIONED PROPERLY, BUT SEVERE LIQUID LEAKAGE OCCURRED. THEREFORE, THE CATHETER WAS REMOVED AND FOUND DAMAGED AT THE SCALE OF 14 CM AND 16 CM. THE DAMAGE AT THE 14-CM SCALE WAS SEVER. CATHETER LEAKAGE, IMPOSSIBLE TO PERFORM CHEMOTHERAPY, LEADING TO DELAY IN TREATMENT CYCLE. SUBSEQUENT THERAPY WAS GIVEN WITH ANOTHER INFUSION TOOL. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02451 | 2023-06-15 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | IT WAS REPORTED THAT UPON UNPACKING, THE VASCULAR SHEATH WAS FOUND SPLIT PERIPHERALLY, UNUSABLE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02450 | 2023-06-15 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PPICC 4F SL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000991 |  | IT WAS REPORTED "RN INSERTED PICC - THEY MET SOME RESISTANCE THEN THE PICC WAS ABLE TO BE INSERTED WITHOUT DIFFICULTY AFTER RE-DIRECTING PICC. WHEN THE RN WAS REMOVING THE 3CG WIRE THE PICC TIP KINKED AND ALMOST DISCONNECTED WHILE IT WAS STILL IN THE PATIENT." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02453 | 2023-06-15 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED VIA MEDWATCH "BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET REFERENCE # S1275108FD5, LOT # REHN2437 & REGY0842 NOTED LEAKAGE WITH FLUSHING AT RUBBER-LIKE STOPPER. REHN ULTIMATELY SUCCESSFULLY PLACED. STYLETS RETAINED." THIS FILE ADDRESSES LOT# REGY0842. ADDITIONAL INFORMATION RECEIVED 6/3/2023: IT WAS REPORTED THERE WERE NO VISIBLE CRACKS OR HOLES THAT CAUSED LEAKAGE. NO HARM NOTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02454 | 2023-06-15 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1295108FD2 |  | IT WAS REPORTED BY THE CUSTOMER "PICC LEAKING AROUND WIRE AND RUBBER PORT." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02466 | 2023-06-15 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER ¿PRESENCE OF BARBS. ON (B)(6) 2023, FIVE PATIENTS WITH LUNG TUMOR NEEDED TO HAVE A PICC CATHETER PLACED TO INFUSE CHEMOTHERAPY DRUGS. THE CATHETERS WERE PLACED AT CATHETER PLACEMENT ROOM OF THE TUMOR DEPARTMENT. THE PUNCTURE WAS CONDUCTED WITH THE GUIDANCE OF ULTRASOUND. AFTER THE PUNCTURE SUCCEEDED, THE GUIDE WIRE WAS INTRODUCED. AT THE TIME OF REMOVING THE INTRODUCER NEEDLE, BARBS WERE NOTED AT THE DISTAL END OF THE GUIDE WIRE. THE PRESENCE OF THE BARBS AFFECTED THE REMOVAL OF THE INTRODUCER NEEDLE AND THE INTRODUCTION OF THE INTRODUCER SHEATH. AFTER THE OPERATOR CUT OFF THE BARBS WITH ASEPTIC SCISSORS, THE INTRODUCER NEEDLE WAS REMOVED SUCCESSFULLY. THIS CAUSED TENSION TO THE PATIENT, INCREASED UNNECESSARY WORKLOAD TO THE OPERATOR, AND CREATED UNFORESEEABLE RISKS FOR CATHETER PLACEMENT.¿ THIS REPORT ADDRESSES THE FOURTH PATIENT AND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02459 | 2023-06-15 | 2023-02-21 | I, n, j, u, r, y | Manufacturer report | BD 20G 8CM POWERGLIDE | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "A 20G 8CM POWER GLIDE PRODUCT UTILIZED FOR MIDLINE INSERTION IN THE PATIENT'S LEFT BASILIC VEIN UTILIZING ULTRASOUND FOR PLACEMENT. THE EQUIPMENT WAS CHECKED BY THE CLINICIAN PRIOR TO USE. ACCESS WAS OBTAINED AND A FLASH OF BLOOD WAS IDENTIFIED IN THE MIDLINE CATHETER; THE WIRE WAS ADVANCED FOLLOWED BY THE CATHETER BEING ADVANCED OVER THE WIRE. UPON ADVANCEMENT OF THE CATHETER, RESISTANCE WAS APPRECIATED WITH APPROXIMATELY 3CM STILL DISLODGED FROM THE INSERTION SITE. ONCE THE RESISTANCE WAS MET, THE ENTIRE DEVICE (CATHETER AND WIRE) WAS ATTEMPTED TO BE REMOVED FROM THE EXTREMITY. RESISTANCE WAS AGAIN MET, AND THE ENTIRE DEVICE WAS UNABLE TO SUCCESSFULLY BE REMOVED. AN X-RAY AND ULTRASOUND WAS COMPLETED OF THE EXTREMITY CONFIRMING A RETAINED FOREIGN BODY. THE PATIENT REQUIRED GENERAL ANESTHESIA FOR THE IR PROCEDURE TO RETRIEVE THE RETAINED PORTION OF THE CATHETER AND GUIDEWIRE. " | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02460 | 2023-06-15 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F DL FT KIT W/GUARDIVA & PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1295108FD2 |  | IT WAS REPORTED BY CUSTOMER ON 5/21/23 "LEAKING AROUND WIRE AND RUBBER PORT". NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02463 | 2023-06-15 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER ¿PRESENCE OF BARBS. ON (B)(6) 2023, FIVE PATIENTS WITH LUNG TUMOR NEEDED TO HAVE A PICC CATHETER PLACED TO INFUSE CHEMOTHERAPY DRUGS. THE CATHETERS WERE PLACED AT CATHETER PLACEMENT ROOM OF THE TUMOR DEPARTMENT. THE PUNCTURE WAS CONDUCTED WITH THE GUIDANCE OF ULTRASOUND. AFTER THE PUNCTURE SUCCEEDED, THE GUIDE WIRE WAS INTRODUCED. AT THE TIME OF REMOVING THE INTRODUCER NEEDLE, BARBS WERE NOTED AT THE DISTAL END OF THE GUIDE WIRE. THE PRESENCE OF THE BARBS AFFECTED THE REMOVAL OF THE INTRODUCER NEEDLE AND THE INTRODUCTION OF THE INTRODUCER SHEATH. AFTER THE OPERATOR CUT OFF THE BARBS WITH ASEPTIC SCISSORS, THE INTRODUCER NEEDLE WAS REMOVED SUCCESSFULLY. THIS CAUSED TENSION TO THE PATIENT, INCREASED UNNECESSARY WORKLOAD TO THE OPERATOR, AND CREATED UNFORESEEABLE RISKS FOR CATHETER PLACEMENT.¿ THIS REPORT ADDRESSES THE FIRST PATIENT AND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02465 | 2023-06-15 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER ¿PRESENCE OF BARBS. ON MAY 30, 2023, FIVE PATIENTS WITH LUNG TUMOR NEEDED TO HAVE A PICC CATHETER PLACED TO INFUSE CHEMOTHERAPY DRUGS. THE CATHETERS WERE PLACED AT CATHETER PLACEMENT ROOM OF THE TUMOR DEPARTMENT. THE PUNCTURE WAS CONDUCTED WITH THE GUIDANCE OF ULTRASOUND. AFTER THE PUNCTURE SUCCEEDED, THE GUIDE WIRE WAS INTRODUCED. AT THE TIME OF REMOVING THE INTRODUCER NEEDLE, BARBS WERE NOTED AT THE DISTAL END OF THE GUIDE WIRE. THE PRESENCE OF THE BARBS AFFECTED THE REMOVAL OF THE INTRODUCER NEEDLE AND THE INTRODUCTION OF THE INTRODUCER SHEATH. AFTER THE OPERATOR CUT OFF THE BARBS WITH ASEPTIC SCISSORS, THE INTRODUCER NEEDLE WAS REMOVED SUCCESSFULLY. THIS CAUSED TENSION TO THE PATIENT, INCREASED UNNECESSARY WORKLOAD TO THE OPERATOR, AND CREATED UNFORESEEABLE RISKS FOR CATHETER PLACEMENT.¿ THIS REPORT ADDRESSES THE THIRD PATIENT AND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02464 | 2023-06-15 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER ¿PRESENCE OF BARBS. ON (B)(6) 2023, FIVE PATIENTS WITH LUNG TUMOR NEEDED TO HAVE A PICC CATHETER PLACED TO INFUSE CHEMOTHERAPY DRUGS. THE CATHETERS WERE PLACED AT CATHETER PLACEMENT ROOM OF THE TUMOR DEPARTMENT. THE PUNCTURE WAS CONDUCTED WITH THE GUIDANCE OF ULTRASOUND. AFTER THE PUNCTURE SUCCEEDED, THE GUIDE WIRE WAS INTRODUCED. AT THE TIME OF REMOVING THE INTRODUCER NEEDLE, BARBS WERE NOTED AT THE DISTAL END OF THE GUIDE WIRE. THE PRESENCE OF THE BARBS AFFECTED THE REMOVAL OF THE INTRODUCER NEEDLE AND THE INTRODUCTION OF THE INTRODUCER SHEATH. AFTER THE OPERATOR CUT OFF THE BARBS WITH ASEPTIC SCISSORS, THE INTRODUCER NEEDLE WAS REMOVED SUCCESSFULLY. THIS CAUSED TENSION TO THE PATIENT, INCREASED UNNECESSARY WORKLOAD TO THE OPERATOR, AND CREATED UNFORESEEABLE RISKS FOR CATHETER PLACEMENT.¿ THIS REPORT ADDRESSES THE SECOND PATIENT AND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02462 | 2023-06-15 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERLOC SAFETY INFUSION SET 20G X 0.75IN WITH Y-SITE | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0672034J |  | IT WAS REPORTED THAT NO LEAKAGE WAS OBSERVED DURING INFUSION. AFTER INFUSION, WHEN FLUSH WAS PERFORMED, LEAKAGE OCCURRED. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CRACKED Y-SITE WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE SUPPLIER RELATED. THE SUPPLIER HAS BEEN NOTIFIED OF THIS EVENT. BD IS WORKING CLOSELY WITH THE SUPPLIER TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02467 | 2023-06-15 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER ¿PRESENCE OF BARBS. ON (B)(6) 2023, FIVE PATIENTS WITH LUNG TUMOR NEEDED TO HAVE A PICC CATHETER PLACED TO INFUSE CHEMOTHERAPY DRUGS. THE CATHETERS WERE PLACED AT CATHETER PLACEMENT ROOM OF THE TUMOR DEPARTMENT. THE PUNCTURE WAS CONDUCTED WITH THE GUIDANCE OF ULTRASOUND. AFTER THE PUNCTURE SUCCEEDED, THE GUIDE WIRE WAS INTRODUCED. AT THE TIME OF REMOVING THE INTRODUCER NEEDLE, BARBS WERE NOTED AT THE DISTAL END OF THE GUIDE WIRE. THE PRESENCE OF THE BARBS AFFECTED THE REMOVAL OF THE INTRODUCER NEEDLE AND THE INTRODUCTION OF THE INTRODUCER SHEATH. AFTER THE OPERATOR CUT OFF THE BARBS WITH ASEPTIC SCISSORS, THE INTRODUCER NEEDLE WAS REMOVED SUCCESSFULLY. THIS CAUSED TENSION TO THE PATIENT, INCREASED UNNECESSARY WORKLOAD TO THE OPERATOR, AND CREATED UNFORESEEABLE RISKS FOR CATHETER PLACEMENT.¿ THIS REPORT ADDRESSES THE FIFTH PATIENT AND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
MW5118432 | 2023-06-15 | 2023-06-13 | I, n, j, u, r, y | Voluntary report | POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | BARD ACCESS SYSTEMS, INC. |  | 1808560 |  | POWER PORT, ORIGINALLY INSERTED IN (B)(6) 2020, HAD SEVERAL PROBLEMS OVER THREE YEARS INCLUDING PAIN WITH INJECTION. PATIENT HAD "PORT STUDY" COMPLETED SHOWING "LEAKING" AT THE DISTAL END OF CATHETER. PORT REMOVED (B)(6) 2023. UPON INSPECTION OF REMOVED PORT SEVERAL CRACKS WERE SEEN AT DISTAL END OF CATHETER. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04378 | 2023-06-15 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT WATER WAS UNABLE TO INJECT THROUGH THE INFLATION VALVE OF THE FOLEY CATHETER. IT WAS STATED THAT THE UPON CHECKING THE BALLOON INFLATION BEFORE THE CATHETER INSERTION, DISTILLED WATER COULD NOT BE INJECTED. IT WAS ALSO STATED THAT THE POSSIBLE VALVE FAILURE. IT WAS NOTED THAT THE POSSIBLE VALVE FAILURE HAD BEEN SUGGESTED AS THE CAUSE OF THE INABILITY TO WATER INJECTION IN THE JAPANESE DESCRIPTION, BUT NO COMPLAINTS OF VALVE FAILURE HAVE BEEN REPORTED. IT WAS ALSO NOTED THAT THIS COMPLAINT WAS JUST AGAINST THE INABILITY TO WATER INJECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02458 | 2023-06-15 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5 F POLYURETHANE CATHETER NURSING BASIC TRAY WITH MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3175118 |  | CUSTOMER REPORTED THAT THE PATIENT WITH PICC CATHETER IMPLANTATION, IMPLANTATION DATE (B)(6)2023. FAMILY REPORTS LEAKAGE SINCE CATHETER IMPLANTATION. WHEN VISITING, THE CATHETER PRESENTS A LEAK AT THE EDGE, REFLUXING ALL INFUSED MEDICATION. CATHETER DISCARDED AND REPLACED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04377 | 2023-06-15 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. IT WAS STATED THAT THERE WAS NO PROBLEM DURING A PRETEST BEFORE THE CATHETER INSERTION TO THE PATIENT IN THE OPERATING ROOM. THE DAY AFTER THE CATHETER WAS PLACED (24 HOURS LATER), THE CATHETER FELL OFF WHEN THE PATIENT WAS MOVED IN A SIDE LYING POSITION DURING A PUBLIC LAVAGE. IT WAS ALSO STATED THAT THE BALLOON WAS DEFLATED AT THAT TIME. THE CATHETER WAS FILLED WITH WATER AGAIN TO SEE IF THE BALLOON WAS DAMAGED, BUT NO WATER LEAKAGE COULD BE CONFIRMED. IT WAS NOTED THAT THE UPON THE SAMPLE RETRIEVAL, THE BALLOON REMAINED INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02449 | 2023-06-15 | 2023-03-24 | I, n, j, u, r, y | Manufacturer report | POWERPICC 4F SL, 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000746 |  | IT WAS REPORTED VIA MEDWATCH ¿DURING PICC (PERIPHERALLY INSERTED CENTRAL CATHETER) LINE PLACEMENT, STAFF REPORTED THE FOLLOWING¿ THE VEIN WAS ACCESSED AND THEN THE WIRE WAS INSERTED THROUGH THE NEEDLE PER PROCEDURE. THE WIRE DID NOT THREAD PROPERLY. DIFFICULTY MET WHEN ATTEMPTING TO REMOVE THE WIRE. STAFF NOTED THE NEEDLE DISLODGED AND THE WIRE REMAINED IN THE TISSUE. WIRE WAS RESISTANT TO REMOVAL AND DURING REMOVAL, IT BEGAN TO UNBRAID. X-RAYS WERE OBTAINED TO CONFIRM THERE WAS NO RETAINED FOREIGN BODY AS A RESULT OF THE WIRE UNRAVELING.¿ | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04375 | 2023-06-15 | 2023-05-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER OF LOT# NGGY5874 WAS INSERTED INTO PATIENT, WHEN PATIENT WENT HOME THE CATHETER HAD FALLEN OUT AS THE BALLOON DEFLATED AND UPON INSPECTION A SMALL PINHOLE WAS DISCOVERED. PATIENT HAD TO RETURN TO THE EMERGENCY ROOM. PER FOLLOW UP VIA PHONE ON 17MAY2023, IT WAS REPORTED THAT THE CUSTOMER HAD AN ADDITIONAL DEFECTIVE SAMPLE THAT THEY WOULD RETURN FOR EVALUATION. PER FOLLOW UP VIA PHONE ON 22MAY2023, IT WAS REPORTED THAT THEY SENT BACK SIX CATHETERS TO THE MANUFACTURER. PER SAMPLE FORM RECEIVED ON 25MAY2023, IT WAS STATED THAT ON (B)(6) 2023, THE PATIENT HAD A FOLEY IN PLACE AND WAS DISCHARGED FROM THE EMERGENCY ROOM. WHILE THE PATIENT WAS AT HOME THE FOLEY FELL OUT. WHEN THEY TROUBLESHOOTED IT, THEY NOTICED A PINHOLE IN THE BALLOON WHERE THE SALINE WAS LEAKING. THEY PUT A NEW ONE AND SENT THEM HOME. THE PATIENT RETURNED ON (B)(6) 2023 FOR THE SAME PROBLEM, AGAIN STAFF TROUBLESHOOTED THE BALLOON AND NOTICED A PIN SIZE HOLE AND THIRD FOLEY WAS INSERTED. BOTH TIMES THEY INFLATED THE BALLOON PRIOR TO THE INSERTION THEY FOUND NO ISSUES, THEN AFTER IT WAS INSERTED, THE HOLE APPEARED. ALSO STATED THAT THE 2ND FOLEY WAS THROWN AWAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04376 | 2023-06-15 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT WATER IN THE BALLOON COULD NOT BE REMOVED. THE USER ATTEMPTED TO DRAIN THE WATER BY PUMPING ACTION, BUT THERE WAS NO RESPONSE AND THE CATHETER WAS LEFT IN PLACE FOR SEVERAL HOURS WITH THE SYRINGE ATTACHED AND EVENTUALLY THE WATER WAS REMOVED. ACCORDING TO THE INQUIRY SHEET PROVIDED, ABOUT 10 ML OF SALINE WAS ASPIRATED BY THE SYRINGE. THE BALLOON WAS TRIED TO DEFLATE BY ASPIRATING MANUALLY WITH THE SYRINGE AND GENTLE DEFLATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02444 | 2023-06-15 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | A710962 | A710962 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, SHEATH ALLEGEDLY BROKE WHEN THE INTRODUCER WAS UNSCREWED AND REMOVED. REPORTEDLY PORT WAS PLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 05/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02445 | 2023-06-15 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT ONE DAY POST PORT PLACEMENT PROCEDURE, A LITTLE BLOOD WAS ALLEGEDLY OOZING FROM THE INSERTION SITE. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING MEDICATION OR TREATMENT FOR BLEEDING. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04333 | 2023-06-15 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT ITEM HAD A HOLE ON THE BOW PART OF THE FOLEY CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO OPERATOR ERROR, MECHANICAL FAILURE, USER RELATED (EG: USE OF SHARP OBJECT). A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: 1) SECURE THE FOLEY CATHETER. USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED. 2) MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. 3) MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. 4) KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. 5) EMPTY THE COLLECTION BAG REGULARLY USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH. FOLEY CATHETER REMOVAL: 1) TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. 2) ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. 3) USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. 4) IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. 5) SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04334 | 2023-06-15 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT ON (B)(6),2023, A THREE-LUMEN URINARY CATHETER WAS INDWELLED AFTER THE OPERATION FOR DRAINAGE AND COMPRESSION TO STOP BLEEDING. ON 30TH MAY, IT WAS FOUND THAT THE BALLOON OF THE URINARY CATHETER WAS RUPTURED AND THE CATHETER FELL OFF. THE ISSUE WAS DISCOVERED ON TIME AND A NEW CATHETER WAS REPLACED IMMEDIATELY. THE ISSUE DID NOT OCCUR AGAIN AND THERE WAS NO ADVERSE IMPACT CAUSED TO THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04332 | 2023-06-15 | 2023-05-24 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION WHILE USING CATHETERS AND BLEEDING. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04335 | 2023-06-15 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS INFLATED DURING TESTING, BUT AFTER INFLATION, THE BALLOON WAS DEFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04338 | 2023-06-15 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | BFF5 | UNK |  | IT WAS REPORTED THAT THE NURSE COMPLAINED THAT THE FOLEY CATHETER SLIPPED OUT AUTOMATICALLY. THE NURSE WENT TO THE BEDSIDE TO CHECK AND FOUND THAT THE BLADDER OF THE FOLEY CATHETER WAS DEFLATED, AND CHECKED THAT THE PATIENT'S URETHRAL OPENING WAS NOT DAMAGED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO BUBBLE VOID ON RUBBERIZE LAYER THAT MAY LEAD TO ISSUE LUMEN TO LUMEN LEAKAGE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04336 | 2023-06-15 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER WAS DIFFICULT TO BE REMOVED FROM THE FOLEY CATHETER DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04345 | 2023-06-15 | 2023-05-22 | I, n, j, u, r, y | Manufacturer report | INTERMITTENT ROCHESTER CATHETERS | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50812G | UNK |  | IT WAS REPORTED THAT THESE INTERMITTENT CATHETERS CAUSED PATIENT AN INFECTION AND WAS HOSPITALIZED AND CAUGHT TETANUS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR INFECTION AND TETANUS. PER FOLLOW UP VIA PHONE ON 22MAY2023, IT WAS REPORTED THAT THE CUSTOMER RECEIVED A VOLUNTARY EMERGENCY NOTICE FROM BD REGARDING A RECALL FOR A NONCONFORMITY WITH THE MAGIC GO 14 FR CATHETERS. THEY WERE REPORTED IN A COMPANY-WIDE NOTICE AS HAVING NONSTERILE COUDE, AND THEY HAPPENED TO BE ONE OF THOSE IMPACTED BY IT. AT THE TIME, THEY WERE SENT OUT FROM 180 MEDICAL. THIS LED TO NUMEROUS BLADDER INFECTIONS AND SEPSIS, AND THIS MADE THINGS DIFFICULT FOR THEM AS THEY COULD NOT WALK. THEY STILL HAD SAMPLES AVAILABLE AND PROVIDED THE LOT# JUGW1445. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-06
2020394-2023-00429 | 2023-06-15 | 2021-12-21 | D, e, a, t, h | Manufacturer report | ULTRAVERSE 018 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U820084 | U820084 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT A BALLOON EXPANDABLE STENT WAS PLACED IN A PATIENT AFTER BEING DIAGNOSED WITH HYPERTENSION, CACHEXIA, UNSTABLE ANGINA, AND HEAVILY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY LESION. AT SOME TIME, POST STENT DEPLOYMENT, IT WAS ALLEGED THAT THE DEVICE WAS DIFFICULT TO BE REMOVED AFTER AN UNSUCCESSFUL PERCUTANEOUS REMOVAL PROCEDURE AND MATERIAL DEFORMED. THE PATIENT REPORTEDLY EXPIRED DUE TO CARDIAC ARREST, POST PERCUTANEOUS CORONARY INTERVENTION VENTRICULAR FIBRILLATION ARREST AND HEART FAILURE. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04364 | 2023-06-15 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT AFTER A SURGERY AND ONLY THE CATHETER WAS RECOVERED. ACCORDING TO THE INQUIRY SHEET ATTACHED, THE EVENT OCCURRED AFTER 2 HOURS OF PLACEMENT AND THERE WAS NO BALLOON RUPTURE/TEAR. THE WATER LEAKAGE WAS UNKNOWN AND NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. THERE WAS NO PATIENT'S URINE STONE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04362 | 2023-06-15 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS SPONTANEOUSLY REMOVED DURING SURGERY. THE EVENT OCCURRED ONE HOUR AFTER THE PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAK LOCATION WAS UNKNOWN BECAUSE THE EVENT WAS CATHETER FELL OUT AND THEY DID NOT TRY TO REPRODUCE THE EVENT. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04363 | 2023-06-15 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM SYRINGE AND VALVE CONNECTION OF THE FOLEY CATHETER. THERE WAS A POSSIBILITY OF CRACKS IN THE VALVE. THE EVENT OCCURRED DURING PRETEST. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE VALVE MENTIONED AS BACKFLOW. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 05JUN2023, STATED THAT THE LEAK WAS BETWEEN THE SYRINGE AND THE VALVE, SO IT WAS HARD TO TELL WHICH SIDE THE LEAK CAME FROM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04366 | 2023-06-15 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INDWELLING, BUT THE CATHETER PART HAD BEEN SEVERED FOR SOME REASON. THE USER REQUESTED TO INVESTIGATE WHY IT WAS DISCONNECTED (WAS IT CUT WITH A SHARP OBJECT SUCH AS SCISSORS, OR WAS IT DUE TO SOME OTHER REASON). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02447 | 2023-06-15 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY BROKEN. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04370 | 2023-06-15 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER (SILICONE) | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER WAS DIFFICULT TO BE REMOVED FROM THE FOLEY CATHETER DURING THE PRETEST. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED THAT THE BALLOON WAS PARTIALLY INFLATED. ATTEMPTED TO DEFLATE BALLOON PASSIVELY WITH RETURNED SYRINGE AND SOLUTION COULD NOT BE WITHDRAWN. ATTEMPTED TO DEFLATE BALLOON PASSIVELY USING IN-HOUSE LUER LOCK SYRINGE AND SOLUTION COULD NOT BE WITHDRAWN. SOLUTION COULD NOT BE WITHDRAWN BY ASPIRATION USING EITHER SYRINGE. DISSECTED BALLOON AND FOUND THAT THE NOTCH WAS PERFORATED COMPLETELY. DISSECTED INFLATION FUNNEL AND PIN GAUGES 0.034¿ AND BELOW COULD FREELY PASS THROUGH LUMEN. 0.035¿ - 0.041¿ TOOK SOME EFFORT TO GET THROUGH AND 0.042¿ AND ABOVE CANNOT PASS THROUGH. BASED ON PHYSICAL SAMPLE RECEIVED THE PRODUCT DOES NOT MEET SPECIFICATIONS WHICH STATES "SHAFT MUST NOT BE DAMAGED OR PINCHED". THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. A POTENTIAL FAILURE MODES ARE COLLAPSED LUMEN AND BLOCKED LUMEN. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿BARD® SILVER LUBRI-SIL® FOLEY TRAY SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY.] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04369 | 2023-06-15 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE CONNECTION WAS DEFECTIVE. IT WAS STATED THAT THE EVENT OCCURED JUST AFTER THE PLACEMENT. IT WAS STATED THAT THE THERE WAS NO BALLOON RUPTURE OR TEAR WAS FOUND. THERE WAS NO MATERIALS POSSIBLY CONTACTED TO THE CATHETER PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 30MAY2023, IT WAS STATED THAT THE CATHETER'S INFLATION VALVE WAS DISCONNECTED. THE DEVICE WAS USED ON A PATIENT. THE CATHETER'S INFLATION VALVE WAS DISCONNECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04371 | 2023-06-15 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER WAS DIFFICULT TO BE REMOVED FROM THE FOLEY CATHETER DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04373 | 2023-06-15 | 2023-05-25 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MALE INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION FROM THE MAGIC 3 GO MALE INTERMITTENT CATHETER. IT WAS UNKNOWN IF THE DEVICE CAUSED THE INFECTION AND UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02448 | 2023-06-15 | 2022-12-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT DURING PREPARATION OF A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE INTRODUCER NEEDLE WAS ALLEGEDLY FOUND TO BE NOT VACCUMED AFTER OPENING THE PUNCTURE KIT. IT WAS FURTHER REPORTED THAT DURING INJECTION BY THE SYRINGE, EXTERNAL LEAKAGE OCCURRED ON NEEDLE TIP AND THE NEEDLE WAS ALLEGEDLY BROKEN. REPORTEDLY, THE GUIDE WIRE GOT STUCK WHEN BEING INSERTED AND UNABLE TO BE ADVANCED OR RETRACTED. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2023). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04413 | 2023-06-16 | 2023-05-26 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MAGIC3 GO MALE COUDE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATED THAT PATIENT GOT URINARY TRACT INFECTION FROM THE CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04411 | 2023-06-16 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MALE EXTERNAL CATHETER | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE MALE EXTERNAL CATHETER COULD NOT SEEM TO GET IT OFF EASILY OR COMPLETELY REMOVE THE ADHESIVE. THEY HAD A BUILD UP OF ADHESIVE OVER THE COUPLE OF DAYS. THEY HAD TRIED TAKING IT OFF IN SHOWER AND USING ALCOHOL TO REMOVE THE ADHESIVE. REPRESENTATIVE ADVISED USING WARM WET CLOTH TO ASSIST IN BREAKING DOWN ADHESIVES. ALSO SUGGESTED ADDING SOAP OF CHOICE TO THE WASH CLOTH, DISCOURAGED USE OF ALCOHOL THAT CLOSE TO MUCOUS MEMBRANE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04407 | 2023-06-16 | 2023-05-26 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MALE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT HAS HAD SEVERE URINARY TRACT INFECTION FOR A LONG PERIOD OF TIME THE NOTE STATED THEY PUT THEM KIDNEYS AND LIFE AS RISK. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, THE PATIENT TOSSED THE REMAINDER OF THE CATHETERS AWAY THAT THEY SAID WAS CONTAMINATED. THE URINARY TRACT INFECTION HAD CLEARED UP WITH SEVERAL DIFFERENT ANTIBIOTICS AND THEIR KIDNEYS WERE WORKING BUT THEY WERE GOING TO CONTACT AN ATTORNEY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04408 | 2023-06-16 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE FUNNEL AREA OF THE FOLEY CATHETER DURING TESTING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04405 | 2023-06-16 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT PERFORMING A LEAK TEST, WATER LEAKAGE OCCURRED FROM THE CONNECTION WITH THE SYRINGE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04412 | 2023-06-16 | 2023-05-26 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MAGIC3 GO MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD A VERY BAD URINARY TRACT INFECTION WHILE USING THE MAGIC 3 GO MALE COUDE INTERMITTENT CATHETER AND WAS HOSPITALIZED FOR 6 DAYS AROUND APRIL AND MAY. PER FOLLOW UP VIA PHONE ON 07JUN2023, IT WAS REPORTED THAT THE CUSTOMER DID DEVELOP A URINARY TRACT INFECTION BUT DID NOT RECEIVE CONFIRMATION THAT THE CATHETERS WERE THE CAUSE. THEY WERE PRESCRIBED ANTIBIOTICS AND SPENT 6 DAYS IN THE HOSPITAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04410 | 2023-06-16 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER EDELMAN MODEL | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 087016 | 087016 |  | IT WAS REPORTED THAT URETERAL CATHETERS HAD A HAIR IN THE STERILE PACKAGE. PER FOLLOW-UP VIA EMAIL ON 01JUN2023, THE PRODUCT NEVER TOUCHED THE STERILE FIELD, SO NO HARM OR POTENTIAL HARM CAME TO THE PATIENT. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. 1 SAMPLE WERE CONFIRMED TO EXHIBIT THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITH ORIGINAL PACKAGING) AND, UNUSED BARDEX EDELMAN CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED ONE LONG STRAND OF HAIR INSIDE OF THE PACKAGING. THIS DOES NOT MEET SPECIFICATION WHICH STATES THAT LOOSE FOREIGN MATTER, ENGRAVED AND SPOTS GREATER-THAN 1/32 IS NOT ALLOWED. THE EXACT ROOT CAUSE COULD NOT BE IDENTIFIED. ALTHOUGH A ROOT CAUSE COULD NOT BE DEFINITIVELY IDENTIFIED, BASED ON THE RISK DOCUMENTATION REVIEW, A POTENTIAL ROOT CAUSE FOR THIS TYPE OF FAILURE COULD BE DEFECTIVE/ CONTAMINATED COMPONENTS FROM THE SUPPLIER. HOWEVER, THERE WAS INSUFFICIENT INFORMATION TO CONFIRM THIS POTENTIAL ROOT CAUSE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. UDI FORMAT UPDATED WITH AVAILABLE PRODUCT INFORMATION PER GUDID. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04421 | 2023-06-16 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE MAGIC 3 GO MALE INTERMITTENT CATHETER WAS NOT LUBRICATED ENOUGH SO THE PATIENT WAS IRRITATED. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02484 | 2023-06-16 | 2023-05-25 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D4 |  | IT WAS REPORTED BY CUSTOMER THAT UPON INSERTING THIS PICC, SOME RESISTANCE WAS FELT SO PICC RN PULLED PICC OUT AND THE GOLD TIP OF THE WIRE CAME OFF AND WAS LODGED IN PATIENT'S TISSUE. PATIENT HAD WIRE IN SOFT TISSUE (NOT VEIN) AND USER COULD NOT PLACE PICC. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02480 | 2023-06-16 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F TL MAX BARRIER PLUS KIT W/GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1395108D2 |  | IT WAS REPORTED BY THE CUSTOMER "LEAKING FROM AROUND WIRE AND RUBBER PORT." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02481 | 2023-06-16 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED THAT THIS PATIENT WAS ADMITTED TO HOSPITAL ON (B)(6)2023. THE NURSE MAINTAINED THE CATHETER BEFORE ADMINISTRATION OF DRUG, BUT NO BLOOD WAS DRAWN. AN IV NURSE IN THE HOSPITAL WAS INVITED FOR CONSULTATION. WHEN FLUSHING THE CATHETER WITH SALINE, LIQUID WAS FOUND LEAKING FROM THE PORTION OF THE CATHETER THAT WAS EXPOSED EXTERNALLY. THE CATHETER WAS REMOVED EVENTUALLY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02483 | 2023-06-16 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED VIA MEDWATCH "BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET REFERENCE # S1275108FD5, LOT # REHN2437 & REGY0842 NOTED LEAKAGE WITH FLUSHING AT RUBBER-LIKE STOPPER. REHN ULTIMATELY SUCCESSFULLY PLACED. STYLETS RETAINED." THIS FILE ADDRESSES LOT# REHN2437. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: IT WAS REPORTED THERE WERE NO VISIBLE CRACKS OR HOLES THAT CAUSED LEAKAGE. NO HARM NOTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02485 | 2023-06-16 | 2023-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED THAT THIS PATIENT WAS ADMITTED TO HOSPITAL DUE TO CEREBRAL INFARCTION ON (B)(6), 2023. ON THE DAY WHEN THE PATIENT WAS TRANSFERRED TO THE REHABILITATION DEPARTMENT, ONE SINGLE-LUMEN SOLO CATHETER WAS INSERTED, WITH 4 CM EXPOSED EXTERNALLY. THE PATIENT WAS THEN WAS TRANSFERRED TO ICU AND THE CATHETER WAS IN USE PROPERLY BY (B)(6). A SMALL AMOUNT OF SEEPAGE WAS NOTED ON (B)(6). THE CATHETER WAS SUSPECTED OF BEING OCCLUDED ON (B)(6). UROKINASE WAS GIVEN FOR THROMBOLYSIS. ON (B)(6), NO IMPROVEMENT WAS MADE IN SEEPAGE AND CATHETER DAMAGE WAS SUSPECTED. THE NURSE ATTEMPTED TO PULL OUT THE CATHETER BY 1 CM FOR FLUSHING AND FOUND THAT THE CATHETER WAS RUPTURED AT THE SCALE OF 5-6 CM. THE CATHETER HAD TO BE REMOVED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A SPLIT IN THE CATHETER IS CONFIRMED. ONE 4 FR S/L POWERPICC SOLO CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUES THROUGHOUT THE RETURNED DEVICE, AND A LONGITUDINALLY ALIGNED SPLIT WAS LOCATED AT A KINK OBSERVED IN THE CATHETER BETWEEN THE 4 CM AND 5 CM DEPTH MARKERS. A MICROSCOPIC OBSERVATION REVEALED A SPLIT AT THE CORNER OF THE KINK OBSERVED BETWEEN THE 4 CM AND 5 CM DEPTH MARKERS. THE SPLIT HAS A GRANULAR FRACTURE SURFACE AND APPEARS TO BE CAUSED BY CYCLIC KINKING OF THE CATHETER. MATERIAL FATIGUE CAN OCCUR WHEN THE CATHETER UNDERGOES CYCLIC KINKING. THE CATHETER WALL THICKNESS WAS MEASURED IN THE VICINITY OF THE SPLIT AND FOUND TO BE WITHIN MANUFACTURING SPECIFICATIONS. AFTER OBSERVATION OF THE RETURNED DEVICE, THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED AS A SPLIT WAS OBSERVED BETWEEN THE 4 CM AND 5 CM DEPTH MARKERS WHICH APPEARED TO BE THE RESULT OF SHARP KINKING OF THE CATHETER SHAFT. BASED ON THE LOCATION OF THE SPLIT, IT IS POSSIBLE THAT INSERTION AND MAINTENANCE TECHNIQUES CONTRIBUTED TO THE FAILURE OF THE DEVICE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11 | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02487 | 2023-06-16 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | NA | 5613200J |  | IT WAS REPORTED THAT THE PART OF THE RED LURE CONNECTOR WAS DETACHED DURING THE CATHETER USE. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02488 | 2023-06-16 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 3CG 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1194108D |  | IT WAS REPORTED BY A DIRECT CONSUMER THAT THEIR FATHER HAS A PICC LINE, TRYING TO GIVE HIM AN INFUSION AT HOME, THE LINE IS NOT WORKING TODAY, AND NOBODY HAS CHECKED IT SINCE (B)(6). CUSTOMER NEEDS SOMEONE TO COME OUT TO THEM TO ASSIST. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02490 | 2023-06-16 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG TPS STYLET 4F MAXIMAL BARRIER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED BY THE CUSTOMER THAT "WHILE PLACING A PICC LINE, I MEASURED OUT THE CATHETER AND WITHDREW THE INNER WIRE. I TRIMMED THE CATHETER AND UPON REPLACING THE INNER WIRE I NOTICED THAT THE COPPER TIP APPEARED TO BE MISSHAPEN. IT APPEARED TO BE BENT, I TOUCHED THE WIRE AND THE COPPER TIP FELL OFF INTO MY GLOVED HAND. WE SET THE PICC LINE ASIDE AND USED A NEW KIT" NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04406 | 2023-06-16 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE SIXTEEN MAGIC3 COUDE INTERMITTENT CATHETER FROM RECENT ORDER DID NOT HAVE ENOUGH WATER IN THE WATER PACK TO HYDRATE THEM AND IT CAUSED DISCOMFORT AND HAD TO BE THROWN AWAY. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04422 | 2023-06-16 | 2023-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PLACED ON (B)(6) 2023, AND SPONTANEOUS WITHDRAWAL WAS SUBSEQUENTLY CONFIRMED. AFTER ONE DAY OF OBSERVATION WITH AIR IN THE BALLOON, IT WAS NOTICED THAT THE BALLOON HAD BECOME SMALLER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02479 | 2023-06-16 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (20G) (10CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120100T |  | IT WAS REPORTED BY THE CUSTOMER, "HAVING ISSUES WITH THE POWERGLIDES AT THE TIP NEAR THE PINK TO GREEN HUB HAVING HOLES IN THEM. ONCE WE FLUSH IT, THE SITE APPEARS INFILTRATED BUT IT'S JUST BECAUSE THERE'S A HOLE IN THE CATHETER AND SO ANY MEDS SIT IN THE SUBQ TISSUE. IT HAS HAPPENED AT LEAST 3 TIMES." WHILE PUSHING FLUID THROUGH THE CATHETER, THERE IS A STREAM OF FLUID COMING OUT OF THE SIDE OF THE CATHETER, WHERE A HOLE APPEARS TO BE. THIS REPORT ADDRESSES THE THIRD OCCURRENCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04394 | 2023-06-16 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE PATIENT WAS IN THE STATE OF CATHETERIZATION AND THE NURSE FOUND THAT THE CATHETER COULD NOT BE REMOVED AFTER THE BALLOON WAS EVACUATED. THE CATHETER WAS REPEATEDLY INFLATED AND DEFLATED SEVERAL TIMES AND WAS FINALLY REMOVED. THE PATIENT WAS KEPT OBSERVED AND NO ABNORMALITY WAS FOUND. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO THIN RUBBERIZE WALL THAT CAUSING COLLAPSE/PINCH INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT". THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
17142665 | 2023-06-16 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWER PORT ISP GROSHONG 8FR TITANIUM CUSTOM | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | BARD ACCESS SYSTEMS, INC. | CP00003 | CP00003 |  | ANGIOGRAM SHOWED LEAKING OF CONTRAST IN THE NECK OF THE PATIENT AND NO FIBRIN SHEATH. MEDIPORT WAS REMOVED AND REPLACED WITH THE OLD CATHETER HAVING A HOLE IN THE SITE WHERE IT TURNED DOWN TO THE IJ IN THE NECK. |  | N | Y |  |  | 2023-06
3006260740-2023-02474 | 2023-06-16 | 2020-08-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE CATHETER ALLEGEDLY SLIPPED FROM THE VEIN. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY CRACKED. REPORTEDLY PATIENT ALLEGEDLY EXPERIENCED BRUISING. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02475 | 2023-06-16 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH HEMODIALYSIS CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT ONE MONTH AND TEN DAYS POST DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY LEAKED IN EXTENSION LINE BELOW THE HUB. IT WAS FURTHER REPORTED THAT THE EXTENSION LINE WAS ALLEGEDLY COMPLETELY TORN AND SEPARATED FROM THE CATHETER. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORD IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04382 | 2023-06-16 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT FOLEY CATHETER PACKAGING IDENTIFIED THAT THE BALLOON WAS 5CC'S. THE CATHETER IDENTIFIED THE BALLOON AS BEING 10CC. PRODUCTS NOT USED ON A PATIENT, NO KNOWN PATIENT HARM. 4 CATHETERS WITH THIS ISSUE HAD BEEN ISOLATED. PER SAMPLE RETURNED ON 13JUN2023, CUSTOMER HAD RETURNED THREE ADDITIONAL SAMPLES WITH THREE DIFFERENT MATERIAL NUMBERS.(0174L18- BATCH# NGFN2804), (165818-BATCH# NGFQ5644), (265716-BATCH# MYDZ3022). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04383 | 2023-06-16 | 2023-05-26 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MAGIC3 GO MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION (JUGT3299) (JUGT3353). PATIENT STATED THAT PATIENT GOT URINARY TRACT INFECTION FROM THE CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04386 | 2023-06-16 | 2023-05-24 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE PATIENT HAD TWO SEVERE URINARY TRACT INFECTIONS. PATIENT STATED THAT WITH ONE PATIENT WAS NOT ABLE TO PEE AND ALMOST STROKED OUT AND DIED. PATIENT STATED THAT PATIENT THEN GOT A URINARY TRACT INFECTION AND ALMOST DIED IN MARCH, PATIENT WAS GIVEN A 2-MONTH SUPPLY OF ANTIBIOTICS AND SEEMS TO BE ON THE MEND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04385 | 2023-06-16 | 2023-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED FOR THE PATIENT AND LATER THE CATHETER SPONTANEOUSLY FELL OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02473 | 2023-06-16 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616000 | 1616000 |  | IT WAS REPORTED THAT DURING A PORT REMOVAL PROCEDURE, THE PORT ALLEGEDLY DIFFICULT TO BE REMOVED. IT WAS FURTHER THAT THE CATHETER WAS CUT BUT ONLY SIX CENTIMETER WAS ABLE TO PULL. REPORTEDLY ADDITIONAL SURGICAL INTERVENTION PERFORMED TO REMOVE THE PORT. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04390 | 2023-06-16 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT THESE FOLEY CATHETER (LATEX) WERE TOO HARD TO GO INSIDE THE PATIENT BODY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04384 | 2023-06-16 | 2023-05-27 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION AND SEPSIS. THE CUSTOMER WAS NOT SURE IF IT WAS CAUSED BY THE PUREWICK FEMALE EXTERNAL CATHETER, AND THEY ARE NO LONGER USING PUREWICK. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04393 | 2023-06-16 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT AFTER INFLATION OF INDWELLING URINARY CATHETER WITH 10MLS WATER WHEN DETACHING THE SYRINGE, THE INFLATION VALVE DETACHED FROM THE INDWELLING URINARY CATHETER. IT WAS STATED THAT THE SURGEON DEFLATED THE INDWELLING URINARY CATHETER WITH GREAT DIFFICULTY. IT WAS STATED THAT THE BALLOON WAS INTAKE WHEN IT WAS TAKEN OUT AND IDC WAS REINSERTED WITH DIFFERENT BATCH NUMBER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04395 | 2023-06-16 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT ON 27TH MAR, 2023, A THREE LUMEN URINARY CATHETER WAS INDWELLED DURING THE OPERATION AND THE CATHETER BALLOON WAS FOUND RUPTURED ON (B)(6) 2023. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04398 | 2023-06-16 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, THE FIXATION WATER COULD NOT BE REMOVED FROM THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04399 | 2023-06-16 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT ON THE NEXT DAY AFTER PLACEMENT, THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. ACCORDING TO THE INQUIRY SHEET, THERE WAS NO BALLOON RUPTURE OR TEAR AND NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER. THERE WAS NO PATIENT'S URINE STONE AND THE CATHETER WAS NOT PULLED. AS PER THE ADDITIONAL INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, THERE WAS A BALLOON RUPTURE OBSERVED ON THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04403 | 2023-06-16 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER (SILICONE) | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER WAS DIFFICULT TO REMOVE FROM THE FOLEY CATHETER. IT WAS STATED THAT THE DURING THE BALLOON DEFLATION, THE WATER WAS NOT COMPLETELY REMOVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04404 | 2023-06-16 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN THE FOLEY CATHETER BALLOON WAS CHECKED, STERILE WATER WAS DIFFICULT TO INJECT AND DIFFICULT TO DRAIN WATER FORM THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02477 | 2023-06-16 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (20G) (10CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120100T |  | IT WAS REPORTED BY THE CUSTOMER, "HAVING ISSUES WITH THE POWERGLIDES AT THE TIP NEAR THE PINK TO GREEN HUB HAVING HOLES IN THEM. ONCE WE FLUSH IT, THE SITE APPEARS INFILTRATED BUT IT'S JUST BECAUSE THERE'S A HOLE IN THE CATHETER AND SO ANY MEDS SIT IN THE SUBQ TISSUE. IT HAS HAPPENED AT LEAST 3 TIMES." WHILE PUSHING FLUID THROUGH THE CATHETER, THERE IS A STREAM OF FLUID COMING OUT OF THE SIDE OF THE CATHETER, WHERE A HOLE APPEARS TO BE. THIS REPORT ADDRESSES THE FIRST OCCURRENCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02478 | 2023-06-16 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (20G) (10CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120100T |  | IT WAS REPORTED BY THE CUSTOMER, "HAVING ISSUES WITH THE POWERGLIDES AT THE TIP NEAR THE PINK TO GREEN HUB HAVING HOLES IN THEM. ONCE WE FLUSH IT, THE SITE APPEARS INFILTRATED BUT IT'S JUST BECAUSE THERE'S A HOLE IN THE CATHETER AND SO ANY MEDS SIT IN THE SUBQ TISSUE. IT HAS HAPPENED AT LEAST 3 TIMES." WHILE PUSHING FLUID THROUGH THE CATHETER, THERE IS A STREAM OF FLUID COMING OUT OF THE SIDE OF THE CATHETER, WHERE A HOLE APPEARS TO BE. THIS REPORT ADDRESSES THE SECOND OCCURRENCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04387 | 2023-06-16 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED FOR THE PATIENT AND LATER THE CATHETER SPONTANEOUSLY FELL OUT. AS PER THE ADDITIONAL INFORMATION RECEIVED ON 05JUN2023, THERE WAS A DAMAGE ON THE BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02492 | 2023-06-19 | 2023-01-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833730 | 5833730 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, AIR WAS ALLEGEDLY DRAWN INTO THE INTRODUCER NEEDLE. IT WAS FURTHER REPORTED THAT WHEN THE TIP OF THE INTRODUCER NEEDLE WAS BLOCKED, SALINE ALLEGEDLY LEAKED OUT OF THE HANDLE OF THE INTRODUCER NEEDLE WHEN INJECTING THE SALINE WITH THE SYRINGE. THE PROCEDURE WAS COMPLETED USING ANOTHER NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A VIDEO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02493 | 2023-06-19 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM LOW-PROFILE IMPLANTABLE PORT, BROVIAC SINGLE-LUMEN, 6.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602190 | 0602190 |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS AND ONE MONTH POST-PORT PLACEMENT, THE PORT ALLEGEDLY BROKE AWAY FROM THE CATHETER DURING THE ATTEMPT TO REMOVE THE DEVICE. IT WAS FURTHER REPORTED THAT THE PORT WAS EASILY DISSECTED FREE OF THE CAPSULE. IT WAS ALSO REPORTED THAT THE CATHETER WAS ALLEGEDLY FIXED AT THE SUBCLAVIAN VEIN ENTRY SITE AND WAS DISSECTED FREE OF SURROUNDING TISSUE DOWN TO THE VEIN UNDER THE CLAVICLE. HOWEVER, THE EXTERNAL ASPECT OF THE CATHETER WAS NOTED TO BE CHANGED FROM SMOOTH TO TEXTURED, LIKELY RELATED TO CALCIFICATIONS, ADHESION, OR INGROWTH WITHIN THE VEIN. REPORTEDLY, THE CATHETER WAS UNABLE TO BE PULLED OUT WITH VEIN TENTING UPON TRACTION, AND IT WAS DECIDED TO LEAVE THE INTERNAL CATHETER IN PLACE DUE TO THE RISK OF VEIN INJURY OR HEMORRHAGE. THE EXTERNAL CATHETER AND THE SUBCUTANEOUS PORT WERE REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRY DATE: 04/2023. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04425 | 2023-06-19 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L20 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE FOLEY CATHETER DURING THE PRE-TEST, ALTHOUGH THE LEAK LOCATION WAS NOT KNOWN. | THE REPORTED EVENT IS CONFIRMED - MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS INFLATED WITH WATER AND OBSERVED WATER LEAKING THRU DRAINAGE EYE. DISSECTED THE CATHETER AND OBSERVED TEAR ON RUBBERIZE LAYER OF INFLATION LUMEN OF CATHETER THAT CAUSE INTERLUMEN LEAK. TEAR WAS EXAMINED UNDER MICROSCOPE AND NOTED TO BE SMOOTH. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, WIRE PRESSING TECHNIQUE, STRIPPING TECHNIQUE ETC. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04427 | 2023-06-19 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ALL SILICONE CATHETER | SILICONE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER COMPLAINED ABOUT THE SILICONE DRAIN BAG FOLEY TRAY WITH ANTI-REFLUX CHAMBER, TAMPER-EVIDENT SEAL. THERE WERE 4 PATIENTS GOT HURT WHEN THEY USE FOLEY CATHETER AND ALSO INFORMED IT WAS SEEMED THICKER AND CUSTOMER WERE WONDERING IF THEY CHANGED THE KIT. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04431 | 2023-06-19 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE USER ATTEMPTED TO INJECT STERILE WATER WHILE TESTING THE FOLEY CATHETER BALLOON INFLATION, BUT WATER DID NOT ENTER. | THE REPORTED EVENT WAS UNCONFIRMED AS PRODUCT MEETS SPECIFICATIONS. VISUAL EVALUATION NOTED RECEIVED A 3-WAY TEMP SENSING CATHETER WITH CUT PORTION OF INLET TUBING. INFLATED BALLOON WITH 10 ML OF SOLUTION SUCCESSFULLY BUT WITH SOME DIFFICULTY USING AN IN-HOUSE LUER LOCK SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. DEFLATED BALLOON PASSIVELY WITH NO CUFFING OR LEAKS NOTED. ALSO RECEIVED 4 PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWS OVERVIEW OF CATHETER. SECOND PHOTO SAMPLE ZOOMS IN ON INFLATION CAP. THIRD PHOTO SAMPLE SHOWS TOP VIEW OF INFLATION CAP WITH NO OBSTRUCTIONS ALONG PATHWAY. FOURTH PHOTO SAMPLE ZOOMS IN ON END OF CATHETER WITH BALLOON NOT INFLATED. BASED ON PHYSICAL SAMPLE RECEIVED PRODUCT MEETS SPECIFICATIONS. THE PRODUCT WAS NOT USED FOR PATIENT DIAGNOSTIC OR TREATMENT. THE PRODUCT WAS NOT INFLUENCED BY THE REPORTED EVENT. THE DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE REPORTED EVENT WAS UNCONFIRMED. THE LABELING REVIEW WAS NOT REQUIRED AS THE REPORTED EVENT WAS UNCONFIRMED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04436 | 2023-06-19 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0166SI16 | 0166SI16 |  | IT WAS REPORTED THAT THE REPRESENTATIVE WANTED TO ENQUIRE AS TO WHETHER OR NOT THERE HAD BEEN ANY CHANGES TO THE DESIGN OF THE BARDEX IC CATHETERS. A CUSTOMER RECENTLY PURCHASED A LARGE AMOUNT AND HAD NOTED THAT THE EYELET WAS TOO LOW ON THE CATHETER TO BE USABLE, AS COMPARED TO HIGHER UP. THE CUSTOMER PURCHASED THE 0166SI16 16FG SIZING. WHILE THIS MIGHT JUST BE A DESIGN DIFFERENCE BETWEEN BARDEX AND OTHER STYLES OF CATHETERS, THE REPRESENTATIVE WOULD LIKE TO CONFIRM THAT THERE WAS NO FAULT OR DESIGN CHANGE, TO ASSIST THE CUSTOMER SERVICE TEAM TO RESOLVE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE SLOT WRONG POSITION/INCORRECT SLOT USED. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]: 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO 26 EVERY 2FR [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY.] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION CATHETER KINKING, DAMAGE, RUPTURE DIFFICULTY OR FAILURE TO REMOVE THE DEVICE OCCLUSION OF CATHETER INNER LUMENS ENCRUSTATION ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS URINARY-TRACT INFECTION HEMORRHAGE, HEMATURIA ALLERGY REACTION TO THE DEVICE CALCULUS FORMATION EDEMA PAIN DISCOMFORT INJURY OF BLADDER OR URETHRAL URETHRITIS, URINARY INCONTINENCE RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02498 | 2023-06-19 | 2023-05-23 | I, n, j, u, r, y | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY W/MICROINTRODUCER 55CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | IT WAS REPORTED THAT, PICC WAS INSERTED INTO PATIENT IN (B)(6) 2023. PATIENT RECEIVED CONTRAST FOR SCAN ON THE (B)(6) AS THE CONTRAST WAS BEING POWER INJECTED EXTRAVASATION WAS NOTICED IN PATIENT'S ARM. MINIMAL PAIN AND MILD SWELLING NOTED. PICC WAS REMOVED AND NOTED TO BE RUPTURED. CT WAS INFUSED VIA NEWLY INSERTED PERIPHERAL CANNULA. EXTRAVASATION WAS TREATED WITH ICE PACK AND PERIPHERAL CANNULA INSTEAD. ADDITIONAL INFORMATION RECEIVED 06/12/2023: THERE WAS NO NEGATIVE OUTCOME. THE PATIENT WAS CHECKED THE NEXT DAY AND THE SITE WAS COMFORTABLE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04458 | 2023-06-19 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE BALLOON PORTION RUPTURED AND WAS WITHDRAWN AFTER THE FOLEY CATHETER WAS PLACED IN THE PATIENT. PER MAIL RECEIVED FROM IBC ON 14JUN2023, STATED THAT THEY RECEIVED INFORMATION FROM SALES THAT THE EVENT WAS WRONG. THE CORRECT EVENT WAS THAT THERE WAS A CRACK IN THE CATHETER SHAFT AREA AND WATER LEAKED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04465 | 2023-06-19 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED, THAT AT 22:00 ON (B)(6) 2023. THE PATIENT UNDERWENT SURGERY UNDER GENERAL ANESTHESIA AND CATHETERIZATION. THEN, WHEN THEY RETURNED TO THE WARD AT 23:30, THE CATHETER COULD DRAIN SMOOTHLY. ON (B)(6) 2023, THE PATIENT REPORTED, THAT THE CATHETER FELL OUT BY ITSELF. AND THE NURSE FOUND, THAT THE FOLEY CATHETER BALLOON RUPTURED. | THE REPORTED EVENT WAS INCONCLUSIVE, BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN, WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN, WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT, EXPOSURE TO PETROLATUM BASED PRODUCTS, MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED, AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX . WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS, SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX ". SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04459 | 2023-06-19 | 2023-05-30 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION FROM USING THE MAGIC 3 FEMALE INTERMITTENT CATHETER AND WAS HOSPITALIZED. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. THE PATIENT REPORTED ISSUES WITH LOT# JUGU0871, LOT# JUGR0160. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04462 | 2023-06-19 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON SPONTANEOUSLY FELL OUT OF THE PATIENT ON THE NEXT DAY AFTER PLACEMENT. THE POSSIBLE CAUSES OF THIS SPONTANEOUS CATHETER REMOVAL INCLUDE BACKFLOW FROM THE INFLATION VALVE OR BALLOON DAMAGE. THE WATER LEAKED FROM THE BALLOON OR THE VALVE MENTIONED AS BACKFLOW. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. PER MAIL RECEIVED FROM IBC ON 06JUN2023, STATED THAT AT THIS TIME, THE LEAKAGE LOCATION WAS ONE, BUT IT WAS NOT JUDGED WHICH IS CORRECT, THE BALLOON OR THE VALVE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, WIRE PRESSING TECHNIQUE, STRIPPING TECHNIQUE ETC. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE 1. MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTIONS: D, H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04466 | 2023-06-19 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER HAD BEEN INSERTED LESS THAN 4 WEEKS PRIOR AND THE STAFF FOUND THE CATHETER IN THE PATIENT'S BED WITH THE BURST BALLOON. THERE HAD BEEN NO TENSION ON THE FOLEY CATHETER WHILE INDWELLING AND THIS HAD NOT HAPPENED WITH OTHER DEVICES THIS PATIENT HAS HAD INSERTED. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. BASED ON THE ATTACHED PHOTO, A BURST BALLOON WAS OBSERVED ON THE SAMPLE. HOWEVER THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. POTENTIAL ROOT CAUSE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04471 | 2023-06-19 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE WATER LEAKED POSSIBLY FROM THE BALLOON PART DURING THE PRETEST. THE FOLEY CATHETER WAS NOT USED BECAUSE OF THE LEAK. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 05JUN2023, STATED THAT THERE WAS A CRACK FOUND AT THE CATHETER SHAFT UPON CHECKING THE RETURNED SAMPLE. POSSIBLY WATER LEAKED FROM THE CRACK. THERE WAS NO DAMAGE OBSERVED TO THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02500 | 2023-06-19 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT DURING WARD ROUNDS AT NIGHT, THE EXTENSION TUBING OF THE CATHETER WAS FOUND RUPTURED. THE CATHETER WAS PLACED IN THIS PATIENT LESS THAN ONE MONTH AGO. THE HEAD NURSE REPLACED THE CATHETER WITH ANOTHER ONE IN SITU, CAUSING NO IMPACT TO THE PATIENT¿S LIFE AND HEALTH. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04474 | 2023-06-19 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS SPONTANEOUSLY WITHDRAWN FROM A PATIENT IN THE HOSPITAL WARD. PER PHOTO SAMPLE FROM IBC ON 30MAY2023, IT WAS FOUND THAT THE INFLATION FUNNEL WAS FOUND TO BE BROKEN. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 05JUN2023, STATED THAT THE INFLATION FUNNEL WAS RUPTURED. THE DETAILS WEREN'T PROVIDED FROM CUSTOMER. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 09JUN2023, STATED THAT THE INFLATION FUNNEL NEVER ENTERED INSIDE THE BODY OF THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02501 | 2023-06-19 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED THAT WHILE DOING A DRESSING CHANGE, NURSE NOTICED THE PICC WAS CRACKED/BROKEN RIGHT AT THE HUB. WHEN DOING A PICC EXCHANGE, SHE NOTICED THERE WAS ANOTHER TEAR AROUND 9CMS IN THE LINE. HAD TO DO A OVERWIRE EXCHANGE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04485 | 2023-06-19 | 2023-05-31 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT GOT 3 URINARY TRACT INFECTIONS USING INTERMITTENT CATHETER AND DID NOT SPECIFY THE TIME FRAME OF THE URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FPR URINARY TRACT INFECTION. AS PER FOLLOW UP VIA PHONE ON (B)(6) 2023, CUSTOMER STATED THAT THEY HAD 3 UTIS WITH USE AND WAS TREATED WITH ANTIBIOTICS FOR THE INFECTIONS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04484 | 2023-06-19 | 2023-05-31 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT RECEIVED URINARY TRACT INFECTIONS AFTER USING THE MAGIC 3 GO INTERMITTENT CATHETER. THE PATIENT HAD URINARY TRACT INFECTION AT THIS MOMENT AND HAD SEEN A PHYSICIAN, THE PATIENT WAS ON ANTIBIOTICS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04488 | 2023-06-19 | 2023-05-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT WAS INDWELLED BY A URINARY CATHETER BY A NURSE. DURING THE OPERATION, THE PATIENT FELT A SOUND IN THE ABDOMEN AND THEN THE CATHETER FELL OUT OF THE URETHRA. UPON CHECKING THE BALLOON, THE BALLOON RUPTURE WAS CONFIRMED WITH NO ADVERSE EFFECT ON THE PATIENT AND THEN REPLACED BY A NEW CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04487 | 2023-06-19 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165818 | 165818 |  | IT WAS REPORTED, THAT THE BALLOON PORT FELL OFF FROM THE FOLEY CATHETER. WHEN THE USER WAS TRYING TO REMOVE SYRINGE. NO APPARENT HARM REACHED THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04490 | 2023-06-19 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT ON (B)(6) 2023, 20ML OF STERILIZED WATER WAS INJECTED INTO THE BALLOON TO CHECK WHETHER THE URETERAL BALLOON FUNCTION WAS NORMAL BEFORE CATHETERIZATION, THEN THE STERILE WATER COULD NOT BE DRAINED OUT. IT WAS REPLACED WITH A NEW CATHETER FOR USE AND THERE WERE NO ISSUES FOUND AT THE TIME OF USE AND DID NOT CAUSE ANY HARM TO THE PATIENT. ACCORDING TO THE INJURY PERFORMANCE, THERE WAS A DELAYED CATHETERIZATION. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE CATHETER WAS NOT USED ON THE PATIENT AND ONCE THE DEFECT WAS FOUND, ANOTHER CATHETER WAS USED. NO INTERVENTION REQUIRED TO REMOVE WATER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04491 | 2023-06-19 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CATHETER FELL OUT AND UPON CHECKING THE CATHETER, THE BALLOON BURST. THERE WERE NO MISSING PIECES OF BALLOON. THE PATIENT WAS THEN REPLACED WITH A NEW 20FR CATHETER WHERE IT DRAINED 500ML OF LIGHT RED URINE, GIVING A GOOD FIXATION, AND THE NURSE INSTRUCTED THE PATIENTS ABOUT THE PRECAUTIONS FOR THE ANTI-SLIPPING OF TUBE. ACCORDING TO THE INJURY AND DEVICE FAILURE PERFORMANCE, THE PATIENT HAD SECONDARY INTUBATION AND THE WATER BLADDER RUPTURED RESPECTIVELY. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON 05JUN2023, CLARIFIED THAT THERE WAS BLEEDING DUE TO THE DEFECTIVE CATHETER AND THERE WAS DELAYED TREATMENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04428 | 2023-06-19 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTERMITTENT ROCHESTER CATHETERS | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE SLEEVE DID NOT SLIDE NOT SAFE TO USE. AFFECTED QUANTITY NOT REPORTED. THE PATIENT NOT USED THIS PRODUCT BEFORE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04455 | 2023-06-19 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE SHARP. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04456 | 2023-06-19 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON GOT BURST DURING THE PRETEST AND THE CATHETER USE WAS DISCONTINUED. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. VISUAL EVALUATION NOTED RECEIVED A 3-WAY TEMP SENSING CATHETER. VISUAL INSPECTION NOTED BALLOON BURST MEASURING 0.4245" WITH NO MISSING PIECES NOTED. ALSO RECEIVED 2 PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWS OVERVIEW OF CATHETER. SECOND PHOTO SAMPLE ZOOMS IN END OF CATHETER WITH BALLOON NOT INFLATED BUT RIP IS PRESENT. BASED ON SAMPLES RECEIVED PRODUCT DOES NOT MEET SPECIFICATIONS ACCORDING TO "BALLOONS SHALL NOT BURST WHEN INFLATED TO MINIMUM BURST VOLUME ". ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE HIGH MODULUS SILICONE. THE PRODUCT WAS NOT USED FOR PATIENT DIAGNOSTIC OR TREATMENT. THE PRODUCT WAS INFLUENCED BY THE REPORTED EVENT. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE 1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. 2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. 3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. 4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. 5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. 9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04473 | 2023-06-19 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 902916 | 0165L16 |  | IT WAS REPORTED THAT MALFUNCTIONING DRAINAGE SYSTEM IN THE FOLEY TRAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04452 | 2023-06-19 | 2023-05-30 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION FROM USING THE MAGIC 3 FEMALE INTERMITTENT CATHETER AND WAS HOSPITALIZED. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. THE PATIENT REPORTED ISSUES WITH LOT# JUGU0871, LOT# JUGR0160. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04457 | 2023-06-19 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE BALLOON UPON INJECTING STERILE WATER INTO THE FOLEY CATHETER BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02494 | 2023-06-19 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716070J |  | IT WAS REPORTED THAT PRIOR TO PORT PLACEMENT PROCEDURE, THE GRAY SHEATH OF THE SHEATH INTRODUCER WAS ALLEGEDLY SOAKING WET AND COULD NOT BE USED. THE AMOUNT OF SILICONE OIL APPLIED MAY HAVE BEEN TOO MUCH. THERE WAS NO PATIENT CONTACT. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 12/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02495 | 2023-06-19 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, IT WAS ALLEGEDLY FOUND THAT THE CATHETER WAS ALLEGEDLY OOZING BLOOD ABOUT SEVENTEEN CENTIMETERS FROM THE NECK. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 12/2023). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04437 | 2023-06-19 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE CATHETER WAS PLANNED TO BE USED FOR THE PATIENT AND DURING A PRETEST, THE BALLOON WAS FOUND LEAKING. IT WAS NOT USED ON THE PATIENT AND WAS THEN REPLACED WITH A NEW ONE TO COMPLETE THE PROCEDURE. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, INCORRECT WIRE PRESSING TECHNIQUE, INCORRECT STRIPPING TECHNIQUE ETC AND MIGHT BE CONTRIBUTE BY USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02496 | 2023-06-19 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER KIT 13 F 15CM TL | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5603150 |  | IT WAS REPORTED BY CUSTOMER THAT THE GUIDEWIRE IN THE TEMP HD CATH VERY DIFFICULT/NOT ABLE TO THREAD THROUGH THE LINE. CLINICIAN HAD TO OPEN A SECOND KIT. EVENT DIRECTLY INVOLVED A PATIENT FOR AN URGENT MEDICAL SITUATION. NO HARM OR INJURY. DELAY IN INSERTION CAUSED DELAY IN DIALYSIS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04446 | 2023-06-19 | 2023-05-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MAGIC3 GO MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT HAD TWO URINARY TRACT INFECTION WHILE USING THE MAGIC 3 GO MALE COUDE INTERMITTENT CATHETER AND THE PATIENT WAS IN HOSPITAL BOTH TIMES. IT WAS STATED THAT PATIENT USED MIGHT BE 10 OF THEM. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED.. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04443 | 2023-06-19 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT A PART OF THE BALLOON WAS DEFORMED IN A BUMPY SHAPE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04439 | 2023-06-19 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT STAFF ATTEMPTED TO PLACE TWO CATHETERS IN PATIENTS AND THE BALLOONS WERE LEAKED. PER FOLLOW UP VIA PHONE ON 07JUN2023, IT WAS REPORTED THAT THE LOTS WERE NGHN4280 AND NGGW0794, THESE WERE THE ONLY LOTS THAT HAVE HAD AN ISSUE AND NO PATIENT IMPACT REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "IMPERFECTION IN BALLOON DUE TO PROCESS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED DIRECTIONS FOR USE WASH HANDS AND DON CLEAN GLOVES EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERIURETHRAL AREA 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL 5. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP 6. DON STERILE GLOVES 7. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE 8. POSITION FENESTRATED DRAPE ON PATIENT 9. SATURATE 3 FOAM SWAB STICKS IN POVIDONE IODINE 10.ATTACH THE WATER FILLED SYRINGE TO THE INFLATION PORT NOTE: IT IS NOT NECESSARY TO PRE-TEST THE FOLEY CATHETER BALLOON 11. REMOVE FOLEY CATHETER FROM WRAP AND LUBRICATE CATHETER 12. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA INORA MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD 13. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND A. WHEN CATHETER TIP HAS NTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE B. INSERT CATHETER TWO MORE INCHES AND INFLATE CATHETER BALLOON 14. INFLATE CATHETER BALLOON USING ENTIRE 10CC OF STERILE WATER PROVIDED IN THE PREFILLED SYRINGE NOTE: USE OF LESS THAN 10CC CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON 15. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK 16. SECURE THE FOLEY CATHETER TO THE PATIENT USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED (SEE STATLOCK® FOLEY STABILIZATION DEVICE IFU) NOTE: PLEASE MAKE SURE PATIENT IS APPROPRIATE FOR USE OF STATLOCK® STABILIZATION DEVICE 17. POSITION HANGER ON BED RAIL AT THE FOOT OF THE BED NOTE: EXERCISE CARE TO KEEP BAG OFF THE FLOOR 18. USE GREEN SHEETING CLIP TO SECURE DRAINAGE TUBE TO THE SHEET. MAKE SURE TUBE IS NOT KINKED 19. INDICATE TIME AND DATE OF CATHETER INSERTION ON PROVIDED LABELS. PLACE DESIGNATED LABELS ON PATIENT CHART AND DRAINAGE SYSTEM 20. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL FOLEY CATHETER REMOVAL 1. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE 2. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY 3. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION 4. IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL 5. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT¿. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04448 | 2023-06-19 | 2023-05-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MAGIC3 GO MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION WHILE USING THE MAGIC 3 GO MALE COUDE INTERMITTENT CATHETER THAT WERE ON RECALL BUT IT WAS STATED THAT THE PATIENT DID NOT HAVE LOT#. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN. WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS. URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04450 | 2023-06-19 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT TWO OF THE NINETY INTERMITTENT CATHETERS WERE NOT LUBRICATING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04453 | 2023-06-19 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE SOME OF THE MAGIC 3 INTERMITTENT CATHETER DID NOT COME WITH ENOUGH WATER TO HYDRATE THE CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "AIR IN LINES". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04451 | 2023-06-19 | 2023-05-30 | I, n, j, u, r, y | Manufacturer report | MAGIC 3 GO 16 FR MALE COUDE CATHETER WITH SUREGRIP | MAGIC 3 GO MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT MENTIONED THAT THEY PREVIOUSLY GOT A URINARY TRACT INFECTION, NOT SURE IF IT WAS FROM MAGIC 3 CATHETER. THE PATIENT DID NOT SPECIFIC THE TIMEFRAME OF THE URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04447 | 2023-06-19 | 2023-05-27 | I, n, j, u, r, y | Manufacturer report | BARD® WOVEN URETERAL CATHETER | URETERAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 36304 | 36304 |  | IT WAS REPORTED THAT THE PATIENT WAS GETTING A RASH ON HIS PENIS, WENT TO THE DOCTOR AND GAVE MEDICATION AND IT IS CLEARING UP. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04454 | 2023-06-19 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165L20 | 0165L20 |  | IT WAS REPORTED THAT ONE OF THESE GOT PLUGGED AND URINE WOULD NOT GO THROUGH INTO THE LATEX FOLEY CATHETER, BUT THE PATIENT FIGURED OUT HOW TO FIX IT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02486 | 2023-06-19 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | IT WAS REPORTED THAT PICC LINE WAS HAVING SMALL LEAKAGE, THEN DOCTOR JUST PULLED THE CATHETER OUT AND HAD TO CUT IT AND PLACE REPAIR KIT AND FIXED IT AGAIN. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAKING CATHETER WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS FOUR PHOTOGRAPHS WHICH DEPICTED A 4FR S/L GROSHONG CATHETER. THE DEPICTED DEVICE WAS INSERTED INTO THE ARM OF A PATIENT. THE VISIBLE PORTION OF THE DEVICE INCLUDED THE DISTAL PORTION OF THE ASSEMBLED CONNECTOR AND THE SEGMENT OF CATHETER BETWEEN THE CONNECTOR AND THE INSERTION SITE. CLEAR LEAKING INFUSATE WAS VISIBLE IN THE PHOTOGRAPHS NEAR THE INSERTION SITE. A CATHETER LEAK WAS EVIDENT IN THE SUBMITTED PHOTOGRAPHS; HOWEVER, INSPECTION OF THE PHOTOGRAPHS WAS INSUFFICIENT TO IDENTIFY THE CAUSE OF THE LEAK. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE MATERIAL FATIGUE AND SHARP INSTRUMENT CONTACT. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02497 | 2023-06-19 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PATIENT ATTENDED FOR BLOODS ON (B)(6) 2023, BLOOD WAS ASPIRATED WITH DIFFICULTY AND WHEN PICC FLUSHED WITH SALINE THE PATIENT REPORTED A NIPPING PAIN AT THE PICC ENTRY SITE WITH SOME CLEAR FLUID OOZING FROM THE SITE ALSO. ADVICE SOUGHT FROM INFUSIONAL SERVICES AND PICC REMOVED. SPLIT NOTED AT 7CM. SECUREMENT DEVICE: SECUREACATH. USED FOR NABPACILITAXEL. NO DELAY IN PATIENT'S TREATMENT. A NEW PICC IS REQUIRED. DATIX COMPETED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04520 | 2023-06-20 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TIEMANN MODEL UROLOGICAL CATHETER COUDE TIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 010116 | 010116 |  | IT WAS REPORTED THAT CUSTOMER GOT 3 BOXES OF 12 INTERMITTENT CATHETERS THAT WERE DEFECTIVE ALL THE SAME LOT. CUSTOMER STATED THAT THE END OF THEM WAS CURL OVER AND WERE VERY HARD. CUSTOMER ALSO MENTIONED THAT UP AT THE TOP THEY WERE SHORTER IN THE PACKAGING, 2 INCH SHORTER THAN USUAL. UP AT THE TOP OF THE CATHETER, THE OUTLET PART THERE WERE A LARGE BUMP ON IT. SO, INSERTION WOULD BE DIFFICULT BECAUSE OF THE HARD CURL END. CUSTOMER REQUEST REPLACEMENT FOR THE DEFECTIVE PRODUCT. PER FOLLOW UP VIA PHONE 08JUN2023, CUSTOMER STATED THAT THEY DID NOT USE THE CATHETERS, THEY SAW THE DEFECT BEFORE USE AND WANTED TO INFORM THE COMPANY OF THE AFFECTED LOTS. PER NOTIFICATION FROM INVESTIGATOR ON 31JUL2023, DURING SAMPLE EVALUATION IT WAS FOUND THAT THE SCRATCH FOUND ON THE SHAFT CLOSE TO THE FUNNEL OF CATHETER. PER ADDITIONAL INFORMATION RECEIVED VIA SAMPLE FORM ON 20JUN2023, IT WAS REPORTED THAT CUSTOMER OPENED ONE BOX AND SAW A DEFECTIVE CATHETER, CHECKED THE OTHERS IN THE BOX, THEY WERE ALL THE SAME, SO HE THREW THAT BOX AWAY. STATED THAT THE NEXT BOX WAS FINE AND THEN THEY GOT ANOTHER BOX OF DEFECTIVE CATHETERS AND STATED THAT THESE WERE SHORTER THAN THEY SHOULD BE. THE TIPS WERE CURLED AND THE HOLE IN THE TIP WAS MUCH FURTHER UP THE CATHETER THAN IT SHOULD BE AND THEY WARE HARDER AND LESS PLIABLE THAN THEY SHOULD BE. THE OUTLET END HAD A BUMP ON THE OUTSIDE. STATED THAT THEY ALSO NOTICED THAT THE PRINT ON THE PACKAGING WAS AN OFF COLOR FROM THE USUAL PACKAGING AND THEY WENT THROUGH ALL THEIR REMAINING BOXES, AND THESE WERE THE ONLY ONES, AND THEY WERE ALL THE SAME LOT NUMBER. STATED THAT THEY WOULD BE TERRIBLY UNCOMFORTABLE TO USE AND SEEM LIKE THEY COULD HURT SOMEONE IF THEY TRIED TO USE IT, AND THIS WAS THE REMAINDER OF TWO BOXES AND COULD SEE THE FLAW THROUGH THE PACKAGING. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE DEVICE MEETS SPECIFICATIONS. GROSS VISUAL EVALUATION: NO NON-CONFORMITIES WERE NOTED WITH THE HARDNESS AT THE TIP OF THE CATHETER. COUDE CURVE ALIGNED WITH THE COUDE INDICATOR NOTCH, SOME DEGRADATION DUE TO OXIDATION WAS NOTED WITHIN THE EYELETS OF CATHETERS 1, 2, 3, 4, 5 (VERY SLIGHT SIGNS OF OXIDATION WITHIN THE EYELETS), 6, 8, 9, AS THIS WAS NOT REPORTED BY THE CUSTOMER, AND IT IS UNKNOWN HOW THE CATHETER WAS BEING STORED OR WHEN THE OXIDATION OCCURRED A NEW COMPLAINT WILL NOT BE CREATED. SOME SCRATCH MARKS WERE NOTED ON THE FUNNEL OF CATHETERS 4, 6 & 9, AND A SMALL SCRATCH WAS NOTED ON THE SHAFT CLOSE TO THE FUNNEL OF CATHETER 10; NO SCRATCHES OR ABNORMALITIES WERE NOTED ON THE PACKAGING. DIMENSIONAL EVALUATION: USING A FRENCH GAUGE, ALL CATHETERS WERE CONFIRMED TO BE 16FR. THE OVERALL LENGTH OF THE CATHETERS WAS MEASURED AND FOUND TO BE WITHIN SPECIFICATION: CATHETER 1- 15.75", CATHETER 2 - 15.75", CATHETER 3 - 15.72", CATHETER 4 - 15.75", CATHETER 5 - 15.81", CATHETER 6 - 15.87", CATHETER 7 - 15.75",CATHETER 8 - 16.00",CATHETER 9 - 16.00", CATHETER 10 - 15.87". A DHR IS NOT REQUIRED SINCE THE REPORTED FAILURE WAS UNCONFIRMED. THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02515 | 2023-06-20 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SLIM-CATH STRAIGHT SHORT-TERM DIALYSIS (12F) (TL) (24CM) | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5853240J |  | IT WAS REPORTED THAT THE GUIDEWIRE WAS DAMAGED AND SLIGHTLY FRAYED. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02514 | 2023-06-20 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F FULL TRAY (DL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2275108 |  | THE CUSTOMER REPORTED THAT THE LUER LOCK CONNECTOR OF THE STYLET DOES NOT CLOSE PROPERLY AND AIR IS GETTING INTO THE TUBE. AIR ENTERS THE SYSTEM. THE PICC IS CORRECTLY IMPLANTED IN THE PATIENT AND WORKS WELL. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02513 | 2023-06-20 | 2023-06-02 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO CATHETER 4F BASIC TRAY (SL) (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6194118 |  | IT WAS REPORTED CATHETER POWERPICC SOLO (B)(6) WAS LEAKING. CATHETER WAS TUNNELED IN THE AXILLARY VEIN. AFTER CATHETER REMOVAL IT WAS NOTED THAT THERE IS A RUPTURE PRESENT ON 15TH CENTIMETER. UNFORTUNATELY, THERE WAS ONGOING CHEMOTHERAPY SO PATIENT HAS NECROSIS AND PATIENT MAY NEED RECONSTRUCTION PROCEDURE, THEY ARE STILL OBSERVING, DECISION NOT YET TAKEN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE S/L 4 FR POWERPICC SOLO. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND TWO SPLITS WERE OBSERVED BETWEEN THE 14 CM AND 15 CM DEPTH MARKERS AND BETWEEN THE 15 CM AND 16 CM DEPTH MARKERS. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED. FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR. 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER. OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING) THE CATHETER WAS SUBSEQUENTLY CUT TO MEASURE THE WALL THICKNESS OF THE CATHETER TO ENSURE THE CATHETER WAS WITHIN SPECIFICATIONS OF THE DRAWING. THE CATHETER WAS MEASURED AND WAS FOUND TO BE WITHIN SPECIFICATION. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04523 | 2023-06-20 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON COULD NOT BE DEFLATED WHEN THE USER ATTEMPTED TO BE DEFLATE IT AFTER INJECTING STERILE WATER INTO THE BALLOON DURING THE PRETEST. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. A REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04524 | 2023-06-20 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILE DISTILLED WATER LEAKED FROM THE VALVE OF THE FOLEY CATHETER WHEN THE USER ATTEMPTED TO INJECT STERILE WATER. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. IT WAS UNKNOWN IF THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE COULD BE LUMEN COMPROMISED BY A PUNCTURE OR CUT. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE SOLUTION IMMEDIATELY WENT INTO THE DRAINAGE LUMEN AT BIFURCATION AND CREATED LUMEN TO LUMEN LEAK. THIS IS OUT OF SPECIFICATION, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04521 | 2023-06-20 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT WHEN THE SYRINGE WAS INSERTED INTO THE INFLATION VALVE OF THE FOLEY CATHETER, THE CLEAR PLASTIC PART OF THE INFLATION VALVE WAS CRACKED (BROKEN). IT SEEMED THAT THE SYRINGE WAS NOT INSERTED WITH A PARTICULARLY STRONG FORCE OR DIAGONALLY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02517 | 2023-06-20 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC HF 5F TL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3385118QJ |  | IT WAS REPORTED THAT THE PATIENT HAD A POWERPICC IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, BLEEDING OCCURRED FROM THE INSERTION SITE. WHEN THE CATHETER WAS PULLED OUT BY APPROXIMATELY 1 CM DURING HEMOSTASIS, IT WAS FOUND THAT BLOOD LEAKED FROM THE PART NEAR THE INSERTION SITE. BLOOD TRANSFUSION WAS PERFORMED THROUGH THE TRIPLE-LUMEN MAIN LUMEN, AND THE OTHER TWO LUMENS WERE LOCKED. BFLUID INJECTION AND ANTIBIOTICS WERE ADMINISTERED AFTER CATHETERIZATION. THE DEVICE WAS REMOVED AND ANOTHER NEW CATHETER WAS REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04522 | 2023-06-20 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TIEMANN MODEL UROLOGICAL CATHETER COUDE TIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 010116 | 010116 |  | IT WAS REPORTED THAT CUSTOMER GOT 3 BOXES OF 12 INTERMITTENT CATHETERS THAT WERE DEFECTIVE ALL THE SAME LOT. CUSTOMER STATED THAT THE END OF THEM WAS CURL OVER AND WERE VERY HARD. CUSTOMER ALSO MENTIONED THAT UP AT THE TOP THEY WERE SHORTER IN THE PACKAGING, 2 INCH SHORTER THAN USUAL. UP AT THE TOP OF THE CATHETER, THE OUTLET PART THERE WERE A LARGE BUMP ON IT. SO, INSERTION WOULD BE DIFFICULT BECAUSE OF THE HARD CURL END. CUSTOMER REQUEST REPLACEMENT FOR THE DEFECTIVE PRODUCT. PER FOLLOW UP VIA PONE (B)(6) 2023, CUSTOMER STATED THAT THEY DID NOT USE THE CATHETERS, THEY SAW THE DEFECT BEFORE USE AND WANTED TO INFORM THE COMPANY OF THE AFFECTED LOTS. PER NOTIFICATION FROM INVESTIGATOR ON 31JUL2023, DURING SAMPLE EVALUATION IT WAS FOUND THAT THE SCRATCH FOUND ON THE SHAFT CLOSE TO THE FUNNEL OF INTERMITTENT CATHETER. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND THAT THERE WAS NO ALLEGATION AGAINST PRODUCT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02518 | 2023-06-20 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER FULL TRAY 13 F (15CM) (TL) | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5605150 |  | IT WAS REPORTED BY THE CUSTOMER THAT A RIJ TRIALYSIS DISLODGED. THEY THINK THE HOLDER WAS DEFECTIVE BECAUSE THE SUTURES ARE COMPLETELY INTACT. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02531 | 2023-06-20 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 18G X 2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1182250 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN NURSE PULLED THE ACCUCATH ACE OUT OF THE PATIENT THE WHOLE GUIDEWIRE WASN'T INTACT LEAVING HIM TO BELIEVE THAT SOME OF THE GUIDEWIRE WAS LEFT IN THE PATIENT. THEY SEND THE PATIENT TO X-RAY AND CT AND THEY HAVEN'T BEEN ABLE TO LOCATE THE GUIDEWIRE. HE SAID THAT HE FELT RESISTANCE AT FIRST, NO LONGER FELT IT AS HE CONTINUED TO ADVANCE AND WHEN HE PULLED THE DEVICE OUT HE NOTICED THE WHOLE GUIDEWIRE WASN'T INTACT. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02522 | 2023-06-20 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT DURING THE PORT PLACEMENT PROCEDURE, THE INTRODUCER WAS ALLEGEDLY SEEMED TO BE LONGER. IT WAS FURTHER REPORTED THAT THE CATHETER KEPT GOING TO THE RIGHT SIDE OF THE SUBCLAVIAN VEIN WHEN THE PORT WAS PLACED ON THE LEFT SUBCLAVIAN VEIN. REPORTEDLY, THE PHYSICIAN HAS TO PULL BACK THE SHEATH A LITTLE BIT TO PLACE THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3, H6 (DEVICE, METHOD). H11: B5, E1, D1, D2, D4 (MEDICAL DEVICE CATALOGUE NUMBER). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02523 | 2023-06-20 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT DURING THE PORT PLACEMENT PROCEDURE, THE INTRODUCER SEEMED TO BE LONGER. IT WAS FURTHER REPORTED THAT THE CATHETER KEPT GOING TO THE RIGHT SIDE OF THE SUBCLAVIAN VEIN WHEN THE PORT WAS PLACED ON THE LEFT SUBCLAVIAN VEIN. REPORTEDLY, THE PHYSICIAN HAS TO PULL BACK THE SHEATH A LITTLE BIT TO PLACE THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3, H6 (DEVICE, METHOD) H11: B5, E1, D1, D2, D4 (MEDICAL DEVICE CATALOGUE NUMBER) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02520 | 2023-06-20 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT DURING THE PORT PLACEMENT PROCEDURE, THE INTRODUCER SEEMED TO BE LONGER. IT WAS FURTHER REPORTED THAT THE CATHETER KEPT GOING TO THE RIGHT SIDE OF THE SUBCLAVIAN VEIN WHEN THE PORT WAS PLACED ON THE LEFT SUBCLAVIAN VEIN. REPORTEDLY, THE PHYSICIAN HAS TO PULL BACK THE SHEATH A LITTLE BIT TO PLACE THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3, H6 (DEVICE, METHOD). H11: B5, E1, D1, D2, D4 (MEDICAL DEVICE CATALOGUE NUMBER). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02521 | 2023-06-20 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT DURING A CHRONIC CATHETER PLACEMENT PROCEDURE, THE INTRODUCER WAS ALLEGEDLY SEEMS TO BE LONGER. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY KEPT GOING TO THE OTHER RIGHT SIDE SUBCLAVIAN VEIN WHEN THE PORT WAS PLACED ON THE LEFT SUBCLAVIAN VEIN. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02528 | 2023-06-20 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER BASIC KIT (4F) (SINGLE-LUMEN) (20CM) | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P6154118 |  | IT WAS REPORTED, "DURING USE, THE MIDLINE LEAKS AT THE JUNCTION OF THE T AND THE PURPLE TUBING. NOTION OF SOAKED DRESSING. CURRENT PATIENT CONDITION: NO CLINICAL CONSEQUENCES REPORTED FOR THE PATIENT. NO ACTIONS TAKEN IN THE CARE FACILITY TO MANAGE THE PATIENT." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02530 | 2023-06-20 | 2023-05-18 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "PATIENT ATTENDED ON (B)(6) 2023 FOR CYCLE 3 DAY 15 OF AVD. DURING ADMINISTRATION PATIENT REPORTED DISCOMFORT IN ARM. PATIENT AND LINE ASSESSED AND SCAT COMPLETED. ON (B)(6) 2023, PATIENT CALLED HELPLINE WITH ARM PAIN AND DISCOLORATION. ATTENDED ON (B)(6) 2023 FOR ASSESSMENT. LINOGRAM ARRANGED FOR (B)(6) 2023 AND PATIENT RETURNED HOME. LINOGRAM CONFIRMED PICC FRACTURE. EXTRAVASATION POLICY FOLLOWED (DMSO, HYDROCORTISONE 1% AND COOL COMPRESS) AND PICC REMOVED. PATIENT HAS ATTENDED ON (B)(6) 2023 FOR ASSESSMENT OF EXTRAVASATION SITE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK IN THE CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR SL POWERPICC SOLO CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED BETWEEN THE 5 CM AND 6 CM MARKINGS ON THE CATHETER BODY. A KINK IMPRESSION WAS OBSERVED IN THE CATHETER TUBING ADJACENT TO THIS SPLIT. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH CATHETER DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02526 | 2023-06-20 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1385108D5 |  | IT WAS REPORTED THAT PICC APPEARS TO BE DROPPING INTO SVC. HEADING DOWNWARD. THEN, MAGNIFYING GLASS FLIPS UP TO TOP RIGHT CORNER OF SHERLOCK. TRIED TO RECALIBRATE MULTIPLE TIMES TO SEE WHETHER THAT WOULD FIX IT BUT CONTINUED TO FLIP UP AWAY FROM SVC. X-RAY CONFIRMS LOW SVC OF PICC TIP. ADDITIONAL INFORMATION RECEIVED 5/31/2023: "CUSTOMER REPORTS THEY HAVEN¿T HAD ANY NEW NURSES. CUSTOMER SUSPECTS ISSUE WITH THE PICC SENSOR; SOMETHING WITH THE TECHNOLOGY WIRE IN THE PICCS." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02512 | 2023-06-20 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT APPROXIMATELY FOUR YEARS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY FELT A SMALL METAL WIRE WAS POKING OUT OF THE PORT. IT WAS FURTHER REPORTED THAT THE NURSE WAS ABLE TO FEEL THE OBJECT WHEN CLEANING THE PORT AREA AND IT ALLEGEDLY CATCHES ON THE SWABS. REPORTEDLY, THE PATIENT ALLEGEDLY HAD A PORT ACCESS NEEDLE RIPPED FROM THE PORT ABOUT A MONTH AGO. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02527 | 2023-06-20 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT THERE WAS A LONG HAIR IN THE CATHETER ONCE THE PACKAGE WAS OPENED FOR THE FIRST TIME. THE OPERATOR DISCARDED THE CATHETER DUE TO THE CONCERN ABOUT CONTAMINATION BY THE USAGE OF THE DEVICE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02525 | 2023-06-20 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRO 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT, ON (B)(6) 2023, BEFORE THE EXTENSION GUIDE WIRE WAS PUSHED TO THE VASCULAR SHEATH, IT WAS FOUND THAT THE END WAS NOT SMOOTH AND THERE WAS A PROTRUSION, WHICH PREVENTED THE VASCULAR SHEATH FROM EXTENDING THE GUIDE WIRE INTO THE BLOOD VESSEL AND COMPLETED THE CATHETER PLACEMENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02510 | 2023-06-20 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER (TPS) STYLET FULL KIT (4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED A PICC INSERTED IN TO PATIENT WITHOUT ANY ISSUE UNTIL THE STYLET WAS BEING REMOVED AND THE STYLET WAS VERY DIFFICULT TO REMOVE LIKE IT WAS STUCK, EVENTUALLY AFTER MANY ATTEMPTS MANAGED TO REMOVE IT. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02519 | 2023-06-20 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, THE CATHETER CRACKED NEAR THE CATHETER CONNECTOR AND LEAKAGE OCCURRED FROM THE CRACK. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02509 | 2023-06-20 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER (TPS) STYLET FULL KIT (4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED A PICC INSERTED IN TO PATIENT WITHOUT ANY ISSUE UNTIL THE STYLET WAS BEING REMOVED AND THE STYLET WAS VERY DIFFICULT TO REMOVE LIKE IT WAS STUCK, EVENTUALLY AFTER MANY ATTEMPTS MANAGED TO REMOVE IT. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04515 | 2023-06-20 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER WAS DIFFICULT TO REMOVE FROM THE FOLEY CATHETER. | THE REPORTED EVENT IS UNCONFIRMED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED RECEIVED TWO (2) PHOTO SAMPLES. BOTH PHOTO SAMPLES SHOWCASES AN OVERVIEW A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING . VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) USING THE RETURNED SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN 1 MINUTE AND 39 SECONDS WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT NECESSARY DUE TO THE REPORTED EVENT BEING UNCONFIRMED. THE REPORTED EVENT IS UNCONFIRMED A RISK AND LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04517 | 2023-06-20 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT AFTER THE FOLEY CATHETER WAS INSERTED INTO THE PATIENT, IT FELL OUT, AND WHEN CHECKED, THE BALLOON HAD CRACKED AND RUPTURED. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. EVALUATION CONFIRMED THE BALLOON WAS BURST. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE LOW LATEX VISCOSITY, SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER IS A BALLOON CATHETER WITH TEMPERATURE- SENSING. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14 TO 18 EVERY 2FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. THE SURFACE OF CATHETER IS COATED FIRST WITH A TRACE AMOUNT OF METALLIC SILVER, AND THEN WITH POLYURETHANE ONTO THAT, TO INHIBIT PROLIFERATION OF BACTERIA THAT MAY ADHERE TO THE CATHETER. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. 3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 5) CONNECT LEAD WIRE TO MONITOR THROUGH EXTENSION CABLE. 6) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 7) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 8) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. 9) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 10) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 11) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 12) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 13) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 14) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 15) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 16) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 17) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02508 | 2023-06-20 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED AT 15:00 ON (B)(6) 2023, THE NURSE FOLLOWED THE DOCTOR'S ORDER TO PLACE PICC IN THE PATIENT. DURING THE PREPARATION PROCESS ACCORDING TO THE PICC PLACEMENT PROCEDURE, IT WAS FOUND THAT THE CENTRAL VENOUS CATHETER INSERTED THROUGH THE PERIPHERY WAS BLOCKED, THE RETRACTION WAS NOT SMOOTH, THE RESISTANCE TO GENTLE PUSHING WAS HIGH, AND THERE WAS NO SALINE OUTFLOW FOR WATER INJECTION. PICC PLACEMENT WAS COMPLETED AFTER REPLACING THE CENTRAL VENOUS CATHETER WITH A NEW ONE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02503 | 2023-06-20 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 4808061 | 4808061 |  | IT WAS REPORTED POST PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY EXPERIENCED THROBBING AT PORT SITE AFTER IMPLANTED. IT WAS FURTHER REPORTED THAT THERE IS NO FURTHER INFORMATION REGARDING ADDITIONAL INTERVENTION OR MEDICATION TO TREAT THE PAIN. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 01/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02524 | 2023-06-20 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT FOUR YEARS FOUR MONTHS AND TWENTY TWO DAYS POST A PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN. IT WAS FURTHER REPORTED THAT THE PORT SYSTEM WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04492 | 2023-06-20 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT COULD NOT GET IT TO DRAIN. PATIENT HAD TESTING AND TEST RESULT CAME BACK NEGATIVE INDICATING THAT THERE WAS NO INFECTION THAT WOULD PREVENT THE PRODUCT NOT TO WORK. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE PATIENT WAS A QUADRIPLEGIC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04494 | 2023-06-20 | 2023-05-01 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT THE FOLEY INDWELLING CATHETER FELL OUT AFTER PATIENT WAS IN PRONE POSITION. UPON INSPECTION OF CATHETER, TIP OF CATHETER RIPPED OFF CAUSING THE BALLOON TO DEFLATE AND CATHETER TO FALL OUT. BALLOON STILL INTACT, APPROXIMATELY 1CM OF THE TIP OF THE CATHETER WAS MISSING AND THERE WAS NO BLOOD NOTED ON CATHETER TO INDICATE ANY TRAUMA. CATHETER WAS NOT TUGGED ON OR PULLED DURING POSITIONING AND TIP NOT FOUND ON INITIAL ASSESSMENT OF AREA. DOCTORS WERE MADE AWARE AND PLAN TO INSPECT THE MEATUS AFTER THE CASE WHEN PATIENT WAS SUPINE PER DOCTOR, NOTIFIED UROLOGY WAS UNABLE TO FOUND BROKEN TIP. NURSE NOTIFIED OF EVENT AND UPDATED ON PLAN OF CARE REGARDING THE MISSING CATHETER TIP. UROLOGY CONSULTED PAGED AND INTRAOPERATIVE CYSTOGRAPHY COMPLETED TO RULE OUT FOREIGN BODY. CYSTOGRAPHY WAS NEGATIVE FOR FOREIGN BODY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04493 | 2023-06-20 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT THE CATHETER FELL OUT SPONTANEOUSLY FROM THE PATIENT. UPON CHECKING THE BALLOON, THERE WAS NO WATER IN THE BALLOON AND THE BALLOON HAD NO DAMAGE. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, INCORRECT WIRE PRESSING TECHNIQUE, INCORRECT STRIPPING TECHNIQUE ETC AND MIGHT BE CONTRIBUTE BY USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04498 | 2023-06-20 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 | FOL0102 | FOL0102 |  | IT WAS REPORTED THAT FOLEY STATLOCK WAS SEPARATED INAPPROPRIATELY AND DETACHED. PER FOLLOW-UP VIA EMAIL ON (B)(6) 2023, THE PATIENT HAD BEEN USING STATLOCK FOR A WHILE AND INDICATED THAT ONLY 1-2 HAD COME APART. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04495 | 2023-06-20 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT ELECTROCISION OF URETHRA AND BLADDER TUMOR PLUS EXCISION OF BENIGN MASSES ON THE HEAD AND FACE UNDER GENERAL ANESTHESIA ON(B)(6) 2023 AND WAS SENT BACK TO THE WARD AT 13:00. AFTER OPERATION, A 20G BARD TRIPLE-LUMEN URETHRAL CATHETER WAS SECURED PROPERLY FOR CONTINUOUS RINSING OF THE BLADDER. THE RINSED LIQUIDS WERE CLEAR. AT 13:10, THE URETHRAL CATHETER FELL OFF ON ITS OWN AND THE BALLOON WAS DEFLATED. NO OBVIOUS DAMAGE WAS OBSERVED WITH THE BALLOON OF THE URETHRAL CATHETER. AFTER THE BALLOON WAS INFLATED WITH A SYRINGE, A THIN SPLIT WAS SEEN LEAKING AIR. THE PHYSICIAN PLACED A NEW 22 G BARD TRIPLE-LUMEN URETHRAL CATHETER FOR CONTINUOUS FLUSHING OF THE BLADDER. THE RINSED LIQUIDS WERE CLEAR. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, INCORRECT WIRE PRESSING TECHNIQUE, INCORRECT STRIPPING TECHNIQUE ETC AND MIGHT BE CONTRIBUTE BY USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04499 | 2023-06-20 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON SPONTANEOUSLY FELL OUT OF THE PATIENT ON THE NEXT DAY AFTER PLACEMENT AND THE BALLOON WAS FOUND TO BE RUPTURED. THE BALLOON RUPTURE OR TEAR WAS OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BALLOON. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04501 | 2023-06-20 | 2023-06-02 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. THE PATIENT RECEIVED URINARY TRACT INFECTION FOR JUGU0870 AND JUFX2028. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04502 | 2023-06-20 | 2023-06-01 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD CONTINUOUS URINARY TRACT INFECTIONS, WAS NOT SURE IF IT WAS RELATED TO THE CATHETERS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02502 | 2023-06-20 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE DEVICE WAS ALLEGEDLY DIFFICULT TO INSERT. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 8.0FR PEEL-APART SHEATH, ONE VESSEL DILATOR AND ONE CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. BUNCHING AND BENDS WERE NOTED THROUGHOUT THE 8.0FR PEEL-APART SHEATH AND THE VESSEL DILATOR APPEARED BENT. THE DISTAL TIP OF THE VESSEL DILATOR WAS NOTED TO BE FRACTURED. AN ATTEMPT TO LOAD THE RECEIVED CATHETER INTO THE PEEL-APART SHEATH WAS PERFORMED AND THE CATHETER SUCCESSFULLY LOADED AFTER RESISTANCE WAS FELT DUE TO THE CONDITION OF THE PEEL-APART SHEATH. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED FRACTURE AND DEFORMATION ISSUE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO INSERT ISSUE, AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 02/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04505 | 2023-06-20 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316 | 119316 |  | IT WAS REPORTED THAT THE CATHETER WAS INSERTED IN THE OPERATING ROOM AND SPONTANEOUSLY WITHDREW AFTER TWO HOURS WHILE RETURNING TO THE WARD. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 07JUN2023, IT WAS STATED THAT THERE APPEARS TO BE NO DAMAGE ON THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RICK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE PUNCH DEPTH TOO LARGE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3-WAY TEMPERATURE SENSING CATHETER. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE SOLUTION IMMEDIATELY LEAKED INTO THE DRAINAGE LUMEN FROM THE BALLOON. THE BALLOON WAS DISSECTED AND FOUND THAT THE NOTCH IS DOUBLE PERFORATED. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02504 | 2023-06-20 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT THREE YEARS THREE MONTHS AND TWENTY FOUR DAYS POST PORT PLACEMENT, THE PORT ALLEGEDLY LEAKED POST TREATMENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. THREE PHOTOS WERE PROVIDED FOR REVIEW. THE PHOTO SHOWS THE TISSUE WITH BLOOD ALL OVER IT AND THE PORT PUNCTURED SITE IN THE PATIENT. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED LEAK ISSUE, AS NO OBJECTIVE EVIDENCE OF LEAK WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3, H6 (METHOD). H11: B5, H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04510 | 2023-06-20 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE DISTILLED WATER WAS UNABLE TO REMOVED FROM THE FOLEY CATHETER BALLOON DURING A PRETEST. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING WITH A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED USING RETURNED SYRINGE IN UNDER 2 MINUTES WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02506 | 2023-06-20 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE PROXIMAL END OF THE SELDINGER ROLLED UP. NO PATIENT IMPACT REPORTED. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF INABILITY TO INSERT THE GUIDEWIRE INTO THE INTRODUCER WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE SEGMENT OF DIGITAL VIDEO WHICH DEPICTED A 4.5FR MICROINTRODUCER DILATOR/SHEATH AND ONE GUIDEWIRE. THE VIDEO DEPICTED UNSUCCESSFUL ATTEMPTS TO INSERT THE GUIDEWIRE INTO THE TIP OF THE VESSEL DILATOR. NO DILATOR OR GUIDEWIRE DAMAGE WAS DISCERNIBLE IN THE VIDEO. THE VIDEO APPEARED TO DEPICT INCOMPATIBILITY BETWEEN THE INTRODUCER AND THE GUIDEWIRE; HOWEVER, INSPECTION OF THE VIDEO WAS INSUFFICIENT TO IDENTIFY THE CAUSE. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE DILATOR OCCLUSION AND GUIDEWIRE DAMAGE. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02507 | 2023-06-20 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE PROXIMAL END OF THE SELDINGER ROLLED UP. NO PATIENT HARM REPORTED. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED GUIDEWIRE TIP IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE PHOTOGRAPH OF A GUIDEWIRE WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED ONE END OF THE GUIDEWIRE PROTRUDING FROM A LUER HUB. THE GUIDEWIRE APPEARS TO HAVE A BURR ON THE WELD TIP. WHAT APPEARS TO BE USE RESIDUE WAS OBSERVED NEAR THE HUB. NO OTHER OBVIOUS DAMAGE COULD BE SEEN ON THE SAMPLE IN THE RETURNED PHOTOGRAPH. WHILE THE EXACT CAUSE OF THE DAMAGED GUIDEWIRE TIP COULD NOT BE DETERMINED FROM THE RETURNED PHOTOGRAPH, POSSIBLE CAUSES INCLUDE DAMAGE DURING MANUFACTURING, PACKAGING, HANDLING, OR USE; HOWEVER, THE REPORTED EVENT COULD NOT BE RELATED TO A DEFICIENCY IN PRODUCT MANUFACTURE OR DEFICIENCY WHILE UNDER CORRECT PRODUCT USE. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04513 | 2023-06-20 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT FOLEY CATHETER HAS AN ELEVATION AT THE BALLOON. THIS MAKES IT NEAR IMPOSSIBLE TO PLACE AS A SUPRAPUBIC CATHETER. THIS MEANS THAT THEY WERE HAVING TO USE A SMALLER DIAMETER CATHETER, I.E., 14FR, TO MAINTAIN THE TRACT OF THE SUPRAPUBIC CATHETER. THESE WORK TO HOLD THE TRACT OPEN BUT DECREASES EFFICACY OF DRAINING. THIS LEAVES THE PATIENT MORE PRONE TO OBSTRUCTIONS. CUSTOMER ASKED WAS THERE ANY OTHER ALTERNATIVES FOR 16FR STRAIGHT CATHETERS THAT THEY CAN USE. PER FOLLOW UP VIA MAIL RECEIVED ON (B)(6)2023, THE ISSUES CUSTOMER HAD BEEN WITH THE NON-LATEX CATHETERS. ON (B)(6)2023 PATIENT WHO HAD A 16FR SUPRAPUBIC CATHETER SINCE (B)(6)2022 WAS UNABLE TO HAVE 16FR REPLACED, HAD TO DOWNSIZE TO 14FR. ON (B)(6)2023PATIENT BEGAN WITH UNC UROLOGY (B)(6)2022 WITH SPT IN PLACE. ATTEMPTED TO PLACE 18FR CATHETER (SAME SIZE SINCE 10/22), UNSUCCESSFULLY. UNABLE TO PLACE 16FR, PROVIDER PAGED. BY THE TIME OF PROVIDER'S ARRIVAL, NO SPT COULD BE PLACED. PATIENT HAD TO GO THROUGH INTERVENTIONAL RADIOLOGY TO HAVE SPT REESTABLISHED. THERE WERE TWO MORE, BUT AFTER OUR INTERNAL REPORTS WERE PLACED, THE EXACT INFORMATION WAS FORGOTTEN. THE PROBLEM WAS THE RIDGE ON THE CATHETER BALLOON. IT DID NOT EASILY PASS INTO THE BLADDER. THE OTHER ISSUE WITH HAVE WITH ALL OF THE BD CATHETERS WAS THE COST. THERE WAS ONLY ONE OPTION THAT WAS LATEX FREE. THIS OPTION WAS 2-3 TIMES THE COST OF PREVIOUS CATHETERS USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04516 | 2023-06-20 | 2023-06-01 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE NURSE ATTEMPTED TO DRAIN THE STERILE WATER DURING A ROUTINE EXCHANGE BUT COULD NOT REMOVE THE STERILE WATER. THE DOCTOR ALSO TRIED TO REMOVE BUT COULD NOT DRAIN THE STERILE WATER, SO THE DOCTOR PUNCTURED THE BALLOON UNDER ULTRASOUND AND RUPTURED TO REMOVE THE BALLOON. THE CATHETER WAS CLAMPED TWICE WEEKLY DURING BLADDER IRRIGATION. | THE REPORTED EVENT IS INCONCLUSIVE DUE TO POOR SAMPLE CONDITION. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS EVALUATED AND OBSERVED INFLATION EYE ARE MEET INTERNAL SPECIFICATION. HOWEVER, DUE TO POOR SAMPLE CONDITION THE COMPLETE INVESTIGATION COULD NOT BE PERFORMED AND THIS COMPLAINT IS CLASSIFIED AS INCONCLUSIVE-POOR SAMPLE CONDITION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE THIN RUBBERIZE WALL (EXAMPLE: CAUSING COLLAPSE/ PINCHED INFLATION LUMEN UNDER THE BALLOON) OR USER RELATED-MISHANDLING DURING USED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE 1. MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04511 | 2023-06-20 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50612 | 50612 |  | IT WAS REPORTED THAT ABOUT 20 CATHETERS DID NOT HAVE SALINE IN THE PACKET, AND PATIENT HAD TO THROW THEM AWAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04509 | 2023-06-20 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WAS A CRACK NEAR THE BASE OF THE THREE CROTCHES OF THE FOLEY CATHETER. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED A 3-WAY TEMP SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED NO CRACKS OR TEARS ON THE SHAFT OR THE CATHETER ITSELF. ATTEMPTED TO INFLATE BALLOON WITH 10 ML OF SOLUTION AND IT IMMEDIATELY LEAKED OUT OF A 0.011" PINHOLE FOUND ON THE SHAFT OF THE CATHETER AND LOCATED 0.2755" FROM THE TRIFURCATION. THERE WAS ALSO A LEAK FROM THE INFLATION LUMEN INTO THE DRAINAGE LUMEN RESULTING IN A LUMEN TO LUMEN LEAK. ALSO RECEIVED (B)(4) PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWS TOP VIEW OF CATHETER. SECOND PHOTO SAMPLE ZOOMS IN LOCATION OF WHERE PINHOLE IS PRESENT ON THE SHAFT. THEREFORE, PRODUCT DOES NOT MEET SPECIFICATIONS ACCORDING TO (B)(4) REV 21 "CAP AND VALVE MUST BE PRESENT AND ASSEMBLED THE TIP OF THE BLACK MATERIAL OF THE VALVE IS VISIBLE ON THE SURFACE OF THE CAP [2] WITHOUT CUTS, HOLES OR TEARS ON THE INFLATION ARM." ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE USER APPLIES EXCESSIVE TENSION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS]. 1. METHOD FOR USE. (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]. 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. SIZES OF CATHETERS. AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ CORRECTIONS: D, F, H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04507 | 2023-06-20 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT FOLEY CATHETER LEAKED THROUGH THE TOP OF THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02505 | 2023-06-20 | 2023-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 20G X 10 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F320108PT |  | IT WAS REPORTED THAT DEVICE DID NOT PROPERLY SAFETY AND NEEDLE WAS EXPOSED. PATIENT AND CLINICIAN WERE NOT HARMED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04512 | 2023-06-20 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT THEY PLACED A FOLEY ON A PATIENT IN PCU, AND AS THE NURSE INSERTED IT, THEY FOUND A PRIMARY PART TO HAVE A CRACK IN IT THAT WAS NOT APPARENT TO THE NAKED EYE UPON INSERTION. THIS CAUSED THEM TO REMOVE, RE-PERFORM STERILE TECHNIQUE AND RE-INSERT A FRESH FOLEY CATHETER. PER FOLLOW UP VIA EMAIL ON (B)(6) 2023, IT WAS STATED THAT THE PATIENT WAS STABLE, PLACED A NEW PRODUCT, NO ISSUES. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE CRACK IN VALVE HOUSING. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), DRAINAGE BAG ATTACHED TO A 2-WAY CATHETER. VISUAL INSPECTION NOTED A CRACK ON THE INFLATION VALVE OF THE CATHETER. THIS IS OUT OF SPECIFICATION PER INSPECTION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE. (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS. PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES]. BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. MATERIAL: BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO 22 EVERY 2FR. 1. BALLOON CATHETER: FOLEY CATHETER. TEMPERATURE-SENSING CATHETER. 2. ACCESSORIES: CLOSED DRAINAGE BAG. (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) 3. MALFUNCTION AND ADVERSE EVENTS. 1) MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE; DIFFICULTY OR FAILURE TO REMOVE THE DEVICE; OCCLUSION OF CATHETER INNER LUMENS; ENCRUSTATION; ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE; DEVICE DAMAGE DUE TO INAPPROPRIATE USE; FAILURE TO MEASURE TEMPERATURE; IMPROPER TEMPERATURE INDICATION. 2) ADVERSE EVENTS: URINARY-TRACT INFECTION; HEMORRHAGE, HEMATURIA; ALLERGY REACTION TO THE DEVICE; CALCULUS FORMATION; EDEMA; PAIN; DISCOMFORT; INJURY OF BLADDER OR URETHRAL; URETHRITIS, URINARY INCONTINENCE; RETAINED BALLOON FRAGMENTS; [STORAGE METHOD AND EXPIRATION DATE]. 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
MW5118702 | 2023-06-21 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | 3 FR. POLY MIDLINE | CATHETER, INTRAVASCULAR, THERAPEUTIC, LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. |  |  |  | DURING INSERTION OF 3 FR. POLY MIDLINE, ONE HALF OF THE HUB OF THE PEEL-AWAY SHEATH POPPED OFF OF THE ALREADY SPLIT BODY OF THE SHEATH WHILE IT WAS STILL IN THE PATIENT. I WAS ABLE TO GRASP THE END WITH MY FINGERS AND REMOVE IT ENTIRELY FROM THE PATIENT'S ARM. THE HUB AND THE BODY ARE NOT SUPPOSED TO SEPARATE FROM EACH OTHER AT ANY TIME DURING THE PROCEDURE. LOT # REGT4350 BARD. SHEATH IS AT THE DESK OF THE IV(INTRAVENOUS) TEAM OFFICE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04552 | 2023-06-21 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT ON (B)(6)2023, AFTER THE OPERATION, THE PATIENT WAS REQUIRED TO INDWELL WITH AN F20 URINARY CATHETER AND THE BALLOON WAS INJECTED FOR 30ML. THEN THE BALLOON BURST AND THE URINARY CATHETER FELL OUT, WHICH CAUSED THE PATIENT DIFFICULT TO URINATE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04564 | 2023-06-21 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI16 |  | IT WAS REPORTED THAT THE INFLATION VALVE CAME OFF IMMEDIATELY AFTER THE FOLEY CATHETER WAS PLACED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04563 | 2023-06-21 | 2023-06-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD A COUPLE OF URINARY TRACT INFECTIONS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTIONS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04561 | 2023-06-21 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE INFLATION VALVE UPON ATTEMPTING TO INJECT STERILE WATER INTO THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04560 | 2023-06-21 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE INFLATION VALVE OF THE FOLEY CATHETER. THIS EVENT OCCURRED IN THE OPERATION ROOM. THE DEVICE WAS NOT USED ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04553 | 2023-06-21 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER COULD NOT BE DRAINED. THE DEVICE WAS NOT USED ON THE PATIENT. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEM OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK®FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE 1. MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04557 | 2023-06-21 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS INFLATED BUT DEFLATED A FEW HOURS LATER. NO PINHOLES WERE OBSERVED. THE EVENT OCCURRED SEVERAL HOURS AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BALLOON. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. BASED ON THE EVALUATION, A TEAR WAS FOUND ON THE RUBBERIZED LAYER CAUSING INTERLUMEN LEAKAGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON UNIFORM RUBBERIZED THICKNESS ETC CAUSING TORN RUBBERIZE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® I.C. SILVER FOLEY TRAY B SINGLE USE ONLY. [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER; WATER SOLUBLE LUBRICANT; ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION; TWEEZERS; GAUZE PADS; WATERPROOF SHEET; COTTON BALLS; GLOVES; STATLOCK® FOLEY. [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS. 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE; - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE; - OCCLUSION OF CATHETER INNER LUMENS; - ENCRUSTATION; - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE; - DEVICE DAMAGE DUE TO INAPPROPRIATE USE; 2) ADVERSE EVENTS - URINARY-TRACT INFECTION; - HEMORRHAGE, HEMATURIA; - ALLERGY REACTION TO THE DEVICE; - CALCULUS FORMATION; - EDEMA; - PAIN; - DISCOMFORT; - INJURY OF BLADDER OR URETHRAL; - URETHRITIS, URINARY INCONTINENCE; - RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE, STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE, INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D, H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04555 | 2023-06-21 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE MIDDLE OF THE FOLEY CATHETER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04554 | 2023-06-21 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HEMATURIA FOLEY CATHETER | HEMATURIA FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 2551H22 |  | IT WAS REPORTED THAT THE FOLEY CATHETER GOT TORN LONGITUDINALLY AND FELL OUT OF THE PATIENT AFTER 5 DAYS OF PLACEMENT. THE USER DID NOT USE FORCEPS DURING THE CATHETER INSERTION. HEMATURIA WAS OBSERVED. IT WAS SUSPECTED THAT THE PATIENT HAD BLADDER STONE. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6)2023, THE CUSTOMER DID NOT CONFIRM IF THE BLADDER STONE WAS PATIENT'S MEDICAL HISTORY. PER MAIL RECEIVED FROM IBC ON (B)(6)2023, STATED THAT THE BALLOON BURST WAS OBSERVED TO THE RETURNED SAMPLE. "CATHETER BREAKAGE" WAS WRONG REPORT, THE EVENT DESCRIPTION SHOULD BE CORRECTED TO "BALLOON BURST". | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04558 | 2023-06-21 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE SUPPLIES WERE NOT PACKAGED CORRECTLY, SOME OF THE SUPPLIES ARE STUCK TOGETHER AND THE LIQUID PACKET DID NOT HAVE THE STERILE SALINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04539 | 2023-06-21 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50812G | 50812G |  | IT WAS REPORTED THAT THE CUSTOMER HAD TROUBLE WITH INTERMITTENT CATHETERS. THEY ORDERED 15 BOXES, AND SOME OF THE CATHETERS DID NOT ENTER THE BLADDER. 3.65 ML DID NOT ENTER, AND 3.43 ML DID ENTER. PER FOLLOW UP VIA EMAIL ON 01MAY2023, IT WAS REPORTED THAT THE CUSTOMER HAD GOOD THEY WORKED AND BAD THEY WOULD NOT ENTER THE BLADDER SAMPLES FROM TWO LOTS (LOT# JUGU0844) AND (LOT# JUGV1290). CUSTOMER CURRENTLY HAD TWELVE BOXES OF LOT # JUGV1290 AND BOUGHT 18 BOXES TOTALLY. PER SAMPLE RECEIVED ON 08MAY2023, IT WAS STATED THAT THE CUSTOMER RETURNED THREE ADDITIONAL USED INTERMITTENT CATHETERS SAMPLES FOR LOT# JUGU0844 AND THREE ADDITIONAL USED SAMPLES FOR LOT# JUGV1290. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04549 | 2023-06-21 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THEY HAD SEVERAL TEMPERATURE SENSING FOLEYS WITH BARD NOT WORKING OVER THE LAST 2-3 WEEKS. THEY TROUBLE SHOOT THE MONITORS, REPLACED CABLES AND STILL WERE UNABLE TO GET TEMPERATURE ON THE MONITOR. THEY TESTED THE CABLE, AND IT WAS NOT THE CABLE OR THE MONITOR WHICH MEANT IT WAS THE FOLEY. IT HAPPENED AGAIN IN NS 5 AND WAS STARTING TO RAISE A RED FLAG. PER ADDITIONAL INFORMATION RECEIVED ON 05JUN2023, ANOTHER FOLEY PLACED LAST NIGHT WAS NOT WORKING UNFORTUNATELY THE PACKAGE WAS NOT KEPT. | PER INVESTIGATOR EVALUATION, BD HAS DETERMINED THAT THIS MDR EVENT IS NOT REPORTABLE. H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04543 | 2023-06-21 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS RUPTURED, AND STERILE WATER WAS INJECTED BUT COULD NOT BE USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04541 | 2023-06-21 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTERMITTENT ROCHESTER CATHETERS | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT CUSTOMER WAS HAVING TROUBLE GETTING THE CATHETER, WHICH IS A MAGIC3 CATHETER TO LUBRICATE PROPERLY. CUSTOMER OPENED THE SILVER PACKET AND POURED IT IN BUT IT DID NOT SEEM TO BE HELPING. CUSTOMER WAS EXPERIENCING DISCOMFORT DUE TO IT NOT BEING LUBRICATED ENOUGH. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND LOT NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04542 | 2023-06-21 | 2023-06-02 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATED PATIENT RECEIVED A NASTY URINARY TRACT INFECTION (JUGU0870) AND (JUFX2028). IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "REVIEW THE SELF CATHETERIZATION: PROCEDURE: WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
17170296 | 2023-06-21 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP INTERMITTENT CATHETER TRAY | CATHETER, UROLOGICAL | C. R. BARD, INC. | INTP14 | INTP14 |  | ATTEMPTING TO STRAIGHT CATH PATIENT. NURSE TECH OPENED STERILE KIT AND STARTED UNFOLDING PAPER/WRAPPING WHEN IT WAS NOTED THAT CATHETER WAS ON THE OUTSIDE OF THE PAPER/WRAPPING. COULD NOT OPEN LAST FOLD OF PAPER WITHOUT CONTAMINATING CATHETER/BREAKING STERILITY. SECOND KIT WAS OBTAINED. SECOND KIT ALSO HAD PART OF CATHETER ON TOP OF PAPER/NOT IN TRAY. THE NURSE TECH WAS ABLE TO CAREFULLY, AND SLOWLY, SLIDE THE PAPER/WRAPPING OPEN WITHOUT CONTAMINATING CATHETER. CONCERNS THAT THERE ARE MORE KITS PACKAGED THIS WAY AND THAT THERE WILL BE WASTE AND OR PATIENTS STRAIGHT CATHED WITH CONTAMINATED CATHETER. |  | N | Y | OTHER |  | 2023-06
1018233-2023-04537 | 2023-06-21 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165L16 | 0165L16 |  | IT WAS REPORTED THAT THE PATIENT WAS IN THE BATHROOM COMPLETING PERI CARE AND UPON NURSE ENTERING BATHROOM PATIENT WAS HOLDING THE END OF THE FOLEY CATHETER AND BALLOON WAS NOT INTACT WITH THE SALINE. PATIENT STATED THAT IT JUST FELL OUT. NURSE ASSESSED THE END OF THE FOLEY BALLOON, AND IT APPEARED THERE WAS A TEAR IN THE BALLOON PART OF THE FOLEY. THE NURSE MANAGER SAID THE NURSE THAT WAS ON TODAY TOLD THEM IT LOOKED LIKE THE BALLOON EXPLODED, IT HAD A BIG HOLE. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® EZ-LOK® SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK (FIG. 1A) OR SLIP TIP SYRINGES (FIG. 1B). OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. SWAB SURFACE OF BARD® EZ-LOK® SAMPLING PORT WITH ANTISEPTIC WIPE. (FIG. 2). USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT (AT APPROXIMATELY 80 - 100 DEGREE ANGLE). PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. (FIG. 3). SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND REMOVE SYRINGE FROM SAMPLE PORT. (FIG. 4) UNKINK TUBING AND TRANSFER URINE SPECIMEN INTO SPECIMEN CUP OR FOLLOW HOSPITAL PROTOCOL. DISCARD SYRINGE ACCORDING TO HOSPITAL PROTOCOL. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB. STERILE: CONTENTS OF INNER WRAP ARE STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RE-STERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04533 | 2023-06-21 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT BALLOON COULD NOT BE DEFLATED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02540 | 2023-06-21 | 2023-05-26 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT EXPERIENCING SWELLING AROUND THE PORT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04530 | 2023-06-21 | 2023-06-01 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETER CAUSED URINARY TRACT INFECTIONS. PER FOLLOW UP VIA PHONE ON 07JUN2023, IT WAS REPORTED THAT THE CUSTOMER WAS IN THE HOSPITAL ON 2 SEPARATE OCCASIONS DUE TO A URINARY TRACT INFECTION SHE ACQUIRED WHILE USING THE PUREWICK EXTERNAL FEMALE CATHETER. THE MOST RECENT HOSPITAL VISIT WAS ON (B)(6) 2023, IN WHICH PATIENT WAS PRESCRIBED ANTIBIOTICS FROM THE DOCTOR FOR THE INFECTION. CUSTOMER WAS REQUESTING THAT PATIENT RECEIVE SOME TYPE OF COMPENSATION OR REIMBURSEMENT DUE TO THE NATURE OF THE INCIDENT AND THE FACT THAT PATIENT CAN NO LONGER USE THE PRODUCT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE COMPLETED DUE TO NO LOT NUMBER PROVIDED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE. " DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. " " REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER." "ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS." (B)(6) REV 0 IS FOUND TO BE ADEQUATE TO FULFILL (B)(6) REVISION 18 AS IT STATES : PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE:- EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE. "¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS." "6. REPLACE THE PUREWICK¿ FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICK¿ FEMALE EXTERNAL CATHETER." "¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02533 | 2023-06-21 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT LONG-TERM HEMODIALYSIS CATHETER, POLYURETHANE, KIT, 14.5F | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5703690 | 5703690 |  | IT WAS REPORTED THAT DURING THE PREPARATION OF A PORT PLACEMENT PROCEDURE, A SMALL HOLE WAS ALLEGEDLY FOUND ON THE CATHETER AT THE CUFF. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 19CM HEMOSPLIT D/L CATHETER, ONE TUNNELER, ONE 15FR PEEL-APART SHEATH AND VESSEL DILATOR, ONE 8FR DILATOR, ONE 12FR DILATOR, TWO END CAPS, ONE INTRODUCER NEEDLE, ONE J-TIP GUIDEWIRE IN A GUIDEWIRE HOOP AND ONE INTRODUCER NEEDLE WERE RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. A SMALL PINHOLE WAS NOTED UNDER THE TISSUE CUFF AND UPON INFUSION, A SMALL LEAK WAS NOTED AROUND THE TISSUE CUFF. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, THREE ELECTRONIC PHOTOS WERE PROVIDED FOR REVIEW. THE PHOTO SHOWS THE CATHETER AND A SMALL PINHOLE WAS NOTED UNDER THE TISSUE CUFF. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED MATERIAL PUNCTURE/HOLE ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 12/2023), H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04550 | 2023-06-21 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THEY HAD SEVERAL TEMPERATURE SENSING FOLEYS WITH BARD NOT WORKING OVER THE LAST 2-3 WEEKS. THEY TROUBLE SHOOT THE MONITORS, REPLACED CABLES AND STILL WERE UNABLE TO GET TEMPERATURE ON THE MONITOR. THEY TESTED THE CABLE, AND IT WAS NOT THE CABLE OR THE MONITOR WHICH MEANT IT WAS THE FOLEY. IT HAPPENED AGAIN IN NS 5 AND WAS STARTING TO RAISE A RED FLAG. PER ADDITIONAL INFORMATION RECEIVED ON 05JUN2023, ANOTHER FOLEY PLACED LAST NIGHT WAS NOT WORKING UNFORTUNATELY THE PACKAGE WAS NOT KEPT. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS INITIALLY REPORTED IN ERROR AS THIS EVENT IS NOT REPORTABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04556 | 2023-06-21 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L20 |  | IT WAS REPORTED THAT THE USER FACED DIFFICULTY WHILE INJECTING STERILE WATER INTO THE FOLEY CATHETER. THE USER TRIED TO INJECT STERILIZED WATER AFTER OPENING THE PACKAGE, BUT IT WAS HARD AND THE WATER WOULD NOT GO INSIDE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS: (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS: (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX. [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR: 1. FOLEY CATHETER. 2. ACCESSORIES. CLOSED DRAINAGE BAG: (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER, WATER SOLUBLE LUBRICANT, ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION, TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS, GLOVES, STATLOCK® FOLEY. [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE: (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE 1. MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING.¿ <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD: WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD: 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17). 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE, DIFFICULTY OR FAILURE TO REMOVE THE DEVICE, OCCLUSION OF CATHETER INNER LUMENS, ENCRUSTATION, ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE, DEVICE DAMAGE DUE TO INAPPROPRIATE USE. 2) ADVERSE EVENTS: URINARYTRACT INFECTION, HEMORRHAGE, HEMATURIA, ALLERGY REACTION TO THE DEVICE, CALCULUS FORMATION, EDEMA, PAIN, DISCOMFORT, INJURY OF BLADDER OR URETHRAL, URETHRITIS, URINARY INCONTINENCE, RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04532 | 2023-06-21 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WAS WATER LEAKAGE FROM THE FUNNEL AREA OF THE FOLEY CATHETER. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. IT WAS UNKNOWN IF THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND IT IMMEDIATELY LEAKED OUT OF A 0.012" PINHOLE FOUND ON THE BALLOON WITH NO ADDITIONAL LEAKS NOTED AROUND THE INFLATION FUNNEL. THIS IS OUT OF SPECIFICATION, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02534 | 2023-06-21 | 2021-04-19 | D, e, a, t, h | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808001 | 1808001 |  | IT WAS REPORTED THAT APPROXIMATELY SEVEN MONTHS AND TWO DAYS POST PORT PLACEMENT, THE CATHETER ALLEGEDLY FRACTURE WITH EXTRAVASATION JUST PROXIMAL TO THE ENTRY POINT OF THE CATHETER. IT WAS FURTHER REPORTED THAT THE PATIENT ALLEGEDLY EXPERIENCED CHEST PAIN IN RELATION TO THE CATHETER. REPORTEDLY THE PATIENT EXPIRED. | H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. HOWEVER, MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THE MEDICAL RECORDS INDICATE THAT THE CATHETER WAS COILED JUST PRIOR TO THE ENTRANCE TO THE VEIN AND ALSO THAT IT HAD A FRACTURE AND ASSOCIATED LEAKAGE/EXTRAVASATION OF CONTRAST OUTSIDE OF THE VEIN. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED CATHETER FRACTURE, COILING, AND LEAKAGE. THIS COMPLAINT IS INVOLVED WITH LITIGATION. AT THIS TIME, THE RELATIONSHIP BETWEEN THE FRACTURE, COILING, AND LEAKAGE AND THE REPORTED DEATH IS UNKNOWN. THE DEFINITIVE ROOT CAUSE OF THESE EVENTS COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 08/2021), G2, G3, H6 (PATIENT, DEVICE, METHOD) H11: B3, B5, H6 (RESULT) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02539 | 2023-06-21 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT DUO M.R.I. IMPLANTABLE PORT, CHRONOFLEX DUAL-LUMEN, 9.5F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1829500 | 1829500 |  | IT WAS REPORTED THAT TWENTY ONE DAYS POST PORT PLACEMENT, THE SEPTUM OF THE PORT ALLEGEDLY DISLODGED OFF THE RIM OF THE PORT. REPORTEDLY. THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: (B)(6) 2024), G3, H6 (METHOD) H11: H6 (RESULT, CONCLUSION) H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02532 | 2023-06-21 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT DURING THE PREPARATION FOR A PORT PLACEMENT PROCEDURE, THE CATH LOCK ALLEGEDLY ADHERED TO THE TYVEK LID. IT WAS FURTHER REPORTED THAT THE CATH LOCK FELL TO THE FLOOR AFTER OPENING THE KIT AND THEREFORE COULD NOT BE USED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: D4 (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02535 | 2023-06-21 | 2023-05-19 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT ONE YEAR FIVE MONTHS AND TWENTY-THREE DAYS POST PORT PLACEMENT IN THE RIGHT INTERNAL JUGULAR VEIN, PORT BODY ALLEGED LEAKED SUBCUTANEOUSLY. IT WAS FURTHER REPORTED THAT CATHETER ALLEGEDLY BROKEN AND CATHETER SEGMENTS WERE REMOVED. REPORTEDLY THE DISTAL CATHETER SEGMENT HAD PERFORATED THE WALL OF THE SUPERIOR VENA CAVA AND MIGRATED TO THE RIGHT ATRIUM AND REMAINS INSIDE OF THE PATIENT¿S BODY. FURTHERMORE, PATIENT ALLEGEDLY DEVELOPED MEDIASTINITIS. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO THE GROSHONG CATHETER WERE RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED AT THE DISTAL END OF THE ATTACHED CATHETER THAT WAS ELLIPTICAL IN SHAPE. WHAT APPEARED LIKE A SPLIT, WAS NOTED ON THE PORT SEPTA. TOOL DAMAGE WAS NOT NOTED TO THE CATH-LOCK. THE EDGES OF THE COMPLETE CIRCUMFERENTIAL BREAK ON THE DISTAL END OF THE ATTACHED CATHETER WERE NOTED TO BE UNEVEN AND THE SURFACE WAS NOTED TO BE ROUND AND SMOOTH IN ONE REGION AND GRANULAR IN THE OTHER REGION. UPON INFUSION, THROMBUS WITH WATER WAS OBSERVED EXITING THE PORT STEM. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, LEAK AND MATERIAL SEPARATION AND IDENTIFIED MATERIAL SPLIT, WEAR AND DEFORMATION ISSUES. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MIGRATION ISSUES AS THERE WAS NO OBJECTIVE EVIDENCE OBTAINED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED, FLEXURAL FATIGUE (REPEATED/CYCLIC KINKING OF THE CATHETER WHICH CAN CAUSE GRADUAL TEARING IN THE CATHETER) COULD HAVE POTENTIALLY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 08/2023), H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02538 | 2023-06-21 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1618000 | 1618000 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY HAD CRACK. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE POWERPORT CLEARVUE SLIM IMPLANTABLE PORT ATTACHED TO A CATHETER WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC VISUAL AND FUNCTIONAL TESTING WERE PERFORMED. TACTILE EVALUATION WAS PERFORMED ON THE CATHETER RETURNED PRIOR TO FUNCTION TESTING AND THE CATHETER WERE NORMALLY ELASTIC AFTER GENTLY STRETCHED. AN IN-HOUSE SYRINGE WAS ATTACHED TO A SAFETY INFUSION SET AND THE NON-CORING NEEDLE WAS INSERTED INTO THE PORT SEPTUM WITH THE ATTACHED CATHETER AND WAS PATENT TO BOTH INFUSION AND ASPIRATION. THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED FRACTURE ISSUES AS THERE WERE NO SPLITS OR CRACKS WERE NOTED ON THE CATHETER AND NO LEAKS WERE OBSERVED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: (B)(6) 2024), G3 H11: H6 (METHOD, RESULT, CONCLUSION) H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02544 | 2023-06-22 | 2023-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8274118 |  | IT WAS REPORTED A PATIENT WHO DURING THE SHIFT PRESENTS MOISTURE IN THE HEALING AREA, IT IS OBSERVED WITH DRAINAGE AND MOISTURE IN THE PATIENT BED, EVIDENCE OF SOLUTION LEAKAGE BETWEEN THE CATHETER AND THE BODY WHEN WASHING, FINDING THE OCCLUDED POWER PATHWAY. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02556 | 2023-06-22 | 2023-06-07 | I, n, j, u, r, y | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. DUAL LUMEN PLACED IN THE UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02554 | 2023-06-22 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED UPON OPENING THE STERILE SET, A HAIR WAS NOTED IN THE STERILE PACKAGE. OCCURRED BEFORE USE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF FOREIGN MATERIAL IS INCONCLUSIVE DUE TO THE STATE OF THE RETURNED SAMPLE. TWO PHOTOGRAPHS OF A KIT TRAY WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPHS SHOWED A HAIR IN WHAT APPEARS TO BE AN OPENED KIT TRAY. THE OPENED STATE OF THE PACKAGING MADE IT DIFFICULT TO ASSESS WHEN AND WHERE THE FOREIGN MATERIAL ENTERED THE PACKAGE. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. MANUFACTURING CONTROLS ARE IN PLACE TO MITIGATE THE OCCURRENCE OF THIS TYPE OF FAILURE. THE COMPLAINT OF FOREIGN MATERIAL COULD NOT BE INDEPENDENTLY CONFIRMED WITHOUT EVALUATION OF A SEALED KIT; HOWEVER, THE REPORT OF FOREIGN MATERIAL HAS BEEN DOCUMENTED AND INCLUDED IN COMPLAINT TRENDING. THIS TYPE OF COMPLAINT WILL CONTINUE TO BE MONITORED AS PART OF ONGOING EFFORTS TO IDENTIFY POTENTIAL MANUFACTURING RELATED ISSUES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04586 | 2023-06-22 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE BALLOON AREA DURING PRE-TEST. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 09JUN2023, THE BALLOON WAS RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02555 | 2023-06-22 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (22G) (8CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | F122080T |  | CUSTOMER REPORTED THAT DURING PLACEMENT, GUIDEWIRE WENT IN SMOOTHLY. WHEN CLINICIAN WENT TO THREAD THE CATHETER OFF OF THE NEEDLE, THEY FELT RESISTANCE. THEY PULLED THE CATHETER BACK AND THEN PULLED BACK THE GUIDEWIRE WITH THE PURPLE ADVANCER. WHEN THEY PULLED THE WHOLE DEVICE OUT OF THE PATIENT, THEY NOTICED THE GUIDEWIRE WAS BENT. AS THEY WERE LOOKING AT THE GUIDEWIRE, THEY NOTICED THAT IT LOOKED UNRAVELED AND IT THEN JUST FELL OFF ONTO THE GROUND. PATIENT HAD TO BE RE-STUCK. SECOND INSERTION WAS SUCCESSFUL. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04587 | 2023-06-22 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316 | 119316 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04584 | 2023-06-22 | 2023-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE BALLOON. UPON ATTEMPTING TO PLACE THE CATHETER ON THE PATIENT, THERE WAS A HOLE IN THE BALLOON AND WATER LEAKED FROM THERE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04583 | 2023-06-22 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175814 |  | IT WAS REPORTED, THAT THE BALLOON OF THE FOLEY CATHETER WAS LEAKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04577 | 2023-06-22 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPECIALTY TUBES & CATHETER | SPECIALTY TUBES & CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE NURSE GOT AN ALARM 14 (PROBE OUT OF RANGE) WHEN TRIED TO CONNECT FOLEY TO THE ARCTIC SUN DEVICE AND TRIED OUTLET CONTROL TEMPERATURE (T2) BUT COULD NOT GET A READING. EXPLAINED OUTLET CONTROL TEMPERATURE (T2) MUST BE ENABLED, HOWEVER IT WOULD NOT RUN THERAPY. EXPLAINED THEY NEED TO CHANGE THE CABLE AND SEE IF THIS RESOLVED THE ISSUE. IF NOT, THE PROBE MIGHT NEED TO BE REPLACED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02543 | 2023-06-22 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY (WITH 70CM NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174355 |  | IT WAS REPORTED THE PATIENT HOSPITALIZED FOR SYLVIAN INFARCTION, HAD A PICC-LINE INSERTED ON (B)(6) 2023, FEBRILE SPIKES APPEARED WITH SKIN INFLAMMATION AT THE PICC-LINE SITE. BLOOD CULTURES TAKEN ON THE DEVICE THE SAME DAY WERE POSITIVE FOR PSEUDOMONAS AERUGINOSA. PICC REMOVED, PATIENT PUT ON ANTIBIOTIC TREATMENT VIA PERIPHERAL CATHETER INSERTED ON (B)(6) 2023, BUT INFLAMMATORY SIGNS, SO WAITING FOR NEW PICC-LINE TO BE INSERTED. DIFFICULTY COLLECTING BLOOD TESTS. ADDITIONAL INFORMATION PROVIDED 6/26. DIFFICULT TO DRAW BLOOD FROM THIS PATIENT. CATHETER MAINTENANCE PRACTICE: IN-HOUSE PROTOCOL DRAWN UP BY THE HOSPITAL HYGIENE DEPARTMENT. CHANGE OF ABSORBENT DRESSING EVERY 72-96H (OR AS SOON AS SOILED/WET/UNSTUCK) WITH STERILE GLOVES AND 4-STAGE ANTISEPSIS OF CATHETER INSERTION AREA: CLEANING WITH MILD LIQUID SOAP, RINSING WITH STERILE WATER, DRYING, APPLICATION OF ALCOHOLIC ANTISEPTIC. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04576 | 2023-06-22 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TIEMANN MODEL UROLOGICAL CATHETER COUDE TIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 010114 | 010114 |  | IT WAS REPORTED THAT THE CATHETERS WERE BENT AND TWISTED MAKING THEM HARD TO INSERT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "DAMAGED PRODUCT DURING SHIPPING". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSE BUT IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04575 | 2023-06-22 | 2023-06-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT USED THE INTERMITTENT MAGIC 3 GO PATIENT HAD NONE LEFT AT ALL BUT WAS REPORTING PATIENT HAD A URINARY TRACT INFECTION IN APRIL OF 2022 WAS IN THE HOSPITAL FOR 8 DAYS AND ENDED UP WITH SEPSIS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04571 | 2023-06-22 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 175812 |  | IT WAS REPORTED THAT THE PATIENT STATED THAT WHEN THE CATHETER WAS INSERTED TO HIM, THE CATHETER FELL OUT AND THE BALLOON WAS DEFLATED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 09JUN2023, STATED THAT THERE WERE NO VISIBLE DAMAGES IN THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04570 | 2023-06-22 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED OUT FROM THE FOLEY CATHETER AND SPONTANEOUSLY FELL OUT OF THE PATIENT. UPON CHECKING THE CATHETER, IT WAS FOUND THAT THERE WAS A PINHOLE. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. BASED ON THE EVALUATION, A TEAR WAS FOUND ON THE RUBBERIZED LAYER CAUSING INTERLUMEN LEAKAGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON UNIFORM RUBBERIZED THICKNESS ETC CAUSING TORN RUBBERIZE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. (THE URETHRA MAY BE INJURED) (2) DO NOT PULL THE CATHETER HARD. (THE BLADDER/URETHRA MAYBE INJURED) 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER (WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED). CONTRAINDICATIONS 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE ON PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04572 | 2023-06-22 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE PATIENT GOT THE URINARY CATHETER ON (B)(6) 2023. ON (B)(6), 2023, WHEN THE NURSE ROUTINELY DISINFECTED THE PATIENT'S URETHRA, THE CATHETER FELL OUT DURING THE DISINFECTION PROCESS AND THE BALLOON WAS DEFLATED. WHEN ASPIRATED THE BALLOON WITH A 20ML SYRINGE, ONLY A SMALL AMOUNT OF FLUID WAS DRAWN OUT. THEN THE PATIENT WAS REPLACED WITH A NEW CATHETER AND THERE WAS NO IMPACT ON THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, INCORRECT WIRE PRESSING TECHNIQUE, INCORRECT STRIPPING TECHNIQUE ETC AND MIGHT BE CONTRIBUTE BY USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04573 | 2023-06-22 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119416M | 119416M |  | IT WAS REPORTED THAT THE TEMPERATURE SENSING FOLEY WAS NOT REGISTERING ANY TEMPERATURE. THERE WERE TWO NON-WORKING TEMPERATURE FOLEYS MATERIAL# 119416M AND MATERIAL# 119216M WHEN THEY WERE ON-SITE. THE FOLEY WAS STILL IN PATIENT AND REQUESTED THAT IT BE KEPT TO BE RETURNED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04566 | 2023-06-22 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED, THAT PATIENT HAD DISCOMFORT AND BLEEDING WHEN USING THE INTERMITTENT CATHETER. AND ALSO BELIEVED, THAT THIS WAS DUE TO THE LACK OF LUBRICATION. THE USER HAS NOTICED, THAT THE LUBRICANT DOESN¿T SEEM TO BE AS EFFECTIVE IN THIS BATCH AS IT APPEARED TO BE OF A DIFFERENT CONSISTENCY. AND WASN¿T COATING AS WELL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04565 | 2023-06-22 | 2023-06-01 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT CUSTOMER STATED THAT PUREWICK URINE COLLECTION SYSTEM WAS NOT SUCTIONING AND ALSO PATIENT HAD A LOT OF URINARY TRACT INFECTION AND WAS NOW ANTIBIOTIC RESISTANT . REPRESENTATIVE ADVISED THE CUSTOMER TO TROUBLE SHOOT WITHOUT THE NEW KIT OR WITH A NEW ONE AND ADVISED ON THE 60-DAY REPLACEMENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. WARNINGS: ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: ¿ AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER ¿ HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. ¿ EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. ¿ PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
2020394-2023-00442 | 2023-06-22 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80142 | ATG80142 |  | IT WAS REPORTED THAT DURING DILATATION IN THE AREA OF THE BRACHIO-CEPHALIC VEIN LEADING INTO THE SUPERIOR VENA CAVA, THE INFLATION OF THE BALLOON ALLEGEDLY CAUSED THE BALLOON TO PARTIALLY SLIP OFF OF THE BALLOON CATHETER. IT WAS FURTHER REPORTED THAT THE BALLOON DID NOT DISLODGE, BUT ALLEGEDLY FAILED TO DEFLATE WHEN NEGATIVE PRESSURE WAS APPLIED VIA THE INFLATION DEVICE. REPORTEDLY, A GUIDEWIRE WAS GUIDED BEHIND AND ANOTHER BALLOON WAS INFLATED BEHIND THE PTA BALLOON IN ORDER TO DEFLATE IT VIA EXTERNAL PRESSURE. THE CATHETER INCLUDING ALL BALLOON MATERIAL WAS REMOVED SUCCESSFULLY. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BREAK AND DEFLATION ISSUE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. D4 (EXPIRY DATE: 12/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02557 | 2023-06-22 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER TRAY 13 F (15CM) (TL) | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5605150 |  | IT WAS REPORTED OUR NURSE SUBMITTED THE PSE AND SENT ME THIS PICTURE. IT LOOKS LIKE THE HOLDER WAS DEFECTIVE BECAUSE THE SUTURES ARE COMPLETELY INTACT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04574 | 2023-06-22 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE BURST PACKETS DID NOT HAVE ENOUGH SOLUTION TO BURST THE PACKET AND ACTIVATE THE HYDROPHILIC COATING TO CREATE LUBRICATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02553 | 2023-06-22 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (20G) (10CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120100T |  | IT WAS REPORTED BY THE CUSTOMER, "HAVING ISSUE WITH THE POWERGLIDE AT THE TIP NEAR THE PINK TO GREEN HUB HAS A HOLE. ONCE WE FLUSH IT, THE SITE APPEARS INFILTRATED BUT IT'S JUST BECAUSE THERE'S A HOLE IN THE CATHETER AND SO ANY MEDS SIT IN THE SUBQ TISSUE. A PATIENT WAS STARTED ON SALINE BOLUS AND NO HARM OCCURRED BUT WE HAD TO PULL THE LINE AND START ANOTHER." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02560 | 2023-06-22 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT,4FR POWERMIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P4154108D |  | IT WAS REPORTED BY THE CUSTOMER THAT THEY ARE UNABLE TO CLAMP ONE SIDE OF STATLOCK ONTO MIDLINE CATHETER, WAS ABLE TO SECURE THE CLAMP DOWN USING A STERISTRIP. STATED THIS HAS HAPPENED OCCASIONALLY IN PAST BUT WAS OBSERVED BY MYSELF ON A MIDLINE INSERTION RESPONSE RECEIVED FROM BD REP ON (B)(6) 2023. IT OCCURRED ONCE DURING A POWERMIDLINE INSERTION. THE NURSES SAID THEY HAD EXPERIENCED THIS OCCASIONALLY IN THE PAST BUT THAT THEY HAD NOT REPORTED IT PREVIOUSLY. A SPECIFIC NUMBER OF AFFECTED DEVICES WAS NOT PROVIDED BY THE COMPLAINANT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02581 | 2023-06-22 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE FULL KIT | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | S4153108 |  | IT WAS REPORTED BY THE CUSTOMER THAT FELT RESISTANCE WHEN PLACING MIDLINE AND WENT TO RETRACT WIRE AND NEEDLE SIMULTANEOUSLY. WHEN WIRE WAS RETRACTED, THEY NOTICED WIRE WAS UNRAVELED AT THE END. WIRE WAS RETRIEVED AND PATIENT HAD TO BE STUCK AGAIN. ADDITIONAL INFO RECEIVED ON 07-JUNE-23 PATIENT HAS A HISTORY OF ALCOHOL ABUSE AND DIABETES. NO URGENT/LIFE THREATENING SITUATION EVENT DISCOVERED WHEN WITHDRAWING THE GUIDEWIRE. THE WIRE DID NOT BREAK BUT DID UNRAVEL. NO PIECES WERE RETAINED IN THE PATIENT. A CHEST, CLAVICLE AND ARM XRAY WAS OBTAINED. PATIENT DID NO EXPERIENCE ANY SYMPTOMS, COMPLICATIONS OR NEGATIVE OUTCOMES. ADDITIONAL INFORMATION RECEIVED 6/8: WIRE UNRAVELED APPROXIMATELY 15-20CM FROM THE TIP. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02579 | 2023-06-22 | 2023-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV 4F DL MAX 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D |  | IT WAS REPORTED BY CUSTOMER THAT PICC WOULDN'T COME OUT OF PATIENT DURING PICC PROCEDURE. THE RN TOOK INTRODUCER OUT AND PULLED WHOLE THING AND THE WIRE WAS BROKEN AND OUT OF SKIN. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02580 | 2023-06-22 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1174108D |  | IT WAS REPORTED BY THE CUSTOMER THAT SECUREMENT CLAMP ON PICC STATLOCK WOULD NOT SECURELY CLOSE ON ONE SIDE. THERE WAS NO PATIENT IMPACT. SECURED CLOSED WITH DRESSING BUT CREATES FRUSTRATION WHEN NOT WORKING PROPERLY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02576 | 2023-06-22 | 2023-06-02 | I, n, j, u, r, y | Manufacturer report | BASIC KIT 20G X 10 CM WITH GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120107T |  | IT WAS REPORTED THAT POWERGLIDE (PG) CAME APART IN PATIENT'S ARM. PG LOCATED IN PATIENT'S ARM. PATIENT GOING TO SURGERY IN THE MORNING FOR REMOVAL. ADDITIONAL INFO RECEIVED 6/6/2023: IT WAS REPORTED BY THE CUSTOMER "PATIENT HAD TO GO TO SURGERY TO REMOVE CATHETER. ALL PIECES WERE REMOVED AND ACCOUNTED FOR." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED MIDLINE CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 20GA X 10CM POWERGLIDE PRO MIDLINE CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH CATHETER DAMAGE CAUSED BY CONTACT WITH A SHARP EDGED INSTRUMENT SUCH AS A SCALPEL. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A BREAK WAS OBSERVED JUST DISTAL OF THE MOLDED JOINT. MICROSCOPIC EXAMINATION OF THE CATHETER BREAK CONFIRMED IT WAS TYPICAL OF CONTACT WITH A SHARP BLADED INSTRUMENT, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: THE FRACTURE SURFACE WAS REFLECTIVE IN NATURE AND CONTAINED A STRIATED PATTERN. THIS CAN OCCUR DUE TO PATTERN TRANSFER OF THE SHARPENED EDGE TYPICALLY FOUND ON A SCALPEL-TYPE INSTRUMENT. SHARPLY FORMED FRACTURE EDGES. PLANAR SHAPE TO THE FRACTURE SURFACE. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02577 | 2023-06-22 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER "THE PATIENT COMES FOR THE 8TH TREATMENT WITH DOCETAXEL, WHEN THE PICCLINE IS FLUSHED WITH NACL, YOU SEE THAT LIQUID COMES FROM THE INSERTION SITE. THE PICCLINE IS THEN PULLED AND YOU SEE A HOLE 5-6CM FROM THE WINGS TOWARDS THE CATHETER TIP. " 6/8/2023: IT WAS REPORTED THERE WAS NO HARM TO THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02575 | 2023-06-22 | 2023-06-01 | I, n, j, u, r, y | Manufacturer report | FULL KIT 20G X 10 CM WITH GUARDIVA AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120108PT |  | IT WAS REPORTED THAT CATHETER EMBOLISM TO PATIENT'S LEFT UPPER ARM, 10 CM CATHETER CAME OUT 9 CM. ADDITIONAL INFO 6/8/2023: IT WAS REPORTED ¿THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. THE PATIENT EXPERIENCED INCREASED LEVEL OF PAIN WITH INSERTION BEYOND INITIAL NEEDLE STICK. CATHETER BREAKAGE WAS RETAINED IN PATIENT'S TISSUE. DOCTORS WERE MADE AWARE, CHEST AND ARM X-RAY WERE ORDERED. TOURNIQUET WAS APPLIED ABOVE THE SITE. PATIENT REMAINED FLAT. AN RRT NURSE WAS MADE AWARE WHO CONSULTED INTENSIVIST AND IR. THE EVENT CAUSED A DELAY IN MEDICATION. PATIENT NEEDED IV ACCESS FOR NAUSEA MEDICATION. SHE WAS A DIFFICULT STICK AND WE WERE CONSULTED TO PLACE A MIDLINE FOR THIS REASON. ADDITIONAL INFORMATION RECEIVED 6/14/2023: IT WAS REPORTED BY THE CUSTOMER ¿THE RETAINED CATHETER SEGMENT WAS NOT REMOVED FROM THE PATIENT AND REMAINS IN THE LEFT UPPER ARM WHERE IT HAD BEEN INSERTED. THERE WERE NOT NOTED ISSUES WITH THE CATHETER DURING INSERTION. CATHETER WAS ONLY IN THE PATIENT FOR SEVERAL SECONDS.¿ | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02570 | 2023-06-22 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 3F SL MBPLUS (3CG, BIOPATCH, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1173108D4 |  | IT WAS REPORTED BY THE CUSTOMER THAT 1 PPICC PROVENA T-LOCK ASSEMBLY W/STYLET SUCKING AIR AT T-PIECE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02571 | 2023-06-22 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER W/ SHERLOCK 3CG (TPS) STYLET (4F) (SL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED A PARTIAL RUPTURE OF THE CATHETER NEAR THE JUNCTION. SPILLAGE OF SUBSTANCE AT THE TIME OF INFUSION IS ALSO REPORTED. THE PATIENT REPORTS THAT HE DID NOT USE THE HIGH-PRESSURE MDC INFUSION CATHETER. ADDITIONAL INFORMATION PROVIDED 6/9 THE BROKEN PICC WAS USED IN THE PAST FOR THE ADMINISTRATION OF CHEMOTHERAPY. NO ADVERSE EVENTS WERE REPORTED FOR EITHER THE OPERATORS OR THE PATIENT. THE DEVICE WAS REMOVED NORMALLY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02568 | 2023-06-22 | 2023-06-07 | I, n, j, u, r, y | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (20G) (8CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120080T |  | IT WAS REPORTED THAT THE GUIDE WIRE MAY HAVE SHEARED OFF. ADDITIONAL INFORMATION RECEIVED 6/13/2023: CUSTOMER REPORTED THAT THE GUIDE WIRE SHEARED OFF DURING INSERTION AND PATIENT REQUIRED A 90 MINUTE SITE TO REMOVE. CLINICIAN THINKS THIS IS IMPART DUE TO THE DIFFICULTY GETTING OFF THE CATHETER. CUSTOMER STATES THAT IF THIS HAS OCCURRED IN THE VEIN - THE PATIENT COULD HAVE POTENTIALLY VERY SERIOUS EVENT. ADDITIONAL INFORMATION GATHERED FROM MD (END-USER) TODAY: ATTEMPTED TO CANNULATE THE BRACHIAL VEIN, THE WIRE WAS DEPLOYED OUTSIDE THE VEIN, JUST BESIDE THE BRACHIAL VEIN, INTO THE INTERSTITIAL TISSUE. THEY NOTICE OVER THE PAST 6 MONTHS, IT IS DIFFICULT TO DEPLOY THE GUIDEWIRE, EVEN OUT OF THE PACKAGE WHEN TESTING IT PRIOR TO INSERTION, IT FEELS LIKE A ¿JOLT OR ABRUPT DEPLOYMENT.¿ THE OTHER MD WAS THE ONE WHO INSERTED THE POWERGLIDEPRO, SHE SAID IT DIDN'T FEEL RIGHT SO SHE PULLED IT OUT AFTER THE WIRE WAS DEPLOYED. WHEN SHE PULLED IT OUT, THE GUIDEWIRE REMAIN IN THE ARM, PATIENT WENT TO THE OR TO HAVE IT REMOVED. SHE STATED, IT FEELS LIKE THEY REALLY HAVE TO ¿REEF¿ ON THE DEVICE TO HAVE IT DEPLOY AND WHEN IT ACTUALLY DOES, IT GOES VERY FAST, PERHAPS CAUSING THE NEEDLE TO POP OUT OF THE VEIN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02569 | 2023-06-22 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, PICC FRACTURE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02572 | 2023-06-22 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER 4F RADIOLOGY BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 3194335 |  | IT WAS REPORTED BY THE CUSTOMER THAT BLOOD WAS FLOWING OUT OF PICC LINE HUB DURING ROUTINE WEEKLY CAP AND DRESSING CHANGE. PICC IS A VALVED LINE WITH NO EXTERNAL CLAMP. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02567 | 2023-06-22 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT,4FR POWERMIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P4154108D |  | CALLER STATES HER DAD WAS IN THE HOSPITAL LAST WEEK. THE HOSPITALIST ORDERED A MIDLINE TO BE PLACED FOR AT HOME ANTIBIOTICS. THE MIDLINE WAS PLACED ON FRIDAY (B)(6) 2023 AND THEY WERE DISCHARGE HOME. CALLER STATES IT WAS LEAKING AFTER IT WAS PLACED. CALLER STATES SHE CHANGED THE DRESSING ON MONDAY BECAUSE IT WAS SATURATED. CALLER STATES ON TUESDAY MORNING IT FLUSHED WITHOUT DIFFICULTY AND SHE GAVE HIM THE ANTIBIOTIC. WHEN SHE FLUSHED IT TUESDAY EVENING, THE SALINE FLUSH AND BLOOD ALL LEAKED OUT INTO THE DRESSING AND DOWN HIS ARM. CALLER STATES SHE TOOK HIM TO THE ER TO HAVE IT LOOKED AT. PER CALLER, THEY DID AN X-RAY AND THEY SAID THE MIDLINE WAS STILL IN PLACE. THEY THEN DID AN ULTRASOUND AND THE TECHNICIAN SAID THERE WAS A BLOOD CLOT, HOWEVER THE RADIOLOGIST SAID THERE WAS NO BLOOD CLOT. CALLER STATES THEY WERE USING HEPARIN LOCKS IN-BETWEEN EACH ANTIBIOTIC INFUSION. WHEN THE NURSES FLUSHED IT, THE FLUID WOULD LEAK BACK AT AS WELL. THE NURSES CHANGED THE DRESSING AND SENT HIM HOME. THE NURSES IN THE ER SAID THEY WERE UNABLE TO REMOVE MIDLINES WITHOUT AUTHORIZATION. HER DAD IS COMPLAINING OF PAIN IN THE ARM WITH THE MIDLINE. CALLER STATES SHE CONTACTED THE HOSPITALIST THAT ORDERED THE MIDLINE TO DISCUSS WITH HIM AND HE STATED HE WOULD FOLLOW UP WITH THEM TODAY TO FIGURE OUT THE NEXT COURSE OF ACTION. CALLER IS NOT SURE WHAT TO DO OR WHERE TO GO TO HAVE THE MIDLINE REMOVED. CALLER STATES SHE PLACED A PERIPHERAL I IN HIS AC TO GIVE THE ANTIBIOTIC IN THE MEANTIME. DISCUSSED WITH CALLER IF THE MIDLINE IS LEAKING, WE DO NOT RECOMMEND CONTINUING TO USE IT. ALTHOUGH THE ER CONFIRMED PLACEMENT WITH AN X-RAY, IT MAY BE LEAKING DUE TO DAMAGE TO THE CATHETER. EXPLAINED A PROVIDER MAY CONSIDER ORDERING A DYE STUDY TO CONFIRM THE CATHETER IS INTACT AND NOT LEAKING. INFORMED CALLER WE WOULD RECOMMEND GOING TO THE PROVIDER OR HOSPITAL WHERE THE MIDLINE IS PLACED. SUGGESTED CALLER CONTINUE TO TRY TO GET IN TOUCH WITH THE HOSPITALIST THAT ORDERED THE MIDLINE TO DISCUSS DYE STUDY OR REMOVAL. SUGGESTED TO CALLER THEY MAY CONSIDER GOING BACK TO THE ER FOR ANOTHER OPINION FROM THE DAYSHIFT PROVIDERS. CALLER STATES SHE DOES NOT WANT TO DO THAT AT THIS TIME. INFORMED CALLER IF HER DAD DOES START EXPERIENCING MORE SEVERE SYMPTOMS SUCH AS CHEST PAIN, HEART PALPITATIONS, SHORTNESS OF BREATH ETC. TO GO TO THE ER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02565 | 2023-06-22 | 2023-06-07 | I, n, j, u, r, y | Manufacturer report | DOT PG PRO 22G 8CM RT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F322088PTX |  | IT WAS REPORTED BY THE CUSTOMER THAT MIDLINE GOT STUCK IN THE PATIENT. THE SURGEON HAD TO COME AND REMOVE IT. ADDITIONAL INFO RECEIVED ON 15-JUNE-2023. SURGERY HAD TO BE CALLED TO REMOVE CATHETER. THE EVENT HAPPENED AT TIME OF PLACEMENT. WHEN INSERTING THE MIDLINE, THE GUIDEWIRE AND CATHETER ADVANCED WITHOUT ANY RESISTANCE BUT UPON TRYING TO RETRACT THE GUIDEWIRE IT WOULD NOT RETRACT. HEAT PACKS WERE APPLIED AND MINIMAL FORCE TO TRY TO PULL ENTIRE MIDLINE OUT WITH NO SUCCESS. IN FEAR OF A PORTION OF THE MIDLINE BREAKING OFF INSIDE THE PATIENT SURGERY WAS CALLED FOR SAFE REMOVAL. COMPLAINANT WAS UNSURE IF THE EVENT INTERRUPT THE ADMINISTRATION OF ANY MEDICATION, OR CAUSE A CLINICALLY SIGNIFICANT DELAY IN MEDICATION. ULTRASOUND WAS USED DURING THE EVENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04602 | 2023-06-22 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRE-TEST WATER WITHDRAWAL, THE STERILE WATER COULD NOT BE DRAINED. IT WAS DRAINED AFTER THE SYRINGE WAS ATTACHED FOR A WHILE, BUT THE ASYMMETRY OF THE BALLOON WAS ALSO NOTICEABLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02561 | 2023-06-22 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F TRAY W/ MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT "ONCE THE PICC WAS INSERTED, THE INSERTER PULLED BACK ON THE SYRINGE TO CHECK FOR BLOOD RETURN. WHEN SHE PULLED BACK TO ATTAIN BLOOD RETURN FROM THE LUMEN WITH THE YELLOW RUBBER STOPPED, AIR BUBBLES COULD BE SEEN ENTERING THE SYRINGE. THIS ONLY HAPPENED THROUGH THE LUMEN WITH THE RUBBER STOPPER THAT WAS HOUSED BY THE STYLET. SHE REMOVED THE SYRINGE AS TO NOT FLUSH AIR BACK INTO THE PICC AND USED A PRE-PRIMED STERILE SALINE SYRINGE TO FLUSH THE LUMEN WITH THE STYLET GOING THROUGH THE YELLOW RUBBER STOPPER. WHEN SHE FLUSHED THE PICC, SMALL DROPLETS OF SALINE APPEARED TO LEAK OUT AROUND THE TOP OF THE RUBBER STOPPER. TO NOTE, THERE WAS NO LEAKING OF FLUID AROUND THE YELLOW RUBBER STOPPING UPON FIRST OPENING THE PICC KIT AND PRIMING THE PICC PRIOR TO INSERTION. SHE SAID THIS HAS HAPPENED BEFORE BUT NOT TODAY. IT IS WITH THE PRESSURE OF FLUSHING THE PICC THAT LEAKING OCCURS AT THE YELLOW STOPPER AND AIR BUBBLES APPEAR IN THE SYRINGE. UPON CLOSE INSPECTION, IT APPEARS THE YELLOW RUBBER STOPPER HAS A SMALL, CLEAR PLASTIC CASING AROUND THE SIDES, RATHER THAN JUST PLAN RUBBER LIKE THE SINGLE LUMEN STOPPERS APPEAR LIKE AND ALL APPEARED LIKE PRIOR TO THIS CHANGE. NO PATIENT HARM. NURSE WORRIES THAT AIR COULD POSSIBLE ENTER THE BLOODSTREAM. ALSO CONCERNED ABOUT THE LEAKING ASPECT TO MAINTAIN ASEPSIS DURING THE PROCEDURE. THIS SAMPLE WAS ACTUALLY IN THE PATIENT AND OBSERVED TO HAVE HAPPENED, BY MYSELF THE REPORTER." ADDITIONAL INFORMATION RECEIVED ON 09-JUN-2023. THERE WAS NO OBSERVED LIFE THREATENING MEDICAL SITUATION RELATED TO THE PICC INSERTION PROCEDURE, AND NO HOLES OR DAMAGE OBSERVED BUT WHEN WE PERFORM A FLUSH OF THE PICC WITH SALINE PRIOR TO INSERTING THE PICC INTO THE PATIENT, WE OBSERVE LEAKING OF SALINE FROM THE RUBBER STOPPER WHERE THE STYLET PASSES THROUGH. THERE HAVE BEEN VARYING DEGREES OF LEAKAGE NOTED DURING THIS SALINE FLUSH. SOME HAVE ONLY 1 DROP OF SALINE EXPRESSED, OTHER TIMES SALINE "SHOOTS" OUT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04599 | 2023-06-22 | 2023-05-20 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT NEEDED TO MIX THEM OWN LUBRICANT ON PLACE IT ON THE MAGIC 3 GO MALE COUDE INTERMITTENT CATHETER. THE REPRESENTATIVE EXPLAINED TO THE CUSTOMER HOW TO USE THE PRODUCT AND HAVE THE LUBRICANT EXTRACTED INSIDE OF THE CATHETER PACKAGING WITHOUT HAVING TO TOUCH THE LUBRICANT. PER FOLLOW UP VIA PHONE (B)(6) 2023, CUSTOMER EXPERIENCED SEPSIS AND HAD TO BE ADMITTED TO THE HOSPITAL FOR 3 DAYS AFTER USING THE CATHETER. CUSTOMER WAS NO LONGER USING THE CATHETERS. | THE REPORTED EVENT WAS CONFIRMED USE RELATED AS THE USER USES THE CATHETERS INCORRECTLY. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "LACK OF TRAINING ON APPROPRIATE TECHNIQUE, USER DOES NOT OPEN PACKAGE AS INDICATED". DHR REVIEW WAS NOT REQUIRED BECAUSE THE INVESTIGATION WAS CONFIRMED - USE RELATED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. INSTRUCTION FOR USE: 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, USE THE ADHESIVE LABEL(S) TO HANG THE PACKAGE ON A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. CLEAN THE OPENING TO THE URETHRA - AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6" FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04600 | 2023-06-22 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE CATHETER FELL OFF SPONTANEOUSLY IN THE WARD AFTER IT WAS PLACED IN THE OPERATING ROOM. THE CATHETER WAS FILLED WITH STERILE WATER AGAIN AND OBSERVATION SHOWED THAT THE BALLOON DEFLATED OVER 7~8 HOURS. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND SOLUTION IMMEDIATELY LEAKED OUT OF A 0.011" PINHOLE FOUND AT THE TRIFURCATION. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02559 | 2023-06-22 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC SOLO 5FR D/L 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 22295118 |  | IT WAS REPORTED THAT AS PER CUSTOMER "I WAS PREPARING TO USE THIS LINE YESTERDAY. I RETRACTED THE GUIDEWIRE WELL PAST THE TRIM POINT. CUT TO LENGTH AND THEN ONCE THE GUIDEWIRE WAS ADVANCED THE KINK IN THE GUIDEWIRE AND ASSOCIATED REGION OF TRIMMED LINE WAS EVIDENT. THIS WAS FROM AN INTACT PACKAGE AND WE HAD NOTED WE NEEDED TO USE THE LINE PRIOR TO SHELF EXPIRY. NO OBVIOUS DAMAGE TO THE PACKAGE OUTER OR INNER, PICC WAS WITH-IN THE CHANNEL. PICC GUIDEWIRE WAS DEFINITELY WITHDRAWN PAST THE TRIM POINT. AS GUIDEWIRE WAS PUSHED BACK TO ENGAGE WITH THE B-H CONNECTOR THE BEND IN THE GUIDEWIRE WAS EVIDENT. MENTALLY RECHECKED THAT IT WAS NOT OPERATOR ERROR, THE WIRE MUST HAVE BEEN 6-8CM BELOW MARKING USED FOR TRIMMING THE LINE. DOCTOR DID NOT PROCEED WITH THIS PICC AND USED ANOTHER PICC KIT AND HAD NO ISSUES WITH THE GUIDEWIRE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02558 | 2023-06-22 | 2023-01-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 3.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668935 |  | CUSTOMER REPORTED THEY OPENED A 3.5 FR MICROINTRODUCER FOR USE BEFORE PLACING THE PATIENT AND FOUND THAT THE GRAY PORTION OF THE SHEATH WAS VERY SOFT AND NOT STIFF ENOUGH TO SUPPORT DELIVERY TO THE PUNCTURE SITE. 5/30/2023 - THE RETURNED DEVICE EXHIBITED MULTIPLE POINTS OF INWARD BUNCHING, FLARING AND DEFORMATION. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED GREY SHEATH IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE 3.5 FR X 5 CM MICROEZ MICROINTRODUCER WAS RETURNED FOR EVALUATION. THE GREY SHEATH HAD NOT BEEN PEELED. DEFORMATION AT THE DISTAL END OF THE SHEATH WAS OBSERVED. BLOOD RESIDUE AND EVIDENCE OF USE WAS NOTED ON THE DEVICE. MICROSCOPIC INSPECTION OF THE SHEATH TIP REVEALED MULTIPLE POINTS OF INWARD BUNCHING. ADDITIONAL DEFORMATION AND OUTWARD FLARING OF THE SHEATH TIP WAS ALSO VISIBLE. A REVIEW OF THE MANUFACTURING RECORDS DID NOT REVEAL ANY EVIDENCE TO INDICATE A MANUFACTURING RELATED ROOT CAUSE. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE ADVANCEMENT AGAINST RESISTANCE, INADEQUATE SKIN INSERTION SITE INCISION, AND STEEP INSERTION ANGLE. SINCE DEFORMATION TO THE SHEATH TIP AND EVIDENCE OF A DIFFICULT INSERTION WAS NOTED ON THE DEVICE, THE COMPLAINT IS CONFIRMED; HOWEVER, THE EXACT FACTORS CONTRIBUTING TO THE REPORTED EVENT REMAIN UNKNOWN. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02542 | 2023-06-22 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FLEXXICON DIALYSIS CATHETER KIT 11F X 20CM (SHORT-TERM) (DUAL-LUMEN) | CATHETER, HEMODIALYSIS, DUAL LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5678200 |  | IT WAS REPORTED THAT THE WIRE WAS EASILY KINKING DURING INSERTION, SO THE PLACER HAD TO UNPACK AN NEW WIRE OF THE OTHER PROVIDER. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GUIDEWIRE WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE PHOTOGRAPH WHICH DEPICTED A J-TIP GUIDEWIRE. USAGE RESIDUES APPEARED TO BE VISIBLE ON THE WIRE. THE GUIDEWIRE EXTENDED FROM ITS PLASTIC HOOP. A BREAK WAS EVIDENT IN THE INNER CORE WIRE AND THE COIL WIRE WAS STRETCHED NEAR THE J-TIP. THE J-TIP APPEARED TO BE INTACT. GUIDEWIRE DAMAGE WAS EVIDENT IN THE SUBMITTED PHOTOGRAPH; HOWEVER, INSPECTION OF THE PHOTOGRAPH WAS INSUFFICIENT TO IDENTIFY THE CAUSE OF THE DAMAGE. CONSEQUENTLY, THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CAUSES INCLUDE WIRE RETRACTION AGAINST THE INTRODUCER NEEDLE BEVEL AND RETRACTION AGAINST RESISTANCE. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04592 | 2023-06-22 | 2023-06-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THESE ARE DEFECTIVE, AND PATIENT HAD A LOT OF URINARY TRACT INFECTIONS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04598 | 2023-06-22 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TIEMANN MODEL UROLOGICAL CATHETER COUDE TIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 010116 | 010116 |  | IT WAS REPORTED THAT CUSTOMER GOT 3 BOXES OF 12 INTERMITTENT CATHETERS THAT WERE DEFECTIVE ALL THE SAME LOT. CUSTOMER STATED THAT THE END OF THEM WAS CURL OVER AND WERE VERY HARD. CUSTOMER ALSO MENTIONED THAT UP AT THE TOP THEY WERE SHORTER IN THE PACKAGING, 2 INCH SHORTER THAN USUAL. UP AT THE TOP OF THE CATHETER, THE OUTLET PART THERE WERE A LARGE BUMP ON IT. SO, INSERTION WOULD BE DIFFICULT BECAUSE OF THE HARD CURL END. CUSTOMER REQUEST REPLACEMENT FOR THE DEFECTIVE PRODUCT. PER FOLLOW UP VIA PHONE (B)(6) 2023, CUSTOMER STATED THAT THEY DID NOT USE THE CATHETERS, THEY SAW THE DEFECT BEFORE USE AND WANTED TO INFORM THE COMPANY OF THE AFFECTED LOTS. PER NOTIFICATION FROM INVESTIGATOR ON 31JUL2023, DURING SAMPLE EVALUATION IT WAS FOUND THAT THE SCRATCH FOUND ON THE SHAFT CLOSE TO THE FUNNEL OF CATHETER. PER ADDITIONAL INFORMATION RECEIVED VIA SAMPLE FORM ON 20JUN2023, IT WAS REPORTED THAT CUSTOMER OPENED ONE BOX AND SAW A DEFECTIVE CATHETER, CHECKED THE OTHERS IN THE BOX, THEY WERE ALL THE SAME, SO HE THREW THAT BOX AWAY. STATED THAT THE NEXT BOX WAS FINE AND THEN THEY GOT ANOTHER BOX OF DEFECTIVE CATHETERS AND STATED THAT THESE WERE SHORTER THAN THEY SHOULD BE. THE TIPS WERE CURLED AND THE HOLE IN THE TIP WAS MUCH FURTHER UP THE CATHETER THAN IT SHOULD BE AND THEY WARE HARDER AND LESS PLIABLE THAN THEY SHOULD BE. THE OUTLET END HAD A BUMP ON THE OUTSIDE. STATED THAT THEY ALSO NOTICED THAT THE PRINT ON THE PACKAGING WAS AN OFF COLOR FROM THE USUAL PACKAGING AND THEY WENT THROUGH ALL THEIR REMAINING BOXES, AND THESE WERE THE ONLY ONES, AND THEY WERE ALL THE SAME LOT NUMBER. STATED THAT THEY WOULD BE TERRIBLY UNCOMFORTABLE TO USE AND SEEM LIKE THEY COULD HURT SOMEONE IF THEY TRIED TO USE IT, AND THIS WAS THE REMAINDER OF TWO BOXES AND COULD SEE THE FLAW THROUGH THE PACKAGING. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE DEVICE MEETS SPECIFICATIONS. GROSS VISUAL EVALUATION: NO NON-CONFORMITIES WERE NOTED WITH THE HARDNESS AT THE TIP OF THE CATHETER. COUDE CURVE ALIGNED WITH THE COUDE INDICATOR NOTCH, SOME DEGRADATION DUE TO OXIDATION WAS NOTED WITHIN THE EYELETS OF CATHETERS 1, 2, 3, 4, 5 (VERY SLIGHT SIGNS OF OXIDATION WITHIN THE EYELETS), 6, 8, 9, AS THIS WAS NOT REPORTED BY THE CUSTOMER, AND IT IS UNKNOWN HOW THE CATHETER WAS BEING STORED OR WHEN THE OXIDATION OCCURRED A NEW COMPLAINT WILL NOT BE CREATED. SOME SCRATCH MARKS WERE NOTED ON THE FUNNEL OF CATHETERS 4, 6 & 9, AND A SMALL SCRATCH WAS NOTED ON THE SHAFT CLOSE TO THE FUNNEL OF CATHETER 10 ; NO SCRATCHES OR ABNORMALITIES WERE NOTED ON THE PACKAGING. DIMENSIONAL EVALUATION: USING A FRENCH GAUGE, ALL CATHETERS WERE CONFIRMED TO BE 16FR. THE OVERALL LENGTH OF THE CATHETERS WAS MEASURED AND FOUND TO BE WITHIN SPECIFICATION: CATHETER 1- 15.75", CATHETER 2 - 15.75", CATHETER 3 - 15.72", CATHETER 4 - 15.75", CATHETER 5 - 15.81", CATHETER 6 - 15.87", CATHETER 7 - 15.75", CATHETER 8 - 16.00", CATHETER 9 - 16.00", CATHETER 10 - 15.87". A DHR IS NOT REQUIRED SINCE THE REPORTED FAILURE WAS UNCONFIRMED. THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. CORRECTIONS: H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02551 | 2023-06-22 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT TWO MONTHS AND FOUR DAYS POST PORT PLACEMENT, THE PATIENT COMPLAINED OF PAIN DURING PRIMING. IT WAS FURTHER REPORTED THAT THE UPON X-RAY EXAMINATION THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS ALSO FURTHER REPORTED THAT THE DISTAL CATHETER SEGMENT WAS REMOVED. REPORTEDLY, THE PORT BODY WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02552 | 2023-06-22 | 2023-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT DUO M.R.I. IMPLANTABLE PORT, CHRONOFLEX DUAL-LUMEN, 9.5F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1829500 | 1829500 |  | IT WAS REPORTED TWO YEARS FOUR MONTHS FOURTEEN DAYS POST PORT PLACEMENT PROCEDURE, THE SEPTUM OF THE PORT ALLEGEDLY DISLODGED OFF THE RIM OF THE PORT. REPORTEDLY. THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2022). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02548 | 2023-06-22 | 2022-12-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT MRI ISP, 8 FR GROSHONG, INT WO SP, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808561 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA RIGHT INTERNAL JUGULAR VEIN, AFTER THE PORT WAS PLACED, WHEN DRAWING BLOOD BEFORE POCKET SUTURE, AIR WAS ALLEGEDLY DRAWN OUT INSTEAD OF BLOOD. IT WAS FURTHER REPORTED THAT THE PORT HUB WAS THEN REMOVED AND FOUND TO BE LEAKING AIR. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTO AND VIDEO WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 09/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02547 | 2023-06-22 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT THREE YEARS FIVE MONTHS AND TWELVE DAYS POST PORT PLACEMENT PROCEDURE VIA THE RIGHT INTERNAL JUGULAR VEIN, THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE PORT BODY AND THE DISTAL CATHETER SEGMENT WERE REMOVED PERCUTANEOUSLY. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, AN IMAGE WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 02/2022). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02545 | 2023-06-22 | 2021-12-04 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1608062 |  | IT WAS REPORTED THAT ONE YEAR, EIGHT MONTHS AND FIVE DAYS POST A PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, THE PATIENT WAS DEVELOPED WITH A FUNGEMIA INFECTION. IT WAS FURTHER REPORTED THAT THE PORT WAS REMOVED AND THE CULTURES FROM THE CATHETER TIP SHOWED THE GROWTH OF YEAST. REPORTEDLY, THE PATIENT WAS PROVIDED WITH ANTIFUNGAL MEDICATION AND THE PORT WAS REPLACED. THERE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. MEDICAL RECORDS WERE PROVIDED AND REVIEWED. APPROXIMATELY ONE YEAR AND EIGHT MONTHS POST PORT PLACEMENT, PATIENT PRESENTED FOR BOWEL OBSTRUCTION FOR WHICH REQUIRED SURGERY. PATIENT STAY WAS COMPLICATED BY DEVELOPMENT OF FUNGEMIA. THE LIKELY SOURCE OF THIS WAS HIS PORT AND PATIENT UNDERWENT REMOVAL OF INFECTED PORT-A-CATH. THE PORT CATHETER TIP SENT FOR CULTURES WHICH SHOWED GROWING YEAST. AROUND THREE DAYS LATER, PATIENT WAS TREATED WITH MICAFUNGIN AND DISCHARGED TO HOME. AROUND ONE DAY LATER, PATIENT UNDERWENT BARD POWERPORT CLEARVUE PLACEMENT PROCEDURE FOR CHEMOTHERAPY TREATMENT FOR BLADDER CANCER, SECONDARY MALIGNANT NEOPLASM OF BONE. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE AS NO OBJECTIVE EVIDENCE HAS BEEN PROVIDED TO CONFIRM ANY ALLEGED DEFICIENCY WITH THE PORT. FURTHERMORE, CLINICAL CONDITION ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 01/2021). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02586 | 2023-06-23 | 2015-12-09 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT FOUR WEEKS POST A PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED CHEST PAIN. IT WAS FURTHER REPORTED THAT THE PATIENT WAS ALLEGEDLY DIAGNOSED WITH STAPHYLOCOCCUS AUREUS BACTEREMIA. REPORTEDLY, THE DEVICE WAS REMOVED. THE PATIENT'S CURRENT STATUS IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. MEDICAL RECORDS WERE PROVIDED AND REVIEWED. APPROXIMATELY TWENTY-EIGHT DAYS POST PORT PLACEMENT, VENOUS ACCESS PORT REMOVAL WAS PERFORMED. PURULENT COLLECTION OF FLUID AROUND THE HUB OF THE VENOUS ACCESS PORT WAS NOTED. ON THE SECOND ROUND OF CHEMOTHERAPY, THE PATIENT DEVELOPED FEVERS, PRONOUNCED NEUTROPENIA, AND STAPH AUREUS BACTEREMIA, SUSPECTED SEEDING FROM ACCESS FROM THE PORT SITE. TENDERNESS OVER THE PORT SITE WAS ALSO NOTED. THE PATIENT WAS TREATED FOR A RIGHT UPPER LOBE PNEUMONIA AND BACTEREMIA. CONCERNING WAS FOR INFECTION OF INDWELLING VENOUS ACCESS PORT. VENOUS ACCESS PORT WAS ENCOUNTERED AND THE FLUID COLLECTION THAT HAS ACCUMULATED AROUND IT. GROSS PURULENT MATERIAL WAS NOTED. CULTURE OF RIGHT CHEST WALL VENOUS ACCESS PORT SITE WAS GIVEN, AND THE VENOUS ACCESS PORT WAS SENT FOR GROSS EXAMINATION. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED BACTEREMIA AND CHEST PAIN AS NO OBJECTIVE EVIDENCE HAS BEEN PROVIDED TO CONFIRM ANY ALLEGED DEFICIENCY WITH THE PORT. FURTHERMORE, CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: B5, D4 (EXPIRATION DATE: 05/2020),. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02609 | 2023-06-23 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 2194108 |  | CUSTOMER REPORTED ¿PATIENT HAD A NEUTRAL SACT DRUG (RITUXIMAB) INFUSING THROUGH HER PICC LINE. WHEN THE RATE OF RITUXIMAB WAS INCREASED AS PER PROTOCOL, PATIENT ALERTED STAFF SHE FELT HER PICC LINE WAS LEAKING. INFUSION PAUSED AND SITE ASSESSED - RITUXIMAB HAD LEAKED OUT OF PICC LINE ONTO PATIENT¿S SKIN (EXTERNAL). THERE WAS A HOLE IN THE PICC LINE (EXTERNALLY). TREATED AS AN EXTRAVASATION AND THE EXTRAVASATION PROTOCOL WAS FOLLOWED - HEAT PACK TO BE APPLIED 20 MINUTES EVERY 4 HOURS FOR 24 HOURS (WAKING HOURS) - PATIENT AWARE TO FOLLOW SAME AT HOME. NO REDNESS/SWELLING/PAIN EVIDENT. PICC LINE WAS REMOVED" NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04630 | 2023-06-23 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT ABOUT 2 HOURS AFTER FOLEY CATHETER PLACEMENT IN THE OPERATING ROOM, THE PATIENT BECAME AWARE OF THE NEED TO URINATE BUT COULD NOT CONFIRM URINE IN THE CATHETER. AFTER CONFIRMING THAT THE PATIENT'S GENERAL CONDITION WAS GOOD, THE USER ATTEMPTED TO REMOVE THE CATHETER. WHEN WATER WAS ADDED TO THE REMOVED CATHETER, THE BALLOON WAS INFLATED, AND NO WATER LEAKAGE WAS CONFIRMED AFTER THAT. IT WAS UNCLEAR IF THERE WERE ANY MICRO RUPTURES THAT COULD NOT BE SEEN. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6)2023, STATED THAT THE EVENT WAS CATHETER FELL OUT AND BOTH INCIDENTS HAPPENED ON SAME CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02608 | 2023-06-23 | 2023-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG MID LINE MINI TRAY 4F | CATHETER, PERCUTANEOUS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7715435CE |  | CUSTOMER REPORTED THAT SPILLAGE OF THERAPY IN DEVICE THAT ALREADY IN THE MORNING DOES NOT WORK ON THE WAY OUT. AT THE TIME OF DRESSING TO INSPECT THE DEVICE HOLE IN CATHETER EXITS IS DETECTED. DEVICE WAS USED, FIRST USE, NO CONSEQUENCES. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02604 | 2023-06-23 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV 4F DL MAX TLS | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1294108D |  | CUSTOMER REPORTED THAT THIS MICROINTRODUCER DID NOT PEEL AWAY PROPERLY AND THE SHEATH APPEARED TO BE NOT COMPLETELY CONNECTED TO THE HANDLE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DIFFICULT PEEL IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. TWO 4.5 FR MICROEZ MICROINTRODUCERS WERE RETURNED FOR EVALUATION. THE FIRST SAMPLE CONSISTED OF THE PEELED TABS WITHOUT THE DILATOR. THE SECOND SAMPLE CONSISTED OF BOTH THE DILATOR AND SHEATH. THE FIRST SAMPLE ONLY HAD ONE PEEL TAB RETURNED. THE SHEATH MATERIAL WAS FOUND TO BE PARTIALLY TORN AT THE ATTACHMENT LOCATION TO THE TAB. THE SECOND TAB WAS NOT RETURNED; HOWEVER, THE GREY SHEATH WAS PRESENT AND APPEARED TO HAVE DETACHED FROM THE TAB. NO APPARENT DAMAGE WAS NOTED ON THE SECOND MICROINTRODUCER. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04632 | 2023-06-23 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT SPONTANEOUS FOLEY CATHETER REMOVAL WAS FOUND AFTER PLACEMENT ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02611 | 2023-06-23 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 18GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1182250 |  | IT WAS REPORTED BY CUSTOMER "THE NEEDLE IS NOT FULLY RETRACTED. THE RN ALSO PROPOSED THE QUESTION IF THE PIGTAIL PORTION MISSING COULD BE IN THE NEEDLE?" | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF NEEDLE FAILURE RETRACTION AND GUIDEWIRE DAMAGE WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS TWO PHOTOGRAPHS WHICH DEPICTED A 20GA X 2.25¿ ACCUCATH ACE PERIPHERAL IV CATHETER ASSEMBLY. THE FIRST PHOTOGRAPH DEPICTED THE PROXIMAL END OF THE HOUSING. THE SAFETY BUTTON WAS DEPRESSED AND THE NEEDLE WAS PARTIALLY WITHDRAWN INTO THE HOUSING. THE TIP OF THE NEEDLE REMAINED EXPOSED, PROTRUDING FROM THE HOUSING. THE GUIDEWIRE WAS ADVANCED AND THE DISTAL END WAS NOT VISIBLE. THE SECOND PHOTOGRAPH DEPICTED THE DISTAL END OF THE GUIDEWIRE. THE TAPERED REGION OF THE WIRE WAS EVIDENT; HOWEVER, THE COIL TIP APPEARED TO BE MISSING/FRACTURED. THE GUIDEWIRE WAS OBSERVED TO DAMAGED AND THE NEEDLE WAS OBSERVED TO HAVE INCOMPLETELY RETRACTED; HOWEVER, INSPECTION OF THE PROVIDED PHOTOGRAPHS WAS INSUFFICIENT TO IDENTIFY THE CAUSE. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. ATTEMPTED INSERTION AGAINST RESISTANCE AND MANIPULATION OF THE GUIDEWIRE MAY RESULT IN BOTH WIRE DAMAGE AND DIFFICULT NEEDLE RETRACTION. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02585 | 2023-06-23 | 2023-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808060 | 1808060 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT WAS ALLEGEDLY UNABLE TO GET BLOOD RETURN AFTER FEW TREATMENTS AND A CT SCAN WAS PERFORMED SHOWING NARROWING OF HER VENA CAVA. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 06/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04636 | 2023-06-23 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE RIGHT SINGLE J CATHETER FELL OUT AND A NEW CATHETER WAS REPLACED. IT WAS FOUND THAT THE BALLOON WAS BURST. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. THEREFORE, NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02614 | 2023-06-23 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194118 |  | IT WAS REPORTED SMALL HOLE/LEAK APPROXIMATELY 4CM FROM TOP EXTERNAL WING. NOTICED THE LINE BULGING WHEN PICC WAS FLUSHED. PATIENT RECEIVED CHEMO ON MAY 2-3, HYDRATION ON MAY 4TH AND RETURNED ON MAY 11TH FOR PICC CARE. THAT IS WHEN THE LEAK WAS NOTED. ADDITIONAL INFORMATION RECEIVED 6/9/2023: IT WAS REPORTED BY THE CUSTOMER ¿EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. PATIENT REQUIRED A REPLACEMENT PICC.¿ | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02615 | 2023-06-23 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED VIA MEDWATCH "BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET LOT # REHN2437 RUBBER-LIKE STOPPER LEAKING WITH FLUSHING.". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04643 | 2023-06-23 | 2023-06-07 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTIONS STARTING IN 2022. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04644 | 2023-06-23 | 2023-06-07 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT DID EXPERIENCE A URINARY TRACT INFECTION IN PREVIOUS MONTH. THE REPRESENTATIVE WAS ABLE TO CONFIRMED THE PATIENT NOT USING THE AFFECTED CATHETER AND EXPLAINED THAT WAS OK TO CONTINUE USING THE MAGIC 3 GO MALE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04645 | 2023-06-23 | 2023-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI24 | 0165SI24 |  | IT WAS REPORTED THAT FOLEY CATHETER STOPPED WORKING AFTER 7 DAYS TUBE CANNOT FLUSH. PER FOLLOW UP VIA EMAIL ON 12JUN2023, THE CUSTOMER STATED THAT THEY WERE STABLE FOR NOW, BUT VERY CONCERNED ABOUT SEVERAL CATHETERS NOT WORKING PROPERLY. THIS WAS NOT THE ONLY ONE TO FAIL. SEVERAL FAILED THIS YEAR, AND SINCE THEIR INSURANCE WILL ONLY PAY FOR ONE A MONTH, SEVERAL WERE PAID OUT OF POCKET. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO BROKEN SPIGOT OR UNUSED SPIGOT WERE NOT PLUGGED CAUSING INSUFFICIENT AIR BLOW THROUGH CATHETER BLOCKAGES. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSERT FOLEY CATHETERS ONLY FOR APPROPRIATE INDICATIONS AND LEAVE IN PLACE ONLY AS LONG AS NEEDED. CDC GUIDELINES FOR APPROPRIATE INDICATIONS FOR INDWELLING URETHRAL CATHETER USE: ¿ PATIENT HAS ACUTE URINARY RETENTION OR BLADDER OUTLET OBSTRUCTION. ¿ NEED FOR ACCURATE URINE OUTPUT MEASUREMENTS. ¿ USE FOR SELECTED SURGICAL PROCEDURES. ¿ TO ASSIST IN HEALING OF OPEN SACRAL OR PERINEAL WOUNDS. ¿ PATIENT REQUIRES PROLONGED IMMOBILIZATION. ¿ TO IMPROVE COMFORT FOR END OF LIFE CARE. PROPER TECHNIQUES FOR URINARY CATHETER INSERTION: - PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION. - INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. - USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. - DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN THE PATIENT RECORD. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: - SECURE THE FOLEY CATHETER. USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED - MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS - MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING - KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES - EMPTY THE COLLECTION BAG REGULARLY USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT * GENERALLY, DRAINAGE IS ACCOMPLISHED BY INSERTING THE CATHETER THROUGH THE URETHRA AND INTO THE BLADDER. HOWEVER, DRAINAGE IS SOMETIMES ACCOMPLISHED BY SUPRAPUBIC OR OTHER PLACEMENT OF THE CATHETER, SUCH AS NEPHROSTOMY TRACT." CORRECTION: H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02601 | 2023-06-23 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | CUSTOMER REPORTED SALINE LEAKAGE FROM THE POINT OF INTERSECTION - BLOOD LEAKAGE NO CONSEQUENCES DEVICE WAS USED. FIRST USE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF LEAKAGE AT THE CONNECTION WAS CONFIRMED AND APPEARED TO BE MANUFACTURING-RELATED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR S/L GROSHONG CATHETER. THE SAMPLE WAS RETURNED ASSEMBLED WITH A TWO-PIECE CONNECTOR. BLOOD RESIDUE WAS ABUNDANT WITHIN THE CATHETER. AN ATTEMPT TO INFUSE WATER THROUGH THE SAMPLE USING A 12ML SYRINGE REVEALED THE CATHETER TO BE PATENT TO INFUSION WITH NO OBSERVED LEAKS; HOWEVER, DURING PRESSURIZATION, THE CATHETER DETACHED FROM THE CONNECTOR. FOLLOWING LONGITUDINAL BISECTION OF THE DISTAL CONNECTOR SEGMENT, INSPECTION OF THE BORE REVEALED THAT NO COMPRESSION SLEEVE WAS PRESENT. THE LACK OF COMPRESSION SLEEVE CAUSED THE CATHETER TO EASILY DETACH FROM THE CONNECTOR DURING SAMPLE EVALUATION. IT IS FURTHER REASONABLE TO CONCLUDE THAT THE MISSING COMPRESSION SLEEVE WOULD ALSO RESULT IN LEAKAGE AT THE CONNECTION AS INDICATED IN THE EVENT DESCRIPTION. THE SLEEVE APPEARED TO HAVE BEEN OMITTED DURING DEVICE ASSEMBLY. PHOTOGRAPHS HAVE BEEN FORWARDED TO THE MANUFACTURING SITE FOR FURTHER EVALUATION. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04648 | 2023-06-23 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE URINE WAS NOT DRAINED THROUGH THE CATHETER AFTER THE FOLEY CATHETER PLACEMENT. WHEN REPLACING IT WITH ANOTHER CATHETER, URINE WAS DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04649 | 2023-06-23 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE CATHETER COULD NOT BE REMOVED. THE FIXATION WATER WAS REMOVED BUT THE CATHETER WAS STUCK AND COULD NOT BE REMOVED. IT WAS NOT CLARIFIED IF THE STERILE WATER WAS COMPLETELY DRAINED AND HOW THE CATHETER WAS EVENTUALLY REMOVED. THE PHYSICAL SAMPLE WAS REMOVED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE 1. MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02617 | 2023-06-23 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1275118 |  | CUSTOMER SUBMITTED PRODUCT AND PATIENT INFORMATION FOR A COMPLAINT BUT DID NOT DESCRIBE THE COMPLAINT ITSELF. CUSTOMER REPORTS THE PRODUCT WAS NOT USED ON A PATIENT AND DID NOT CAUSE HARM. ADDITIONAL INFORMATION RECEIVED ON 08-JUNE-23. CUSTOMER REPORTED THAT WHILE INSERTING THE 5FR DOUBLE LUMEN PICC, THE RUBBER STOPPER WHERE THE INSERTION WIRE/STYLET PASSES THROUGH LEAKS SALINE AND DRAWS AIR INTO THE SYRINGE ON THE EXTENSION PIECE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02618 | 2023-06-23 | 2023-06-06 | I, n, j, u, r, y | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. DUAL LUMEN PLACED IN THE UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02619 | 2023-06-23 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT: FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. THIS EVENT INVOLVED A SINGLE LUMEN PPIC IN THE UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02602 | 2023-06-23 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | CK000516 |  | CUSTOMER REPORTED LEAKS AND LIQUID SPILLAGE DURING INFUSION. A HOLE ON THE FITTING WAS NOTICED. CATHETER WAS REMOVED AND A NEW PICC PLACED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02584 | 2023-06-23 | 2023-05-30 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808060 | 1808060 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY HAD BLOOD CLOT IN ARM. IT WAS FURTHER REPORTED THAT PATIENT EXPERIENCED SWELLING AROUND THE PORT SITE. REPORTEDLY, ADDITIONAL MEDICATION WAS PRESCRIBED FOR THE TREATMENT. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 07/2024). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04610 | 2023-06-23 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, THE CATHETER BALLOON WAS INFLATED BUT THE WATER COULD NOT BE DRAINED. | THE REPORTED EVENT WAS UNCONFIRMED AS THE PRODUCT MET SPECIFICATIONS. NO ROOT CAUSE WAS PROVIDED AS THE REPORTED EVENT WAS UNCONFIRMED. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE REPORTED EVENT WAS UNCONFIRMED HOWEVER, A DEVICE HISTORY RECORD REVIEW WAS COMPLETED PRIOR TO SAMPLE EVALUATION. THE DHR FOUND NO MANUFACTURING DISCREPANCIES IN THE REVIEW THAT COULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. LABELLING REVIEW WAS NOT REQUIRED AS THE REPORTED EVENT WAS UNCONFIRMED. H11: SECTION A THROUGH F THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02582 | 2023-06-23 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT TI L/P 6 CF INT WOSP ATT SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8716001 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA RIGHT SUBCLAVIAN VEIN, THE GUIDE WIRE OF THE PORT WAS ALLEGEDLY FOUND TO HAVE DRAWING PROBLEM. IT WAS FURTHER REPORTED THAT THE GUIDEWIRE ALLEGEDLY WAS NOT TIGHT TWINED AND LOSS TENSION. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 08/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00444 | 2023-06-23 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH DIALYSIS CATHETER KIT 12F X 20CM | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | NA | 5553200 |  | CUSTOMER REPORTED THAT WHEN A DIALYSIS CATHETER WAS USED IN HEMODIALYSIS ROOM TO ESTABLISH VASCULAR ACCESS FOR DIALYSIS PATIENTS, THE GUIDE WIRE IN THE KIT WAS INSERTED INTO THE CATHETER AND THE DEFORMED GUIDE WIRE SPREAD OUT. THE NURSE IMMEDIATELY REPLACED THE CATHETER KIT, WHICH DID NOT HARM THE PATIENT BUT CAUSED DISCOMFORT TO THE PATIENT FOR A SHORT TIME. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02600 | 2023-06-23 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER "WHEN USING PICC THE CATHETER WAS LEAKING. THEY PULLED OUT THE CATHETER AND SAW A HOLE 10 CM INSIDE THE SKIN INTO THE VESSEL. THEY TREATED THE PATIENT AS EXTRAVASATION OF CHEMO. THE PACLITAXEL. THE PATIENT DID NOT SUFFER ANY HARM." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02599 | 2023-06-23 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 20G X 8 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F320088PT |  | IT WAS REPORTED BY THE CUSTOMER THAT POWERGLIDE WIRE FRAYED WHILE ATTEMPTING PLACEMENT. NO HARM, A NEW CATHETER WAS INSERTED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GUIDEWIRE WAS CONFIRMED AND THE CAUSE APPEARED TO BE USE-RELATED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 20GA X 8CM POWERGLIDE PRO MIDLINE CATHETER ASSEMBLY. THE CATHETER HAD BEEN ADVANCED AND WAS NOT RETURNED. THE SAFETY MECHANISM WAS ADVANCED AND FULLY COVERED THE NEEDLE TIP. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. A COMPLETE BREAK WAS OBSERVED IN THE GUIDEWIRE CORE WIRE. THE COIL WIRE WAS ELONGATED BUT INTACT. MICROSCOPIC INSPECTION OF THE BREAK IN THE CORE WIRE REVEALED A GRANULAR FRACTURE SURFACE. A REGION OF INCREASED LUSTER WAS OBSERVED. CURVED SHAPE MEMORY WAS OBSERVED IN THE VICINITY OF THE BREAK. INSPECTION OF THE NEEDLE REVEALED DEFORMATION ALONG THE PROXIMAL EDGE OF THE BEVEL. THE GUIDEWIRE FRACTURE FEATURES AND NEEDLE BEVEL DEFORMATION WAS CONSISTENT WITH DAMAGE CAUSED BY CONTACT BETWEEN THE GUIDEWIRE AND THE NEEDLE TIP. SUCH DAMAGE MAY OCCUR IF THE WIRE IS WITHDRAWN AGAINST THE NEEDLE OR THE NEEDLE IS INSERTED AGAINST THE WIRE FOLLOWING ADVANCEMENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02598 | 2023-06-23 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 4F SINGLE-LUMEN CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6194335 |  | CUSTOMER REPORTED THAT WHEN DILATING THE VEIN, THERE WAS A PROBLEM WITH THE DILATOR MAKING IT IMPOSSIBLE TO INSERT. AFTER SEVERAL ATTEMPTS, THE DISTAL PART OF THE DILATOR WAS FOUND TO BE DAMAGED, MAKING IT IMPOSSIBLE TO INSERT. GREY PART AS IF "TORN". THE DEVICE WAS DISCARDED. DESCRIPTION OF CONSEQUENCES: GREATER INCISION IN THE PATIENT'S ARM TO ALLOW THE DILATOR TO BE INSERTED, UNTIL THE MATERIAL DETERIORATED. IN ADDITION, 1 PICC-LINE WAS OPENED ONLY TO TAKE A NEW DILATOR. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04625 | 2023-06-23 | 2023-06-06 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT CURRENTLY GOT A URINARY TRACT INFECTION AND THEY THOUGHT IT WAS CAUSED BY THE MAGIC 3 GO MALE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, CUSTOMER STATED THAT HE WAS TREATED WITH AN ANTIBIOTIC, AND NO SAMPLE IS AVAILABLE. THE PRODUCT LOT #: JUGB0666, REFERENCE NUMBER (B)(4). CUSTOMER STATED HE HAD BEEN SPEAKING WITH A LAWYER IN REGARD TO HIS USE OF THESE CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04626 | 2023-06-23 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER DID NOT HAD THE WATER SACHET IN THEM. THE CUSTOMER HAD BEEN BOILING THE KETTLE FOR LETTING IT COOL AND THEN ADDING TO THE CATHETER. PER MAIL RECEIVED FROM IBC ON 21JUN2023, STATED THAT THE PATIENT CALLED TO STATE THAT THE NATURE OF THE COMPLAINT WAS WRONG. INITIALLY STATED THAT THEY DIDN'T RECEIVE THE WATER SACHETS. PATIENT STATED THAT THEY DID RECEIVE THEM, THEY JUST DIDN'T HAD ANY WATER IN THEM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04624 | 2023-06-23 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE BELOW PART OF THE FUNNEL UPON THE FOLEY CATHETER REMOVAL PROCEDURE. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 13JUN2023, STATED THAT THE LEAKING WAS FOUND DURING CATHETER REMOVAL. | HE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04619 | 2023-06-23 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119112 | 119112 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT AFTER TWO DAYS OF INDWELLING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04607 | 2023-06-23 | 2023-06-02 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER (SILICONE) | C.R. BARD, INC. (COVINGTON) -1018233 | 123516A | 123516A |  | IT WAS REPORTED THAT THE CUSTOMER WANTS TO RETURN THREE FOLEY CATHETERS FROM ORDER NUMBER# (B)(4). THE CUSTOMER STATED THAT THE DOCTOR SWITCHED THE PATIENT TO A DIFFERENT CATHETER BECAUSE THE USER WAS HAVING AN ALLERGIC REACTION. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, CUSTOMER WAS STILL EXPERIENCED BLADDER PAIN AND URGENCY. CUSTOMER WAS BEING TREATED WITH UNKNOWN ANTIBIOTICS FROM THEIR DOCTOR AND THEY WERE NO LONGER USING PRODUCT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE MECHANICAL FAILURE, OPERATOR ERROR, USER HYPERSENSITIVITY TO LATEX, BACTERIAL INFECTION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON TO BURST. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX.. SINGLE PATIENT USE ONLY. DO NOT REUSE. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SYRINGE. DO NOT USE NEEDLE. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. DO NOT USE IF PACKAGE IS OPENED OR DAMAGED. RECOMMENDED INFLATION CAPACITIES: 5CC BALLOON: USE 10CC STERILE WATER; 30CC BALLOON: USE 35CC STERILE WATER; DO NOT EXCEED RECOMMENDED CAPACITIES. NOTE: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LAWS AND REGULATIONS. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04613 | 2023-06-23 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123420C |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT INDWELLING CATHETERIZATION ON (B)(6) 2023 AND CONTINUED BLADDER IRRIGATION AFTER SURGERY. THE COLOR OF DRAINAGE WAS LIGHT RED. AT 10:05 ON MAY 1. THE PATIENT'S URINARY CATHETER BALLOON RUPTURED, AND THE DOCTOR RE-INDWELLED THE THREE-LUMEN URINARY CATHETER FOR THE PATIENT, CONTINUED TO CONNECT THE BLADDER FOR CONTINUOUS IRRIGATION, AND OBSERVED THE URINE COLOR. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04615 | 2023-06-23 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT STERILE WATER COULD NOT BE DRAINED DURING PRE-TEST. | UPON FURTHER REVIEW OF INVESTIGATION, BD HAS DETERMINED THAT THIS MDR IS NOT REPORTABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04609 | 2023-06-23 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE PATIENT DID NOT FEEL WATER SACHET LUBES WHILE USING THE INTERMITTENT CATHETERS, AND GOOD ENOUGH AND WANTED TO TEST DIFFERENT ONES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04611 | 2023-06-23 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE BASE OF THE CATHETER DURING PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02606 | 2023-06-23 | 2023-05-19 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT ATTEND (B)(6) FOR CHEMO TREATMENT TODAY HAS DIAGNOSED DVT YESTERDAY BUT PICC LOOK LIKE NORMAL AND CONSULTANT IS HAPPY TO START CHEMOTHERAPY AND TODAY CHECK THE PICC LINE THAT IS FLUSHED WELL AND VENOUS RETURN IS GOOD AND STARTED WITH DOXORUBICIN HE COMPLAINT ABOUT KIND OF HEAVY FEEL AND PRESSURE INSIDE STILL VENOUS RETURN WAS GOOD INFORM THE DOCTOR HE IS HAPPY TO CONTINUE AND HAD CHAT WITH PATIENT INFORMED HIM TO DISCOMFORT DUE TO THE THROMBOSIS HE IS ON INJECTION ENOXAPARIN THAT CAN SORTED ISSUES THEN CONTINUE WITH THE CHEMOTHERAPY AND DOUBLE CHECKED THE LINE WITH ANOTHER STAFF NURSE VENOUS RETURN WAS PRESENT FLUSHED WELL THEN STARTED ANOTHER ONE VINBLASTIN FOR 5 MINUETS HE COMPLAINT ABOUT THE PAIN AGAIN SO EVERYTHING STOPPED THEN CONSULTED SEND HIM FOR RADIOGRAPHY AND TREATED AS EXTRAVASATION. PATIENT REQUIRED INTERVENTION BY PLASTIC SURGERY. NO OTHER INFORMATION WAS PROVIDED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A SPLIT IN THE CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR S/L POWERPICC SOLO. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED BETWEEN THE 9 CM AND 10 CM DEPTH MARKER. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED. FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR. 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL. BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING). AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02588 | 2023-06-23 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THIS SENSOR HAS A ECG SIGNAL ON THE YELLOW INTRAVASCULAR LINE EVEN WITHOUT A PICC CONNECTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL BE EVALUATED. A SUPPLEMENTAL WILL BE SUBMITTED WITH EVALUATION RESULTS. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00443 | 2023-06-23 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80144 | ATG80144 |  | IT WAS REPORTED DURING AN ANGIOPLASTY PROCEDURE, BALLOON ALLEGEDLY SEPARATED FROM THE SHAFT. IT WAS FURTHER REPORTED THAT THE BALLOON ALLEGEDLY STARTED MIGRATING THUS IT WAS STENTED IN PLACE. THERE WAS NO REPORTED PATIENT INJURY. | ADDITIONAL INFORMATION WAS RECEIVED, AND THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE REPORTABLE AS MALFUNCTION. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. D4 (EXPIRATION DATE: 02/2026). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04605 | 2023-06-23 | 2023-06-01 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT CUSTOMER DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. IT WAS STATED THAT THE PATIENT WAS HOSPITALIZED TWICE BECAUSE OF URINARY TRACT INFECTION AND DID NOT KNOW IF THE CAUSE WAS FROM THE RECALL MAGIC 3 GO INTERMITTENT CATHETERS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION WAS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04606 | 2023-06-23 | 2023-06-01 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51812 | 51812 |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT HAD MULTIPLE URINARY TRACT INFECTIONS, WAS USING RECALLED INTERMITTENT CATHETERS EVEN AFTER RECEIVING THE LETTER, ADVISED TO THROW THEM AWAY IMMEDIATELY, RUSHED ORDER TO BE OVERNIGHTED. PER FOLLOW UP VIA PHONE (B)(6) 2023, CUSTOMER HAD TO HAVE CONSTANT TREATMENT DUE TO THE URINARY TRACT INFECTIONS. ALSO STATED THE PATIENT WAS ON NONSTOP ANTIBIOTICS AND SEEING DIFFERENT SPECIALISTS BECAUSE NOTHING WAS CURING THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-06
17189915 | 2023-06-23 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | PUREWICK FEMALE EXTERNAL CATHETER | COLLECTOR, URINE, POWERED, NON INDWELLING CATHETER | C. R. BARD, INC. | PWF030 | PWF030 |  | WHILE SETTING UP A NEW PUREWICK, IT WAS NOTED THE PUREWICK DIDN'T HAVE SUCTION. ANOTHER PUREWICK WAS OBTAINED AND WORKED WITHOUT DIFFICULTY. MANUFACTURER RESPONSE FOR FEMALE EXTERNAL CATHETER, PUREWICK FEMALE EXTERNAL CATHETER (PER SITE REPORTER). EQUIPMENT FAILURE REPORTED TO BARD CUSTOMER SERVICE. EQUIPMENT IS AVAILABLE FOR RETURN TO MANUFACTURER. |  | N | Y |  |  | 2023-06
3006260740-2023-02583 | 2023-06-23 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT THREE MONTHS AND TWO DAYS POST DIALYSIS CATHETER PLACEMENT, THE CUFFS APPEARED PAPERY WITH NO FIBROSIS. IT WAS FURTHER REPORTED THAT CATHETER WAS ALLEGEDLY CAME OUT FROM PATIENT SKIN SPONTANEOUSLY. REPORTEDLY THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02616 | 2023-06-23 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT POWERPICC 4F SL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000228B |  | IT WAS REPORTED BY THE CUSTOMER THAT CONNECTOR WOULD NOT ATTACH TO THE HUB OF THE PICC PROPERLY AND CAUSED LEAKING BETWEEN THE HUB AND THE CONNECTOR. HAD TO CHANGE OUT THE PICC. ADDITIONAL INFO RECEIVED ON 06-JUNE-23 THE REPORTED EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION, OR CAUSE A CLINICALLY SIGNIFICANT DELAY IN TREATMENT. NO MEDICATIONS WERE INFUSING AT THE TIME OF THE EVENT. CATHETER WAS SECURED WITH STATLOCK. NO HARM OR INJURY TO THE PATIENT OCCURRED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04631 | 2023-06-23 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS SPONTANEOUSLY WITHDRAWN IN THE ICU AFTER INSERTION OF THE CATHETER IN THE OPERATING ROOM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02605 | 2023-06-23 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405J |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE. THEREFORE, THE CATHETER WAS CUT LITTLE BY LITTLE IN 3 TIMES, BUT THERE WAS NO IMPROVEMENT. THE PROCEDURE WAS COMPLETED BY USING ANOTHER CATHETER CONNECTOR. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LOOSE CONNECTION BETWEEN A PROXIMAL CONNECTOR PIECE AND A CATHETER IS INCONCLUSIVE DUE TO THE STATE OF THE RETURNED SAMPLE. ONE 4 FR PROXIMAL GROSHONG NXT CONNECTOR PIECE WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED THE CONNECTOR PIECE WAS RETURNED WITHOUT ANY CATHETER SEGMENTS OR THE DISTAL CONNECTOR PIECE. A MICROSCOPIC OBSERVATION REVEALED SOME USE RESIDUE ON THE RETURNED SAMPLE. NO DAMAGE WAS OBSERVED ON THE RETURNED SAMPLE. THE OUTER DIAMETER OF THE METAL CANNULA/STEM OF THE PROXIMAL CONNECTOR PIECE WAS MEASURE AND WAS FOUND TO BE WITHIN SPECIFICATION. AN ATTEMPT WAS MADE TO CONNECT THE RETURNED PROXIMAL CONNECTOR PIECE WITH A NON-COMPLAINANT DISTAL CONNECTOR PIECE. AN AUDIBLE CLICK WAS HEARD WHEN THE CONNECTOR PIECES WERE ASSEMBLED AND THE CONNECTION WAS FOUND TO BE FIRM. THE CONNECTION BETWEEN THE PROXIMAL CONNECTOR PIECE AND THE CATHETER COULD NOT BE CONFIRMED WITHOUT THE RETURN OF THE ALLEGED CATHETER. THE CANNULA/STEM OF THE RETURNED PROXIMAL CONNECTOR PIECE WAS FOUND TO BE WITHIN SPECIFICATION AND NO ISSUES WERE FOUND DURING TESTING OF THE CONNECTION WITH A NON-COMPLAINANT CATHETER AND DISTAL CONNECTOR PIECE. THEREFORE, THIS COMPLAINT IS INCONCLUSIVE AT THIS TIME. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02607 | 2023-06-23 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG MID LINE MINI TRAY 4F | CATHETER, PERCUTANEOUS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7715435CE |  | IT WAS REPORTED THAT THE CATHETER IS ANGLED AND ON 27/04 IT SHOWS FISSURE SPILLAGE. CATHETER WAS REMOVED. LOCAL CHEMOTHERAPY SPREAD AND MIDLINE REPOSITIONING (OTHER VENOUS ACCESS). NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
MW5118848 | 2023-06-23 | 2023-06-20 | I, n, j, u, r, y | Voluntary report | POWERGLIDE PRO MIDLINE CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | F118108P | F118108PT |  | THE PATIENT HAD AN 18G RIGHT UPPER ARM MIDLINE (BASILIC VEIN) CATHETER PLACED ON (B)(6) 2023. ON (B)(6) 2023 WHILE AT OUTPATIENT THERAPY TO GET HIS IV ANTIBIOTICS THE LINE WAS NOTED TO BE LEAKING. THE RN PULLED DOWN THE DRESSING AND NOTICED THE CATHETER HUB AND PART OF THE CATHETER HAD BROKEN OFF, LEAVING A PART OF THE CATHETER IN THE PATIENT'S ARM. THE PATIENT WAS TAKEN TO THE ED(EMERGENCY DEPT) AND THEN WENT TO THE OPERATING ROOM FOR A VENOTOMY WHERE THE BROKEN PIECE IN THE PATIENT'S ARM WAS SUCCESSFULLY REMOVED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04662 | 2023-06-26 | 2023-06-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 51814 |  | IT WAS REPORTED THAT CUSTOMER RECEIVED THE MAGIC GO RECALL NOTICE AND HAVE THE AFFECTED PRODUCT. THEY HAD BEEN EXPERIENCING A SEVERE BLADDER INFECTION AND WAS UNDER DOCTOR CARE AND WAS ON ANTIBIOTICS METHENAMINE AND NITROFURTANTOIN MONO MCR. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD WAS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02623 | 2023-06-26 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, THE CONTRAST MEDIUM LEAKED NEAR THE VASCULAR INSERTION SITE. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY HAD A CRACK. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 12/2022). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02626 | 2023-06-26 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT SEVEN YEARS SIX MONTHS AND ONE DAY POST PORT PLACEMENT PROCEDURE VIA THE RIGHT SUBCLAVIAN VEIN, LEAKAGE WAS ALLEGEDLY FOUND WHEN CT EXAMINATION WAS PERFORMED. IT WAS FURTHER REPORTED THAT THE PORT SYSTEM WAS REMOVED AND A CRACK WAS FOUND ON THE REMOVED CATHETER. REPORTEDLY, CT IMAGE REVEALED THAT A KINK WAS FOUND ON CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04678 | 2023-06-26 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT SURGERY YESTERDAY, AND A FOLEY CATHETER WAS PLACED IN THE OPERATING ROOM AFTER THE OPERATION. AT 5 O'CLOCK IN THE MORNING, THE CATHETER BALLOON RUPTURED AND THE URINARY CATHETER PROLAPSED. PER FOLLOW-UP INFORMATION RECEIVED FROM (B)(6) ON 14JUN2023, STATED THAT THE EVENT WAS ABOUT CATHETER BALLOON RUPTURE. THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02636 | 2023-06-26 | 2023-06-08 | I, n, j, u, r, y | Manufacturer report | POWERPICC PROVENA CATHETER W/ SOLO VALVE & SHERLOCK 3CG (TPS) STYLET (5F) (TL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1395108D |  | IT WAS REPORTED BY THE PROVIDER THAT THEY HAD A PATIENT WITH A POWERPICC PROVENA THAT WAS PLACED A FEW WEEKS AGO. WHEN REMOVING THE PICC FROM THE PATIENT, THE PIC WAS STUCK AND HAD ADHERED TO THE VESSEL WALL. PROVIDER STATES THE FIBRIN HAD BUILT UP AROUND THE CATHETER CAUSING IT TO GET STUCK TO THE VEIN WALL. INFORMED PROVIDER THE PICC CATHETER ADHERING TO THE VESSEL WALL IS UNCOMMON. THERE ARE FEW REPORTS OF VAD ADHERING TO THE VESSEL WALLS, I AM UNSURE OF ANY OCCURRENCE WITH THE POWERPICC PROVENA SPECIFICALLY. WE DO NOT HAVE SPECIFIC RECOMMENDATIONS FOR HOW TO REMOVE A PICC WHEN STUCK OR ADHERED TO THE VESSEL WALL. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02637 | 2023-06-26 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT W/BIOPATCH, PROBE COVER, & BELDAR NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED THAT THE SALINE LEAKING OUT OF T-LOCK. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02633 | 2023-06-26 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE CATHETER COULD NOT BE ENTERED DURING CATHETER PLACEMENT. IT KEPT JUMPING TOWARDS THE INSIDE OF THE NECK ON ITS OWN. UPON REMOVAL OF THE STYLET, IT WAS FOUND THAT THE STYLET WAS BENT IN THE MIDDLE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02632 | 2023-06-26 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG (TPS) STYLET (4F) (SL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED THE NURSE NOTICED RED LIQUID COMING FROM THE EXIT SITE WHILST ADMINISTERING EPIRUBICIN. IMMEDIATELY STOPPED - FLUSHED 20MLS. PICC REMOVED ¿INTERNAL FRACTURE AT 36CM ¿ 3CM FROM EXIT SITE. SACT ADMINISTERED PERIPHERALLY. EXIT SITE MONITORED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04689 | 2023-06-26 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED, THAT THE FOLEY TIP WAS DIFFERENT. AND SEEMED A BIT BIGGER. THEY WERE CONTINUING TO HAVE PATIENT COMPLAINTS WITH REMOVAL OF THE FOLEYS. THEY CURRENTLY, HAD A STOCK OF 5 CASES, THAT THEY WOULD LIKE TO RETURN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04688 | 2023-06-26 | 2023-06-09 | I, n, j, u, r, y | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 995081 | 995081 |  | IT WAS REPORTED THAT IN AN ONLINE SURVEY A PHYSICIAN STATED A PERFORATION ADVERSE EVENT, DIRECTLY ATTRIBUTABLE TO THE DEVICE 995081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, STERILE. ALSO STATING THAT THE PATIENT WITH THE DEVICE 996081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, W/ INFLATION DEVICE, STERILE PLACEMENT DID RESULT IN PATIENT INJURY REQUIRING MEDICAL OR SURGICAL INTERVENTION BECAUSE FORCED DILATATION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW UP INFORMATION RECEIVED VIA MAIL ON 14JUN2023, IN AN ONLINE SURVEY, A PHYSICIAN STATED A PERFORATION ADVERSE EVENT, DIRECTLY ATTRIBUTABLE TO THE DEVICE 995081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, STERILE. PHYSICIAN ALSO INDICATED 'PERFORATION' WAS PRESENT PRIOR TO DEVICE PLACEMENT. ALSO STATING THAT THE PATIENT WITH THE DEVICE 996081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, W/ INFLATION DEVICE, STERILE PLACEMENT DID RESULT IN PATIENT INJURY REQUIRING MEDICAL OR SURGICAL INTERVENTION BECAUSE FORCED DILATATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04687 | 2023-06-26 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE TEMPERATURE SENSING FOLEYS WERE NOT ACCURATE IN PATIENTS WHO WERE UNABLE TO URINATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02627 | 2023-06-26 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE. THEREFORE, THE CATHETER WAS CUT AND COMPLETED TO THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04684 | 2023-06-26 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED, THAT THE FOLEY CATHETER WAS WITHDRAWN SPONTANEOUSLY SOON AFTER USE. WHEN CHECKED, THE BALLOON WAS FOUND TO BE RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04685 | 2023-06-26 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED, THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT. ABOUT 2 HOURS, AFTER FOLEY CATHETER PLACEMENT IN THE OPERATING ROOM. THE PATIENT BECAME AWARE OF THE NEED TO URINATE, BUT COULD NOT CONFIRM URINE IN THE CATHETER. AFTER CONFIRMING, THAT THE PATIENT'S GENERAL CONDITION WAS GOOD, THE USER ATTEMPTED TO REMOVE THE CATHETER. WHEN WATER WAS ADDED TO THE REMOVED CATHETER, THE BALLOON WAS INFLATED, AND NO WATER LEAKAGE WAS CONFIRMED. AFTER THAT, IT WAS UNCLEAR IF THERE WERE ANY MICRO RUPTURES THAT COULD NOT BE SEEN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04683 | 2023-06-26 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS USED FOR CONTINUOUS BLADDER IRRIGATION, WHICH COULD DRAIN FLUENTLY AND WAS FIXED GOOD ON (B)(6)2023. ON THE MORNING OF (B)(6)2023, WHEN THE BLADDER WAS FLUSHED AND COULD ONLY GO IN AND OUT, AND THEN THE THREE-LUMEN CATHETER CAME OUT BY ITSELF, AND THE BALLOON WAS FOUND TO BE RUPTURED. AFTER REPORTING IT TO THE DOCTOR, A NEW CATHETER WAS REPLACED. AS PER INJURY PERFORMANCE, STATED IT CAUSED PAIN TO THE PATIENT. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04665 | 2023-06-26 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT STAFF ATTEMPTED TO PLACE TWO CATHETERS IN PATIENTS AND THE BALLOONS WERE LEAKED. PER FOLLOW UP VIA PHONE ON 07JUN2023, IT WAS REPORTED THAT THE LOTS WERE NGHN4280 AND NGGW0794, THESE WERE THE ONLY LOTS THAT HAVE HAD AN ISSUE AND NO PATIENT IMPACT REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02620 | 2023-06-26 | 2023-05-29 | I, n, j, u, r, y | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0602830T |  | IT WAS REPORTED THAT FOUR YEARS TEN MONTHS AND ELEVEN DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, THE CATHETER WAS ALLEGEDLY RUPTURED UPON REMOVAL. IT WAS FURTHER REPORTED THE PART OF THE CATHETER ALLEGEDLY FELL INTO THE HEART AND THEN REMOVED BY INTERVENTIONAL MANNER. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE BARDPORT MRI IMPLANTABLE PORT AND A GROSHONG CATHETER IN TWO SEGMENTS WERE RETURNED AND ONE X-RAY IMAGE WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED TO THE PORT STEM AND CRACKS WERE NOTED TO THE CATH-LOCK. A PARTIAL LONGITUDINAL SPLIT WAS NOTED ON THE PROXIMAL CATHETER SEGMENT APPROXIMATELY 5.6CM FROM THE DISTAL END OF THE CATH-LOCK AND A COMPLETE COMPOUND CIRCUMFERENTIAL BREAK WAS ALSO NOTED AT THE DISTAL END OF THE PROXIMAL CATHETER SEGMENT. THE X-RAY DEMONSTRATES A PORT CATHETER IN THE RIGHT ATRIUM, FULLY SEPARATED FROM THE PORT. THE PORT CATHETER HAS SEPARATED FROM THE EXTERNAL PORT. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION AND MIGRATION ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04663 | 2023-06-26 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53618 | 53618 |  | IT WAS REPORTED THAT PATIENT STATED, THERE WAS HARDLY ANY LUBRICANT. IT WAS HARDLY THERE AT ALL, AND THE CATHETER WAS ACTUALLY SQUEAKING. PATIENT WANTED SOMETHING DIFFERENT. NO ADDITIONAL INFORMATION PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
2020394-2023-00450 | 2023-06-26 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQ5074 | CQ5074 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY BURST WHEN THE PRESSURE WAS LESS THAN 30 ATM. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. ONE PHOTO WAS REVIEWED. THE PHOTO SHOWS THE BALLOON WAS NOTED TO BE STUCK WITHIN AN UNKNOWN SHEATH, THE BALLOON WAS NOTED TO BE IN A DEFLATED CONDITION, AND THE CATHETER WAS NOTED TO BE DETACHED FROM THE BALLOON, EXPOSING THE INNER GUIDE WIRE LUMEN. NO EVIDENCE OF BALLOON RUPTURE COULD BE NOTED DUE TO THE DEVICE'S CONDITION. NO OTHER ANOMALIES WERE NOTED. BASED ON THE PHOTO REVIEW, IT WAS OBSERVED THAT THE CATHETER WAS NOTED TO BE DETACHED, EXPOSING THE INNER GUIDE WIRE LUMEN. THE BALLOON WAS NOTED TO BE STUCK WITHIN AN UNKNOWN SHEATH, BUT NO OBJECTIVE EVIDENCE OF BALLOON RUPTURE WAS NOTED. THEREFORE, THE INVESTIGATION WAS CONFIRMED FOR THE IDENTIFIED BALLOON REMOVAL ISSUE, CATHETER DETACHMENT AND THE REPORTED BALLOON RUPTURE REMAINS INCONCLUSIVE. A DEFINITIVE ROOT CAUSE FOR THE IDENTIFIED BALLOON REMOVAL ISSUE, CATHETER DETACHMENT AND REPORTED BALLOON RUPTURE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELLING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRY DATE: 01/2025). THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04669 | 2023-06-26 | 2023-06-08 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT CUSTOMER KEPT GETTING URINARY TRACT INFECTION (UTIS) CONSTANTLY WHILE USING INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, CUSTOMER WAS TREATED WITH ANTIBIOTICS FOR THE URINARY TRACT INFECTION (UTIS) AND NO LONGER USES THE CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04664 | 2023-06-26 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT PATIENT STATES CATHETERS WERE MESSY AND TOOK TOO LONG. REPRESENTATIVE ADVISED PATIENT ON HOW TO USE THESE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INSTRUCTIONS FOR USE: INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS OR PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04674 | 2023-06-26 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED, THAT WHEN THE STERILE WATER WAS INJECTED INTO THE BALLOON DURING PRETEST. IT FLOWED BACK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04668 | 2023-06-26 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILE WATER DID NOT RETURN TO THE SYRINGE DURING THE DEFLATION PROCEDURE OF THE PRETEST. THEN THE FOLEY CATHETER BALLOON WAS DEFLATED. ABOUT 10 ML OF WATER WAS ASPIRATED BY SYRINGE. THE USER TRIED TO DEFLATE BALLOON BY SPONTANEOUS OR GENTLE DEFLATION BY SYRINGE. USUALLY, BALLOON CAN BE DEFLATED WITHIN FEW SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04670 | 2023-06-26 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI18 |  | IT WAS REPORTED THAT THE CATHETER FELL OUT OF THE PATIENT AND WAS FOUND STERILE WATER LEAKING FROM THE SHAFT AREA. ACCORDING TO THE INQUIRY SHEET, THE EVENT OCCURRED JUST AFTER CATHETER PLACEMENT AND THERE WAS NO BALLOON RUPTURE/TEAR. NO MATERIALS CAME IN CONTACT WITH THE CATHETER AND THERE WAS NO PATIENT'S URINE STONE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04671 | 2023-06-26 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT THE CUSTOMER HAD ISSUE WITH OUR SILICONE FOLEY IN FOLEY TRAY. FROM WHAT THEY DESCRIBED, IT SOUNDED LIKE A SILICONE CUFFING ISSUE UPON FOLEY REMOVAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04660 | 2023-06-26 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED, THAT DURING A PRETEST. THE CATHETER BALLOON WAS INFLATED, BUT THE STERILE WATER COULD NOT BE REMOVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04661 | 2023-06-26 | 2023-06-03 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT DEVELOPED A URINARY TRACT INFECTION AND SPENT 8 DAYS IN THE HOSPITAL. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
2020394-2023-00449 | 2023-06-26 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DORADO | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | DR8074 | DR8074 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE BALLOON ALLEGEDLY GOT STUCK IN THE SHEATH. IT WAS FURTHER REPORTED THAT THE BALLOON WAS REMOVED WITH THE SHEATH. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 09/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02624 | 2023-06-26 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN CV CATHETER, TRIPLE-LUMEN, 12.5F | CHRONIC CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 0600650 | 0600650 |  | IT WAS REPORTED THAT SOMETIME POST A CHRONIC CATHETER PLACEMENT, THE WHITE LEG WAS ALLEGEDLY RUPTURED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
17197854 | 2023-06-26 | 2023-04-29 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | DECATHLON DF LONG-TERM HEMODIALYSIS CATHETER | CATHETER, HEMODIALYSIS, IMPLANTED | BARD ACCESS SYSTEMS, INC. | DF19SH24K | DF19SH24K |  | WHILE TUNNELING THE DIALYSIS CATHETER, IT BROKE AWAY FROM THE TUNNELER AND WOULD NOT STAY ON. THERE WAS PATIENT CONTACT, BUT NO HARM DONE. A NEW CATHETER KIT WAS USED TO COMPLETE THE PROCEDURE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04657 | 2023-06-26 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING PAIN WITH SUCTION FROM USING PUREWICK FEMALE EXTERNAL CATHETER. REPRESENTATIVE ADVISED TO STOP USAGE AND SPEAK WITH DOCTOR AND THE LOT INFORMATION NOT GIVEN. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04752 | 2023-06-27 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | INTL15 | INTL15 |  | IT WAS REPORTED THAT THE SURESTEP INTERMITTENT CATHETER TRAY WERE NOT COMPLETELY DRAINING THE PATIENT'S BLADDER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿HUMAN ERROR / LACK OF TRAINING¿. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® EZ-LOK¿ SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK (FIG. 1A) OR SLIP TIP COLLECTION DEVICES (FIG. 1B). BD VACUTAINER® URINE PRODUCTS CAN BE USED FOR COLLECTION, STORAGE, AND TRANSPORT OF URINE SPECIMENS FROM A URINARY CATHETER. 1. SWAB SURFACE OF BARD® EZ-LOK¿ SAMPLING PORT WITH ANTISEPTIC (FIG. 2) E.G., SITE-SCRUB¿ IPA DEVICE WHICH IS AN EFFECTIVE DISINFECTING AGENT FOR LUER HUBS. 2. USING ASEPTIC TECHNIQUE, POSITION THE COLLECTION DEVICE IN THE CENTER OF THE SAMPLING PORT. PRESS THE DEVICE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. (FIG. 3) 3. IF USING BD VACUTAINER® LUER-LOK¿ ACCESS DEVICE (FIG. 4), COLLECT URINE INTO THE BD VACUTAINER® TUBES PER HOSPITAL PROTOCOL. IF USING SYRINGE, SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND TRANSFER TO COLLECTION CUP OR FOLLOW HOSPITAL PROTOCOL. 4. AFTER SAMPLE IS OBTAINED, DISCARD COLLECTION DEVICE ACCORDING TO HOSPITAL PROTOCOL. 5. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02648 | 2023-06-27 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945J |  | IT WAS REPORTED THAT IMMEDIATELY AFTER OPENING THE PACKAGE, IT WAS NOTICED THAT THE TIP OF THE GUIDEWIRE WAS BENT DIFFERENTLY THAN USUAL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04745 | 2023-06-27 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILE WATER WAS INJECTED INTO THE BALLOON FOR A BALLOON TEST AND UPON TRYING TO REMOVE THE STERILE WATER FROM THE BALLOON, THE WATER COULD NOT BE DRAINED. THERE WAS NO VISUAL DAMAGE OBSERVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04748 | 2023-06-27 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT 8 CC OF STERILE WATER WAS DRAINED, BUT THE REMAINING 2 CC WAS NOT DRAINED FROM THE FOLEY CATHETER DURING THE PRETEST. THE SAME PHENOMENON OCCURRED WHEN THE TEST WAS REPEATED AGAIN. AFTER LEAVING IT FOR A WHILE, THE REMAINING 2 CC OF STERILE WATER CAME OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04749 | 2023-06-27 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE INFLATION VALVE DURING STERILE WATER INJECTION AND THE THERMISTOR WIRE WAS CRIMPED INSIDE THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04750 | 2023-06-27 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING THE BALLOON TEST, AFTER FILLING THE BALLOON WITH THE STERILE WATER, UPON TRYING TO REMOVE WATER USING A SYRINGE, THE WATER COULD NOT BE DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04753 | 2023-06-27 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER COUDE TIP | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0100L16 | 0100L16 |  | IT WAS REPORTED THAT THERE WAS A CONSTANT LEAK, THE TIP OF PRODUCT ONLY HAD ONE HOLE, SUPPOSED TO HAVE 2. THE BATCH OF 6 RECEIVED HAD ALL OF THE SAME ISSUE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02658 | 2023-06-27 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT AFTER INSERTION INTO THE PATIENT¿S BODY, LIQUID LEAKAGE OCCURRED WHEN FLUSHING THE CATHETER. THE CATHETER WAS IMMEDIATELY PULLED OUT ON THE OPERATING TABLE AND A SMALL BREACH WAS FOUND ON THE CATHETER. THE NURSE THUS REMOVED THE CATHETER FROM THE PATIENT¿S BODY. THE PATIENT IS IN GOOD CONDITION AND NO IMPACT WAS CAUSED TO THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02653 | 2023-06-27 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE OVERSLEEVE PORTION CONNECTOR COULD NOT BE ATTACHED WITH THE EXTENSION TUBING TIGHTLY. A GAP WAS LEFT AFTER THE ATTACHMENT AND BOTH WERE EASILY SEPARATED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02655 | 2023-06-27 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG 3F SINGLE-LUMEN PICC BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7717305 |  | IT WAS REPORTED THAT WHEN PLACING THE PICC CATHETER, THE CUSTOMER OPENED THE PACKAGE AND PRE-FLUSHED THE CATHETER BY DRAWING SALINE WITH A 10 ML SYRINGE TO CONFIRM THE INTEGRITY OF THE CATHETER. BUT DURING PRE-FLUSHING, LIQUIDS LEAKED FROM THE PROXIMAL END OF THE CATHETER. THEN SOME OF THE GUIDE WIRE WAS REMOVED AND ONE FINE HOLE WAS NOTED AT THE PROXIMAL END OF THE CATHETER. WATER SPRAYING OCCURRED DURING PRE-FLUSHING. THE CATHETER WAS IMMEDIATELY REPLACED AND NO IMPACT WAS CAUSED TO THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04733 | 2023-06-27 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT FOLEY CATHETER STOPPED WORKING AFTER 7 DAYS TUBE CANNOT FLUSH. PER FOLLOW UP VIA EMAIL ON (B)(6)2023, THE CUSTOMER STATED THAT THEY WERE STABLE FOR NOW, BUT VERY CONCERNED ABOUT SEVERAL CATHETERS NOT WORKING PROPERLY. THIS WAS NOT THE ONLY ONE TO FAIL. SEVERAL FAILED THIS YEAR, AND SINCE THEIR INSURANCE WILL ONLY PAY FOR ONE A MONTH, SEVERAL WERE PAID OUT OF POCKET. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02661 | 2023-06-27 | 2023-06-14 | I, n, j, u, r, y | Manufacturer report | BARD MIDLINE | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT THE CUSTOMER HAS A PIECE OF THE WIRE AND CATHETER THAT WAS REMOVED FROM THE PATIENT THAT NEEDS TO BE SENT BACK. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02660 | 2023-06-27 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THE PICC CATHETER HAS BEEN FUNCTIONING WELL RIGHT UP UNTIL THE END DATE WHEN THE NURSE SUDDENLY NOTICED A LEAK AT THE CONNECTION ON THE OUTER CATHETER PART AND WING. ADMITTED 22.12.21 RIGHT BASILICA WITHOUT COMPLICATIONS DURING ADMISSION. END 23.06.07 DUE TO LEAKAGE AT THE WING. RISK OF PATIENT SKIN DAMAGE AND RISK FOR STAFF WITH LEAKAGE OF MEDICATION. ONGOING TREATMENT WITH PEMBROLIZUMAB + PEMETREXED EVERY THREE WEEKS. CATHETER SECURED WITH STATLOCK. CATHETER WAS REMOVED, CATHETER NOT REPLACED. EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. THE PATIENT ARE OK AND DIDN´T GET ANY HARM, INJURY OR NEGATIVE OUTCOME MORE THEN THE NEED OF PULLING THE PICCLINE AWAY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02662 | 2023-06-27 | 2023-06-14 | I, n, j, u, r, y | Manufacturer report | BARD MIDLINE | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED DURING CLINICAL ASSESSMENT "REDNESS OR SWELLING AT OR AROUND THE MIDLINE SITE." DEVICE INVOLVED IS A POWER MIDLINE IN THE UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02663 | 2023-06-27 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC SV 4F DL BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 |  | 3274118J |  | IT WAS REPORTED THAT THE PATIENT HAD A POWERPICC IMPLANTED ON (B)(6) 2023. ON JUNE 2, 2023, FLUSH COULD NOT BE PERFORMED. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02651 | 2023-06-27 | 2023-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED ON (B)(6) 2022, THE PATIENT WAS ADMITTED TO HOSPITAL FOR BREAST CANCER, AND A PICC CATHETER WAS INSERTED. ON THE AFTERNOON OF (B)(6) 2023, THE NURSE FOUND CONTINUOUS FLUID LEAKAGE AT THE PUNCTURE POINT DURING MAINTENANCE OF THE PICC CATHETER, WHICH COULD NOT BE USED NORMALLY. IMMEDIATELY, THE CATHETER WAS PULLED OUT, AND A CREASE WAS FOUND 2CM AWAY FROM THE PUNCTURE POINT, AND A BROKEN MOUTH WAS FOUND 4CM AWAY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04730 | 2023-06-27 | 2023-06-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD BEEN HAVING ON AND OFF URINARY TRACT INFECTION AND WAS UNSURE OF THEY ARE FROM THE MAGIC3 GO INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. THE PATIENT REPORTED ISSUES WITH LOT#: JUGT3350, LOT#: JUGQ0914 AND LOT#: JUGQ2913. PER FOLLOW UP VIA PHONE ON 15JUN2023, IT WAS REPORTED THAT THE CUSTOMER DEVELOPED UTI AND SOUGHT MEDICAL INTERVENTION. CUSTOMER WAS PRESCRIBED ANTIBIOTICS AND LIBERATOR REPLACED CATHETERS OF DIFFERENT TYPE WHICH HAVE CAUSED NO ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04715 | 2023-06-27 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123618A | 123618A |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WOULD NOT DEFLATE. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE CATHETER WAS CUT OFF. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON BURST. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE PATIENT USE ONLY. DO NOT REUSE AND RESTERILIZE. FOR UROLOGICAL USE ONLY. USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PROTOCOL, THE MAY BE CUT OFF. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04723 | 2023-06-27 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT MANY OF THE NURSES BELIEVE THAT THE TEMPERATURE SENSING FOLEYS WERE NOT ACCURATE IN PATIENTS WHO WERE UNABLE TO URINATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02622 | 2023-06-27 | 2022-12-09 | I, n, j, u, r, y | Manufacturer report | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0602680 |  | IT WAS REPORTED THAT TEN MONTHS AND THIRTY DAYS POST A PORT PLACEMENT, THE PORT WAS ALLEGEDLY INACCESSIBLE AND THE PORT CATHETER WAS ALLEGEDLY FRACTURED AT THE FIRST RIB PINCH POINT CONSISTENT WITH PINCH OFF SYNDROME. IT WAS FURTHER REPORTED THAT A PIECE OF THE CATHETER MOVED INTO THE RIGHT ATRIUM AND THE RIGHT VENTRICLE. REPORTEDLY, THE CATHETER FRAGMENTS WERE REMOVED BY THE SNARE IN ITS ENTIRETY AND THE PORT WAS REMOVED AND REPLACED. THE PATIENT'S CURRENT STATUS IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. MEDICAL RECORDS WERE PROVIDED AND REVIEWED. APPROXIMATELY TEN MONTHS AND THIRTY DAYS POST PORT PLACEMENT, PATIENT PRESENTED FOR CHEST PORT DYSFUNCTION AND UNDERWENT TRANSCATHETER FOREIGN BODY RETRIEVAL, IMPLANTED CHEST PORT REMOVAL AND PLACEMENT OF NEW CHEST PORT. INITIAL FLUOROSCOPIC IMAGES SHOW FRACTURE OF THE CHEST PORT CATHETER AT THE FIRST RIB PINCH POINT CONSISTENT WITH PINCH OFF SYNDROME. THE DISTAL ASPECT OF THE PORT CATHETER WAS SEEN TO BE FLOATING IN THE RIGHT ATRIUM AND RIGHT VENTRICLE. THROUGH THE SELDINGER TECHNIQUE AN 8 FRENCH SIDEARM SHEATH WAS PLACED TO THE SUPERIOR VENA CAVA. SUBTRACTION VENOGRAMS WERE PERFORMED SHOWING NO EVIDENCE OF THROMBUS IN THE SVC, RIGHT ATRIUM AND CHEST PORT CATHETER EXTENDING FROM THE RIGHT ATRIUM INTO THE RIGHT VENTRICLE, WITH CATHETER TIP NEAR THE APEX OF THE RIGHT VENTRICLE. A SNARE WAS ADVANCED, AND THE CATHETER WAS REMOVED THROUGH THE SHEATH. THE FRACTURED PORTION OF THE CATHETER WAS REMOVED IN ITS ENTIRETY. NO ADDITIONAL CATHETER FRAGMENTS WERE IDENTIFIED IN THE RIGHT HEART BY FLUOROSCOPY. CHEST PORT REMOVAL OF THE EXISTING DYSFUNCTIONAL PORT WAS PERFORMED. FINALLY, NEW CHEST PORT WAS PLACED SUCCESSFULLY. NEW PORT FUNCTIONING WELL AND READY FOR USE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MIGRATION, MATERIAL SEPARATION AND OBSTRUCTION OF FLOW. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED BLOCKED CONNECTION AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. BASED UPON AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02638 | 2023-06-27 | 2022-09-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833350 | 5833350 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE INTRODUCER NEEDLE WAS ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT THE INTRODUCER NEEDLE ALLEGEDLY LEAKED AIR. FURTHERMORE, SALINE ALLEGEDLY LEAKED IN THE PINK JOINT OF THE INTRODUCER NEEDLE WHEN INJECTED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2023) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02641 | 2023-06-27 | 2022-12-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN PORT SINGLE LUMEN | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0603880C |  | IT WAS REPORTED THAT DURING THE PREPARATION OF A PORT PLACEMENT PROCEDURE IN RIGHT SIDE OF SIXTH THORACIC VERTEBRA VIA RIGHT INTERNAL JUGULAR VEIN, THE PEEL-AWAY SHEATH WAS INSERTED AFTER THE GUIDE WIRE WAS PASSED THROUGH THE RIGHT INTERNAL JUGULAR VEIN. IT WAS FURTHER REPORTED THAT RESISTANCE WAS ALLEGEDLY FELT WHEN DELIVERING THE SHEATH OVER THE GUIDEWIRE AND ONE BULGE AND SIGNS OF CRACK WERE ALLEGEDLY OBSERVED AT THE TIP. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 12/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02642 | 2023-06-27 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602680 | 0602680 |  | IT WAS REPORTED THAT DURING PREPARATION OF A PORT PLACEMENT PROCEDURE, A HAIR WAS ALLEGEDLY FOUND IN THE PACKAGE UPON OPENING. IT WAS FURTHER REPORTED THAT THERE WAS A HOLE ALL THE WAY THROUGH THE INNER PAPER PACKAGE OF THE DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 09/2026), G3. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04698 | 2023-06-27 | 2023-06-12 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0102L16 | 0102L16 |  | IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED WITH URINARY TRACT INFECTION DUE TO FOLEY CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04699 | 2023-06-27 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE TEMP SENSING FOLEY CATHETER | BARDEX® LUBRI-SIL® ALL-SILICONE TEMP SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119216M | 119216M |  | IT WAS REPORTED THAT THE TEMPERATURE SENSING FOLEY WAS NOT REGISTERING ANY TEMPERATURE. THERE WERE TWO NON-WORKING TEMPERATURE FOLEYS (MATERIAL# 119416M) (MATERIAL# 119216M) WHEN THEY WERE ON-SITE. THE FOLEY WAS STILL IN PATIENT AND REQUESTED THAT IT BE KEPT TO BE RETURNED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00457 | 2023-06-27 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U3513076 | U3513076 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE VIA THE BRACHIAL ACCESS, THE CATHETER SHAFT ALLEGEDLY HAD A BREAK. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE ULTRAVERSE 035 PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, NO KINKS WERE NOTED TO THE CATHETER AND NO ANOMALIES NOTED TO THE Y-BODY/STRAIN RELIEF. DURING FUNCTIONAL EVALUATION, AN IN-HOUSE PRESTO INFLATION DEVICE WAS USED TO INFLATE THE BALLOON, AND A PINHOLE RUPTURE WAS NOTED TO THE PROXIMAL END OF THE BALLOON. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. ONE VIDEO WAS PROVIDED AND REVIEWED. THE VIDEO SHOWS THE ULTRAVERSE 035 DEVICE HELD BY A HEALTH PROFESSIONAL. THE BALLOON IS SEEN BEING INFLATED, BUT LIQUID IS SEEN STREAMING FROM THE PROXIMAL END OF THE BALLOON. THE SOURCE OF THE LEAK CANNOT NOT BE DETERMINED. NO OTHER ANOMALIES NOTED. THEREFORE, THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED BREAK AS THERE WAS NO BREAK NOTED ON THE DEVICE. HOWEVER, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED PINHOLE BALLOON RUPTURE AS THERE WAS A PINHOLE RUPTURE WAS NOTED TO THE PROXIMAL END OF THE BALLOON DURING INFLATION OF THE RETURNED DEVICE. HENCE, THE SOURCE OF THE LEAK SEEN IN THE PROVIDED VIDEO IS DETERMINED TO BE FROM THE PINHOLE RUPTURE. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED BREAK AND IDENTIFIED PINHOLE BALLOON RUPTURE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 03/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04725 | 2023-06-27 | 2023-06-09 | I, n, j, u, r, y | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 996101 | 996101 |  | IT WAS REPORTED, THAT IN AN ONLINE SURVEY. A PHYSICIAN STATED, A PERFORATION ADVERSE EVENT, DIRECTLY ATTRIBUTABLE TO THE DEVICE 996101 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 10MM X 15CM, W/ INFLATION DEVICE, STERILE. ALSO STATING, THAT THE PATIENT WITH THE DEVICE 996101 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 10MM X 15CM, W/ INFLATION DEVICE, STERILE PLACEMENT DID RESULT IN PATIENT INJURY, REQUIRING MEDICAL OR SURGICAL INTERVENTION, BECAUSE SURGICAL REPAIR. MEDICAL INTERVENTION WAS UNKNOWN. PER FOLLOW UP INFORMATION RECEIVED VIA EMAIL ON 14 JUN 2023, IN AN ONLINE SURVEY, A PHYSICIAN STATED, A PERFORATION ADVERSE EVENT, DIRECTLY ATTRIBUTABLE TO THE DEVICE 996101 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 10MM X 15CM, W/ INFLATION DEVICE, STERILE. PHYSICIAN ALSO INDICATED, PERFORATION WAS PRESENT PRIOR TO PLACEMENT OF THE DEVICE. ALSO STATING, THAT THE PATIENT WITH THE DEVICE 996101 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 10MM X 15CM, W/ INFLATION DEVICE, STERILE PLACEMENT DID RESULT IN PATIENT INJURY REQUIRING MEDICAL OR SURGICAL INTERVENTION, BECAUSE SURGICAL REPAIR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04707 | 2023-06-27 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI20 |  | IT WAS REPORTED THAT THE DIRT WAS FOUND ON THE TIP OF THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04708 | 2023-06-27 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316 | 119316 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DUE TO A BALLOON RUPTURE. ON THE NEXT DAY AFTER SURGERY, UPON REMOVING THE TAPE FOR FIXING THE CATHETER AND FOLLOWING THE CATHETER TUBE TO CLEAN UP THE INSERTION SITE, IT WAS FOUND THAT THE CATHETER HAD FALLEN OUT TO THE TIP. WHEN THE BALLOON WAS CHECKED, IT WAS RUPTURED. AS PER THE NOTIFICATION RECEIVED ON (B)(6) 2023, STATED THAT THERE WAS NO BALLOON RUPTURE OBSERVED ON THE RETURNED SAMPLE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 16JUN2023, IT WAS STATED THAT THE BALLOON RUPTURE WAS NOT OBSERVED TO THE RETURNED SAMPLE. A PINHOLE WAS FOUND ON THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04704 | 2023-06-27 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER FROM THE BEGINNING, THERE WAS A HOLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04719 | 2023-06-27 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER COUDE TIP | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0100L16 | 0100L16 |  | IT WAS REPORTED THAT URINE WAS POURING FROM THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04710 | 2023-06-27 | 2023-06-11 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE THESE MAGIC 3 INTERMITTENT CATHETER FELT LIKE THEY CUT THEM AND THE PATIENT NOW HAD A URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04697 | 2023-06-27 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED URETHRAL CATHETER KIT | URETHRAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE PATIENT STATED INTERMITTENT CATHETER WERE CAUSING PAIN AND BLEEDING. PER FOLLOW UP VIA PHONE ON (B)(6)2023, PATIENT HAD TO GO TO THE DOCTOR AND WAS WAITING TO SEE HIS UROLOGIST. PATIENT WAS NOT ABLE TO FULLY INSERT NOR VOID WITH THE CATHETERS. PATIENT HAD SAMPLES TO RETURN AND WILL SEND 3 UNUSED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04728 | 2023-06-27 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER GOT DEFORMED. THE LUMEN OF THE CATHETER BECAME AS IF IT HAD GOT GLUED AND URINE WAS UNABLE TO DRAIN. ALTHOUGH THERE WERE SOME TRACES OF URINATION, THE SYMPTOMS OF THE ACTUAL PRODUCT OCCURRED AT AN EARLY STAGE AFTER INDWELLING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04731 | 2023-06-27 | 2023-06-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED WITH BLEEDING AND A URINARY TRACT INFECTION WHILE USING THE MAGIC 3 GO MALE COUDE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. PER CUSTOMER VIA PHONE ON 19JUNE2023, THEIR ORIGINAL CATHETER WAS DISCONTINUED AND THEY WAS SENT 50814G AS REPLACEMENTS. HE STATED THAT THE PATIENT WAS HOSPITALIZE AND HAD TO HAVE THEM PROSTATE CAUTERIZED DUE TO BLEEDING AFTER INSERTION AND PUNCTURING THEIR PROSTATE. CUSTOMER STATED THAT THEY BLEEDING EVERY TIME THEY TRIED TO INSERT THE CATHETERS. THEY STATED THAT THIS ISSUE HAD DECREASED WITH TREATMENT FROM THEIR UROLOGIST OVER THE PAST MONTH AND THE PATIENT WAS ABLE TO FULLY INSERT AND VOID WITH THE CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04702 | 2023-06-27 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE THESE MAGIC 3 INTERMITTENT CATHETERS DID NOT HAVE ENOUGH LUBRICANT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04703 | 2023-06-27 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175814 |  | IT WAS REPORTED THAT THE PATIENT PULLED ON THE CATHETER AND IT SNAPPED IN HALF LEAVING HALF OF THE CATHETER IN SITU. THE NURSES MANAGED TO REMOVE THE REMAINING PART OF THE CATHETER. THE PATIENT WAS NOT HURT. NO MEDICAL INTERVENTION WAS REQUIRED. AS PER THE FOLLOW UP INFORMATION RECEIVED VIA IBC ON 13JUN2023, SNAPPED DESCRIBED AS "THE ARROW POINTS TO THE PART OF THE CATHETER WHERE THE PORT TO INFLATE THE BALLOON AND THE LEG BAG CONNECTION PORT, IS ATTACHED TO THE MAIN SHAFT OF THE INDWELLING PART OF THE CATHETER. IT IS AT THIS POINT THE CATHETER SEPARATED. IF THE BALLOON HAD NOT EMPTIED ITSELF THE USER WOULD NOT HAVE BEEN ABLE TO REMOVE THE CATHETER AS THEY WOULD NOT HAVE BEEN ABLE TO DEFLATE THE BALLOON BECAUSE THE LEG BAG PORT AND THE INFLATION PORT HAD DETACHED FROM REST OF THE CATHETER.". | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02654 | 2023-06-27 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REPLACEMENT CONNECTOR FOR 4F GROSHONG NXT CLEARVUE PICC W/ EXTENSION LEG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7812400 |  | IT WAS REPORTED THAT WHEN CONNECTING THE EXTENSION TUBING KIT FOR GROSHONG PICC PUNCTURE, IT COULD NOT BE ENGAGED PROPERLY WITH THE CATHETER. A NEW KIT WAS USED AND CONNECTED SUCCESSFULLY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02640 | 2023-06-27 | 2018-01-18 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1608062 |  | IT WAS REPORTED THAT SIX MONTHS AND THIRTEEN DAYS POST PORT PLACEMENT VIA THE RIGHT JUGULAR VEIN WITH THE CATHETER INTO THE MID RIGHT ATRIUM, THE PORT WAS ALLEGEDLY INFECTED. IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED IRRITATION AND REDNESS SURROUNDING THE INSERTION AREA. FURTHERMORE, PATIENT UNDERWENT ANESTHESIA AND THE PORT REMOVAL PROCEDURE DUE TO INFECTION. REPORTEDLY, THE PORT WAS REMOVED. THE PATIENT'S CURRENT STATUS WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2018). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02664 | 2023-06-27 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (18G) (10CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F118100T |  | IT WAS REPORTED BY CUSTOMER THAT WHILE PRACTICING ON THE VEIN BLOCK IN DIDACTIC CLASS, NURSE THREAD THE CATHETER AND WHEN SHE WENT TO PULL BACK THE DEVICE AND SAFETY THE NEEDLE, THE SAFETY PORTION STAYED ATTACHED TO THE DEVICE AND THE NEEDLE DID NOT SAFETY. ADDITIONAL INFORMATION RECEIVED 06/18/2023: IT WAS REPORTED "DEVICES WERE NOT USED ON A PATIENT, THIS HAPPENED ON THE VEIN BLOCK ON DIDACTIC CLASS." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DEFECTIVE NEEDLE SAFETY MECHANISM IS CONFIRMED AND WAS DETERMINED TO BE RELATED TO MANUFACTURING. ONE 18 GA POWERGLIDE PRO WAS RETURNED FOR EVALUATION. AFTER FUNCTIONAL TESTING AND OBSERVATIONS OF THE RETURNED POWERGLIDE PRO, IT WAS OBSERVED THAT THE SAFETY MECHANISM OF THE DEVICE WAS BROKEN. THE METAL COMPONENT INSIDE THE SAFETY MECHANISM APPEARED SHARP AND SLIGHTLY GROUND. THE RED RING INSIDE THE SAFETY MECHANISM WAS BROKEN, AND THE CAUSE WAS TRACED TO MANUFACTURING. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. THE INVESTIGATION WAS FORWARDED TO THE MANUFACTURING FACILITY FOR FURTHER EVALUATION, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02665 | 2023-06-28 | 2022-09-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT APPROXIMATELY TWO MONTHS POST DIALYSIS CATHETER PLACEMENT, BLOOD EXUDATION ALLEGEDLY OCCURRED AT THE PORTION OF THE CATHETER WHERE THE FIXED WINGS WERE ATTACHED. IT WAS FURTHER REPORTED THAT THERE WAS BREAKAGE AT THE TRANSPARENT CONNECTING TUBE OUTSIDE THE BODY. REPORTEDLY, THE CATHETER WAS REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 03/2022). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04755 | 2023-06-28 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI12 |  | IT WAS REPORTED THAT THERE WAS DIFFICULTY DRAINING WATER DURING PRE-TEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02639 | 2023-06-28 | 2022-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833350 | 5833350 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, INTRODUCER NEEDLE ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT INTRODUCER NEEDLE ALLEGEDLY LEAKED AIR. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04756 | 2023-06-28 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER CREEVY MODEL | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 01267520 | 01267520 |  | IT WAS REPORTED THAT THE BALLOON LEAKED AND THE FOLEY CATHETER FELL OUT OF THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02671 | 2023-06-28 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT WHEN FLUSHING THE CATHETER UPON UNPACKING, THE CUSTOMER SAW A FOREIGN MATTER LOOKING LIKE A FINE HAIR INSIDE THE CATHETER. THE CUSTOMER IMMEDIATELY REPLACED THE CATHETER AND PLACED A NEW ONE IN THE PATIENT¿S BODY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04757 | 2023-06-28 | 2023-06-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD BEEN HAVING ON AND OFF URINARY TRACT INFECTION AND WAS UNSURE OF THEY ARE FROM THE MAGIC3 GO INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. THE PATIENT REPORTED ISSUES WITH LOT# JUGT3350, LOT# JUGQ0914 AND LOT# JUGQ2913. PER FOLLOW UP VIA PHONE ON 15JUN2023, IT WAS REPORTED THAT THE CUSTOMER DEVELOPED UTI AND SOUGHT MEDICAL INTERVENTION. CUSTOMER WAS PRESCRIBED ANTIBIOTICS AND LIBERATOR REPLACED CATHETERS OF DIFFERENT TYPE WHICH HAVE CAUSED NO ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
1018233-2023-04758 | 2023-06-28 | 2023-06-06 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE CATHETER WAS BROKEN. THE HEAD NURSE OF THE WARD STATED THAT THE PATIENT PULLED THE CATHETER AND THAT WAS THE FIRST TIME THE CATHETER BROKE. SINCE THE EVENT WAS A PATIENT DEPRIVED PHENOMENON, THE SALES REPRESENTATIVE EXPLAINED THE HANDLING PRECAUTIONS AND ASKED FOR UNDERSTANDING BUT THE CUSTOMER REQUESTED TO RECOVER THE SAMPLE AND INVESTIGATE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, IT WAS STATED THAT THE A BROKEN PARTS HAD LEFT INSIDE THE PATIENT'S BODY. HOWEVER, AFTER THE FOLEY CATHETER IS TORN OUT, THE REMAINING PORTION OF THE BROKEN CATHETER IN THE BLADDER WAS REMOVED BY A UROLOGIST USING A CYSTOSCOPE. THE BLADDER AND URETHRA WERE JUDGED TO BE UNDAMAGED SINCE THERE WAS NO HEMATURIA, AND THE PATIENT'S CONDITION DID NOT DETERIORATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02672 | 2023-06-28 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 8194118 |  | CUSTOMER REPORTED THAT WHEN PRE-FLUSHING THE POWER PICC SOLO CATHETER DURING PLACEMENT ON (B)(6) 2023, BLACK MATTERS WERE FLUSHING OUT OF THE CATHETER. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04762 | 2023-06-28 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE INFLATION VALVE DURING STERILE WATER INJECTION AND THE THERMISTOR WIRE WAS CRIMPED INSIDE THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04763 | 2023-06-28 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT CUSTOMER HAD AN ACCOUNT WITH COMPLAINTS ABOUT THE MRI SAFE PLASTIC FOLEY CLAMP HAVING DIFFICULTIES CLAMPING AND URINE SPILLED OUT. LASTLY, THIS SAME ACCOUNT SWITCHED FROM OUR SILICONE SURESTEP FOLEY TRAY TO STANDARD LATEX. THEY WERE HAVING DIFFICULTIES WITH THE BLUE SLEEVE ON THE CATHETER. REPRESENTATIVE ADVISED IN REGARD TO THE MRI SAFE PLASTIC CLAMP ON THE FOLEY DRAINAGE BAG, THE VIDEO INSTRUCTIONS ON HOW TO OPEN & CLOSE THE CLAMP WAS THE MOST SUFFICIENT RESOURCE. AS FAR AS THE BLUE SLEEVE IN THE SURESTEP TRAY, IT WAS THERE TO PROTECT THE FOLEY CATHETER DURING THE MANUFACTURING PROCESS AND WAS AN ADDITIONAL STERILE MATERIAL WITHIN THE KIT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INSUFFICIENT SEALING SPACE OF POLYSLEEVE" OR USER MISHANDLING. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04764 | 2023-06-28 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE 4 -5 MAGIC 3 GO MALE HYDROPHILIC INTERMITTENT CATHETER DID NOT HAVE HYDROPHILIC COATING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04765 | 2023-06-28 | 2023-06-12 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD BEEN HAVING ON AND OFF URINARY TRACT INFECTION AND WAS UNSURE OF THEY ARE FROM THE MAGIC3 GO INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. THE PATIENT REPORTED ISSUES WITH LOT# JUGT3350, LOT# JUGQ0914 AND LOT# JUGQ2913. PER FOLLOW UP VIA PHONE ON (B)(6)2023, IT WAS REPORTED THAT THE CUSTOMER DEVELOPED UTI AND SOUGHT MEDICAL INTERVENTION. CUSTOMER WAS PRESCRIBED ANTIBIOTICS AND LIBERATOR REPLACED CATHETERS OF DIFFERENT TYPE WHICH HAVE CAUSED NO ISSUES. . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02673 | 2023-06-28 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1295108D |  | IT WAS REPORTED BY CUSTOMER THAT PICC HUB IS DESTROYED. ADDITIONAL INFORMATION RECEIVED 06/14/2023. THE PATIENT WAS ALREADY VERY ANXIOUS ABOUT HAVING THE LINE PLACED IN THE FIRST PLACE. IT IS UNFORTUNATE THAT SHE HAD TO UNDERGO THE PROCEDURE AGAIN. MEDICATION DELIVERY WAS DELAYED DUE TO THE FACT THAT THE DAMAGED LINE HAD TO BE REMOVED AND A NEW LINE REPLACED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02674 | 2023-06-28 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE. THEREFORE, THE CATHETER WAS CUT AND RECONNECTED, THE PROCEDURE WAS COMPLETED. THERE WAS NO REPORTED PATIENT INJURY. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02675 | 2023-06-28 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE. THEREFORE, THE CATHETER WAS CUT AND RECONNECTED, THE PROCEDURE WAS COMPLETED. THERE WAS NO REPORTED PATIENT INJURY. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02670 | 2023-06-28 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED DURING THE CATHETER OUTPATIENT MAINTENANCE, IT WAS FOUND THAT THE CATHETER WAS PARTIALLY DETACHED, AFFECTING THE USE OF THE CATHETER AND THE POSSIBILITY OF FLUID OVERFLOW. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02677 | 2023-06-28 | 2023-06-14 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 5F BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1295118 |  | IT WAS REPORTED THE PICC LEAKED, AS IT HAD A HOLE IN THE INVERTED CONE REGION. IT WAS NECESSARY TO REMOVE THE CATHETER FROM THE PATIENT AND PERFORM ANOTHER INSERTION PROCEDURE. 06/21/2023: ADD INFO RECEIVED. ON THE NIGHT SHIFT ON (B)(6) 2023 AT SIX O'CLOCK IN THE MORNING AT THE SHIFT CHANGEOVER, THE NURSE OBSERVED THE PATIENT VERY AGITATED DESPITE RECEIVING SEDATION IN CONTINUOUS INFUSION PUMP. WHEN THE NURSE WAS ON DUTY AND WENT TO ASSESS THE PATIENT, SHE OBSERVED THE WET SHEETS, AND THE PATIENT WAS VERY AGITATED. WHEN PERFORMING THE SAS (AGITATION AND SEDATION SCALE), SHE OBSERVED MEDICATION LEAKAGE WETTING THE SHEET. THE FRACTURE WAS EXTERNAL IN THE MIDDLE OF THE DORMONID AND FENTANYL AND THE NORADRENALINE WAS IMMEDIATELY TURNED OFF. PATIENT EXPERIENCED PSYCHOMOTOR AGITATION BECAUSE HE WAS UNDER OROTRACHEAL INTUBATION AND SEDATED. TREATMENT WAS PERFORMED IMMEDIATELY DUE TO THE RISK OF DRUG EXTRAVASATION AND THE LACK OF DELIVERY OF THE DRUG THERAPY OFFERED, WE DECIDED TO REMOVE THE CATHETER AND PERFORM A NEW PASSAGE. REPLACEMENT PROCEDURE WAS PERFORMED AT THE BEDSIDE BY THE INFUSION THERAPY TEAM NURSE. TEGADERM CHG 3M WAS USED TO SECURE THE DEVICE. THE PATIENT HAS AN ULNAR FRACTURE WITH SURGICAL REAPPROACH THAT INACTIVATES THE IMPLANTATION OF A NEW CATHETER IN THE UPPER RIGHT LIMB, SO A NEW PUNCTURE WAS PERFORMED IN THE SAME MEMBER OF THE UPPER LEFT LIMB ADVERSE EVENT, EXPOSING THE CRITICAL PATIENT TO A NEW PROCEDURE. NO OTHER INFORMATION WAS PROVIDED. | H11: A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02669 | 2023-06-28 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT TWELVE MONTHS POST PORT PLACEMENT, THE LINE WAS ALLEGEDLY KINKED AND THE PORT HAD BEEN WASHED BUT CONTAMINATED WITH BLOOD. IT WAS FURTHER REPORTED THAT THE DEVICE ALLEGEDLY HAD A LEAK AT THE PORT SITE. REPORTEDLY, THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02668 | 2023-06-28 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY BROKEN. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04769 | 2023-06-28 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE FOLEY CATHETER BALLOON. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE INFORMATION ABOUT WHETHER DEVICE WAS USED OR NOT WAS NOT PROVIDED FROM CUSTOMER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04772 | 2023-06-28 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63616 |  | IT WAS REPORTED THAT THE PATIENT HAD AN INDWELLING CATHETER FOR 3 WEEKS AS THE PATIENT HAD A FALSE PASSAGE WITH THEIR INTERMITTENT SELF-CATHETERIZATION CAUSING TRAUMA. THE PATIENT WAS USING 63616M HYDROSIL CATHETER AND THEY ADVISED THAT THEY DID NOT FEEL THE LUBRICANT ENTERING THE URETHRA AND IT SCRAPED OFF WHILE THE CATHETER WAS INSERTED. THAT CAUSED A TIGHT FEELING AND, SINCE THE PATIENT HAD A STRICTURE DUE TO PROSTATE CANCER, IT CAUSED TRAUMA. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 15JUN2023, IT WAS STATED THAT THE PATIENTS TREATMENT WAS TO HAVE AN INDWELLING CATHETER FOR 3 WEEKS TO HEAL THE FALSE PASSAGE. PATIENT HAD TRAINING ON HOW TO INSERT CATHETERS SOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04775 | 2023-06-28 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE CUSTOMER USED FOLEY CATHETER. AT HOME, THE HOME NURSE WOULD INFLATE BALLOON WITH 10CC OF WATER. THE HOSPITAL FACILITY INFLATED BALLOON WITH 5CC OF WATER WHICH CAUSED THE CATHETER TO LEAK. PATIENT STATED LTC CHANGED CATHETER TO A 30CC BUT IT CAUSED PAIN. PATIENT WANTED TO KNOW IF THERE DIRECTIONS FOR INFLATING CATHETERS THAT CAN BE SHARED WITH THE LTC. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-06
1018233-2023-04776 | 2023-06-28 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT FOLEY CATHETER HAS AN ELEVATION AT THE BALLOON. THIS MAKES IT NEAR IMPOSSIBLE TO PLACE AS A SUPRAPUBIC CATHETER. THIS MEANS THAT THEY WERE HAVING TO USE A SMALLER DIAMETER CATHETER, I.E., 14FR, TO MAINTAIN THE TRACT OF THE SUPRAPUBIC CATHETER. THESE WORK TO HOLD THE TRACT OPEN BUT DECREASES EFFICACY OF DRAINING. THIS LEAVES THE PATIENT MORE PRONE TO OBSTRUCTIONS. CUSTOMER ASKED WAS THERE ANY OTHER ALTERNATIVES FOR 16FR STRAIGHT CATHETERS THAT THEY CAN USE. PER FOLLOW UP VIA MAIL RECEIVED ON (B)(6)2023, THE ISSUES CUSTOMER HAD BEEN WITH THE NON-LATEX CATHETERS. ON (B)(6) PATIENT WHO HAD A 16FR SUPRAPUBIC CATHETER SINCE (B)(6) WAS UNABLE TO HAVE 16FR REPLACED, HAD TO DOWNSIZE TO 14FR. ON (B)(6) PATIENT BEGAN WITH UNC UROLOGY (B)(6)22 WITH SPT IN PLACE. ATTEMPTED TO PLACE 18FR CATHETER (SAME SIZE SINCE (B)(6)), UNSUCCESSFULLY. UNABLE TO PLACE 16FR, PROVIDER PAGED. BY THE TIME OF PROVIDER'S ARRIVAL, NO SPT COULD BE PLACED. PATIENT HAD TO GO THROUGH INTERVENTIONAL RADIOLOGY TO HAVE SPT REESTABLISHED. THERE WERE TWO MORE, BUT AFTER OUR INTERNAL REPORTS WERE PLACED, THE EXACT INFORMATION WAS FORGOTTEN. THE PROBLEM WAS THE RIDGE ON THE CATHETER BALLOON. IT DID NOT EASILY PASS INTO THE BLADDER. THE OTHER ISSUE WITH HAVE WITH ALL OF THE BD CATHETERS WAS THE COST. THERE WAS ONLY ONE OPTION THAT WAS LATEX FREE. THIS OPTION WAS 2-3 TIMES THE COST OF PREVIOUS CATHETERS USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04777 | 2023-06-28 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ALL SILICONE CATHETER | ALL SILICONE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT FOLEY CATHETER HAS AN ELEVATION AT THE BALLOON. THIS MAKES IT NEAR IMPOSSIBLE TO PLACE AS A SUPRAPUBIC CATHETER. THIS MEANS THAT THEY WERE HAVING TO USE A SMALLER DIAMETER CATHETER, I.E., 14FR, TO MAINTAIN THE TRACT OF THE SUPRAPUBIC CATHETER. THESE WORK TO HOLD THE TRACT OPEN BUT DECREASES EFFICACY OF DRAINING. THIS LEAVES THE PATIENT MORE PRONE TO OBSTRUCTIONS. CUSTOMER ASKED WAS THERE ANY OTHER ALTERNATIVES FOR 16FR STRAIGHT CATHETERS THAT THEY CAN USE. PER FOLLOW UP VIA MAIL RECEIVED ON 12JUN2023, THE ISSUES CUSTOMER HAD BEEN WITH THE NON-LATEX CATHETERS. ON 5/26 PATIENT WHO HAD A 16FR SUPRAPUBIC CATHETER SINCE 12/21 WAS UNABLE TO HAVE 16FR REPLACED, HAD TO DOWNSIZE TO 14FR. ON (B)(6) PATIENT BEGAN WITH UNC UROLOGY (B)(6)2022 WITH SPT IN PLACE. ATTEMPTED TO PLACE 18FR CATHETER (SAME SIZE SINCE 10/22), UNSUCCESSFULLY. UNABLE TO PLACE 16FR, PROVIDER PAGED. BY THE TIME OF PROVIDER'S ARRIVAL, NO SPT COULD BE PLACED. PATIENT HAD TO GO THROUGH INTERVENTIONAL RADIOLOGY TO HAVE SPT REESTABLISHED. THERE WERE TWO MORE, BUT AFTER OUR INTERNAL REPORTS WERE PLACED, THE EXACT INFORMATION WAS FORGOTTEN. THE PROBLEM WAS THE RIDGE ON THE CATHETER BALLOON. IT DID NOT EASILY PASS INTO THE BLADDER. THE OTHER ISSUE WITH HAVE WITH ALL OF THE BD CATHETERS WAS THE COST. THERE WAS ONLY ONE OPTION THAT WAS LATEX FREE. THIS OPTION WAS 2-3 TIMES THE COST OF PREVIOUS CATHETERS USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00459 | 2023-06-28 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CQ7562J |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON WAS ALLEGEDLY STUCK WITH THE TIP OF THE STENT. IT WAS FURTHER REPORTED THAT THE BALLOON ALLEGEDLY FAILED TO CROSS THE LESION. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02680 | 2023-06-28 | 2023-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG (TPS) STYLET FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000375 |  | IT WAS REPORTED PATIENT ATTENDED [REDACTED] FOR PLANNED MAGNESIUM INFUSION ADMINISTERED VIA PICC LINE. WHILE FLUSHING THE LINE NOTICED THE CATHETER IS LEAKING FROM A FINE FRACTURED LINE. DISCUSSED WITH DOCTOR AND SAME REMOVED. ADDITIONAL INFORMATION PROVIDED 6/8. 4FR PICC WAS INSERTED ON (B)(6) 2023. INTERNAL LENGTH 52CM, EXTERNAL LENGTH 21CM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02681 | 2023-06-28 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV SP 4F DL MAX | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S3274108PD |  | IT WAS REPORTED BY CUSTOMER THAT LEAK NOTED TO PURPLE LUMEN BY BEDSIDE NURSE ON (B)(6)2023 AT 1857. ADDITIONAL INFORMATION RECEIVED 06/19/2023 PATIENT IS A PREVIOUSLY HEALTHY MALE WITH PANCYTOPENIA AND PERIPHERAL BLASTS ADMITTED TO THE PICU WITH CONTINUED BLEEDING AND PANCYTOPENIA, WORK UP CONSISTENT WITH ACUTE PROMYELOCYTIC LEUKEMIA. INITIATED ON CHEMOTHERAPY (B)(6) 2023 PER AAML1331 PROTOCOL. THIS EVENT DIRECTLY INVOLVED A PATIENT. THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. PATIENT REQUIRED NEW PICC LINE PLACEMENT IMMEDIATELY TO AVOID ANY INTERRUPTION IN CHEMOTHERAPY PLAN. PATIENT REQUIRED AN ADDITIONAL SEDATION EVENT FOR ANOTHER PICC LINE INSERTION. THIS RN DOES IS NOT AWARE OF ANY MEDICATION INTERRUPTION. NEW PICC LINE PLACED AS SOON AS POSSIBLE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
MW5118983 | 2023-06-28 | 2023-06-06 | I, n, j, u, r, y | Voluntary report | BARDEX ANTI-INFECTIVE FOLEY CATHETER | CATHETER, URETHRAL | C.R. BARD, INC. |  |  |  | LUER LOCK PORT BROKE OFF OF URINARY CATHETER WHILE APPLYING AN EQUASHIELD. HAD TO REMOVE URINARY CATHETER AND PLACE A NEW ONE TO ADMINISTER IVES CHEMOTHERAPY. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04766 | 2023-06-28 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ALL SILICONE CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT FOLEY CATHETER HAS AN ELEVATION AT THE BALLOON. THIS MAKES IT NEAR IMPOSSIBLE TO PLACE AS A SUPRAPUBIC CATHETER. THIS MEANS THAT THEY WERE HAVING TO USE A SMALLER DIAMETER CATHETER, I.E., 14FR, TO MAINTAIN THE TRACT OF THE SUPRAPUBIC CATHETER. THESE WORK TO HOLD THE TRACT OPEN BUT DECREASES EFFICACY OF DRAINING. THIS LEAVES THE PATIENT MORE PRONE TO OBSTRUCTIONS. CUSTOMER ASKED WAS THERE ANY OTHER ALTERNATIVES FOR 16FR STRAIGHT CATHETERS THAT THEY CAN USE. PER FOLLOW UP VIA MAIL RECEIVED ON 12JUN2023, THE ISSUES CUSTOMER HAD BEEN WITH THE NON-LATEX CATHETERS. ON (B)(6) 2023 PATIENT WHO HAD A 16FR SUPRAPUBIC CATHETER SINCE 12/21 WAS UNABLE TO HAVE 16FR REPLACED, HAD TO DOWNSIZE TO 14FR. ON (B)(6) 2023 PATIENT BEGAN WITH UNC UROLOGY (B)(6) 2022 WITH SPT IN PLACE. ATTEMPTED TO PLACE 18FR CATHETER (SAME SIZE SINCE 10/22), UNSUCCESSFULLY. UNABLE TO PLACE 16FR, PROVIDER PAGED. BY THE TIME OF PROVIDER'S ARRIVAL, NO SPT COULD BE PLACED. PATIENT HAD TO GO THROUGH INTERVENTIONAL RADIOLOGY TO HAVE SPT REESTABLISHED. THERE WERE TWO MORE, BUT AFTER OUR INTERNAL REPORTS WERE PLACED, THE EXACT INFORMATION WAS FORGOTTEN. THE PROBLEM WAS THE RIDGE ON THE CATHETER BALLOON. IT DID NOT EASILY PASS INTO THE BLADDER. THE OTHER ISSUE WITH HAVE WITH ALL OF THE BD CATHETERS WAS THE COST. THERE WAS ONLY ONE OPTION THAT WAS LATEX FREE. THIS OPTION WAS 2-3 TIMES THE COST OF PREVIOUS CATHETERS USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00458 | 2023-06-28 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA80104 | VA80104 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY BURST. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 02/2026). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04761 | 2023-06-28 | 2023-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316M | 119316M |  | IT WAS REPORTED THAT OPERATION ROOM NURSE INSERTED FOLEY CATHETER WITHOUT ISSUES ONCE PATIENT WAS PUT UNDER ANESTHESIA. DURING THE SURGICAL CASE PERFUSION AND ANESTHESIA NOTED THERE WAS NO URINE OUTPUT. ONCE OPERATION ROOM NURSE HAD ACCESS TO CHECK UNDER THE DRAPES, THE FOLEY CATHETER WAS FOUND LYING ON THE BED, NOT INSIDE THE PATIENT. FOLEY CATHETER WAS TAKEN OFF THE BED BY OPERATION ROOM NURSE AND INSPECTED. THE BALLOON WAS REINFLATED WITH 10 CC OF SALINE AND OPERATION ROOM NURSE OBSERVED THE FOLEY BALLOON SLOWLY DEFLATE OVER SERVAL MINUTES. PER FOLLOW UP VIA EMAIL ON (B)(6)2023, THE PATIENT INVOLVED WAS DISCHARGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
17216052 | 2023-06-28 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPORT ISP IMPLANTABLE PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | BARD ACCESS SYSTEMS, INC. | CP00003 | CP00003 |  | PATIENT ARRIVED FOR TREATMENT. MEDIPORT WAS ACCESSED, HOWEVER BLOOD RETURN WAS UNABLE TO BE OBTAINED. UPON FLUSHING THE CATHETER, SWELLING WAS NOTED AN INCH ABOVE THE PORT SITE. PATIENT NOTED PAIN WITH FLUSH. CHEST X-RAY REVEALED RIGHT-SIDED CENTRAL VENOUS PORT CATHETER WITH FRACTURED CATHETER TUBING. PATIENT UNDERWENT EXTRACTION OF PORT AND RETRIEVAL OF CATHETER TUBING AND PLACEMENT OF NEW PORT. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02676 | 2023-06-28 | 2023-06-01 | I, n, j, u, r, y | Manufacturer report | POWERPORT VUE, ISP, MR, 8FR, MICRO | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608052 | 1608052 |  | IT WAS REPORTED THAT APPROXIMATELY EIGHT TO TEN YEARS POST A PORT PLACEMENT, THE PORT WAS ALLEGEDLY DIFFICULT TO BE ACCESSED. IT WAS FURTHER REPORTED THAT THE PORT WAS VERY DIFFICULT TO BE ACCESSED AND SOMETIMES WITH NO ACCESS, AND USED HEPARIN WITH EVERY OTHER SYRINGE TO FLUSH. REPORTEDLY, THE PATIENT ALLEGEDLY EXPERIENCED FIBRIN SHEATH FORMATION. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02667 | 2023-06-28 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1606062 | 1606062 |  | IT WAS REPORTED THAT APPROXIMATELY TEN MONTHS POST PORT PLACEMENT, ULCERATION WAS ALLEGEDLY FOUND AROUND THE PORT. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY SUSPECTED TO BE DAMAGED. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 04/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
17216116 | 2023-06-28 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPORT ISP IMPLANTABLE PORT | PORT CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | BARD ACCESS SYSTEMS, INC. | CP00003 | CP00003 |  | PATIENT ARRIVED IN OUTPATIENT CLINIC FOR CHEMOTHERAPY INFUSION. MEDIPORT WAS ACCESSED AND FLUSHED WITHOUT DIFFICULTY, ALTHOUGH PATIENT REPORT SLIGHT PAIN WITH FLUSHING. HOWEVER, BLOOD RETURN WAS UNABLE TO BE OBTAINED. PATIENT UNDERWENT CHEST X-RAY AND VENOGRAM WHICH REVEALED RIGHT CHEST PORT CONNECTED TO A PORTION OF THE TUBING THAT EXTENDS TOWARDS THE INTERNAL JUGULAR VEIN. THE MAJORITY OF THE TUBING IS DISPLACED DISTALLY REJECTING OVER THE EXPECTED LOCATION OF THE RIGHT ATRIUM AND IVC. PATIENT UNDERWENT REMOVAL OF MEDIPORT, RETRIEVAL OF CATHETER FRAGMENT, AND INSERTION OF A NEW MEDIPORT. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02683 | 2023-06-29 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, AIR WAS ALLEGEDLY ASPIRATED FROM THE INTRODUCER NEEDLE. IT WAS FURTHER REPORTED THAT HUB OF INTRODUCE NEEDLE WAS CRACKED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE INTRODUCER NEEDLE ATTACHED TO ONE 6 ML TERUMO SYRINGE WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. MULTIPLE CRACKS WERE NOTED ON THE PROXIMAL NEEDLE HUB. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. ALTHOUGH, THERE WAS NO LEAKS WERE NOTED UPON INFUSION (AFTER MULTIPLE ATTEMPTS) AND THE NEEDLE WAS PATENT TO ASPIRATION WITHOUT ISSUE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED AIR LEAK ISSUES AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED AND THE SAMPLE EVALUATION RESULTS INDICATING NO DIFFICULTY UNDER LABORATORY CONDITIONS ARE NOT BY THEMSELVES SUFFICIENT TO CONFIRM THIS EVENT OCCURRED UNDER CLINICAL CONDITIONS. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 01/2025), G3. H11: E1, H6 (METHOD, RESULT, CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02691 | 2023-06-29 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS AND SIX MONTHS POST PORT PLACEMENT, A COMPUTED TOMOGRAPHY REVEALED CYLINDRICAL CALCIFIC DENSITY IN THE LEFT UPPER LOBE SEGMENTAL PULMONARY ARTERY, AND WAS SUSPICIOUS FOR AN EMBOLIZED CATHETER FRAGMENT. IT WAS FURTHER REPORTED THAT THE PORT WAS REMOVED AND IT WAS CONFIRMED TO BE INTACT UPON INSPECTION. REPORTEDLY, RADIOLOGICAL ANALYSIS REVEALED THAT THE RADIO DENSITY OF THE OPACITY IN THIS PATIENT'S COMPUTED TOMOGRAPHIC SCAN WAS NOT COMPATIBLE WITH HUMAN TISSUE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2020). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02701 | 2023-06-29 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT TITANIUM ISP 6CF INT WSP ATT SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8706060 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, PATIENT EXPERIENCED SWELLING UNDERNEATH THE SUBCLAVIAN AREA. IT WAS FURTHER REPORTED THAT CATHETER WAS ALLEGEDLY FOUND RUPTURED AND THE CATHETER WAS ALLEGEDLY OBSTRUCTED. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02704 | 2023-06-29 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT, THE CATHETER ALLEGEDLY OCCLUDED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04780 | 2023-06-29 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT LONG CURLY BLACK HAIR FOUND IN STERILE 18 FR FOLEY UPON OPENING PACKAGE. PER FOLLOW UP VIA EMAIL ON (B)(6) 2023, THE PATIENT INVOLVED WAS DISCHARGED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO UNCLEAN WORKING AREA. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02706 | 2023-06-29 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616070 | 1616070 |  | IT WAS REPORTED THAT SOME TIME POST PORT PLACEMENT, WHEN THE PORT SYSTEM WAS ATTEMPTED TO BE REMOVED AND THE SKIN WAS INCISED TO EXPOSE THE PORT, IT WAS FOUND THAT THE CATHETER LOCK AND THE CATHETER WERE DETACHED FROM THE PORT BODY. REPORTEDLY, THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 05/2023) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00465 | 2023-06-29 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U415074RX | U415074RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE SUPERFICIAL FEMORAL ARTERY VIA AN IPSILATERAL ANTEGRADE APPROACH, A LARGE MISALIGNMENT WAS ALLEGEDLY CONFIRMED BETWEEN THE BALLOON LENGTH AND THE RADIOPAQUE MARKER. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04782 | 2023-06-29 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
1018233-2023-04783 | 2023-06-29 | 2023-06-07 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK URINE COLLECTION SYSTEM HAD NO SUCTION. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE CUSTOMER WAS IN THE HOSPITAL ON 2 SEPARATE OCCASIONS DUE TO A URINARY TRACT INFECTION THAT PATIENT ACQUIRED WHILE USING THE PUREWICK EXTERNAL FEMALE CATHETER. THE MOST RECENT HOSPITAL VISIT WAS ON (B)(6) 2023- (B)(6) 2023, IN WHICH PATIENT WAS PRESCRIBED ANTIBIOTICS FROM THE DOCTOR FOR THE INFECTION. CUSTOMER REQUESTED COMPENSATION OR REIMBURSEMENT DUE TO THE NATURE OF THE INCIDENT AND THE FACT THAT THEY COULD NO LONGER USE THE PRODUCT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-06
3006260740-2023-02685 | 2023-06-29 | 2023-05-29 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED ¿ON (B)(6)2022, A CENTRAL VENOUS CATHETER WAS INSERTED INTO THE PATIENT. THE CATHETER VESSEL WAS THE LEFT BRACHIAL VEIN, THE CATHETER DEPTH WAS 42CM, AND THE CATHETER TIP WAS LOCATED IN THE SUPERIOR VENA CAVA. THE NURSE CAME TO OUR HOSPITAL FOR MAINTENANCE AT 8:05 AM ON (B)(6)2023, AND FOUND THAT THERE WAS LIGHT BLOODY LIQUID IN THE CONNECTION TUBE, BUT NO BLOOD WAS FOUND AFTER THE FIRST WITHDRAWAL. WHEN THE TUBE WAS SLIGHTLY FLUSHED, THE CATHETER IMMEDIATELY SLIPPED OUTSIDE AND SHOWED DISCONNECTION, THE PORT WAS FLAT, AND THE SCALE AT THE SEGMENT WAS 39.1CM, WHICH WAS 2CM AWAY FROM THE PUNCTURE POINT, AND SHE IMMEDIATELY STUCK THE CATHETER 10CM ABOVE THE PUNCTURE POINT WITH HER HAND. ANOTHER STAFF MEMBER IMMEDIATELY TOOK THE PRESSURE PULSE BAND AND TIED IT IN THE ARMPIT. ULTRASOUND EXAMINATION SHOWED THAT THERE WAS NO CATHETER SHADOW IN THE LEFT UPPER ARM VEIN, AND IMMEDIATELY CONTACTED THE DSA DOCTOR IN OUR HOSPITAL FOR FLUOROSCOPY, WHICH SHOWED THAT THE CATHETER WAS IN THE HEART, THE OUTLET OF THE PULMONARY ARTERY WAS REVERSED, AND THE END WAS IN THE SUPERIOR VENA CAVA, SUGGESTING DSA TO ARREST AND REMOVE IT. REPORT TO THE NURSING DEPARTMENT LEADERSHIP, AND IMMEDIATELY CONTACT WITH OTHER HOSPITAL, THE PATIENT AND FAMILY MEMBERS DECIDED TO GO TO OTHER HOSPITAL FOR TREATMENT. THE PATIENT WAS SENT TO THE OTHER HOSPITAL BY AMBULANCE, AND THE DEPARTMENT OF VASCULAR SURGERY OF THE OTHER HOSPITAL PERFORMED ENDOVASCULAR FOREIGN BODY REMOVAL UNDER LOCAL ANESTHESIA DSA + SUPERIOR VENA CAVA + PULMONARY ARTERY ANGIOGRAPHY. THE PATIENT WAS SUCCESSFULLY RETURNED TO THE WARD WITHOUT COMPLAINING OF DISCOMFORT. THE INCIDENT CAUSED THE PATIENT TO BE ANESTHETIZED AGAIN, CAUSING A SECONDARY INJURY." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02687 | 2023-06-29 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 22 GAUGE POWERGLIDE PRO MIDLINE CATHETER 8CM BASIC KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6F122080 |  | IT WAS REPORTED BY THE CUSTOMER THAT THEY WENT TO SLIDE THE MIDLINE IN, THE DELIVERY DEVICE SORT OF BUCKLED/FELL APART AND THE LINE HAD NOT SLID OVER THE WIRE AND SHEARED THE LINE. ADDITIONAL INFORMATION RECEIVED 6/22/2023: CUSTOMER REPORTED THAT DEVICE BUCKLING, LINE CONCERTINA¿D EXCEPT, THE LINE SHEARED OFF THE HUB AND ONLY 1-2 MM OF IT WERE OUTSIDE THE PATIENT WHICH CLINICIAN QUICKLY GRABBED WITH AN ARTERY FORCEPS, SURGEON CAME TO ASSIST BY DISSECTING AND REMOVING LINE AND WIRE. PATIENT IS FINE - MINOR BRUISE ONLY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02686 | 2023-06-29 | 2023-06-09 | I, n, j, u, r, y | Manufacturer report | 22 GAUGE POWERGLIDE PRO MIDLINE CATHETER 8CM BASIC KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6F122080 |  | IT WAS REPORTED THAT, CLINICAN INSERTED THE CANNULA INTO THE R BASILIC VEIN WHICH WAS QUITE DEEP ON THIS PARTICULAR PT (SO CANNULA AT A REASONABLE ANGLE), THE WIRE SEEMED TO SLIDE IN OK. CLINICIAN WENT TO SLIDE THE MIDLINE IN, THE DELIVERY DEVICE SORT OF BUCKLED/FELL APART AND THE LINE HAD NOT SLID OVER THE WIRE ¿ IT HAD IN FACT CONCERTINA¿D AND THEN THE LINE AND WIRE WERE STUCK IN PTS ARM. CLINICIAN GOT SURGEON TO DISSECT DOWN TOWARDS THE VESSEL UNTIL HE COULD REMOVE THEM BOTH. CLINICIAN THOUGHT THIS WAS ¿ME¿ EG ANGLE TOO GREAT, PERHAPS NEARER TO WHERE TOURNIQUET WAS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02684 | 2023-06-29 | 2023-02-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE IN RIGHT INTERNAL JUGULAR VEIN, THE ACCESS PROCESS WAS DONE AND GUIDE WIRE WAS INSERTED SUCCESSFULLY. IT WAS FURTHER REPORTED THAT THE PEEL-APART SHEATH WAS INSERTED OVER THE GUIDE WIRE SMOOTHLY AND THE RESISTANCE WAS ALLEGEDLY EXPERIENCED WHEN REMOVING THE GUIDE WIRE FROM PATIENT. REPORTEDLY, AFTER REMOVAL FROM THE VESSEL, WIREDRAWING WAS NOTED WITH THE TIP OF THE GUIDE WIRE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 02/2023). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02690 | 2023-06-29 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "I HAVE A PATIENT THAT THE BEDSIDE NURSES WERE CONCERNED THAT THE PICC WAS INFILTRATING AT THE INSERTION SITE BECAUSE THE PATIENT'S ARM IS SWOLLEN, RED, AND WARM. HOWEVER, BOTH ARMS ARE LIKE THIS. THEY WERE CONCERNED BECAUSE THE FLUID ON THE BIOPATCH WAS YELLOW AND PATIENT IS ON MINOCYCLINE (OUR PREPARATIONS ARE A NEON YELLOW COLOR). HOWEVER, I LOOKED UNDER ULTRASOUND WHILE FLUSHING AND COULDN'T SEE ANY INFILTRATION HAPPENING AND WAS ABLE TO ASPIRATE AND FLUSH ALL 3 LUMENS. I COULDN'T SEE ANY FLUID LEAKING WHEN I FLUSHED". ADDITIONAL INFORMATION (B)(6)2023: IT WAS REPORTED BY THE CUSTOMER ¿THE PICC CATHETER WASN'T INFILTRATING AT THE SITE. IT WAS A MISUNDERSTANDING BY THE BEDSIDE STAFF. THE PATIENT WAS RECEIVING MINOCYCLINE, WHICH WHEN PREPARED AT OUR HOSPITAL IS A NEON YELLOW COLOR. THE FLUID WEEPING FROM THE PATIENT WAS ALSO YELLOW BUT A SEROUS YELLOW. PATIENT WAS FLUID OVERLOADED. EDUCATION WAS PROVIDED TO BEDSIDE STAFF¿. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
MW5119034 | 2023-06-29 | 2023-06-20 | I, n, j, u, r, y | Voluntary report | ACCUCATH ACE IV CATHETER WITH COILED GUIDEWIRE | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. |  | AC1202250 |  | VASCULAR ACCESS RN NOTICED THAT THE TIP OF THE WIRE WAS MISSING AFTER INSERTING A 20G X 2.25 IN BARD ACCUCATH ACE IV WITH NITINOL GUIDEWIRE WITH COILED TIP. A STAT X-RAY IDENTIFIED THAT THE WIRE WAS STILL IN THE PATIENT'S ARM. THE GUIDEWIRE TIP WAS REMOVED THE FOLLOWING DAY UNDER FLUOROSCOPY GUIDANCE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02689 | 2023-06-29 | 2023-06-09 | I, n, j, u, r, y | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT HER DAUGHTER GOT A PICC LINE PLACED AND THAT DEVICE HAD TO BE REMOVED BECAUSE SHE DEVELOPED A BLOOD CLOT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02694 | 2023-06-29 | 2023-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 20G X 10 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F320108PT |  | IT WAS REPORTED BY THE CUSTOMER THAT THEY HAD TWO SEPARATE CATHETERS THAT HAD HOLES WHERE THE CATHETER AND THE HUB MEET. THEY HAD TO PULL THE CATHETER AND PLACE ANOTHER ONE. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. INITIAL MEDWATCH REPORT WAS SUBMITTED, UPON FURTHER REVIEW IT WAS FOUND THAT THIS MEDWATCH REPORT 3006260740-2023-02694 IS A DUPLICATE FILE AND HAS BEEN VOIDED. THE ORIGINAL EVENT WAS SUBMITTED ON MEDWATCH REPORT 3006260740-2023-02918. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02700 | 2023-06-29 | 2023-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV SP 4F DL MAX | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S3274108PD |  | IT WAS REPORTED BY CUSTOMER THAT LEAK NOTED FROM PURPLE PLASTIC HUB BY BEDSIDE NURSE (B)(6) 23. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023 PATIENT HAS HISTORY OF AML WHO PRESENTS DUE TO FATIGUE AND HEMOPTYSIS WITH LABS SIGNIFICANT FOR SEVERE THROMBOCYTOPENIA. HE RECEIVED ALL PREVIOUS CARE IN MEXICO BESIDES RECENT ADMISSION AT OSH AND FURTHER RESEARCH WILL NEED TO BE MADE INTO PRIOR TREATMENTS. S/P BRONCHOSCOPY, BIOPSY, LP, INTRATHECAL CYTARABINE. FEVER S/P PROCEDURES ON 6/9, SPIKED FEVER 6/10 AM. CONTINUES CEFEPIME AND LINEZOLID. PLAN TO REPEAT BLOOD AND URINE CULTURES AFTER 24 HOURS SINCE LAST DRAW IF PATIENT REMAINS FEBRILE. PLAN TO REPLACE PICC TOMORROW DUE TO HOLE IN ONE OF THE LINES, FUNCTIONALLY PATIENT ONLY HAS SINGLE LUMEN PICC. THIS EVENT DIRECTLY INVOLVED A PATIENT. IT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. PATIENT REQUIRED PLACEMENT OF ANOTHER PICC LINE WHICH ALWAYS RUNS THE RISK OF INTRODUCTION INFECTION. THIS RN IS NOT AWARE OF ANY MEDICATION INTERRUPTION. NEW PICC LINE WAS PLACED IMMEDIATELY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF LEAKAGE AT THE LUER ADAPTER WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR D/L POWERPICC PROVENA CATHETER. LIGHT USAGE RESIDUES WERE OBSERVED AND NEEDLELESS INJECTION CAPS WERE ATTACHED TO BOTH LUER ADAPTERS. AN ATTEMPT TO INFUSE WATER THROUGH THE SAMPLE USING A 12ML SYRINGE REVEALED BOTH LUMENS TO BE PATENT TO INFUSION AND ASPIRATION; HOWEVER, DURING PRESSURIZATION, A LEAK WAS OBSERVED AT THE INTERFACE BETWEEN THE PURPLE LUER ADAPTER AND THE NEEDLELESS INJECTION CAP. NO LEAKS WERE OBSERVED WHEN PRESSURIZING THE PURPLE LUER ADAPTER USING VARIOUS NON-COMPLAINANT ACCESSORIES. MICROSCOPIC INSPECTION OF THE PURPLE LUER ADAPTER DID NOT REVEAL ANY DAMAGE OR DEFORMITY. INSPECTION OF THE LEAKING INJECTION CAP ALSO DID NOT REVEAL ANY DAMAGE OR DEFORMITY. THE OBSERVED LEAK APPEARED TO CAUSED BY INCOMPATIBILITY BETWEEN THE CATHETER AND THE INJECTION CAP; HOWEVER, NO DAMAGE OR DEFORMITY WAS DISCOVERED. CONSEQUENTLY, THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02697 | 2023-06-29 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH STRAIGHT DUAL-LUMEN DIALYSIS KIT 12F X 15CM | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5503150 |  | IT WAS REPORTED BY CUSTOMER THAT FOUND HAIR IN TRAY. ADDITIONAL INFORMATION RECEIVED 6/13/2023: IT WAS REPORTED THAT IR RADIOLOGY OPENED STERILE PACKAGE IN THEIR DEPARTMENT. THERE WAS NO PATIENT INVOLVEMENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02702 | 2023-06-29 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT W/BIOPATCH PROBE COVER & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1194108D5 |  | IT WAS REPORTED BY THE CUSTOMER THAT PICC LINE HAD SOMETHING ON THE CATHETER. THEY HAD TO BREAK STERILITY TO OPEN A NEW KIT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF FOREIGN MATERIAL ON THE CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS 4FR S/L POWERPICC SOLO CATHETER. LIGHT USAGE RESIDUES WERE OBSERVED ON THE SAMPLE. WHITE MATERIAL WAS ADHERED TO THE CATHETER SHAFT BETWEEN THE 3CM AND 4CM DEPTH MARKINGS. MICROSCOPIC INSPECTION OF THE MATERIAL REVEALED IT TO BE GLOSSY. THE MATERIAL DID NOT EXHIBIT UNIFORM THICKNESS. STREAKS AND MATERIAL FLOW WERE OBSERVED, WHICH SUGGESTED THAT THE MATERIAL WAS WIPED OR STREAKED ONTO THE CATHETER. ATTEMPTS TO SCRAPE OR REMOVE THE MATERIAL WERE UNSUCCESSFUL. THE SURFACE OF THE MATERIAL APPEARED SEMI-RIGID AND THE MATERIAL WAS FIRMLY ADHERED TO THE CATHETER SHAFT. THE FOREIGN MATERIAL OBSERVED WAS BONDED TO THE CATHETER SHAFT, SUGGESTING THAT IT MAY HAVE BEEN ADHESIVE OR ANOTHER MATERIAL THAT CURED FOLLOWING DEPOSITION ON THE CATHETER SHAFT. PHOTOGRAPHS HAVE BEEN FORWARDED TO THE MANUFACTURING SITE FOR FURTHER EVALUATION. BASED ON THE EVALUATION PERFORMED AT THE MANUFACTURING SITE, THE WHITE MATERIAL WAS IDENTIFIED AS THE INK USED TO PRINT MARKINGS ON THE MOLDED JOINT OF THE CATHETER. THEREFORE, THE ROOT CAUSE WAS DETERMINED TO BE RELATED TO DEVICE MANUFACTURE. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02705 | 2023-06-29 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 22 GAUGE POWERGLIDE PRO MIDLINE CATHETER 8CM BASIC KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6F122080 |  | IT WAS REPORTED ATTEMPT TO INSERT A MIDLINE IN A CHILD WITH RECURRENT PYELONEPHRITIS CAUSED BY MULTI-RESISTANT GERMS, ACCOMPANIED BY HIS MOTHER AND HER NURSE. ULTRASOUND IDENTIFICATION OF THE VEIN. LOCAL ANESTHETIC WITH LIDOCAINE 10 MG/ML (1ML). ULTRASOUND-GUIDED SKIN PUNCTURE, VISUALIZATION OF THE NEEDLE, DIFFICULT APPROACH TO THE VEIN DUE TO THE MOVEMENTS OF THE YOUNG CHILD. ULTRASOUND-GUIDED VEIN PUNCTURE, ULTRASOUND-GUIDED GUIDE EXIT (THE GUIDE IS PLACED ON THE POSTERIOR WALL OF THE VEIN). ATTEMPT TO DEPLOY THE CATHETER ALONG THE VEIN. RESISTANCE TO DEPLOYMENT. ATTEMPT TO WITHDRAW THE GUIDE, BLOCKAGE WHEN THE GUIDE IS WITHDRAWN AFTER A FEW CENTIMETERS. PERCEPTION OF A "BALL" IN THE SUPERFICIAL SUBCUTANEOUS. THE WHOLE DEVICE CANNOT RETRACT AND EXIT. NEED TO MAKE A 2 MM INCISION AT THE PUNCTURE POINT USING A COLD KNIFE. GENTLE TRACTION OF THE DEVICE WHICH EXITS THROUGH THE INCISION, VERY DAMAGED, WITH NO OBVIOUS DAMAGE TO THE SUBCUTANEOUS TISSUE OF THE CHILD CONCERNED. WOUND CLEANED WITH ALCOHOLIC CHLORHEXIDINE. NO NEED FOR SUTURES. A DRESSING IS APPLIED (COMPRESS + TEGADERM). RETURN OF THE CHILD ACCOMPANIED BY HIS MOTHER AND NURSE TO HIS PEDIATRIC WARD AND ADVICE TO ORGANIZE PLACEMENT OF MIDLINE UNDER GA. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF COMPLICATIONS DURING THE INSERTION PROCESS WAS CONFIRMED AND WAS DETERMINED TO BE ASSOCIATED WITH USE. ONE 22G POWERGLIDE CATHETER WAS PROVIDED FOR INVESTIGATION. THE CATHETER WAS RECEIVED IN TWO PIECES OVER THE NEEDLE OF THE DEPLOYMENT SYSTEM. THE CATHETER AND GUIDEWIRE WERE FOUND TO BE SEVERED AND WERE LIKELY DAMAGED BY THE NEEDLE BEVEL. THE NEEDLE WAS BENT AT THE FLASH NOTCH. IT WAS REPORTED THAT MOVEMENTS OF THE YOUNG PATIENT COMPLICATED THE PROCEDURE. NO DAMAGE ASSOCIATED WITH THE MANUFACTURING PROCESS COULD BE CONFIRMED FROM THE CONDITION OF THE RETURNED SAMPLE. A REVIEW OF THE INSPECTION RECORDS AND QUALITY/MANUFACTURING CONTROLS FOR THE IMPLICATED LOT INDICATED NO ISSUES WITH THE MANUFACTURING PROCESS. THE COMPLAINT HAS BEEN DOCUMENTED AND WILL CONTINUE TO BE MONITORED AS PART OF ONGOING EFFORTS TO IDENTIFY POTENTIAL MANUFACTURING RELATED ISSUES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02707 | 2023-06-29 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV 4F DL MAX 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D |  | IT WAS REPORTED BY THE CUSTOMER, CLINICIAN GOT INTO THE VEIN AND PUT THE INTRODUCER IN BUT COULD NOT THREAD THE PICC. SHE WAS ABLE TO REMOVE THE PICC WITHOUT ANY PROBLEM BUT AFTER REMOVING THE PICC SHE CHECKED THE WIRE AND NOTED IT WAS BENT IT 2 PLACES AND WAS DANGLING FROM THOSE 2 PLACES. PIECES DID NOT BREAK OFF. WE DID NOT USE ANY GREAT FORCE TO PULL OR PUSH IN THE PICC AS WE NEVER FORCEFULLY WITH PICCS OR WIRES. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02693 | 2023-06-29 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 20G X 10 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F320108PT |  | IT WAS REPORTED BY THE CUSTOMER THAT THEY HAD TWO SEPARATE CATHETERS THAT HAD HOLES WHERE THE CATHETER AND THE HUB MEET. THEY HAD TO PULL THE CATHETER AND PLACE ANOTHER ONE. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02692 | 2023-06-29 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER (W/ MICROINTRODUCER)(55CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | IT WAS REPORTED ¿DURING INSERTING PICC TO PATIENT CAP IS LEAKING (WHERE GUIDEWIRE IS COMING) AND WHILE ASPIRATING THEY SAW AIR COMING.¿ ADDITIONAL INFORMATION RECEIVED 6/9: IT WAS REPORTED THEY DID NOT CONTINUE TO INSERT THE PICC INTO THE PATIENT, NO HARM DONE. THE CUSTOMER STATED THE GUIDEWIRE IS ¿DIFFERENT FEELING¿ THAN PREVIOUS MATERIAL, THE CAP IS DIFFERENT SIZE THAN PREVIOUS, AND THE GUIDEWIRE IS NOT IN THE MIDDLE OF THE CAP (AS IT USED TO BE). NO CHEMO. GUIDEWIRE IS DIFFERENT FEELING THAT PREVIOUS MATERIAL. CAP IS DIFFERENT SIZE THAN PREVIOUS. GUIDEWIRE IS NOT IN THE MIDDLE OF THE CAP (AS IT USED TO BE). | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00469 | 2023-06-30 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U357576 | U357576 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY BURST. IT WAS FURTHER REPORTED THAT PIECES OF THE BALLOON REMAINED IN THE PATIENT. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 11/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02710 | 2023-06-30 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0607530 | 0607530 |  | IT WAS REPORTED THAT TEN MONTHS AND SEVENTEEN DAYS POST PORT PLACEMENT, THE CATHETER ALLEGEDLY HAD OPACIFICATION. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY LEAKED. IT WAS ALSO FURTHER REPORTED THAT THE CATHETER ALLEGEDLY FRACTURED. REPORTEDLY, THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2023). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
2020394-2023-00471 | 2023-06-30 | 2023-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRUE DILATATION CATHETER | BALLOON VALVULOPLASTY CATHETER | BARD PERIPHERAL VASCULAR, INC. | 0234512 | 0234512 |  | IT WAS REPORTED THAT DURING A VALVULOPLASTY PROCEDURE, THE DEVICE ALLEGEDLY HAD LEAKED AT 2 ATM. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 02/2026). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02711 | 2023-06-30 | 2023-06-06 | I, n, j, u, r, y | Manufacturer report | M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0603880 | 0603880 |  | IT WAS REPORTED THAT FOUR MONTHS AND NINETEEN DAYS POST PORT PLACEMENT PROCEDURE, THE DOCTOR ALLEGEDLY FELT A DIFFICULTY IN INJECTING THE PRODUCT WHILE INSERTING THE NEEDLE FOR THE INJECTION OF AN IMMUNOTHERAPY TREATMENT. IT WAS FURTHER REPORTED THAT THE IMPLANTABLE CHAMBER CATHETER ALLEGEDLY DISCONNECTED FROM THE CHAMBER BOX. REPORTEDLY, THE CATHETER WAS FOUND TO BE MIGRATED BETWEEN THE SUPERIOR VENA CAVA AND THE RIGHT VENTRICLE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 05/2025). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02720 | 2023-06-30 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER "THE GUIDE WIRE CAN'T GET THROUGH THE NEEDLE. REMOVE THE GUIDE WIRE, WIPE THE BLOOD ON THE SURFACE OF THE GUIDE WIRE WITH GAUZE, AND THE GUIDE WIRE IS BROKEN." IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02717 | 2023-06-30 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT SINGLE LUMEN PICCS ARE HAVING ISSUES WITH THE LITTLE RUBBER PIECE WHERE GUIDEWIRE IS PULLED THROUGH AFTER WE CUT THE CATHETER, IT IS NOT ALLOWING US TO FLUSH THE LINE (NO PRESSURE SUPPORT) NOR ARE WE ABLE TO GET BLOOD RETURN FROM IT UNTIL WE PLACED A FINGER OVER THAT RUBBER HUB. NO OTHER INFORMATION PROVIDED. THE CUSTOMER REPORTED MULTIPLE EVENTS HOWEVER THE EXACT NUMBER OF EVENTS WAS NOT PROVIDED TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02718 | 2023-06-30 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER "THE GUIDE WIRE CAN'T GET THROUGH THE NEEDLE. REMOVE THE GUIDE WIRE, WIPE THE BLOOD ON THE SURFACE OF THE GUIDE WIRE WITH GAUZE, AND THE GUIDE WIRE IS BROKEN." IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02721 | 2023-06-30 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER "THE GUIDE WIRE CAN'T GET THROUGH THE NEEDLE. REMOVE THE GUIDE WIRE, WIPE THE BLOOD ON THE SURFACE OF THE GUIDE WIRE WITH GAUZE, AND THE GUIDE WIRE IS BROKEN." IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE FOURTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02722 | 2023-06-30 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER "THE GUIDE WIRE CAN'T GET THROUGH THE NEEDLE. REMOVE THE GUIDE WIRE, WIPE THE BLOOD ON THE SURFACE OF THE GUIDE WIRE WITH GAUZE, AND THE GUIDE WIRE IS BROKEN." IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE FIFTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02724 | 2023-06-30 | 2023-06-14 | I, n, j, u, r, y | Manufacturer report | BARD TL PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. TRIPLE LUMEN BD POWERPICC UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02723 | 2023-06-30 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THE RESPONSIBLE NURSE CHECKED THE INTEGRITY OF THE CATHETER BEFORE PICC PLACEMENT AND FOUND THE CATHETER WAS BROKEN, TUBING RUPTURED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02725 | 2023-06-30 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD MIDLINE | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED DURING CLINICAL ASSESSMENT "LEAKAGE AT OR AROUND THE MIDLINE SITE."DEVICE INVOLVED IS A POWER MIDLINE IN THE UPPER ARM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02726 | 2023-06-30 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED ON (B)(6) 2022, A CENTRAL VENOUS CATHETER WAS INSERTED INTO THE PATIENT. THE CATHETER VESSEL WAS THE LEFT BRACHIAL VEIN, THE CATHETER DEPTH WAS 42CM, AND THE CATHETER TIP WAS LOCATED IN THE SUPERIOR VENA CAVA. AFTER CATHETERIZATION, THE PATIENTS WENT TO THE PICC OUTPATIENT CLINIC OF THE HOSPITAL FOR MAINTENANCE ON (B)(6) 2022, (B)(6) 2023, (B)(6), (B)(6), AND (B)(6) , AND THE CATHETER FUNCTION WAS NORMAL. DURING MAINTENANCE ON (B)(6), THE CATHETER SLIPPED 0.5CM OUTSIDE AND THE LENGTH IN THE BODY WAS 41.5CM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02728 | 2023-06-30 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER "APPROXIMATELY TWO MONTHS AFTER THE IMPLANTATION, A RUPTURE OF THE CATHETER OCCURRED IN THE INTRAVASCULAR PART ABOUT 5CM FROM THE BIFURCATION OF THE CONNECTORS. THERE IS A REPORTED CRACK OF APPROXIMATELY 1CM. THIS RESULTED IN AN OVERFLOW OF THE DRUG INTO THE SUBCUTIS, WHICH WAS APPROPRIATELY IDENTIFIED AND RESOLVED BY REIMPLANTATION OF THE CATHETERS. EVENT OCCURRED DURING USE. NO SERIOUS INJURY, ERRONEOUS RESULTS, NO CHANGE IN COURSE OF TREATMENT, NO EXPOSURE TO BLOOD OR BODY FLUIDS, NO TREATMENT OTHER THAN REPLACEMENT OF CATHETER." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02730 | 2023-06-30 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PROVEN A TM MIDLINE CATHETER 3F SL | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000917 |  | IT WAS REPORTED BY CUSTOMER THAT 3FR DILATORS ARE BUCKLING WHEN TRYING TO TRANSITION THROUGH SKIN. ADDITIONAL INFORMATION 6/19: IN INABILITY TO ADVANCE DILATOR AND SHEATH THROUGH SKIN REQUIRED ADDITIONAL SKIN NICK AND OBTAINING NEW 3 FR. DILATOR. NO URGENT/LIFE THREATING MEDICAL SITUATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02732 | 2023-06-30 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER "APPROXIMATELY TWO MONTHS AFTER THE IMPLANTATION, A RUPTURE OF THE CATHETER OCCURRED IN THE INTRAVASCULAR PART ABOUT 5CM FROM THE BIFURCATION OF THE CONNECTORS. THERE WAS POINT DAMAGE. THIS RESULTED IN AN OVERFLOW OF THE DRUG INTO THE SUBCUTIS, WHICH WAS APPROPRIATELY IDENTIFIED AND RESOLVED BY REIMPLANTATION OF THE CATHETERS. EVENT OCCURRED DURING USE. NO SERIOUS INJURY, ERRONEOUS RESULTS, NO CHANGE IN COURSE OF TREATMENT, NO EXPOSURE TO BLOOD OR BODY FLUIDS, NO TREATMENT OTHER THAN REPLACEMENT OF CATHETER." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02698 | 2023-06-30 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT W/BIOPATCH, PROBE COVER, & BELDAR NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CUSTOMER REPORTED THAT THE SALINE LEAKING OUT OF T-LOCK. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02735 | 2023-06-30 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA CATHETER W/ SOLO VALVE TECHNOLOGY, SHERLOCK 3CG (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1395108D |  | IT WAS REPORTED, ACCOUNT STATES THEY HAVE PLACED FOUR 5FR DL PICCS THAT <12 HOURS OF BEING PLACED, NONE OF THE LUMENS WILL DRAW AND CAN'T BE FLUSHED. STATES THERE HAVE BEEN DIFFERENT PLACERS, AND EACH PLACERS HAS STATED ZERO ISSUES WITH PLACEMENTS. STATES AFTER PLACEMENT EACH OF THE LUMENS DREW BLOOD AND HAD NO ISSUES. TROUBLE SHOOTING WAS PERFORMED: CATHFLO WITH NO IMPROVEMENT. AND STILL WAS UNABLE TO GET BLOOD RETURN AND HAD DIFFICULTY FLUSHED. AN XRAY WAS PERFORMED ON ALL FOUR PICCS AND THE XRAY SHOWED CORRECT PLACED OF THE LINE. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02736 | 2023-06-30 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA CATHETER W/ SOLO VALVE TECHNOLOGY, SHERLOCK 3CG (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1395108D |  | IT WAS REPORTED, ACCOUNT STATES THEY HAVE PLACED FOUR 5FR DL PICCS THAT <12 HOURS OF BEING PLACED, NONE OF THE LUMENS WILL DRAW AND CAN'T BE FLUSHED. STATES THERE HAVE BEEN DIFFERENT PLACERS, AND EACH PLACERS HAS STATED ZERO ISSUES WITH PLACEMENTS. STATES AFTER PLACEMENT EACH OF THE LUMENS DREW BLOOD AND HAD NO ISSUES. TROUBLE SHOOTING WAS PERFORMED: CATHFLO WITH NO IMPROVEMENT. AND STILL WAS UNABLE TO GET BLOOD RETURN AND HAD DIFFICULTY FLUSHED. AN XRAY WAS PERFORMED ON ALL FOUR PICCS AND THE XRAY SHOWED CORRECT PLACED OF THE LINE. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02739 | 2023-06-30 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA CATHETER W/ SOLO VALVE TECHNOLOGY, SHERLOCK 3CG (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1395108D |  | IT WAS REPORTED, ACCOUNT STATES THEY HAVE PLACED FOUR 5FR DL PICCS THAT <12 HOURS OF BEING PLACED, NONE OF THE LUMENS WILL DRAW AND CAN'T BE FLUSHED. STATES THERE HAVE BEEN DIFFERENT PLACERS, AND EACH PLACERS HAS STATED ZERO ISSUES WITH PLACEMENTS. STATES AFTER PLACEMENT EACH OF THE LUMENS DREW BLOOD AND HAD NO ISSUES. TROUBLE SHOOTING WAS PERFORMED: CATHFLO WITH NO IMPROVEMENT. AND STILL WAS UNABLE TO GET BLOOD RETURN AND HAD DIFFICULTY FLUSHED. AN XRAY WAS PERFORMED ON ALL FOUR PICCS AND THE XRAY SHOWED CORRECT PLACED OF THE LINE. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE FOURTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02742 | 2023-06-30 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "THE STATLOCKS ARE BREAKING AT THE CONNECTION ON THE A-LINES." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH SIX DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02743 | 2023-06-30 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "THE STATLOCKS ARE BREAKING AT THE CONNECTION ON THE A-LINES." ADDITIONAL INFORMATION RECEIVED ON 5 JULY 2023: CUSTOMER REPORTS THE PATIENT DID EXPERIENCE BLOOD LOSE. INFORMATION ON WHETHER THE BLOOD LOSS CAUSED COMPLICATIONS, PATIENT SYMPTOMS, OR REQUIRED TREATMENT WAS REQUESTED BUT NOT PROVIDED. IT WAS REPORTED THIS OCCURRED WITH SIX DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02741 | 2023-06-30 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "THE STATLOCKS ARE BREAKING AT THE CONNECTION ON THE A-LINES." ADDITIONAL INFORMATION RECEIVED ON 5 JULY 2023: CUSTOMER REPORTS THE PATIENT DID EXPERIENCE BLOOD LOSE. INFORMATION ON WHETHER THE BLOOD LOSS CAUSED COMPLICATIONS, PATIENT SYMPTOMS, OR REQUIRED TREATMENT WAS REQUESTED BUT NOT PROVIDED. IT WAS REPORTED THIS OCCURRED WITH SIX DEVICES. THIS REPORT ADDRESSES THE SIXTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02740 | 2023-06-30 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "THE STATLOCKS ARE BREAKING AT THE CONNECTION ON THE A-LINES." ADDITIONAL INFORMATION RECEIVED ON 5 JULY 2023: CUSTOMER REPORTS THE PATIENT DID EXPERIENCE BLOOD LOSE. INFORMATION ON WHETHER THE BLOOD LOSS CAUSED COMPLICATIONS, PATIENT SYMPTOMS, OR REQUIRED TREATMENT WAS REQUESTED BUT NOT PROVIDED. IT WAS REPORTED THIS OCCURRED WITH SIX DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A STATLOCK EXTENSION TUBING BREAK IS INCONCLUSIVE DUE TO POOR SAMPLE CONDITION. ONE PHOTO SAMPLE OF A STATLOCK ARTERIAL PLUS KIT WAS PROVIDED FOR EVALUATION. THE STATLOCK AND EXTENSION SET TUBING WERE VISIBLE IN THE IMAGE. A RED BOX IS DRAWN OVER THE PROXIMAL END OF THE EXTENSION TUBING AT THE INTERFACE WITH THE LUER HUB. NO APPARENT DAMAGE COULD BE OBSERVED FROM THE IMAGE PROVIDED. SINCE A BREAK IN THE CONNECTION COULD NOT BE OBSERVED FROM THE IMAGE PROVIDED, THE COMPLAINT COULD NOT BE CONFIRMED AND REMAINS INCONCLUSIVE AT THIS TIME. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02744 | 2023-06-30 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "THE STATLOCKS ARE BREAKING AT THE CONNECTION ON THE A-LINES." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH SIX DEVICES. THIS REPORT ADDRESSES THE FOURTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02745 | 2023-06-30 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "THE STATLOCKS ARE BREAKING AT THE CONNECTION ON THE A-LINES." NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH SIX DEVICES. THIS REPORT ADDRESSES THE FIFTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02734 | 2023-06-30 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIG JAPAN HANG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405J |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS IMPLANTED AND CONNECTED TO THE CATHETER CONNECTOR, WHEN THE CATHETER WAS CONNECTED TO A SYRINGE AND ATTEMPTED TO BE INFUSED WITH HEPARIN SALINE SOLUTION, THE INFUSION COULD NOT BE PERFORMED. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02719 | 2023-06-30 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER "THE GUIDE WIRE CAN'T GET THROUGH THE NEEDLE. REMOVE THE GUIDE WIRE, WIPE THE BLOOD ON THE SURFACE OF THE GUIDE WIRE WITH GAUZE, AND THE GUIDE WIRE IS BROKEN." IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02715 | 2023-06-30 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR POST A PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, LEAKAGE WAS ALLEGEDLY FOUND. IT WAS FURTHER REPORTED THAT THE PATIENT COMPLAINED OF DISCOMFORT. REPORTEDLY. THE PORT SYSTEM WAS REMOVED AND THERE IS NO INFORMATION REGARDING THE TREATMENT FOR THE DISCOMFORT. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02737 | 2023-06-30 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA CATHETER W/ SOLO VALVE TECHNOLOGY, SHERLOCK 3CG (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1395108D |  | IT WAS REPORTED, ACCOUNT STATES THEY HAVE PLACED FOUR 5FR DL PICCS THAT <12 HOURS OF BEING PLACED, NONE OF THE LUMENS WILL DRAW AND CAN'T BE FLUSHED. STATES THERE HAVE BEEN DIFFERENT PLACERS, AND EACH PLACERS HAS STATED ZERO ISSUES WITH PLACEMENTS. STATES AFTER PLACEMENT EACH OF THE LUMENS DREW BLOOD AND HAD NO ISSUES. TROUBLE SHOOTING WAS PERFORMED: CATHFLO WITH NO IMPROVEMENT. AND STILL WAS UNABLE TO GET BLOOD RETURN AND HAD DIFFICULTY FLUSHED. AN XRAY WAS PERFORMED ON ALL FOUR PICCS AND THE XRAY SHOWED CORRECT PLACED OF THE LINE. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02714 | 2023-06-30 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT WAS ALLEGEDLY LEAKED. IT WAS FURTHER REPORTED THAT THERE WAS A CRACK IN THE LINE. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02716 | 2023-06-30 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT APPROXIMATELY EIGHT YEARS POST PORT PLACEMENT, PATIENT ALLEGEDLY EXPERIENCED SOME BLEEDING AFTER HIS LAST TREATMENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-06
3006260740-2023-02750 | 2023-07-03 | 2023-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0606150J |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS AND TEN MONTHS POST PORT PLACEMENT VIA LEFT INTERNAL JUGULAR VEIN, THE CATHETER ALLEGEDLY HAD A CRACK. IT WAS FURTHER REPORTED THAT THE PATIENT ALLEGEDLY HAD PAIN WHEN SALINE SOLUTION WAS INFUSED. REPORTEDLY THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE TITANIUM IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION AND TWO IMAGES WERE PROVIDED FOR REVIEW. VISUAL, MICROSCOPIC, FUNCTIONAL AND TACTILE EVALUATIONS WAS PERFORMED. A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED FROM THE DISTAL END OF THE CATH-LOCK. THE EDGES OF THE CIRCUMFERENTIAL BREAK PROXIMAL END OF THE ATTACHED CATHETER WERE NOTED TO BE UNEVEN AND THE SURFACE APPEARED TO BE ROUGH. THE IMAGES SHOWS NO OBJECTIVE EVIDENCE OF FRACTURE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRY DATE: 07/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00477 | 2023-07-03 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 PTA DILATATION CATHETER | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U357588 | U357588 |  | IT WAS REPORTED THAT DURING THE PROCEDURE, THE ENTIRE SHAFT OF THE PTA BALLOON ALLEGEDLY EXPLODED DUE TO OVER-INFLATION. IT WAS FURTHER REPORTED THAT THE PTA BALLOON AND THE BROKEN PARTS WERE REMOVED. THERE WAS NO REPORTED INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 01/2026). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00475 | 2023-07-03 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DORADO | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | DR4084 | DR4084 |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE, THE RE-SHEATHER WAS ALLEGEDLY FUSED TO THE BALLOON AND PULLED THE BALLOON OFF THE SHAFT. IT WAS FURTHER REPORTED THAT THE BALLOON WAS COMPLETELY DETACHED FROM THE CATHETER SHAFT. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2024) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02753 | 2023-07-03 | 2023-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE VIA THE RIGHT INTERNAL JUGULAR VEIN, THE INTRODUCER NEEDLE WAS ALLEGEDLY UNABLE TO DRAW BLOOD. IT WAS FURTHER REPORTED THAT THE NEEDLE ALLEGEDLY LEAKED AIR. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE SAMPLE WAS NOT RETURNED FOR EVALUATION. HOWEVER, ONE VIDEO WAS PROVIDED FOR REVIEW. THE VIDEO SHOWS A CLINICIAN TRYING TO ASPIRATE THE WATER THROUGH THE INTRODUCER NEEDLE. CLINICIAN REMOVING AND CONNECTING THE NEEDLE AND SYRINGE FOR MULTIPLE TIMES. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED SUCTION PROBLEM AND AIR GAS IN DEVICE ISSUES, AS THE PROVIDED VIDEO IS NOT SUFFICIENT TO CONFIRM THE ISSUE AND ALSO THE EXACT CIRCUMSTANCE AT THE TIME OF THE REPORTED EVENT ARE UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: (B)(6) 2024), G3, H6 (METHOD) H11: H6 (RESULT, CONCLUSION) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02754 | 2023-07-03 | 2022-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT FOUR MONTHS AND THREE DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, THE TIP OF THE CATHETER ALLEGEDLY MISPOSITIONED IN THE HEAD SIDE INSTEAD OF THE HEART SIDE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: E1, H6 (METHOD). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00474 | 2023-07-03 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8088 | VA8088 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY GOT STUCK IN THE SHEATH. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: (B)(6) 2025) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02752 | 2023-07-03 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0606150J |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT, LEAKAGE WAS ALLEGEDLY FOUND. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00472 | 2023-07-03 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA80124 | VA80124 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY BURST. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BALLOON RUPTURE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. D4 (EXPIRATION DATE: 03/2026). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02751 | 2023-07-03 | 2023-05-17 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT TWO YEARS SEVEN MONTHS AND TWELVE DAYS POST PORT PLACEMENT, A COMPUTED TOMOGRAPHY EXAMINATION WAS PERFORMED AND REVEALED THAT THE CATHETER WAS ALLEGEDLY BROKEN AND THE DISTAL CATHETER SEGMENT MIGRATED TO THE PULMONARY ARTERY. IT WAS FURTHER REPORTED THAT ONLY A PART OF THE CATHETER AND THE PORT BODY WERE REMOVED. REPORTEDLY, CONSIDERING THE RISK OF REMOVAL, THE DOCTOR DECIDED THAT THE DISTAL PART OF CATHETER REMAINS INSIDE OF THE BODY AND CONTINUES THE FOLLOW-UP OBSERVATION. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRY DATE: 12/2022. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02748 | 2023-07-03 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT TWO YEARS EIGHT MONTHS AND TWENTY ONE DAYS POST PORT PLACEMENT PROCEDURE VIA THE SUBCLAVIAN VEIN, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY LEAKED. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02747 | 2023-07-03 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6905240 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER WAS UNABLE TO BE INTRODUCED THROUGH THE GUIDE. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY NOT COMPLETELY HOLLOW INSIDE AND RESISTANCE FELT UPON ADVANCEMENT. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 02/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02746 | 2023-07-03 | 2023-04-10 | I, n, j, u, r, y | Manufacturer report | POWERPORT TI L/P 6 CF INT WOSP ATT SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8716001 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE IN RIGHT INTERNAL JUGULAR VEIN, THE DILATATION SHEATH WAS ALLEGEDLY BENT. IT WAS FURTHER REPORTED THAT TIP OF THE SHEATH ALLEGEDLY CRACKED. REPORTEDLY, THE PATIENT EXPERIENCED SLIGHT VESSEL DAMAGE AND HAS BEEN DEALT WITH. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 08/2024). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00483 | 2023-07-04 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | AT75144 | AT75144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN ILIAC VEIN, THE PTA BALLOON ALLEGEDLY BURST AT 18 ATM WHEN PERFORMING MULTIPLE INFLATIONS. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ATLAS PTA DILATATION CATHETER WAS RETURNED FOR EVALUATION. PEELED PEBAX AND UNRAVELING FIBERS WERE ABLE TO BE OBSERVED ON THE BALLOON; NO OTHER ANOMALIES WERE NOTED DURING THE VISUAL EVALUATION. ON THE FUNCTIONAL TESTING, THE BALLOON WAS INFLATED UP TO 6 ATM WITH AN IN-HOUSE PRESTO INFLATION DEVICE AND NO LEAK WAS NOTED. THEN BALLOON WAS INFLATED TO 18 ATM AS REPORTED BY USER BALLOON BURST WAS OBSERVED. UNDER MICROSCOPIC OBSERVATION, A LARGE LONGITUDINAL RUPTURE WAS NOTED IN THE BALLOON. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. TWO PHOTOS WERE PROVIDED AND REVIEWED. BOTH THE PHOTOS SHOW THE BALLOON WITHIN A PLASTIC BAG AND THE BALLOON APPEARED TO BE BLOODY AND IN DEFLATED CONDITION, PEELED PEBAX, UNRAVELLING FIBERS WERE ABLE TO BE OBSERVED ON THE BALLOON. NO OTHER OBJECTIVE EVIDENCE FOR BALLOON RUPTURE COULDN¿T BE NOTED ON THE SUBMITTED PHOTO. FROM THE SUBMITTED PHOTOS, NO EVIDENCE OF BALLOON RUPTURE WAS OBSERVED, BUT PEELED PEBAX AND UNRAVELING FIBERS CAN BE OBSERVED, AND DURING THE RETURNED SAMPLE ANALYSIS, LONGITUDINAL BALLOON RUPTURE WAS NOTED UNDER MICROSCOPIC OBSERVATIONS. HENCE, THE INVESTIGATION CONFIRMS THE REPORTED BALLOON RUPTURE AND THE IDENTIFICATION OF PEELED PEBAX AND UNRAVELING FIBERS, RESPECTIVELY. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BALLOON RUPTURE AND IDENTIFIED PEELED PEBAX AND UNRAVELING FIBERS COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 02/2026). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00481 | 2023-07-04 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U357584 | U357584 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON WAS ALLEGEDLY GOT STUCK WITH THE STENT AND WAS UNABLE TO CROSS THE BENT SITE OF THE STENT. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00484 | 2023-07-04 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8088 | VA8088 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON CATHETER ALLEGEDLY GOT STUCK WITHIN THE INTRODUCER SHEATH. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED SHEATH REMOVAL DIFFICULTY COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: B5, D4 (EXPIRATION DATE: 11/2025), G3, H6 (METHOD) . H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04842 | 2023-07-05 | 2023-06-13 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 947318 | 175818 |  | IT WAS REPORTED THAT THE CUSTOMER REPORTED THAT 5 PATIENTS HAD EXPERIENCED URINARY RETENTION DESPITE CATHETERIZATION. THE WORST AFFECTED PATIENT HAD MORE THAN 500ML OF URINE RETAINED AS PER THE BLADDER SCAN WHICH WAS PERFORMED. THE CUSTOMER WAS UNSURE WHAT WAS THE CAUSE FOR THE RETENTION BUT THE USER OBSERVED THAT THE STAFF HAD BEEN SECURING THE CATHETERS INCORRECTLY WITH STATLOCK AND THEREFORE LEAVING TOO MUCH SLACK. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "TUBING DESIGN (LUMEN ID UNDERSIZED)/ LUMEN WALL THICKNESS UNDERSIZED/ INADEQUATE MATERIAL SELECTION". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "READ CAREFULLY BEFORE USE. APPLICATION TECHNIQUE: PREP 1. PLACE FOLEY CATHETER INTO RETAINER. DIRECTIONAL ARROW SHOULD POINT TOWARDS CATHETER TIP 2. CLOSE LID, BEING CAREFUL TO AVOID PINCHING THE CATHETER." CORRECTION: B, D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04806 | 2023-07-05 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT THE FOLEY TIP WAS DIFFERENT AND SEEMED A BIT BIGGER. THEY WERE CONTINUING TO HAVE PATIENT COMPLAINTS WITH REMOVAL OF THE FOLEYS. THEY CURRENTLY HAD A STOCK OF 5 CASES THAT THEY WOULD LIKE TO RETURN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT LINE CLEARANCE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04850 | 2023-07-05 | 2023-06-13 | I, n, j, u, r, y | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123416C |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS DIFFICULT TO REMOVE. CURRENTLY, THE HOSPITAL REPLACED THE PRODUCTS WITH UNITS FROM OTHER BATCHES. STATED THAT THE PATIENT HAD TO BE HOSPITALIZED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THEY DID NOT HAD ANY OTHER ISSUES OTHER THAN DIFFICULT TO REMOVE. | THE REPORTED EVENT IS INCONCLUSIVE DUE TO CONDITION OF THE SAMPLE RECEIVED. BASED ON THE ATTACHED VIDEO, IT WAS OBSERVED THE CATHETER HAS BEEN CUT AND THE BALLOON BURST AND DEFLATED. HOWEVER, THE REPORTED EVENT COULD NOT BE CONFIRMED DUE TO POOR SAMPLE CONDITION. NO FUNCTIONAL TESTING COULD BE PERFORMED AS ONLY PHOTOS ARE PROVIDED AS SAMPLE. IT IS UNKNOWN IF THE PRODUCT MEETS SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINCH LUMEN/COLLAPSE LUMEN/THICK SAC THICKNESS/VALVE DAMAGE/SHORT INFLATION LUMEN. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04821 | 2023-07-05 | 2023-06-13 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK URINE COLLECTION SYSTEM WAS NOT SUCTIONING, AND THE UNIT HAD BEEN RE TAPPED. PATIENT THOUGHT THAT THEY SENT THE ONE THAT WAS NOT WORKING. PER CUSTOMER VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THEY RECEIVED THE NEW DEVICE AND WAS STILL HAVING SUCTION ISSUES. PER THE FOLLOW UP VIA PHONE 21JUNE2023, THE PURE WICK FEMALE EXTERNAL CATHETER WAS CAUSING THE CUSTOMER TO BREAK OUT IN A RASH WITHIN THE LAST WEEK. PATIENTS DOCTOR GAVE ZINC CREAM TO USE. PATIENT STATED THAT SOMETIMES THE PUREWICK URINE COLLECTION SYSTEM SUCTIONS AND SOMETIMES IT DOES NOT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE OR MATERIAL SURFACE IS ROUGH, ABRASIVE OR UNCOMFORTABLE ". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER. "TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL". "DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS". "EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE". RECOMMENDATIONS: ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS". NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. PERI-CARE AND PLACEMENT: PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-04831 | 2023-07-05 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L22 |  | IT WAS REPORTED THAT AFTER CATHETER INSERTION IN THE OPERATING ROOM, THE BALLOON WAS FOUND DAMAGED. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. EVALUATION CONFIRMED THE BALLOON WAS RUPTURE. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE LOW LATEX VISCOSITY, SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1.METHOD FOR USE (1)DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2)DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BEINJURED.] 2.APPLICABLE PATIENTS (1)PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1)DO NOT REUSE. (2)DO NOT RESTERILIZE. (3)BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4)DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1)PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBERLATEX (2)PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX®SILVER LUBRICATH®FOLEYCATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD®SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO26EVERY 2FR 1. BALLOONCATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT 1. (NEITHER OF THEM MAYBE INCLUDED IN SOME PRODUCT VARIATIONS.) [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1)CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2)LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3)INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4)PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5)TO DEFLATE BALLOON AND REMOVE CATHETER, INSERTA LUERTIP(NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1)WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2)WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGEBY HANDS.[THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETERCANNOT BE REMOVED.] (3)WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4)NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON.[IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY.] (5)DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6)DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7)SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8)WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1.PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1)EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1)WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2)WHEN ANY PART OFTHEBALLOON AND/OR THE CATHETERIS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3)WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAYREMAIN IN THE BLADDER. THEREFORE,TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1)ATTACH LUER TIP SYRINGE TO THE INFLATIONVALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2)IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) 3)IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) 4)IF SITUATION WOULDN'T BEIMPROVED WITH3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) 5)IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) B. BALLOON-RUPTURE METHOD 1)INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) 2)IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A)UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER.(FIG.6) B)IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUMUNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) C)IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) 3.MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION -CATHETER KINKING, DAMAGE, RUPTURE -DIFFICULTY OR FAILURE TO REMOVE THE DEVICE -OCCLUSION OF CATHETER INNER LUMENS -ENCRUSTATION -ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE -DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS -URINARY-TRACT INFECTION -HEMORRHAGE, HEMATURIA -ALLERGY REACTION TO THE DEVICE -CALCULUSFORMATION -EDEMA -PAIN -DISCOMFORT -INJURY OF BLADDER OR URETHRAL -URETHRITIS, URINARY INCONTINENCE -RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: F,H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04825 | 2023-07-05 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE CATHETER FELL OUT OF THE PATIENT IMMEDIATELY AFTER PLACEMENT. UPON CHECKING THE CATHETER AFTER REMOVAL, IT WAS FOUND THAT THE BALLOON WAS DEFLATED. THERE WAS NO BALLOON RUPTURE OR TEAR. WATER LEAKAGE WAS UNKNOWN. THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04844 | 2023-07-05 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX®LUBRICATH® FOLEY CATHETER ALCOCK MODEL | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0134L20 | 0134L20 |  | IT WAS REPORTED THAT STERILE WATER COULD NOT BE INJECTED WHEN ATTEMPTED TO INJECT AFTER THE CATHETER INSERTION. ACCORDING TO THE INQUIRY SHEET ATTACHED, THE PRETEST WAS NOT DONE AND THERE WAS NO PROBLEM OBSERVED DURING INFLATION OF THE CATHETER. THE INFLATION FUNNEL DID NOT SWELL AND THERE WAS NO ABNORMALITY OBSERVED WITH THE SHAFT. USUALLY THE BALLOON COULD BE INFLATED WITHIN SECONDS, BUT SINCE THE PLUNGER COULD NOT BE PUSHED, THE BALLOON COULD NOT BE DEFLATED. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 19JUN2023, IT WAS REPORTED THAT WATER WAS UNABLE TO INJECT INTO THE CATHETER. THE PLUNGER ISSUE WAS NOT REPORTED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR PATIENT TREATMENT. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), USED THREE-WAY LATEX CATHETER FOLEY. VISUAL INSPECTION OF THE SAMPLE NOTED NO VISIBLE NONCONFORMITIES AND ALSO NOTED RESIDUE IN THE DRAINAGE EYE POSSIBLY FROM BEING USED. THE FOLEY CATHETER BALLOON WAS INFLATED WITH 35 ML OF D.I WATER WITH NO ISSUES. WITH THE SYRINGE ATTACHED THE CATHETER BALLOON DEFLATED PASSIVELY AS INTENDED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04858 | 2023-07-05 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CATHETER WAS PLACED BY INJECTING 10 CC OF WATER AND IN AN HOUR, IT FELL OUT OF THE PATIENT. UPON INJECTING 15 CC OF WATER TO CHECK FOR WATER LEAKAGE, THE WATER LEAKED FROM THE BALLOON IN 10 MINUTES. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® SILICONE-COATED FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILICONE COATING <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 30 EVERY 2FR 1. BALLOON CATHETER [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04823 | 2023-07-05 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED THAT THE PATIENT HAD 9 INTERMITTENT CATHETER OUT OF THIS BATCH THAT DID NOT CONTAIN THE SACHET OF LUBRICATION IN THE PACKET AND THEREFORE WERE NOT USABLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-04830 | 2023-07-05 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE CATHETER FELL OFF AND WHEN CHECKED, THE BALLOON WAS FOUND DAMAGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02760 | 2023-07-05 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0607540CE |  | IT WAS REPORTED THAT SEVEN MONTHS AND SEVEN DAYS POST PORT PLACEMENT, THE CATHETER ALLEGEDLY LEAKED. IT WAS FURTHER REPORTED THAT THE DEVICE ALLEGEDLY HAD A SUCTION PROBLEM. FURTHERMORE, THE PATIENT ALLEGEDLY EXPERIENCED PAIN AND DISCOMFORT DURING THE INJECTION. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 04/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02758 | 2023-07-05 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED ITCHING. IT WAS FURTHER REPORTED THAT THERE IS NO INFORMATION PROVIDED REGARDING THE INTERVENTION PERFORMED OR MEDICATION PRESCRIBED FOR THE TREATMENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 01/2025) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04892 | 2023-07-05 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRETEST WHEN THE BALLOON WAS TRIED TO INFLATE, THERE WAS A RESISTANCE. AFTER INFLATION WHEN THE STERILE WATER WAS TRIED TO REMOVE, IT COULD NOT BE DRAINED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "LUMEN PARTIALLY COLLAPSED; PARTIALLY KINKED / PINCHED LUMEN". A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04840 | 2023-07-05 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS DIFFICULT TO DEFLATE DURING PRETEST. ABOUT 5 ML OF WATER WAS ASPIRATED BY SYRINGE. STATED THAT DUE TO PRE-TESTING, THE WATER WAS NOT COMPLETELY DRAINED. USUALLY, BALLOON COULD BE DEFLATED WITHIN 20-40 SECONDS. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. PHOTO: THERE ARE 2 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE PHOTOS SHOW AN OVERVIEW OF THE PARTIALLY INFLATED 3-WAY CATHETER. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED THAT THE BALLOON WAS PARTIALLY INFLATED UPON RECEIVAL OF SAMPLE. DEFLATED THE BALLOON (USING THE RETURNED SYRINGE) PASSIVELY WITH NO CUFFING OR LEAKS NOTED. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 2 MINUTES WITH NO CUFFING OR LEAKS NOTED, RETURNING 10ML OF SOLUTION. THIS IS WITHIN SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04880 | 2023-07-05 | 2023-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE BALLOON COULD BE INFLATED DURING THE PRE-TEST, BUT THE WATER COULD NOT BE DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02761 | 2023-07-05 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1716000 | 1716000 |  | IT WAS REPORTED THAT SOMETIMES POST PORT PLACEMENT PROCEDURE, PATIENT ALLEGEDLY EXPERIENCED ITCHING. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: (EXPIRATION DATE: 10/2021). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02785 | 2023-07-05 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 3.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668935 |  | IT WAS REPORTED THAT THE VASCULAR SHEATH IN THE MICRO-INTRODUCER WAS NOT STRAIGHT. THERE WAS ALSO AN ANGLE BETWEEN THE VASCULAR SHEATH AND THE DILATOR. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02784 | 2023-07-05 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 3.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668935 |  | WAS REPORTED THAT THE VASCULAR SHEATH IN THE MICRO-INTRODUCER WAS NOT STRAIGHT. THERE WAS ALSO AN ANGLE BETWEEN THE VASCULAR SHEATH AND THE DILATOR. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A BENT MICROINTRODUCER IS INCONCLUSIVE DUE TO THE STATE OF THE RETURNED SAMPLE. ONE PHOTOGRAPH OF A 3.5 FR MICROEZ MICROINTRODUCER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED A MICROEZ MICROINTRODUCER KIT. ONLY THE PROXIMAL END OF THE MICROINTRODUCER COULD BE SEEN IN THE RETURNED PHOTOGRAPH. THE GRAY HANDLE PIECE WAS NOT FULLY SEATED IN THE COLLAR OF THE DILATOR AND WAS SITTING AT AN ANGLE. NO DAMAGE COULD BE SEEN ON THE PRODUCT IN THE RETURNED PHOTOGRAPH. ALTHOUGH THE SHEATH HANDLE WAS NOT SEATED IN THE COLLAR OF THE DILATOR AND WAS AT AN ANGLE, IT COULD NOT BE DETERMINED IF THE MICROINTRODUCER WAS BENT DUE TO THE KIT LABEL COVERING PART OF THE SAMPLE IN THE PHOTOGRAPH. THEREFORE, THE COMPLAINT IS INCONCLUSIVE AT THIS TIME. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02783 | 2023-07-05 | 2023-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THE BLUE CORE TUBE IN THE PICC EXTENSION TUBE IS DISCONNECTED AND THE BROKEN END IS IRREGULAR. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02782 | 2023-07-05 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER PICC LINE TUBING HAPPENED TO COME OUT DURING INPATIENT HOSPITAL STAY AND CUSTOMER BEGAN TO BLEED WHICH ALSO CAUSED ARM SWELLING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02780 | 2023-07-05 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PG PRO 22G 8CM RT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F322088PTX |  | IT WAS REPORTED BY THE CUSTOMER THAT MIDLINE GOT STUCK IN PATIENT ON INSERTION ADDITIONAL INFO RECEIVED: CUSTOMER REPORTED THAT PT. HAD PAIN AT INSERTION SITE. DIAGNOSED WITH STREP A, I COULD NOT PULL OUT GUIDEWIRE AND DEVICE ONCE DEPLOYED. I HAD TO REMOVE THE WHOLE DEVICE AND CATHETER WITH GREAT RESISTANCE. NO COMPLICATIONS, I WAS ABLE TO PULL ALL OF IT OUT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF DIFFICULTY REMOVING THE GUIDEWIRE OF A POWERGLIDE PRO IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. ONE 22 GA POWERGLIDE PRO WAS RETURNED FOR EVALUATION. THE GUIDEWIRE IN THE RETURNED POWERGLIDE PRO WAS OBSERVED TO HAVE DISJOINTED COILS THROUGHOUT THE DISTAL END OF THE GUIDEWIRE, AND THE GUIDEWIRE WAS OBSERVED TO HAVE SHARP BENDS COMING OUT OF THE DISTAL END OF THE NEEDLE. BASED ON THE SHARP BENDS AND DISJOINTED COILS OBSERVED ALONG THE RETURNED GUIDEWIRE OF THE POWERGLIDE PRO, THE COMPLAINT OF DIFFICULT DEVICE REMOVAL IS CONFIRMED. INSERTION INTO RESISTANCE SUCH AS TISSUE CAN CAUSE DAMAGE TO THE GUIDEWIRE THUS CAUSING DIFFICULTY DURING REMOVAL OF THE GUIDEWIRE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02759 | 2023-07-05 | 2023-06-09 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808060 | 1808060 |  | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT, THE PATIENT ALLEGEDLY DEVELOPED ITCHING; HOWEVER, THE ITCHING STOPPED. IT WAS FURTHER REPORTED THAT THE PATIENT ALLEGEDLY EXPERIENCED A STAPH INFECTION. REPORTEDLY, THE PORT WAS REMOVED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 07/2023), G3, H6 (METHOD). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04910 | 2023-07-05 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT THE SAME MAGIC 3 MALE CATHETERS DID NOT HAVE HOLES TO DRAIN. REPLACEMENT PLACED FOR 20 CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-04906 | 2023-07-05 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI20 |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS INSERTED, IT CAME OUT ON ITS OWN. THE BALLOON WAS RE-INFLATED IN AIR AND OBSERVED FOR A WHILE, BUT NO ABNORMALITY WAS FOUND. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON UNIFORM RUBBERIZED THICKNESS ETC CAUSING TORN RUBBERIZE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE(S) THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARDEX® SILVER LUBRICATH® FOLEY CATHETER SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D, H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04901 | 2023-07-05 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 33420 | 33420 |  | IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED DUE TO A FAULTY BARD FOLEY CATHETER. STATED THAT THE FOLEY CATHETER EXPLODED WHILE INSIDE PATIENT AND THE CATHETER WAS SURGICALLY REMOVED. PATIENT ALSO STATED THAT THEY DID NOT OVERFILL THE BALLOON AND WENT TO HOSPITAL IN MAY. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "HIGH MODULUS LATEX". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER. ENGLISH UNITS THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. FOR UROLOGICAL USE ONLY. DO NOT USE IF PACKAGE IS DAMAGED BARD, BARDEX, AND LUBRICATH ARE TRADEMARKS AND/OR REGISTERED TRADEMARKS OF C. R. BARD, INC. SINGLE USE DO NOT RE-STERILIZE MANUFACTURER. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. CONTAINS OR PRESENCE OF NATURAL RUBBER LATEX. CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. KEEP AWAY FROM SUNLIGHT VALVE TYPE: USE LUER SLIP TIP SYRINGE. DO NOT USE NEEDLE. RECOMMENDED INFLATION CAPACITIES 3CC BALLOON: USE 5ML STERILE WATER 5CC BALLOON: USE 10ML STERILE WATER 15CC BALLOON: USE 20ML STERILE WATER 20CC BALLOON: USE 25ML STERILE WATER 30CC BALLOON: USE 35ML STERILE WATER 40CC BALLOON: USE 45ML STERILE WATER 75CC BALLOON: USE 80ML STERILE WATER. DO NOT EXCEED RECOMMENDED CAPACITIES. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04843 | 2023-07-05 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PLACED FOR THE PATIENT ON MAY 31, BUT THE LIQUID IN THE BALLOON COULD NOT BE WITHDRAWN. THEREFORE, A NEW CATHETER WAS REPLACED IMMEDIATELY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿DO NOT USE THE PRODUCT OF HAVE LATER ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02757 | 2023-07-05 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED CHEST PAIN, NECK PAIN AND UNDERARM PAIN ON THE RIGHT SIDE. IT WAS FURTHER REPORTED THAT THERE IS NO ADDITIONAL INFORMATION PROVIDED REGARDING THE INTERVENTION PERFORMED OR MEDICATION PRESCRIBED FOR THE TREATMENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04904 | 2023-07-05 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 995081 | 995081 |  | IT WAS REPORTED THAT THE DILATION CATHETER WAS RUPTURE. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR THIS INVESTIGATION. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE POOR BALLOON DESIGN. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSPECTION PRIOR TO USE THE X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER IS A STERILE, SINGLE USE DEVICE. CAREFULLY INSPECT THE CATHETER AND THE STERILE PACKAGING FOR SIGNS OF DAMAGE THAT MAY HAVE OCCURRED DURING SHIPMENT. DO NOT USE THE PRODUCT IF DAMAGE IS EVIDENT. INFLATING THE BALLOON CATHETER 1. FILL THE INFLATION DEVICE (EAGLE¿ INFLATION DEVICE) WITH STERILE, LIQUID MEDIA. 2. ATTACH THE INFLATION DEVICE TO THE BALLOON LUMEN. 3. OPEN THE STOPCOCK AND INFLATE THE BALLOON. 4. ONCE DILATION HAS BEEN ATTAINED, ADVANCE THE SHEATH OVER THE BALLOON. WARNING: DO NOT USE AIR OR ANY GASEOUS SUBSTANCES AS A BALLOON INFLATION MEDIA, ALWAYS USE STERILE LIQUID MEDIA. NOTE: THE USE OF AN INFLATION DEVICE WITH A PRESSURE GAUGE IS HIGHLY RECOMMENDED TO MAKE SURE ADEQUATE PRESSURE IS APPLIED AND THE MAXIMUM LIMITS OF THE BALLOON ARE NOT EXCEEDED. NOTE: SLIGHT REDUCTION IN BALLOON PRESSURE MAY BE REQUIRED IF RESISTANCE IS ENCOUNTERED WHEN ADVANCING THE SHEATH OVER THE BALLOON. CAUTION: DO NOT EXCEED THE RECOMMENDED RATED BURST PRESSURE (RBP) FOR THIS DEVICE. BALLOON RUPTURE MAY OCCUR IF THE RBP RATING IS EXCEEDED. PLEASE REFER TO THE DEVICE LABEL FOR THE RECOMMENDED RBP." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04859 | 2023-07-05 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE BALLOON AREA OF THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02762 | 2023-07-05 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5733690 | 5733690 |  | IT WAS REPORTED THAT ONE MONTH AND THREE DAYS POST DRAINAGE CATHETER PLACEMENT, WHILE FLUSHING WITH SALINE, SWELLING WAS ALLEGEDLY NOTED. IT WAS FURTHER REPORTED THAT LEAKAGE WAS ALLEGEDLY FOUND IN THE CATHETER UNDER FLUOROSCOPY. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A VIDEO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2024). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04876 | 2023-07-05 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THIS WAS THE 3RD FOLEY CATHETER THAT THEY HAD TO CHANGE FOR THE PATIENT AS THE CATHETER GOT BLOCKED. STATED THAT THE THEATRE STOCK CONTROLLER CALLED TO STATED THAT THEY HAD A FAULTY SUPRAPUBIC CATHETER. UPON ARRIVAL, THEY MANAGED TO SPEAK TO DOCTOR AND ADVISED THE CATHETER WAS BLOCKED (EVENT OCCURRENCE DATE: 12JUN2023). THE URINE WAS NOT FLOWING THROUGH THE CATHETER AND WAS EXITING FROM OTHER AVENUES, THIS WAS INSERTED APPROXIMATELY 3 WEEKS PRIOR. THE DOCTOR STATED THAT THEY TRIED TO FLUSH THE CATHETER WITH A TUMI SYRINGE AND WAS NOT ABLE TO, THEY THEN REMOVED THE BLOCKED CATHETER AND REPLACED WITH ANOTHER SIZE 16 BARD SUPRAPUBIC CATHETER. STATED THAT THIS HAPPENED WITH THE SAME PATIENT TWICE BEFORE (EVENT OCCURRENCE DATE: UNK). | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INCORRECT PARAMETER SETTING BLOCKED DRAINAGE EYE/LUMEN OR NO DRAINAGE EYE OR USER RELATED (SALT ACCUMULATION). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: THIS IS A SINGLE USE DEVICE. DO NO RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH LEAD TO INJURY, ILLNESS, OR DEATH OF THE PATIENT. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE THE BALLOON TO BURST." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04872 | 2023-07-05 | 2023-06-14 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53618 | 53618 |  | IT WAS REPORTED THAT THESE CATHETERS ARE CAUSING URINARY TRACT INFECTION TO THE PATIENT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS, MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. THE CATHETER BECOMES SLIPPERY WHEN WETTED WITH WATER, ELIMINATING THE NEED FOR A SEPARATE LUBRICANT. FOR YOUR ADDED CONVENIENCE, THIS CATHETER IS PACKAGED WITH ITS OWN STERILE WATER. SIMPLY RELEASE THE WATER FROM ITS FOIL PACKET AND THEN TIP THE UN-OPENED CATHETER PACKAGE END-TO-END. THE CATHETER ACTS LIKE A MAGNET TO ATTRACT THE WATER AND ACTIVATE ITS SLIPPERY COATING. FOLLOW THESE STEPS FOR BEST RESULTS. 1. RELEASE THE WATER PRIOR TO OPENING THE SEALED CATHETER POUCH: A. APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. B. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. 2. WET THE CATHETER A. HOLD PACKAGE WITH PRINTED SIDE UP. B. TIP PACKAGE END-TO-END THREE TO SIX TIMES TO WET CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. 3. USE THE CATHETER A. PEEL BACK PACKAGE TO EXPOSE FUNNEL END OF CATHETER. B. IF DESIRED, USE THE SELF-ADHESIVE TAPE ON THE PACKAGE TO TEMPORARILY ATTACH THE POUCH TO ANY DRY VERTICAL SURFACE. C. REMOVE CATHETER AND USE ACCORDING TO PHYSICIAN¿S INSTRUCTIONS. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE, DO NOT RESTERILIZE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-07
3006260740-2023-02768 | 2023-07-05 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR (RADEL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770352 |  | IT WAS REPORTED THE CORD HAS PULLED OUT OF THE SENSOR EXPOSING WIRING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04897 | 2023-07-05 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRETEST WHEN THE BALLOON WAS TRIED TO INFLATE, THERE WAS A RESISTANCE. AFTER INFLATION WHEN THE STERILE WATER WAS TRIED TO REMOVE, IT COULD NOT BE DRAINED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "LUMEN PARTIALLY COLLAPSED; PARTIALLY KINKED / PINCHED LUMEN". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04914 | 2023-07-05 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE PATIENT GOT A CATHETERIZATION ON (B)(6)2023, AND ON (B)(6) 2023, IT WAS FOUND THAT THE FOLEY CATHETER BALLOON HAD BURST, BUT THE PATIENT DID NOT FEEL ANY DISCOMFORT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/EXPOSURE TO PETROLATUM BASED PRODUCTS/MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04895 | 2023-07-05 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316M | 119316M |  | IT WAS REPORTED, THAT MULTIPLE REPORTS OF TEMPERATURE SENSING FOLEY CATHETERS LEAKING. AFTER INSERTION (DIFFERENT PATIENTS). FOLEY CATHETER LEAKING URINE AROUND THE FOLEY DEVICE, AS WELL AS THROUGH THE FOLEY TUBING. BALLOON FULLY INFLATED. PRODUCT AND PACKAGING SAVED. WHEN THE LEAKING DID NOT STOP, NURSE REMOVED THE TEMPERATURE SENSING FOLEY. AND PLACED A STANDARD FOLEY, WITHOUT ANY FURTHER LEAKING OR ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04918 | 2023-07-05 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THERE WAS AN ISSUE ON AT LEAST 2 OCCASIONS WHERE THE FOLEY CATHETER BALLOON DID NOT INFLATE PROPERLY. ONE FELL OUT OF A PATIENT, THE OTHER WAS A TEST (NOT USED) WHERE THE BALLOON INFLATED HOWEVER THE BALLOON HAD A HOLE AND LEAKED. PRODUCT WAS DEFORMED, FOLEY CATHETER BALLOON FILLED UP WITH 10CC WAS LOPSIDED. PER FOLLOW UP VIA EMAIL ON 28JUN2023, IT WAS STATED THAT THE ONLY SAMPLE AVAILABLE WAS THE FOLEY THAT HAD A HOLE IN THE TUBING, NOT THE FOLEY WHERE THE BALLOON WAS LOPSIDED BECAUSE IT WAS USED, THEY COULD NOT SAVE IT. PATIENT WITH LOPSIDED FOLEY WAS OKAY, BUT WAS VERY UNCOMFORTABLE THEREFORE THEY REMOVED THE FOLEY AND THAT WAS WHEN THEY REALIZED THAT THE PRODUCT WAS DEFORMED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE MANUFACTURING RELATED DUE TO COLLAPSED / PINCHED INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04902 | 2023-07-05 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT A PINHOLE WAS OBSERVED ON THE BALLOON DURING THE PRE-TEST. THERE WAS NO BALLOON RUPTURE OR TEAR. WATER LEAKED FROM THE BALLOON. THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE BALLOON. THERE WAS NO PATIENT'S URINE STONE. THE CATHETER WAS NOT PULLED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04900 | 2023-07-05 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER COULD BE INJECTED DURING PRETEST BUT COULD NOT BE DRAINED FROM THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02778 | 2023-07-05 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER ¿APPROXIMATELY TWO MONTHS AFTER DEVICE IMPLANTATION, RUPTURE OF THE CATHETER IN THE INTRAVASCULAR PART. APPROXIMATELY 5 CM FROM THE CONNECTOR FORK THERE IS A HOLE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04915 | 2023-07-05 | 2023-05-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT ON (B)(6) 2023, A FR16 DOUBLE LUMEN CATHETER WAS INDWELLED FOR THE PATIENT FOR URINE DRAINAGE. ON (B)(6) 023, THE CATHETER FELL OUT DUE TO BALLOON BURST. AFTER REPORTED TO THE DOCTOR, IT WAS SUGGESTED TO KEEP OBSERVATION AND THE PATIENT HAD BEEN ABLE TO URINATE BY SELF. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04917 | 2023-07-05 | 2023-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE FR20 CATHETER WAS INDWELLED FOR THE PATIENT AFTER THE OPERATION ON (B)(6) . ON (B)(6) 2023, THE CATHETER FELL OUT DUE TO BALLOON BURST. AFTER REPORTING TO THE DOCTOR, A NEW CATHETER WAS USED FOR THE PATIENT WHICH GAVE GOOD FIXATION AND DRAINED SMOOTHLY BUT THE EVENT INCREASED PATIENT'S PHYSICAL PAIN, PSYCHOLOGICAL PRESSURE AND ECONOMIC COST. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04911 | 2023-07-05 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® URETHRAL CATHETER | URETHRAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 120614 | 120614 |  | IT WAS REPORTED THAT THE TIEMANN CATHETER OPENED AND READY TO USE AND A WHITE POWDER WAS FOUND THROUGHOUT THE LENGTH OF THE CATHETER IN THE PACKAGING. STATED THAT THE PACKAGE OPENED PRIOR TO USE AND NOT ABLE TO USE. | THE REPORTED EVENT IS CONFIRMED, MANUFACTURING RELATED. THE RED RUBBER LATEX CATHETER WAS RETURNED IN THE OPENED ORIGINAL PACKAGING. WHITE POWDERY SUBSTANCE WAS NOTED TO BE PRESENT WITHIN ONE OF THE PACKAGES. THE POWDERY SUBSTANCE WAS CONFIRMED TO BE SNOWFLOSS. UPON FURTHER EVALUATION, NO EYELETS WERE PRESENT ON THE CATHETER, THIS IS OUT OF SPECIFICATIONS. THE LACK OF DRAINAGE EYES WAS CONFIRMED TO BE THE CAUSE OF THE SNOWFLOSS AS THIS PREVENTED WATER FROM FLOWING THROUGH THE CATHETER DURING THE WASH/CHLORINATION MACHINE THAT WOULD NORMALLY REMOVE THE RELEASE AGENT. THE COUDE INDICATOR ALIGNED WITH THE CURVE AND THE CATHETER WAS CONFIRMED TO BE 14FR USING A FRENCH A GAUGE. A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE OPERATOR ERROR. AS NO PATIENT INVOLVEMENT WAS REPORTED THE DEVICE WAS NOT USED, HOWEVER IS INTENDED FOR TREATMENT PURPOSES. NO MANUFACTURING ISSUES OR NON-CONFORMANCES WERE NOTED DURING REVIEW OF THE DHR THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABEL/PACK REVIEW IS NOT REQUIRED AS A REVIEW OF THE LABEL COULD NOT HAVE PREVENTED THE REPORTED EVENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02755 | 2023-07-05 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0606150J |  | IT WAS REPORTED THAT THREE MONTHS AND SEVEN DAYS POST PORT PLACEMENT VIA THE RIGHT SUBCLAVIAN VEIN FOR CHEMOTHERAPY, THE PATIENT COMPLAINED OF DISCOMFORT IN THE NECK DURING CHEMOTHERAPY. IT WAS FURTHER REPORTED THAT WHEN THE REMOVED PORT SYSTEM WAS INFUSED, NO DAMAGE TO THE CATHETER WAS OBSERVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE TITANIUM IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DISCOMFORT ISSUE, AS THERE ARE NO OBJECTIVE EVIDENCE PROVIDED TO CONFIRM THE ISSUE. THEREFORE, THE CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT IS INCONCLUSIVE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04888 | 2023-07-05 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN THE BALLOON WAS TESTED, THE INJECTED STERILE WATER COULD NOT BE DRAINED. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. THE POTENTIAL FAILURE MODES ARE COLLAPSED LUMEN OR BLOCKED LUMEN. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED THAT THE BALLOON WAS PARTIALLY INFLATED UPON RECEIVAL OF SAMPLE. DEFLATED BALLOON PASSIVELY USING RETURNED SYRINGE WITH NO CUFFING OR LEAKS NOTED. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) USING RETURNED SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. ATTEMPTED TO DEFLATE BALLOON PASSIVELY USING RETURNED SYRINGE AND ONLY 2 ML COULD BE WITHDRAWN. ATTEMPTED TO DEFLATE BALLOON WITH IN-HOUSE LUER LOCK SYRINGE AND SOLUTION COULD NOT BE WITHDRAWN. REPLACED RETURNED INFLATION VALVE WITH AN IN-HOUSE VALVE AND REATTEMPTED INFLATION AND DEFLATION WITH BOTH SYRINGES. NEITHER SYRINGE COULD WITHDRAW MORE THAN 1 TO 2 ML OF SOLUTION PASSIVELY. SOLUTION WAS ASPIRATED WITH RETURNED SYRINGE. DISSECTED BALLOON AND FOUND THAT THE NOTCH WAS PERFORATED COMPLETELY. DISSECTED INFLATION FUNNEL AND PIN GAUGES 0.024¿ AND BELOW COULD FREELY PASS THROUGH LUMEN. 0.025¿ - 0.035¿ TOOK SOME EFFORT TO GET THROUGH AND 0.036¿ AND ABOVE CANNOT PASS THROUGH. BASED ON PHYSICAL SAMPLE RECEIVED THE PRODUCT DOES NOT MEET SPECIFICATIONS WHICH STATES "SHAFT MUST NOT BE DAMAGED OR PINCHED." A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE: ¿ INCORRECT DRAIN FORMER SELECTION. ¿ GLUE FROM TEMP SENSING ASSEMBLY. ¿ PHYSIOLOGICAL OBSTRUCTION. ¿ PATIENT POSITION. ¿ KINKED CATHETER. ¿ BLOCKED BY CLOTS, ETC. ¿ TUBING DESIGN (LUMEN ID UNDERSIZED) . ¿ LUMEN WALL THICKNESS UNDERSIZED . ¿ INADEQUATE MATERIAL SELECTION. ¿ LUMEN COLLAPSE. ¿ FOREIGN MATTER IN LUMEN. ¿ KINKED OR PINCHED SHAFT. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02756 | 2023-07-05 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN, WHITE, 4.2F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 0601610 | 0601610 |  | IT WAS REPORTED THAT PRIOR TO A CHRONIC CATHETER REPAIR PROCEDURE, THE GLUE WAS ALLEGEDLY HARD. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: A VOLUNTARY RECALL HAS BEEN INITIATED FOR CATHETER REPAIR KITS WHICH ARE PRODUCT CATALOG/LOT NUMBER SPECIFIC. REPORTEDLY THE CATHETER REPAIR KITS HAVE HARDENED/COAGULATED ADHESIVE. HARDENED OR COAGULATED ADHESIVE HAS THE POTENTIAL TO CAUSE DELAYS WHILST AN ALTERNATIVE REPAIR KIT, ADHESIVE OR REPLACEMENT CATHETER IS OBTAINED; THEREFORE, POTENTIALLY PROLONGING SURGERY OR NECESSITATING EXCHANGE. TO DATE THERE HAVE BEEN NO ADVERSE EVENTS WORLDWIDE REPORTED RELATED TO THIS ISSUE. AS RESULT OF THE FIELD ACTION, THIS EVENT IS BEING REPORTED AS A MALFUNCTION REPORTABLE EVENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 09/2023) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04841 | 2023-07-05 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PLACED IN THE PATIENT ON (B)(6) AND ON (B)(6), IT WAS FOUND THAT THE BALLOON RUPTURED AND THE CATHETER FELL OUT DURING A ROUTINE INSPECTION. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 16JUN2023, IT WAS STATED THAT THERE WERE NO MISSING PIECES OF THE BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04815 | 2023-07-05 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. THE CATHETER WAS INSERTED IN THE MORNING AND SPONTANEOUSLY FELL OUT AT AROUND 3:30 P.M. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04839 | 2023-07-05 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE USER ATTEMPTED TO INJECT STERILE WATER BUT THE STERILIZED WATER DID NOT ENTER INTO THE FOLEY CATHETER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. IT WAS UNKNOWN IF THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3-WAY TEMPERATURE SENSING CATHETER. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) USING AN IN-HOUSE LUER LOCK SYRINGE WITH NO ISSUES AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00487 | 2023-07-05 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U3513076 | U3513076 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE BALLOON ALLEGEDLY CAME WITH A SMALL HOLE AT THE JUNCTION WITH THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04822 | 2023-07-05 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT AFTER INJECTING 15ML SALINE, A SOUND OF BURST WAS HEARD AND THE CATHETER FELL OUT OF THE PATIENT. THE CATHETER WAS PUNCTURED REPEATEDLY WHICH WAS EASY TO CAUSE INJURY TO THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04829 | 2023-07-05 | 2023-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 995081 | 995081 |  | IT WAS REPORTED THAT IN AN ONLINE SURVEY A PHYSICIAN STATED THAT THEY WERE NOT ABLE TO USE THE REFOLDING TOOL TO REFOLD THE BALLOON DILATION CATHETER FOR REINSERTION. THE PHYSICIAN WAS NOT ABLE TO USE THE REFOLDING TOOL TO REFOLD THE 995081- X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, STERILE BECAUSE "UNKNOWN". | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE IMPROPER MATERIAL SELECTION. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿DESCRIPTION: THE BALLOON DILATION CATHETER COMES PACKAGED WITH A REFOLDING TOOL AND A WORKING SHEATH. IT IS AVAILABLE WITH OR WITHOUT AN INFLATION DEVICE. IT COMES STERILE AND IS FOR SINGLE USE ONLY. USING THE REFOLDING TOOL: THE BALLOON CATHETER IS SUPPLIED WITH A REFOLDING TOOL. THIS AIDS IN PREPARING THE BALLOON FOR ANOTHER INSERTION DURING THE PROCEDURE. TO USE: MANUALLY COMPRESS THE BALLOON. POSITION THE REFOLDING TOOL AT ONE END OF THE BALLOON. TWIST THE REFOLDING TOOL COUNTERCLOCKWISE AND PUSH DOWN ON THE BALLOON UNTIL THE REFOLDING TOOL TRAVERSES THE ENTIRE LENGTH OF THE BALLOON. ONCE THE BALLOON IS FOLDED, REMOVE THE REFOLDING TOOL AND PROCEED WITH THE PROCEDURE. DISPOSE OF PROPERLY.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04846 | 2023-07-05 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER FELL OUT OF THE PATIENT SPONTANEOUSLY AND THE BALLOON WAS FOUND RUPTURED. URINE WITH BLOOD WAS FOUND IN THE DIAPER. DURING INSERTION, THE CATHETER WAS INSERTED SMOOTHLY AND YELLOW COLORED URINE FLOW WAS OBSERVED DURING INSERTION. DURING CATHETERIZATION, BLEEDING WAS OBSERVED FROM THE INSERTION SITE AND FAINT BLOODY URINE WAS FOUND IN THE TUBING. AT THE TIME OF DISCOVERY, IT WAS FOUND THAT THE INSIDE PART OF THE BALLOON WAS FLIPPED INSIDE OUT AND NO DAMAGE WAS OBSERVED. PRE-TEST CONFIRMED THAT THERE WAS NO DAMAGE TO THE BALLOON. THE BALLOON WAS INSERTED WITHOUT FORCEPS. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 19JUN2023, IT WAS CONFIRMED THAT THE FLIPPING OUT OF THE BALLOON WAS DUE TO BALLOON RUPTURE. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. VISUAL INSPECTION NOTED IRREGULAR BURST WITHOUT MISSING PIECES. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02766 | 2023-07-05 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG* TCS SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770131 |  | IT WAS REPORTED THE CORD HAS BEEN PULLED FROM THE SENSOR/EXPOSING WIRES | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04852 | 2023-07-05 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT WHEN THE USER ACCESSED THE VALVE OF THE FOLEY CATHETER WITH A SYRINGE DURING THE PRETEST, STERILIZED WATER LEAKED FROM THE VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04816 | 2023-07-05 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT ONE DAY AFTER PLACEMENT. | THE REPORTED EVENT WAS CONFIRMED-MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS INFLATED WITH WATER AND OBSERVED WATER LEAKING THRU DRAINAGE EYE. DISSECTED THE CATHETER AND OBSERVED TEAR ON RUBBERIZE LAYER OF INFLATION LUMEN. TEAR WAS EXAMINED UNDER MICROSCOPE AND NOTED TO BE ROUNDED AND NOT SMOOTH AND IDENTIFIED THIS IS RELATED TO MANUFACTURING. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO TORN RUBBERIZE ON RUBBERIZE LAYER THAT MAY LEAD TO LUMEN-TO-LUMEN LEAK. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE(S) THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER SHAPE, CONFIGURATION AND PRINCIPLES: BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY INTENDED USE & EFFECT- EFFICACY: THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. DIRECTIONS FOR USE: 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. CORRECTION- D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02764 | 2023-07-05 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE ¿LOLLIPOP¿ OF THE PICC TIP LOCATION SHOWS GOING IN DIFFERENT DIRECTIONS, EVEN THOUGH X-RAY CONFIRMS TIP LOCATION IS IN THE CORRECT POSITION. ALSO, WHEN USING THIS SENSOR FOR 3CG CONFIRMATION THEY HAVE TROUBLE GETTING CONSISTENT 3CG AND HAVE TO GET X-RAYS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04909 | 2023-07-05 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THERE WAS A STERILE WATER LEAK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04864 | 2023-07-05 | 2023-06-09 | I, n, j, u, r, y | Manufacturer report | BARD® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 365714 | 365714 |  | IT WAS REPORTED, THAT THE FOLEY CATHETER BALLOON BURST WHILE IN USE. THE EVENT OCCURRED DURING PATIENT USE, WITHIN 1-3 HOURS OF USE. THERE WAS NO REPORTED INJURY. HOWEVER, CATHETER BALLOON BURST, WHILE IN USE AND REMNANTS REMAINED IN BLADDER. THUS, PROLONGING ANESTHESIA WAS GIVEN TO RETRIEVE REMNANTS OF THE BURST BALLOON, STATED THAT THE SURGICAL STAFF NOTICED, THAT THE CATHETER WAS ON THE FLOOR UNDER THE PATIENT. PATIENT WAS GIVEN INDWELLING URINARY CATHETER, DURING GYNECOLOGICAL SURGICAL PROCEDURE, BEFORE POSITIONING READY FOR LAPAROSCOPIC PROCEDURE. APPROXIMATELY 2 HOURS INTO THE PROCEDURE, THE CATHETER WAS FOUND ON THE FLOOR UNDER THE PATIENT. CYSTOSCOPY WAS PERFORMED AND CONFIRMED, RETAINED/BURST BALLOON REMNANTS INSIDE BLADDER. GRASPER WAS REQUIRED TO RETRIEVE BROKEN/BURST LATEX BALLOON REMNANTS AND ALL REMNANTS WERE REMOVED. NO FURTHER SURGERY REQUIRED. | THE REPORTED EVENT WAS INCONCLUSIVE, BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT), EXPOSURE TO PETROLATUM BASED PRODUCTS, MECHANICAL FAILURE, OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED, AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. DO NOT USE IF PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON TO BURST. CAUTION: DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. NOTE: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LAWS AND REGULATIONS. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. VALVE TYPE: USE LUER SYRINGE. DO NOT USE NEEDLE. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. RECOMMENDED INFLATION CAPACITIES: 3CC BALLOON: USE 5CC STERILE WATER, 5CC BALLOON: USE 10CC STERILE WATER, 15CC BALLOON: USE 20CC STERILE WATER, 20CC BALLOON: USE 25CC STERILE WATER, 30CC BALLOON: USE 35CC STERILE WATER, 40CC BALLOON: USE 45CC STERILE WATER, 75CC BALLOON: USE 80CC STERILE WATER. DO NOT EXCEED RECOMMENDED CAPACITIES. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04845 | 2023-07-05 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | SILICON FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123618A |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED FOR THE PATIENT AT 15:00 PM ON (B)(6) 18 ML OF SALINE WAS INJECTED INTO THE FOLEY CATHETER BALLOON FOR FIXATION AND THE CATHETERIZATION WAS SMOOTH. ON (B)(6), AT 11:20 AM, IT WAS FOUND THAT THE FOLEY CATHETER HAD FALLEN OUT. UPON CHECKING, IT WAS FOUND THAT THE FOLEY CATHETER HAD SLIPPED OUT DUE TO RUPTURE OF THE BALLOON. A NEW CATHETER WAS REPLACED IMMEDIATELY AND REINSERTED, AND THE CATHETERIZATION WENT SMOOTHLY. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, IT WAS STATED THAT THERE WERE NO MISSING PIECES OF THE BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04903 | 2023-07-05 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT DURING USE IN THE OPERATING ROOM, THE CATHETER FELL OFF. UPON CHECKING, WATER LEAKED FROM THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04899 | 2023-07-05 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HEMATURIA FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 2551H22 |  | IT WAS REPORTED THAT THE FOLEY CATHETER COULD NOT BE REMOVED. THE UROLOGIST CUT OFF THE INFLATION FUNNEL AND AFTER OBSERVATION, THE CATHETER WAS REMOVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04856 | 2023-07-05 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT ON (B)(6) 2023, A FOLEY CATHETER WAS INDWELLED FOR URINE DRAINAGE AND ON THE SAME DAY, THE CATHETER WAS FOUND FELL OUT DUE TO BALLOON BURST. IT WAS STATED THAT THE AFTER REPORTING TO THE DOCTOR, A NEW FOLEY CATHETER WAS USED FOR THE PATIENT, WHICH INCREASE PATIENT'S PHYSICAL PAIN, PSYCHOLOGICAL PRESSURE AND ECONOMIC COST. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04810 | 2023-07-05 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123620A | 123620A |  | IT WAS REPORTED THAT UPON INJECTING STERILE WATER, THE WATER LEAKED FROM THE FOLEY BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE BLISTER PLY (SAC TEARING) CAUSED BY HIGH MECHANICAL STABILITY TIME OF INCOMING LATEX. A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04934 | 2023-07-06 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS ATTEMPTED TO BE REMOVED, THE WATER DID NOT DRAIN OUT, SO IT WAS FORCIBLY PULLED OUT. WHEN THE CATHETER WAS REMOVED, THE BALLOON PORTION WAS DONUT-SHAPED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02796 | 2023-07-06 | 2023-05-10 | I, n, j, u, r, y | Manufacturer report | POWERPORT MRI 8GS, INT. W SP, ATT S/L | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1878001 | 1878001 |  | IT WAS REPORTED THAT TWO DAYS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED RASHES. IT WAS FURTHER REPORTED THAT THE PATIENT ALLEGEDLY EXPERIENCED ITCHING AND BURNING SENSATION. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 12/2025), G3, H6 (METHOD). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02819 | 2023-07-06 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC HF CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F MAXIMAL BARRIER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1385108QD |  | IT WAS REPORTED, WE HAVE HAD A RECENT SCENARIO WHERE THE CLAMP BROKE SO WE HAD TO REPLACE THE LINE. THEY WERE KEEPING THEIR PICC LONG TERM FOR CHEMO. THEY WERE TURNING THE PATIENT AND THE TECH GOT THE PICC CAUGHT SOMEHOW BETWEEN THE BED RAIL. MON 07/03/2023 I AM NOT ABLE TO PROVIDE ANY INFORMATION ON THIS AS THIS WAS REPORTED TO ME FROM ONE OF THE FLOOR NURSES. I AM NOT SURE WHICH PATIENT IT WAS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04933 | 2023-07-06 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE BALLOON COULD BE INFLATED BY INJECTING STERILE WATER DURING THE TEST, BUT IT COULD NOT BE DEFLATED SMOOTHLY DUE TO RESISTANCE WHEN THE WATER WAS DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04938 | 2023-07-06 | 2023-06-08 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED FOLEY TRAY THAT WAS FAULTY. WHEN INFLATING BALLOON, CATHETER FELL OUT PATIENT'S VAGINA. PHYSICIAN EXAMINED CATHETER, TIP OF FOLEY CATHETER WAS MISSING/BROKEN. PHYSICIAN ASKED IF STAFF PRE-INFLATED BALLOON, STAFF SAID THEY DID NOT. PHYSICIAN STATED THEY ALWAYS DO AT MOR SO IT COULD HAVE BEEN PREVENTED. SUPERVISOR PRINTED POLICY STATING BALLOON WAS NOT TO BE INFLATED PRIOR TO USE. SUPERVISOR ASKED PHYSICIAN IF CYSTOSCOPY SHOULD BE PERFORMED FIRST TO LOOK FOR BROKEN PIECE OF CATHETER. PHYSICIAN WISHED TO INSERT SECOND FOLEY CATHETER AND COMPLETED LAPAROSCOPY, PRIOR TO PERFORMING CYSTOSCOPY. CYSTOSCOPY PERFORMED AT END OF CASE, BUT NO CATHETER PIECES FOUND. ROOM AND ASSOCIATED GARBAGE SEARCHED, BUT NOTHING FOUND. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT CAUSED THE REPORTED FAILURE. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED A 2-WAY CATHETER. VISUAL INSPECTION NOTED THAT THE TIP OF THE CATHETER WAS BROKEN OFF. THIS IS OUT OF SPECIFICATION PER INSPECTION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INADEQUATE MATERIAL SELECTION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04944 | 2023-07-06 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE BALLOON WAS BURST, AND THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT JUST AFTER THE PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04945 | 2023-07-06 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT NOT ALL THE STERILE WATER IN THE BALLOON WAS REMOVED. CHECKING THE BALLOON AFTER THE CATHETER WAS REMOVED, A LITTLE OF WATER REMAINED INSIDE THE BALLOON. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE DEVICES MEET THE SPECIFICATION. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04937 | 2023-07-06 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, 10 CC OF WATER INJECTED INTO THE BALLOON WAS UNABLE TO REMOVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02792 | 2023-07-06 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT PRIOR TO A PORT PLACEMENT PROCEDURE IN THE RIGHT INTERNAL JUGULAR VEIN, UPON OPENING OF THE PACKAGE, A LITTLE BUG WAS FOUND INSIDE AND THE PRODUCT WAS CONTAMINATED. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 02/2023). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02814 | 2023-07-06 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F BASIC W/ MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | THE CUSTOMER REPORTED A LEAKING STOPPER FROM THEIR DL PICC LINE. THE RUBBER STOPPER WHERE THE STYLET PASSES THROUGH WAS LEAKING. THEY HAD TO INTERVENE TO ENSURE NOT AIR WAS INTRODUCED INTO THE SYSTEM. IT WAS USED ON A PATIENT AND THERE WAS NO HARM TO THE PATIENT. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: PER THE CUSTOMER, THE EVENT DID NOT INVOLVE AN URGENT LIFE THREATENING MEDICAL SITUATION. THE LEAK WAS DISCOVERED WHEN THE CLINICIAN FLUSHED THE PICC LINE PRIOR TO INSERTION (FLUSHED THE SIDE WITH THE STYLET). THE EVENT DID NOT INTERRUPT TREATMENT. THE PATIENT DID NOT EXPERIENCE ANY ABNORMAL SIGNS OR SYMPTOMS. THERE WAS NO MEDICATION INVOLVED. A STAT LOCK WAS USED TO SECURE THE DEVICE IN THE END. NO PATIENT HARM, INJURY OR NEGATIVE OUTCOME THAT HAS NOT ALREADY BEEN DISCUSSED. CLINICIAN STATES PATIENT INFORMATION IS NOT RELEVANT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04943 | 2023-07-06 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT BODY TEMPERATURE WAS NOT PROPERLY DISPLAYED. THE TEMPERATURE WAS DISPLAYED AS 22°C, WHICH DEVIATED FROM THE AXILLARY TEMPERATURE OF 34.1°C. THE LEAD WIRE WAS REPLACED TWICE, AND THE MONITOR WAS RESTARTED, BUT IT SHOWED ABNORMAL VALUES. AFTER CONSULTATION WITH THE PHYSICIAN, THE CATHETER WAS REPLACED WITH A NEW CATHETER, 5530014LW, WHICH DISPLAYED AN APPROPRIATE TEMPERATURE THAT DID NOT DEVIATE FROM THE AXILLARY TEMPERATURE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04951 | 2023-07-06 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT URINARY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 51812 |  | IT WAS REPORTED THAT THE BD NURSE COMMENTED THAT A COMMUNITY PATIENT VISITED FOR INTERMITTENT CATHETERIZATION ON (B)(6) 2023. WHEN A HYDROSIL ROSE FEMALE 12CH WAS SHOWN, THE CATHETER FELT LIKE THERE WAS NO LUBRICANT, DESPITE WHICH THE WATER SACHET WAS POPPED TO ACTIVATE BUT IT WASN'T WET. THEREFORE, THEY DID NOT USE THE HYDROSIL PRODUCT ON THE PATIENT AND INSTEAD A COLOPLAST PRODUCT WAS USED. THE NURSE HAD A SECOND BOX AT HOME FROM THE SAME BATCH AND THE CATHETER UPON CHECKING ON (B)(6) 2023, ALSO FELT THE SAME. THE NURSE FELT THAT THE LUBRICANT ON THE HYDROSIL ROSE AND GRIPPER CATHETERS HADN'T BEEN THE SAME FOR SOME TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02813 | 2023-07-06 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED SIGN OF EXTRAVASATION. 1 MM PICCLINE BREACH AT MARK 28. ADDITIONAL INFORMATION 06/19/2023. TO DATE, THE PATIENT HAS HAD NO COMPLICATIONS WITH THE ADMINISTRATION OF HIS CHEMOTHERAPY; HOWEVER, HE WAS EXPOSED TO ADDITIONAL LOCAL ANAESTHESIA WHEN A NEW FUNCTIONAL PICCLINE WAS INSERTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02817 | 2023-07-06 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY THE CUSTOMER THAT "DURING PLACEMENT, CLINICIAN EXPERIENCED TUGGING OF THE WIRE WHEN TRYING TO THREAD. PRIOR TO THIS, HE HAD A SEMI AGGRESSIVE PULL BACK OF THE WIRE. WE REMOVED THE ENTIRE CATHETER AND REALIZED THE GUIDE WIRE WAS THREADING THROUGH THE INSTAFLASH NOTCH VERSUS THE NEEDLE TIP OPENING. ADDITIONAL INFORMATION (B)(6) 2023: THIS EVENT TOOK PLACE WITH CUSTOMER AND A TRAINEE. THEY WERE ATTEMPTING TO INSERT A CATHETER FOR MEDICATION ADMINISTRATION. NO URGENT SITUATION OCCURRED, BUT THERE WAS A SECOND VENIPUNCTURE FOR THE PATIENT. NO NEGATIVE IMPACT WAS REPORTED AS ANOTHER LINE WAS PLACED RIGHT AWAY. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. (B)(6) 2023 MICROSCOPIC INSPECTION OF THE GUIDEWIRE CONFIRMED A BREAK WITHIN THE COIL REGION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF AN ACCUCATH GUIDEWIRE PROTRUDING FROM THE FLASH NOTCH WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 20GA X 2.25¿ ACCUCATH ACE PERIPHERAL IV CATHETER ASSEMBLY. THE CATHETER HAD BEEN ADVANCED AND WAS NOT RETURNED FOR EVALUATION. USAGE RESIDUES WERE OBSERVED BOTH ON AND WITHIN THE NEEDLE SHAFT. THE GUIDEWIRE PROTRUDED FROM THE BLOOD FLASH NOTCH. THE COILED TIP APPEARED DEFORMED AND FRACTURED. MICROSCOPIC INSPECTION OF THE GUIDEWIRE CONFIRMED A BREAK WITHIN THE COIL REGION. THE BREAK EXHIBITED SHARPLY DEFINED FRACTURE FEATURES AND A REGION OF INCREASED LUSTER. INSPECTION OF THE NEEDLE BEVEL REVEALED SLIGHT DEFORMATION ALONG THE PROXIMAL EDGE. THE BREAK IN THE GUIDEWIRE COIL RESULTED IN A SHORTER WIRE, WHICH CAUSED IT TO EXTEND FROM THE FLASH NOTCH, RATHER THAN THE NEEDLE TIP. THE FRACTURE FEATURES AND NEEDLE BEVEL DAMAGE WERE CONSISTENT WITH DAMAGE INITIATED BY CONTACT BETWEEN THE NEEDLE AND THE WIRE. BOTH THE EVENT DESCRIPTION AND THE BIOLOGICAL RESIDUES ON THE SAMPLE INDICATED THAT OCCURRED DURING ATTEMPTED DEVICE PLACEMENT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04947 | 2023-07-06 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT AFTER OPERATION, THE CATHETERIZATION WAS PERFORMED. AFTER ONE DAY, THE CATHETER FELL OUT AND WAS FOUND BURST. THEN A NEW CATHETER WAS USED FOR THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE LOW LATEX VISCOSITY, SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. THEREFORE, NO ADDITIONAL ACTION REQUIRED AT THIS TIME. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04936 | 2023-07-06 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175818 | 175818 |  | IT WAS REPORTED THAT THE FOLEY CATHETER HAD FELL OUT. CUSTOMER REPORTED THAT THE THREE DIFFERENT PATIENTS HAD THEIR LUBRISIL FOLEY CATHETERS FALL OUT DESPITE BEING SECURED BY STATLOCKS AND THE CATHETER BALLOONS BEING FILLED. IT WAS STATED THAT TWO OF THE PATIENTS HAD NEW LUBRISIL CATHETERS FITTED, BUT THESE ALSO FELL OUT. THE CUSTOMER WAS UNSURE WHAT MAY HAVE CAUSED THESE ISSUES, BUT THE USER WORKS ON WARD 3, AND THEY HAVE IDENTIFIED THAT THE MAJORITY OF THE STAFF WERE POSITIONING THE STATLOCK INCORRECTLY AND THEREFORE LEAVING TOO MUCH SLACK. IT WAS ALSO STATED THAT THEY HAVE NOW PROVIDED TRAINING ON THIS. AS FAR AS I KNOW, NO OTHER AREAS IN THE HOSPITAL WERE EXPERIENCING THESE ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02818 | 2023-07-06 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY THE CUSTOMER THAT "DURING PLACEMENT, CLINICIAN EXPERIENCED TUGGING OF THE WIRE WHEN TRYING TO THREAD. PRIOR TO THIS, HE HAD A SEMI AGGRESSIVE PULL BACK OF THE WIRE. WE REMOVED THE ENTIRE CATHETER AND REALIZED THE GUIDE WIRE WAS THREADING THROUGH THE INSTAFLASH NOTCH VERSUS THE NEEDLE TIP OPENING. ADDITIONAL INFORMATION (B)(6) 2023: THIS EVENT TOOK PLACE WITH CUSTOMER AND A TRAINEE. THEY WERE ATTEMPTING TO INSERT A CATHETER FOR MEDICATION ADMINISTRATION. NO URGENT SITUATION OCCURRED, BUT THERE WAS A SECOND VENIPUNCTURE FOR THE PATIENT. NO NEGATIVE IMPACT WAS REPORTED AS ANOTHER LINE WAS PLACED RIGHT AWAY. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. (B)(6) 2023 MICROSCOPIC INSPECTION OF THE GUIDEWIRE CONFIRMED A BREAK WITHIN THE COIL REGION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02790 | 2023-07-06 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR POST PORT PLACEMENT VIA THE RIGHT SUBCLAVIAN VEIN, THE INFUSION RATE BY THE PUMP OF THE ANTICANCER AGENT WAS CLEARLY SLOW DURING CHEMOTHERAPY. IT WAS FURTHER REPORTED THAT APPROXIMATELY TEN CENTIMETERS OF RESIDUE WERE FOUND IN THE CUT CATHETER WHEN THE PORT SYSTEM WAS REMOVED. REPORTEDLY, A NEW PORT SYSTEM WAS REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 04/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02794 | 2023-07-06 | 2023-06-13 | I, n, j, u, r, y | Manufacturer report | POWER PORT MRI ISP, 8 FR. GROSHONG, INT W SP, ATTACHABLE SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808560 |  | IT WAS REPORTED THAT THREE MONTHS AND TEN DAYS POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY SEPARATED FROM THE POWER PORT. REPORTEDLY, THE BROKEN PART OF THE CATHETER MIGRATED INSIDE THE HEART. THE PORT WAS REMOVED, BUT THE BROKEN CATHETER WAS NOT YET REMOVED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: ADDITIONAL INFORMATION WAS RECEIVED, AND THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE REPORTABLE AS SERIOUS INJURY. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO AND AN IMAGE WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 09/2023), G3, H6 (DEVICE). H11: B5, H1, H6 (PATIENT). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04935 | 2023-07-06 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT AFTER INFLATING THE BALLOON DURING PRETEST, THE STERILE WATER LEAKED FROM THE BALLOON AREA OF THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02816 | 2023-07-06 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER FULL TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174108 |  | IT WAS REPORTED A PICC LINE INSITU IN PATIENTS ARM WAS FOUND TO BE LEAKING. LINE REMOVED AND HOLE NOTED. 42CM LONG. 3CM OF LINE EXTERNAL. HOLE NOTED AT 4.5CM SERURECATH SECUREMENT DEVICE INSITU. TWO DEVICES WERE RETURNED FOR INVESTIGATION. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04946 | 2023-07-06 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE BALLOON AND THE CATHETER SPONTANEOUSLY FELL OUT FROM THE PATIENT. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED.1018233-2023-00263-1. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02808 | 2023-07-06 | 2023-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 18G X 10 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F318108PT |  | CUSTOMER REPORTED THAT A POWER GLIDE HAD THE SAFETY STUCK INSIDE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02731 | 2023-07-06 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK CV PLUS SLIDING POST, TRICOT HYDROCOLLOID CRESCENT PAD, PIGTAIL PAD | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | CV0220CE |  | IT WAS REPORTED ON (B)(6) 2023, THE PATIENT WAS HOSPITALIZED FOR ADVANCED LUNG CANCER ON THE LEFT SIDE. AT 9:00 AM, THE NURSE PERFORMED DEEP VEIN INFUSION THERAPY FOR THE PATIENT AS ADVISED BY THE DOCTOR. DURING THE MAINTENANCE OF THE DEEP VEIN CATHETER, IT WAS FOUND THAT THE PERIPHERAL INTUBATION CENTRAL VENOUS CATHETER LACKED A CATHETER FIXATOR, SO THE NEW CATHETER WAS IMMEDIATELY REPLACED AND SUCCESSFULLY COMPLETED THE OPERATION WITHOUT CAUSING ADVERSE EFFECTS ON THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02787 | 2023-07-06 | 2023-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | UNK EQUISTREAM SPLIT TIP |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT, THE MATERIAL OF THE CATHETER ALLEGEDLY DEGRADED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE SAMPLE WAS NOT RETURNED FOR EVALUATION. TWO ELECTRONIC PHOTOS WERE PROVIDED FOR REVIEW. THE IMAGES SHOW A PARTIAL VIEW OF EQUISTREAM SPLIT TIP CATHETER PLACED IN A PATIENT. THE CATHETER WAS NOTED TO BE STRETCHED AND TURNED MILKY WHITE. THE PROPERTIES OF THE MATERIAL HAVE BEEN CHANGED IN THE REGION OF WHITENING. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED DEGRADATION AND IDENTIFIED OPACIFICATION, STRETCHED AND MATERIAL PROTRUSION ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02815 | 2023-07-06 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER FULL TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174108 |  | IT WAS REPORTED THE PATIENT WAS RECEIVING AN INFUSION AND FOUND THAT THERE WAS FLUID LEAKING OUT. LINE REMOVED AND FOUND TO HAVE A HOLE IN IT. 42CM LONG 3CM EXTERNAL. HOLE AT 8CM. SECUREACATH SECUREMENT DEVICE USED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02788 | 2023-07-06 | 2023-06-12 | I, n, j, u, r, y | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT TWO DAYS POST DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY HAD A LEAK FROM THE EXTENSION LINE BELOW THE HUB. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY HAD A CRACK IN THE EXTENSION LINE. FURTHERMORE, A SEGMENT OF THE CATHETER WAS ALLEGEDLY MIGRATED TOWARDS THE HEART. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE 19CM GLIDEPATH D/L CATHETER WAS RETURNED AND ONE VIDEO WAS RECEIVED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE AND MICROSCOPIC VISUAL EVALUATIONS WAS PERFORMED. A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED 0.4CM FROM THE DISTAL END OF THE RED LUER AND A COMPLETE COMPOUND CIRCUMFERENTIAL BREAK WAS NOTED AT THE DISTAL END OF THE CATHETER. THE EDGES OF THE DISTAL END OF THE COMPOUND BREAK WERE NOTED TO BE UNEVEN AND ROUGH. AN IN-HOUSE SYRINGE WAS ATTACHED TO THE RED LUER AND WATER WAS NOTED LEAKING FROM THE PARTIAL CIRCUMFERENTIAL BREAK AS WELL AS EXITING FROM THE COMPOUND BREAK. THE VIDEO SHOWS AN IMPLANTED CATHETER WHICH WAS NOTED TO HAVE A CRACK IN THE RED LUER BELOW THE HUB. BLOOD WAS NOTED TO BE LEAKING FROM THE CRACKED AREA OF THE RED LUER EXTENSION LEG. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION AND FLUID LEAK ISSUE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MIGRATION ISSUE AS NO OBJECTIVE EVIDENCE OF MIGRATION WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02791 | 2023-07-06 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602830 | 0602830 |  | IT WAS REPORTED THAT SOMETIMES, POST PORT PLACEMENT, THE DEVICE WAS ALLEGEDLY DIFFICULT TO FLUSH. IT WAS FURTHER REPORTED THAT THE DEVICE WAS ALLEGEDLY UNABLE TO OBTAIN A BLOOD RETURN. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 06/2018). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02793 | 2023-07-06 | 2023-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT SOMETIME POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY HAD A LEAK FROM THE EXTENSION LINE BELOW THE HUB. IT WAS FURTHER REPORTED THAT THE BREAKAGE HAPPENED AT THE EXTENSION LINE BELOW THE HUB. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE GLIDEPATH D/L 23CM CATHETER AND ONE VIDEO WERE RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC AND FUNCTIONAL EVALUATIONS WAS PERFORMED. THE DISTAL CATHETER SEGMENT WAS NOT RETURNED. A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED TO THE DISTAL END OF THE RED LUER WHERE HUB MEETS. IN THE VIDEO, THE CLINICIAN WAS HOLDING AN IMPLANTED CATHETER WHICH WAS NOTED TO HAVE A CRACK IN THE BIFURCATION AREA AND BLOOD WAS NOTED TO BE LEAKING FROM THE CONNECTION AREA OF THE BIFURCATION SITE AND THE EXTENSION LEG. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND FLUID LEAK ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 05/2024), G3, H2, H6 (METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02795 | 2023-07-06 | 2023-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0607555 | 0607555 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CONNECTOR WAS ALLEGEDLY BROKE WHILE INSERTION. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04921 | 2023-07-06 | 2023-06-13 | I, n, j, u, r, y | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123416C |  | IT WAS REPORTED THAT THE BALLOON OF THE FOLEY CATHETER WAS BURST. STATED THAT THERE WERE FOUR CASES OF BALLOON BURST. CURRENTLY, THE HOSPITAL REPLACED THE PRODUCTS WITH UNITS FROM OTHER BATCHES. STATED THAT THE PATIENT HAD TO BE HOSPITALIZED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04919 | 2023-07-06 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THIS WAS THE 3RD FOLEY CATHETER THAT THEY HAD TO CHANGE FOR THE PATIENT AS THE CATHETER GOT BLOCKED. STATED THAT THE THEATRE STOCK CONTROLLER CALLED TO STATED THAT THEY HAD A FAULTY SUPRAPUBIC CATHETER. UPON ARRIVAL, THEY MANAGED TO SPEAK TO DOCTOR AND ADVISED THE CATHETER WAS BLOCKED EVENT OCCURRENCE DATE: (B)(6), 2023). THE URINE WAS NOT FLOWING THROUGH THE CATHETER AND WAS EXITING FROM OTHER AVENUES, THIS WAS INSERTED APPROXIMATELY 3 WEEKS PRIOR. THE DOCTOR STATED THAT THEY TRIED TO FLUSH THE CATHETER WITH A TUMI SYRINGE AND WAS NOT ABLE TO, THEY THEN REMOVED THE BLOCKED CATHETER AND REPLACED WITH ANOTHER SIZE 16 BARD SUPRAPUBIC CATHETER. STATED THAT THIS HAPPENED WITH THE SAME PATIENT TWICE BEFORE (EVENT OCCURRENCE DATE: UNK). | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE POOR SAMPLE CONDITION. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. BASED ON THE ATTACHED PHOTO, OBSERVED THE BALLOON OF CATHETER WAS DEFLATED. THE CONDITIONS OF SAMPLE COULD NOT BEEN IDENTIFIED SINCE THERE ARE ONLY PHOTO OF CATHETER PROVIDED FOR INVESTIGATION. FUNCTIONAL TESTING OF THE REPORTED SAMPLE COULD NOT BE PERFORMED DUE TO ONLY PHOTO PROVIDED. THEREFORE, THE REPORTED EVENT IS INCONCLUSIVE DUE TO POOR SAMPLE CONDITION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PLC FAILURE OR WRONG DURATION SETTING" OR CAN CONTRIBUTED AS USER RELATED DUE TO MISHANDLING THE PRODUCT. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02786 | 2023-07-06 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT ONE MONTH POST PORT PLACEMENT PROCEDURE, THE DEVICE ALLEGEDLY HAD A DIFFICULTY IN ACCESSING THE PORT. IT WAS FURTHER REPORTED THAT BLOOD ASPIRATION WAS HARD AND DIFFICULT TO FLUSH THE DEVICE. REPORTEDLY, THE PATIENT EXPERIENCED HEMATOMA OVER THE PORT. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04923 | 2023-07-06 | 2023-06-13 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 947318 | 175818 |  | IT WAS REPORTED, THAT THE CUSTOMER REPORTED THAT 5 PATIENTS HAD EXPERIENCED URINARY RETENTION DESPITE CATHETERIZATION. THE WORST AFFECTED PATIENT HAD MORE THAN 500ML OF URINE RETAINED. AS PER THE BLADDER SCAN, WHICH WAS PERFORMED. THE CUSTOMER WAS UNSURE WHAT WAS THE CAUSE FOR THE RETENTION, BUT THE USER OBSERVED, THAT THE STAFF HAD BEEN SECURING THE CATHETERS INCORRECTLY WITH STATLOCK. AND THEREFORE, LEAVING TOO MUCH SLACK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00494 | 2023-07-06 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, OF CHRONIC TOTAL OCCLUSION IN THE POPLITEAL ARTERY VIA AN IPSILATERAL ANTEGRADE APPROACH, THE CATHETER ALLEGEDLY SEPARATED APPROXIMATELY AFTER TEN SECONDS OF ACTIVATION. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 06/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04927 | 2023-07-06 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER LEAKED FROM AROUND THE BALLOON DURING PRE-TEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04924 | 2023-07-06 | 2023-06-09 | I, n, j, u, r, y | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 996081 | 996081 |  | IT WAS REPORTED THAT IN AN ONLINE SURVEY A PHYSICIAN STATED A PERFORATION ADVERSE EVENT, DIRECTLY ATTRIBUTABLE TO THE DEVICE 995081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, STERILE . ALSO STATING THAT THE PATIENT WITH THE DEVICE 996081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, W/ INFLATION DEVICE, STERILE PLACEMENT DID RESULT IN PATIENT INJURY REQUIRING MEDICAL OR SURGICAL INTERVENTION BECAUSE FORCED DILATATION. PER FOLLOW UP INFORMATION RECEIVED VIA MAIL ON 14JUN2023, CRITICAL CARE NURSES REPORTED IN AN ONLINE SURVEY THAT A PERFORATION ADVERSE EVENT, DIRECTLY ATTRIBUTABLE TO THE DEVICE 995081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, STERILE. PHYSICIAN ALSO INDICATED 'PERFORATION' WAS PRESENT PRIOR TO DEVICE PLACEMENT. ALSO STATING THAT THE PATIENT WITH THE DEVICE 996081 - X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER, 8MM X 15CM, W/ INFLATION DEVICE, STERILE PLACEMENT DID RESULT IN PATIENT INJURY REQUIRING MEDICAL OR SURGICAL INTERVENTION BECAUSE FORCED DILATATION. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE INADEQUATE PRESSURE APPLIED TO INFLATE BALLOON. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "ONLY A PHYSICIAN WHO HAS AN UNDERSTANDING OF THE CLINICAL APPLICATIONS, TECHNICAL PRINCIPLES AND ASSOCIATED RISKS ASSOCIATED WITH BALLOON DILATION OF THE NEPHROSTOMY TRACT SHOULD USE THIS DEVICE. THE X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER IS A STERILE, SINGLE USE DEVICE. CAREFULLY INSPECT THE CATHETER AND THE STERILE PACKAGING FOR SIGNS OF DAMAGE THAT MAY HAVE OCCURRED DURING SHIPMENT. DO NOT USE THE PRODUCT IF DAMAGE IS EVIDENT. PREPARATION OF THE CATHETER: ALL X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETERS CONTAIN AIR IN THE BALLOON LUMEN. THE AIR MUST BE REMOVED TO ALLOW LIQUID TO FILL THE BALLOON WHEN IT IS INFLATED. REMOVE THE PROTECTIVE SHEATH FROM THE BALLOON. ATTACH THE INFLATION DEVICE TO THE CONNECTOR ON THE BALLOON LUMEN. OPEN THE STOPCOCK, AND DRAW BACK ON THE INFLATION DEVICE TO REMOVE THE AIR FROM THE BALLOON CATHETER. CLOSE THE STOPCOCK, REMOVE THE INFLATION DEVICE, DEPRESS THE PLUNGER TO REMOVE ANY AIR AND REATTACH TO THE BALLOON CATHETER. REPEAT STEPS 3-4 UNTIL ALL AIR IS REMOVED FROM THE BALLOON LUMEN." INFLATING THE BALLOON CATHETER. FILL THE INFLATION DEVICE (EAGLE¿ INFLATION DEVICE) WITH STERILE, LIQUID MEDIA. ATTACH THE INFLATION DEVICE TO THE BALLOON LUMEN. OPEN THE STOPCOCK AND INFLATE THE BALLOON." ONCE DILATION HAS BEEN ATTAINED, ADVANCE THE SHEATH OVER THE BALLOON. WARNING: DO NOT USE AIR OR ANY GASEOUS SUBSTANCES AS A BALLOON INFLATION MEDIA, ALWAYS USE STERILE LIQUID MEDIA. NOTE: THE USE OF AN INFLATION DEVICE WITH A PRESSURE GAUGE IS HIGHLY RECOMMENDED TO MAKE SURE ADEQUATE PRESSURE IS APPLIED AND THE MAXIMUM LIMITS OF THE BALLOON ARE NOT EXCEEDED. NOTE: SLIGHT REDUCTION IN BALLOON PRESSURE MAY BE REQUIRED IF RESISTANCE IS ENCOUNTERED WHEN ADVANCING THE SHEATH OVER THE BALLOON. CAUTION: DO NOT EXCEED THE RECOMMENDED RATED BURST PRESSURE (RBP) FOR THIS DEVICE. BALLOON RUPTURE MAY OCCUR IF THE RBP RATING IS EXCEEDED. PLEASE REFER TO THE DEVICE LABEL FOR THE RECOMMENDED RBP." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04928 | 2023-07-06 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE PATIENT GOT CATHETERIZATION AFTER OPERATION, AFTER INFUSING 20ML LIQUID INTO THE BALLOON, THE BALLOON COULD NOT BE DEFLATED COMPLETELY. THEN A NEW FR16 TRIBBLE-LUMEN CATHETER WAS REPLACED. AS PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 27JUN2023, THERE WAS RESISTANCE WHILE INFLATING THE BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02799 | 2023-07-06 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE BREAST DEPARTMENT OF FACILITY FOUND THAT THERE WERE NO SCALES ON THE PICC CATHETER. NO OTHER INFORMATION WAS PROVIDED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF MISSING DEPTH MARKINGS ON THE CATHETER WAS CONFIRMED AND APPEARS TO BE MANUFACTURING RELATED. A VIDEO OF AN INSERTED 4FR GROSHONG NXT CLEARVUE PICC WAS RETURNED FOR EVALUATION. THE VIDEO SHOWED THE CATHETER BEING ROTATED IN THE PATIENT. NO DEPTH MARKINGS WERE OBSERVED ON THE EXPOSED SECTION OF THE DEVICE. IT APPEARED THAT THE CATHETER WAS NOT PRINTED APPROPRIATELY DURING MANUFACTURING. IN RESPONSE TO THIS COMPLAINT, AN AWARENESS TRAINING WAS PROVIDED TO ALL MANUFACTURING PERSONNEL THAT CONDUCT THIS OPERATION. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02805 | 2023-07-06 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT SOMETIME POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY BROKE. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY LEAKED BLOOD. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE SAMPLE WAS NOT RETURNED FOR EVALUATION. ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS A PARTIAL VIEW OF 19CM GLIDEPATH DIALYSIS CATHETER PLACED IN PATIENT. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND FLUID LEAK AS PARTIAL BREAK WAS NOTED ABOVE THE STRAIN RELIEF AND BLOOD LEAK WAS NOTED AROUND THE BREAK FROM THE PROVIDED PHOTO. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRY DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04922 | 2023-07-06 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT WHEN THE STERILE WATER WAS INJECTED, THE WATER LEAKED FROM THE BALLOON. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS EVALUATED AND OBSERVED THERE WAS A PINHOLE AT THE UPPER PART OF SAC COLLAR THAT ALLOWED WATER TO LEAK OUT. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO BUBBLE DURING DIPPING PROCESS/UNDER SIZE AND LOW SAC WEIGHT-TOO SHORT PICK-UP. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY. [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02807 | 2023-07-06 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY CUSTOMER APPROXIMATELY TWO MONTHS AFTER IMPLANTATION OF THE DEVICE, CATHETER RUPTURE OCCURRED IN THE INTRAVASCULAR PART ABOUT 5 CM FROM THE BIFURCATION OF THE CONNECTORS. THERE WAS EVIDENCE OF ALTERATION OF THE CHARACTERISTICS OF THE DEVICE IN THAT A CRACK OF A FEW MILLIMETERS WAS PRESENT. INCIDENT THAT OCCURRED ON (B)(6) 2023. CONSEQUENCE: SPECIFIC INTERVENTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02809 | 2023-07-06 | 2023-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BROKEN DURING USE. THE DEVICE WAS DAMAGED NEAR THE CLAMP. WITH LEAKAGE OF INFUSED SUBSTANCE, IT BECAME NECESSARY TO REPLACE THE DEVICE. ADDITIONAL INFORMATION RECEIVED 27 JUNE: THE PATIENT OBVIOUSLY HAD TO IMPLANT A NEW DEVICE WITH SOME DIFFICULTY AS SHE IS VERY OBESE. THE CHEMOTHERAPY SHE RECEIVES IS TAXOL AND CARBOPLATIN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04972 | 2023-07-07 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT STERILE WATER WAS DIFFICULT TO REMOVE FROM THE BALLOON DURING A PRETEST. | PER INVESTIGATOR EVALUATION, BD HAS DETERMINED THAT THIS MDR EVENT IS NOT REPORTABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02830 | 2023-07-07 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK PICC PLUS FIXED POST, CRESCENT PAD | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | VPPDFP |  | IT WAS REPORTED BY THE CUSTOMER THAT THE STATLOCK CLIP BROKE OFF WHEN ATTEMPTING TO LATCH. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02832 | 2023-07-07 | 2022-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT EIGHT MONTHS AND TWENTY-SIX DAYS POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER HUB ALLEGEDLY BENT. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY EXPANDED. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 09/2022). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02829 | 2023-07-07 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC HF CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1385108QD |  | IT WAS REPORTED BY "THE CUSTOMER THE CLAMP BROKE SO WE HAD TO REPLACE THE LINE. USED FOR LONG TERM CHEMO." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04973 | 2023-07-07 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE CATHETER WAS DIFFICULT TO INSERT TO THE PATIENT. THERE WAS NO PROBLEM INJECTING WATER INTO THE BALLOON DURING PRETEST. HOWEVER, ATTEMPTING TO INSERT THE CATHETER TO THE PATIENT, IT COULD NOT BE INSERTED. REPLACING WITH ANOTHER CATHETER (7760016LR), IT WAS ABLE TO INSERT AND BE PLACED TO THE PATIENT WITHOUT NO ISSUES. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 22JUN2023, CLARIFIED THAT THE ISSUE WAS NOT ABOUT INFLATION AND IT WAS AN INSERTION ISSUE. WHEN THE USER ATTEMPTED TO INSERT THE CATHETER TO THE PATIENT BUT FAILED. THE CAUSE OF THE ISSUE WAS NOT REPORTED FROM THE CUSTOMER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02831 | 2023-07-07 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 2194108 |  | CUSTOMER REPORTED PICC BLOCKED, CXR TAKEN AND UROKINASE ADMINISTERED AS PER PROTOCOL. UPON FLUSHING PICC LEAKING AT INSERTION SITE. PICC REMOVED. PICC LEAKING AT HUB. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER LEAK WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR POWERPICC SOLO CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED AT THE INTERFACE BETWEEN THE CATHETER AND MOLDED WINGED HUB. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING). AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02821 | 2023-07-07 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT SOMETIME POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY BENT. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY DEFORMED. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 09/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02826 | 2023-07-07 | 2022-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833730 | 5833730 |  | IT WAS REPORTED THAT TWO MONTHS AND TWENTY-THREE DAYS POST A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY APPEARED AS WHITE OPAQUE AND MILKY BELOW THE HUB. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 08/2023) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04963 | 2023-07-07 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING THE BALLOON INFLATION TEST PRIOR TO INSERTION, WATER COULD BE INJECTED BUT COULD NOT BE DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04964 | 2023-07-07 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN THE WATER WAS DRAINED, THERE WAS RESISTANCE AND USE WAS DISCONTINUED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04959 | 2023-07-07 | 2023-06-15 | I, n, j, u, r, y | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 33420 | 33420 |  | IT WAS REPORTED THAT CUSTOMER WAS HOSPITALIZED, UNSURE OF HOW FOLEY CATHETERS FAILED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
3006260740-2023-02825 | 2023-07-07 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT ONE YEAR AND TWENTY-NINE DAYS POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY BENT. THERE WAS NO REPORTED PATIENT INJURY | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 09/2023) . H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
17277248 | 2023-07-07 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDEX ALL-SILICONE FOLEY CATHETER | CATHETER, URETHRAL | C. R. BARD, INC. | 165814 | 165814 |  | INDWELLING SILICONE CATHETER 14F WAS NOT DRAINING AND WAS LEAKING, NEEDED TO BE DISCONTINUED. |  | N | Y |  |  | 2023-07
3006260740-2023-02823 | 2023-07-07 | 2023-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT LONG-TERM HEMODIALYSIS CATHETER, POLYURETHANE, KIT, 14.5F | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5703730 | 5703730 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE PLASTIC AT THE JUNCTION OF THE PUNCTURE NEEDLE ALLEGEDLY CRACKED IN MANY PLACES DURING THE TUBE INSERTION, WHICH CAUSED THE NEEDLE TO FAIL TO DO NEGATIVE PRESSURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 06/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02824 | 2023-07-07 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393270 | 5393270 |  | IT WAS REPORTED THAT TWO YEARS FIVE MONTHS AND FIVE DAYS POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY BENT. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY IN MILKY APPEARANCE BELOW THE HUB. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 08/2021). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04967 | 2023-07-07 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THESE MAGIC3 HYDROPHILIC CATHETER WOULD NOT GO IN ALL THE WAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-04952 | 2023-07-07 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE CONNECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04970 | 2023-07-07 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE STERILE PURIFIED WATER WAS UNABLE TO BE INJECTED INTO THE FOLEY CATHETER. UPON INJECTING THE STERILE PURIFIED WATER TO INFLATE THE BALLOON, THE WATER DID NOT ENTER THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02847 | 2023-07-10 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP IMPLANTABLE PORT, GROSHONG, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | CP00003 | CP00003 |  | IT WAS REPORTED THAT ONE YEAR EIGHT MONTHS AND NINE DAYS POST A PORT PLACEMENT, THE CONTRAST MEDIUM LEAKED SUBCUTANEOUSLY. IT WAS FURTHER REPORTED THAT A HOLE WAS ALLEGEDLY NOTED ON THE CATHETER UPON REMOVAL. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10:D4 (EXPIRY DATE: 12/2022). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04977 | 2023-07-10 | 2023-06-13 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION PRODUCT ALLEGATION. REPRESENTATIVE ADVISED TO STOP USING AND CHECK WITH DOCTOR. PER CUSTOMER VIA PHONE ON 19JUN2023 IT WAS REPORTED THAT THE CUSTOMER SOUGHT MEDICAL INTERVENTION AND WAS PRESCRIBED ANTIBIOTICS TO TREAT THE URINARY TRACT INFECTION. THEY DID NOT RECEIVE CONFIRMATION THAT THE CATHETERS CAUSED THE URINARY TRACT INFECTION BUT DECIDED TO STOP USE FOR NOW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-04976 | 2023-07-10 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS ATTEMPTED TO BE REMOVED, THE WATER DID NOT DRAIN OUT, SO IT WAS FORCIBLY PULLED OUT. WHEN THE CATHETER WAS REMOVED, THE BALLOON PORTION WAS DONUT-SHAPED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
17286513 | 2023-07-10 | 2023-05-22 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. |  |  |  | THE CATHETER WAS LEAKING IN POST-ANESTHESIA CARE UNIT (PACU). WHEN THE PATIENT GOT TO THE FLOOR, THERE WAS A HOLE IN THE Y SILICONE SECTION OF THE CATHETER. |  | N | Y |  |  | 2023-07
1018233-2023-04974 | 2023-07-10 | 2023-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE BALLOON WAS BURST, AND THE CATHETER FELL OUT FOR A PART DURING THE PROCESS OF CONTINUOUS BLADDER IRRIGATION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00506 | 2023-07-10 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | AT75144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE STRING FROM THE PTA BALLOON ALLEGEDLY SNAGGED ON A STENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02843 | 2023-07-10 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | CUSTOMER REPORTED THAT WHEN THE NURSE PLACED THE PICC CATHETER FOR THE PATIENT ON (B)(6), IT WAS FOUND THAT THERE WAS NO SCALE AT THE END OF THE CATHETER DURING THE CATHETER PLACEMENT, AND A NEW CATHETER WAS IMMEDIATELY REPLACED. NO HARM TO PATIENTS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02841 | 2023-07-10 | 2023-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7274118 |  | USER REPORTS THAT THE PICC RUPTURED IN THE INITIAL PORTION, WITH MEDICATION LEAKAGE. THE CATHETER BROKE IN ITS INITIAL PORTION, BEFORE THE CM); IN THE ATTEMPT TO INFUSE THE SERUM EXTRAVASATION WAS OBSERVED. REMOVED THE DEVICE FROM THE PATIENT. RE-PUNCTURE AND DELAY IN MEDICATION. NO OTHER INFORMATION PROVIDED. ADDITIONAL INFORMATION PROVIDED (B)(6) 2023: CUSTOMER ADVISED THERE WAS NO DAMAGE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A CATHETER RUPTURE WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS A 4FR DL POWERPICC SV CATHETER. THE DAMAGE OBSERVED IN THE RETURNED SAMPLE WAS CHARACTERISTIC OF OVER-PRESSURIZATION (BURST) DAMAGE. THIS CAN OCCUR THROUGH THE USE OF SYRINGES SMALLER THAN 10ML, BY FLUSHING AGAINST AN OCCLUSION, OR DUE TO EXCESSIVE FORCE APPLIED DURING INFUSION PROCEDURES. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED ADJACENT TO 0CM MARK ON THE CATHETER BODY. THIS CATHETER DAMAGE WAS TYPICAL OF A BURST, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ 'S' SHAPED SPLIT ¿ TENSILE WEAKNESS AT THE FRACTURE SITE (DUE TO MATERIAL BALLOONING PRIOR TO BURST) ¿ GRANULAR FRACTURE SURFACE TEXTURE (TYPICAL OF TEARING FAILURE MODES) ¿ THE CATHETER MATERIAL WAS VISIBLY LIGHTER IN COLOR AT THE FRACTURE SITE AN OCCLUSION WAS OBSERVED DISTAL TO THE BURST SITE, AND THIS MAY HAVE CONTRIBUTED TO THE FAILURE. FORCEFUL FLUSHING AGAINST AN OCCLUSION CAN CAUSE THIS TYPE OF FAILURE. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02842 | 2023-07-10 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT WHEN PLACING THE PICC CATHETER, THE CUSTOMER OPENED THE EXTERNAL PACKAGE AND FOUND THAT THE STATLOCK IN THE CATHETER WAS DAMAGED. IT WAS SUSPECTED WHETHER IT WAS STERILE. THE CATHETER WAS NOT USED AND NO IMPACT WAS IMPOSED TO THE PATIENT. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02845 | 2023-07-10 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, AIR WAS ALLEGEDLY ASPIRATED WHEN PUNCTURING WITH THE INTRODUCER NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE INTRODUCER NEEDLE WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC AND FUNCTIONAL TESTING WERE PERFORMED. CRACKS WERE NOTED THROUGHOUT THE HUB OF THE INTRODUCER NEEDLE. UPON INFUSION, A LEAK WAS OBSERVED FROM THE INTRODUCER NEEDLE HUB. ASPIRATION OF WATER INTO THE SYRINGE WAS ATTEMPTED AND WAS UNSUCCESSFUL AS WATER DID NOT FULLY RETURN TO THE IN-HOUSE SYRINGE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED AIR OR GAS IN DEVICE AND THE IDENTIFIED FRACTURE ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 01/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02844 | 2023-07-10 | 2022-01-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5743270 | 5743270 |  | IT WAS REPORTED THAT PRIOR TO A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE PACKAGE WAS ALLEGEDLY DEFORMED AND NOT SEALED TIGHTLY. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 03/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02846 | 2023-07-10 | 2023-06-05 | I, n, j, u, r, y | Manufacturer report | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602660 | 0602660 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT ALLEGEDLY BROKE AND TWO PIECES WERE MIGRATED TO HEART AND ONE PIECE WAS MIGRATED TO LIVER. REPORTEDLY, SURGERY WAS PERFORMED TO REMOVE THE PARTS. PORT SYSTEM WAS REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 10/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00505 | 2023-07-10 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U3513076 | U3513076 |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE, THE INTERIOR BOX WAS ALLEGEDLY CRUSHED AT ONE END AND SOAKED IN DETERGENT. THERE WAS NO PATIENT CONTACT. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 10/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02837 | 2023-07-10 | 2022-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT ONE YEAR NINE MONTHS POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY APPEARS OPAQUE RIGHT UNDERNEATH HUB. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY DEFORMED. REPORTEDLY, CATHETER WAS DISCOLORED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 11/2021). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00504 | 2023-07-10 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8088 | VA8088 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE BALLOON ALLEGEDLY FAILED TO INFLATE. IT WAS FURTHER REPORTED THAT THE BALLOON ALLEGEDLY DIFFICULTY REMOVING IT. REPORTEDLY, THE BALLOON WAS BUNCHED UP AT THE TIP. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 12/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00501 | 2023-07-10 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U3575104 | U3575104 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE DEVICE CATHETER SHAFT ALLEGEDLY RUPTURED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 10/2025). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04978 | 2023-07-10 | 2023-06-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE PATIENT HAD BLADDER OR URINARY TRACT INFECTION EVERY 2 MONTHS FOR THE PAST 12 MONTHS AND HAD MEDICATION PRESCRIBED BY DOCTOR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-04979 | 2023-07-10 | 2023-05-25 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0167L20 | 0167L20 |  | IT WAS REPORTED THAT THE PATIENT HAD A THREE-WAY INDWELLING FOLEY CATHETER (20 F) THAT WAS BEING MANUALLY IRRIGATED EVERY 1 TO 2 HOURS. THE ORDERS INDICATED TO IRRIGATE WITH 30-60 ML EVERY 1 TO 2 HOURS. AT 13:00 THE NURSE WAS GOING TO IRRIGATE THE CATHETER WHEN NURSE NOTICED THE CATHETER WAS ALMOST TOTALLY OUT OF THE URETHRA. THE NURSE CONTACTED HOSPITALIST WHO FURTHER REVIEWED THE CATHETER, WHICH WAS NOT TOTALLY INTACT, WITH A MISSING PORTION OF THE CATHETER IN AN AREA OF 3 TO 6 CM FROM THE TIP OF THE CATHETER. THE PORTION OF THE CATHETER MISSING APPEARS TO BE THE BALLOON PORTION OF CATHETER. THERE WAS QUESTION OF WHETHER ANY OF THE MANUAL IRRIGATION WAS UNDERTAKEN IN THE PORT WHERE THE BALLOON WAS FILLED TO 10 ML TO HELP SECURE THE CATHETER PLACEMENT. THE CATHETER WAS REPLACED WITH ANOTHER 20 F, 3-WAY CATHETER. PATIENT REQUIRED CYSTOSCOPY COMPLETED 5/26 - RETAINED FOLEY CATHETER BALLOON WAS VISUALIZED AND GRASPED USING FORCEPS AND REMOVED IN ITS ENTIRETY. PATIENT TOLERATED THE PROCEDURE WELL. SCOPE WAS REMOVED. NO IMMEDIATE COMPLICATIONS IDENTIFIED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04984 | 2023-07-10 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 51812 |  | IT WAS REPORTED THAT THE BD NURSE COMMENTED THAT A COMMUNITY PATIENT VISITED FOR INTERMITTENT CATHETERIZATION ON (B)(6) 2023. WHEN A HYDROSIL ROSE FEMALE 12CH WAS SHOWN, THE CATHETER FELT LIKE THERE WAS NO LUBRICANT, DESPITE WHICH THE WATER SACHET WAS POPPED TO ACTIVATE BUT IT WASN'T WET. THEREFORE, THEY DID NOT USE THE HYDROSIL PRODUCT ON THE PATIENT AND INSTEAD A COLOPLAST PRODUCT WAS USED. THE NURSE HAD A SECOND BOX AT HOME FROM THE SAME BATCH AND THE CATHETER UPON CHECKING ON (B)(6) 2023, ALSO FELT THE SAME. THE NURSE FELT THAT THE LUBRICANT ON THE HYDROSIL ROSE AND GRIPPER CATHETERS HADN'T BEEN THE SAME FOR SOME TIME. | THE REPORTED EVENT IS CONFIRMED - MANUFACTURING RELATED. VISUAL INSPECTION NOTED 56 HYDROSIL ISCS IN CLOSED PACKAGING AND ONE (1) HYDROSIL ISC IN OPEN PACKAGING WITH NO NICKS, BENDS, FLASHES, BUMPS OR SHARPS PRESENT. FURTHER TESTING REQUIRED. FUNCTIONAL EVALUATION NOTED ORGANOLEPTIC TEST WAS CONDUCTED. 13 SAMPLES WERE TESTED. ALL 13 SAMPLES DID NOT HAVE ANY LUBRICIOUS COATING MATERIAL PRESENT. THEREFORE, PRODUCT DOES NOT MEET SPECIFICATIONS. THE POTENTIAL ROOT CAUSE COULD BE MECHANICAL FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04986 | 2023-07-10 | 2023-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PATIENT GOT CATHETERIZATION ON (B)(6) 2023, AND ON THE FIRST DAY AFTER THE OPERATION, IT WAS FOUND THAT THE CATHETER FELL OUT AND THE BALLOON BURST. AS PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, THERE WAS DAMAGE ON THE BALLOON AND THE WATER LEAK FROM IT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05022 | 2023-07-10 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123514A | 123514A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSPECTED POST INCIDENT AND NOTICED THAT THE BALLOONING AREA HAD A TEAR. IT WAS STATED THAT THE UPON INVESTIGATION, IT HAD BEEN CONFIRMED THAT THE DOCTOR INJECTED 10ML WATER FOR INJECTION DURING THE INSERTION AS RECOMMENDED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON BURST. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE PATIENT USE ONLY. DO NOT REUSE AND RESTERILIZE. FOR UROLOGICAL USE ONLY. USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PROTOCOL, THE MAY BE CUT OFF. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05020 | 2023-07-10 | 2023-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT URINARY CATHETER WAS INSERTED ON (B)(6) 2023. ON (B)(6) 2023 PATIENT'S URINARY CATHETER WAS OBSERVED TO BE INTACT AND SECURED DURING ROUTINE NURSING CARE. IT WAS STATED THAT THE NURSE NOTICED THAT THE PATIENT'S URINARY WAS DISLODGED. THE NURSING OFFICER (NO) NOTED THAT THE BALLOON WAS DEFLATED AND TIP OF THE URINARY CATHETER WAS INTACT. IT WAS STATED THE URINARY CATHETER WAS TESTED AND NOTICED THAT WATER WAS LEAKING FROM THE DISTAL END OF THE CATHETER AND THE BALLOON WAS UNABLE TO INFLATE FULLY. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 24JUN2023, IT WAS STATED THAT THE WATER LEAKAGE DID NOT FOUND DUE TO DISLODGE. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, INCORRECT WIRE PRESSING TECHNIQUE, INCORRECT STRIPPING TECHNIQUE ETC AND MIGHT BE CONTRIBUTE BY USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON BURST. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE PATIENT USE ONLY. DO NOT REUSE AND RESTERILIZE. FOR UROLOGICAL USE ONLY. USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PROTOCOL, THE MAY BE CUT OFF. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05017 | 2023-07-10 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER WAS DIFFICULT TO BE INJECTED INTO THE FOLEY CATHETER. WHEN THE BALLOON CUFF TEST WAS PERFORMED, THE BALLOON DID NOT INFLATE AND STERILIZED WATER LEAKED FROM THE CONNECTION BETWEEN THE VALVE AND THE SYRINGE. SINCE THE SAME PHENOMENON HAD OCCURRED IN THE SAME LOT, THE USER ASKED TO INVESTIGATE WHETHER IT WAS CAUSED BY THE LOT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05016 | 2023-07-10 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123516 |  | IT WAS REPORTED THAT THE INFLATION VALVE OF THE FOLEY CATHETER WAS DEFECTIVE. WHEN DISTILLED WATER WAS INJECTED THROUGH THE INFLATION VALVE WITH A SYRINGE, THE BALLOON INFLATED BUT DISTILLED WATER LEAKED OUT OF THE VALVE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO CRACK / CRUSHED VALVE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CONTRAINDICATIONS 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2.APPLICABLE PATIENTS (1) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER" SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05015 | 2023-07-10 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE INFLATION VALVE OF THE FOLEY CATHETER. UPON INJECTING THE STERILE WATER INTO THE BALLOON, WATER LEAKAGE WAS FOUND FROM THE INFLATION VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05009 | 2023-07-10 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 |  | FOL0102 |  | IT WAS STATED, THAT THE PATIENT ADVISED, THAT THE CLIP ON THE FOLEY STATLOCK DID NOT CLIP OVER THE CATHETER. IT WAS STATED, THAT THE CUSTOMER HAD TO CLIP IT OVER THE TUBE FOR BALLOON INFLATION, RATHER THAN WHERE THE LEG BAGS CONNECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE, A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05011 | 2023-07-10 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165816 |  | IT WAS REPORTED THAT THE USER BROUGHT 2 PACKAGES OF PROBES FROM THE SAME LOT WHERE THERE WAS A SMALL HOLE IN THE BALLOON EACH TIME WHICH CAUSED THEM TO FAIL. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 23JUN2023, IT WAS STATED THAT THERE WAS NO BALLOON BURST FOUND IN THE CATHETER. IT WAS STATED THAT ACCORDING TO THE NURSE, THERE WAS NO CRACK, BUT THERE WAS A HOLE. THE PROBES HAD FAILED, REQUIRING A NEW ONE TO BE PLACED EACH TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿IMPERFECTION IN BALLOON DUE TO PROCESS ¿. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02836 | 2023-07-10 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN CV CATHETER, DUAL-LUMEN REPAIR KIT, RED/WHITE, 9F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 |  | 0601700 |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS ONE MONTH AND TWENTY ONE DAYS POST CHRONIC CATHETER PLACEMENT, UPON PERFORMING THE CATHETER REPAIR, A METAL STENT SEGMENT WAS ALLEGEDLY FOUND TO BE PROTRUDING OUT OF THE CATHETER AT AN ANGLE. REPORTEDLY, THE CATHETER WAS REPAIRED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05007 | 2023-07-10 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT THE PATIENT GOT CATHETERIZATION ON JUNE 1 AFTER THE OPERATION. ON JUNE 7, THE FOLEY CATHETER WAS PLANNED TO BE REMOVED BUT DURING THE PROCESS, IT WAS FOUND THAT THE LIQUID IN THE BALLOON COULD NOT BE COMPLETELY DRAINED OUT BY THE SYRINGE. THEREFORE, THE CATHETER WAS DIFFICULT TO BE PULLED OUT. WITH REPEATED ATTEMPTS OF WITHDRAWAL AND USING THE LUBRICANT, THE CATHETER WAS FINALLY REMOVED. NO DAMAGE WAS CAUSED ON THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05001 | 2023-07-10 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE BALLOON HAD BURST. ON JUNE 4TH, THE CATHETER WAS REPLACED DUE TO BAG DAMAGE. ON JUNE 6TH, DURING A DIAPER CHANGE, THE CATHETER WAS FOUND TO BE WITHDRAWN AND UPON CHECKING IT WAS FOUND THAT THE BALLOON WAS DAMAGED. THERE WAS NO EVIDENCE THAT THE CATHETER WAS REMOVED BY THE PATIENT AND ON JUNE 6TH, THE CATHETER WAS REINSERTED BY THE ON-DUTY DOCTOR. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] METHOD FOR USE DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS. PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX.. PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES] BARD I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD SILVER ALLOY COATING. SIZES OF CATHETERS AVAILABLE IN SIZES 8 TO 24 EVERY 2FR FOLEY CATHETER ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER, WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD 10% POVIDONE-IODINE SOLUTION TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS , GLOVES, STATLOCK FOLEY, [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD EZ-LOK SAMPLING PORT OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) PRECAUTIONS FOR USE WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 1, SEVER THE INFLATION FUNNEL OF VALVE. FIG. 10 IF SITUATION WOULDN¿T BE IMPROVED WITH 2, SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. FIG. 11 IF SITUATION WOULDN'T BE IMPROVED WITH, INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. FIG.12 IF SITUATION WOULDN'T BE IMPROVED WITH , INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. FIG.13 B. BALLOON-RUPTURE METHOD INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. FIG. 14 IF SITUATION WOULDN'T BE IMPROVED WITH 1, ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) MALFUNCTION AND ADVERSE EVENTS MALFUNCTION ,CATHETER KINKING, DAMAGE, RUPTURE ,DIFFICULTY OR FAILURE TO REMOVE THE DEVICE , OCCLUSION OF CATHETER INNER LUMENS ,ENCRUSTATION ,ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE ,DEVICE DAMAGE DUE TO INAPPROPRIATE USE. ADVERSE EVENTS ,URINARY-TRACT INFECTION ,HEMORRHAGE, HEMATURIA ,ALLERGY REACTION TO THE DEVICE ,CALCULUS FORMATION , EDEMA , PAIN ,DISCOMFORT , INJURY OF BLADDER OR URETHRAL ,URETHRITIS, URINARY INCONTINENCE ,RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] STORAGE, STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04998 | 2023-07-10 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE PATIENT GOT CATHETERIZATION AFTER OPERATION, AFTER INFUSING 20ML LIQUID INTO THE BALLOON, THE BALLOON COULD NOT BE DEFLATED COMPLETELY. THEN A NEW FR16 TRIBBLE-LUMEN CATHETER WAS REPLACED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, THERE WAS RESISTANCE WHILE INFLATING THE BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINCH LUMEN/COLLAPSE LUMEN/THICK SAC THICKNESS/VALVE DAMAGE/SHORT INFLATION LUMEN. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04992 | 2023-07-10 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | INTS14 | INTS14 |  | IT WAS REPORTED THAT NURSES WERE COMPLAINED ABOUT SURE STEP SILICONE INTERMITTENT TRAY'S. STAFF WERE STATED THAT THERE WERE TIMES WHEN THEY HAVE TROUBLE DETERMINED IF PATIENT STILL HAD URINE IN BLADDER. AS AN ALTERNATIVE THEY WERE USED FOLEY KITS TO STRAIGHT CATCH BECAUSE OF ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04999 | 2023-07-10 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THERE WAS AN ISSUE ON AT LEAST 2 OCCASIONS WHERE THE FOLEY CATHETER BALLOON DID NOT INFLATE PROPERLY. ONE FELL OUT OF A PATIENT, THE OTHER WAS A TEST (NOT USED) WHERE THE BALLOON INFLATED HOWEVER THE BALLOON HAD A HOLE AND LEAKED. PRODUCT WAS DEFORMED, FOLEY CATHETER BALLOON FILLED UP WITH 10CC WAS LOPSIDED. PER FOLLOW UP VIA EMAIL ON (B)(6)2023, IT WAS STATED THAT THE ONLY SAMPLE AVAILABLE WAS THE FOLEY THAT HAD A HOLE IN THE TUBING, NOT THE FOLEY WHERE THE BALLOON WAS LOPSIDED BECAUSE IT WAS USED, THEY COULD NOT SAVE IT. PATIENT WITH LOPSIDED FOLEY WAS OKAY, BUT WAS VERY UNCOMFORTABLE THEREFORE THEY REMOVED THE FOLEY AND THAT WAS WHEN THEY REALIZED THAT THE PRODUCT WAS DEFORMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04994 | 2023-07-10 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | SURESTEP INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | INTP14 | INTP14 |  | IT WAS REPORTED THAT ATTEMPTING TO STRAIGHT CATHETER PATIENT. NURSE TECH OPENED INTERMITTENT STERILE KIT AND STARTED UNFOLDING PAPER OR WRAPPING WHEN IT WAS NOTED THAT CATHETER WAS ON THE OUTSIDE OF THE PAPER OR WRAPPING. COULD NOT OPEN LAST FOLD OF PAPER WITHOUT CONTAMINATING CATHETER OR BREAKING STERILITY (LOT# NGGZ1044). SECOND KIT WAS OBTAINED. SECOND KIT ALSO HAD PART OF CATHETER ON TOP OF PAPER OR NOT IN TRAY (LO# UNK). THE NURSE TECH WAS ABLE TO CAREFULLY, AND SLOWLY, SLIDE THE PAPER OR WRAPPING OPEN WITHOUT CONTAMINATING CATHETER. CONCERNED THAT THERE WERE MORE KITS PACKAGED THIS WAY AND THAT THERE WOULD BE WASTE AND OR PATIENTS STRAIGHT CATHED WITH CONTAMINATED CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04997 | 2023-07-10 | 2023-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT A THREE-LUMEN FOLEY CATHETER WAS INDWELLED IN THE PATIENT DURING THE OPERATION ON (B)(6). THE URINE COLOR WAS CLEAR, AND THE BLADDER WAS CONTINUOUSLY FLUSHED WITH NORMAL SALINE, AND THE FLUSHED LIQUID WAS CLEAR, BUT THERE WAS LEAKAGE AT THE BALLOON, AND THE LIQUID LEAKED OUT BY TOUCHING THE AIR BAG WITH A PAPER TOWEL. IMMEDIATELY REPORTED TO THE DOCTOR, THE DOCTOR PUMPED OUT THE FLUID FROM THE BALLOON AND INJECTED 30ML OF NORMAL SALINE, CLAMPED THE AIR BAG WITH A HEMOSTAT, AND NO LEAKAGE WAS FOUND. AS PER INJURY PERFORMANCE, STATED THAT THE PATIENT TREATMENT WAS AFFECTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT) MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-04989 | 2023-07-10 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER INJECTION WAS DIFFICULT DURING FOLEY CATHETER PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05004 | 2023-07-10 | 2023-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT WAS GIVEN CATHETERIZATION DUE TO DYSURIA DURING THE HOSPITALIZATION. DURING THE WARD ROUND ON THE THIRD DAY OF ADMISSION, THE PATIENT COMPLAINED OF URINE LEAKAGE FROM THE URETHRAL OPENING, AND THE BALLOON OF THE FOLEY CATHETER WAS FOUND TO BE RUPTURED, CAUSING THE CATHETER TO SLIP OFF BY ITSELF. THEREFORE, ANOTHER CATHETER WAS IMMEDIATELY REPLACED FOR CATHETERIZATION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02822 | 2023-07-10 | 2023-06-13 | I, n, j, u, r, y | Manufacturer report | POWER PORT MRI ISP, 8 FR. GROSHONG, INT W SP, ATTACHABLE SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808560 |  | IT WAS REPORTED THAT APPROXIMATELY TWO MONTHS POST PORT PLACEMENT, SUTURES WERE ALLEGEDLY NOT INTACT. IT WAS FURTHER REPORTED THAT THE CATHETER COILED WAS OUT OF VEIN. FURTHERMORE, SUTURES ALLEGEDLY DISLODGED AND THE CHAMBER FLAPPED. REPORTEDLY, THE DEVICE WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO AND A VIDEO WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 05/2024) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02854 | 2023-07-10 | 2023-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERLOC MAX POWER-INJECTIBLE INFUSION SET 20G X 0.75IN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 0132075 |  | IT WAS REPORTED VIA MEDWATCH "THAT TWO SEPARATE INSTANCES OF LEAKS NOTED FROM THE MEDIPORT NEEDLE (SAME LOT NUMBER) TUBING DURING INTRAVENOUS CONTINUOUS INFUSION (IVCI) OF CHEMOTHERAPY MEDICATION. AFTER INSPECTION OF BOTH EVENTS, A CRACK WAS FOUND AT THE BIFURCATION ON THE MEDIPORT TUBING AT THE PROXIMAL SITE. REF REPORT: MW5118289.". ADDITIONAL INFORMATION 6/15/23. CRACK WAS EXTERNAL TO THE PATIENT. MEDIPORT NEEDLES ARE SECURED WITH A TEGADERM DRESSING THAT COMES IN OUR MEDIPORT ACCESS KITS. CHEMO TUBING HAD TO BE CLAMPED AND PAUSED SO THAT PATIENT COULD BE CLEANED UP AND CHEMO BE RE-ADMINISTERED IN A DIFFERENT UNIT. EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. THE CUSTOMER REPORTED TWO SEPARATE DATES OF EVENTS, (B)(6) 2023 AND (B)(6) 2023. THIS REPORT ADDRESSES THE EVENT ON (B)(6) 2023. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02848 | 2023-07-10 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE EXPIRED PORT WAS PLACED KNOWINGLY. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED IMPROPER OR INCORRECT PROCEDURE OR METHOD USED ISSUE SINCE THE LOT EXPIRATION DATE WAS BEFORE THE SAMPLE IMPLANTATION DATE. THE ROOT CAUSE WAS DETERMINED TO BE USER RELATED FACTOR CONTRIBUTED TO THE EVENT. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INSTRUCTIONS FOR USE WAS REVIEWED AND IT WARNS: I. PRIOR TO PLACEMENT: ¿ EXAMINE PACKAGE CAREFULLY BEFORE OPENING TO CONFIRM ITS INTEGRITY AND THAT THE EXPIRATION DATE HAS NOT PASSED. H10: D4 (EXPIRY DATE: 05/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02866 | 2023-07-11 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER W/ SHERLOCK 3CG (TPS) STYLET FULL KIT (4F) (SL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED, CALLED TO SEE PATIENT AS PICC LINE NOT WORKING. R/V BY SENIOR ANAESTHETIC NURSE X2. PICC FLUSHING BUT NOT ASPIRATING. DECISION MADE TO REMOVE PICC LINE, PLACE A PERIPHERAL IV TO MINIMIZE TREATMENT DELAYS AND REBOOK FOR PICC INSERTION ON (B)(6) 2023. ON REMOVAL OF PICC LINE, LINE WAS INSPECTED. LINE NOTED TO BE FRACTURED (APPROX. 5MM IN LENGTH) AT THE DISTAL END. SECURED WITH STATLOCK. INTERNAL LENGTH 38CM. USED FOR IVAB. NO DELAY TO PATIENT'S TREATMENT AS PERIPHERAL IV PLACED AND REBOOKED FOR PICC REINSERTION ON (B)(6) 2023. IT WAS ALSO NOTED THAT THE NEEDLELESS CONNECTOR (CLAVE) HAD BEEN CHANGED ¿ THIS IS STANDARD PRACTICE ON WARDS ¿ CONTACTED BD FOR ADVICE RE: CHANGING THESE CLAVES. IT WAS NOTED THAT WHEN THE NEEDLESS CONNECTOR IS REMOVED, THE PRESSURE INSIDE THE LINE DROPS AND ANY FLUID WITHIN THE LINE LEAKS OUT DISTALLY. AWAITING CLARIFICATION OF THE SAME. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05046 | 2023-07-11 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER WAS NOT BECOMING LUBRICATED ENOUGH, CAME OUT QUITE DRY AND HAD PROBLEMS WITH INSERTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05044 | 2023-07-11 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165818 | 165818 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS GENERALLY DEFECTIVE. PATIENT DOES NOT HAVE THE PACKAGING. PER ADDITIONAL INFORMATION RECEIVED VIA PHONE ON (B)(6) 2023, STATED THAT THE CUSTOMER RECEIVED A REPLACEMENT CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05039 | 2023-07-11 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT PRIOR TO USE, BIFURCATION AREA OF THE FOLEY CATHETER WAS FOUND TORN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05036 | 2023-07-11 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119416M | 119416M |  | IT WAS REPORTED THAT THE CATHETERS WERE NOT DRAINING URINE ONCE INSERTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05037 | 2023-07-11 | 2023-06-10 | I, n, j, u, r, y | Manufacturer report | BARDEX® PEDIATRIC ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165808 | 165808 |  | IT WAS REPORTED TO THE REGISTER NURSE, WENT TO BEDSIDE TO REMOVE FOLEY CATHETER PER ORDERS. BALLOON DEFLATED (5CC). DURING REMOVAL ATTEMPT FOLEY CATHETER BECAME STUCK NEAR THE END OF MEATUS. CALLED CHARGE REGISTER NURSE TO BEDSIDE TO BRING LUBRICANT AND SYRINGE WITH NORMAL SALINE TO ATTEMPT REMOVAL PER UROLOGY¿S RESIDENT RECOMMENDATION. AFTER AN ADDITIONAL ATTEMPT BY BOTH. THIS REGISTER NURSE AND CHARGE REGISTER NURSE, UROLOGY RESIDENT MEDICINE OF DOCTOR CALLED TO BEDSIDE. CATHETER WAS PULLED TO THE TIP OF THE MEATUS. PART OF BALLOON WAS VISIBLE AT MEATUS. ON PALPATION, THERE WAS EVIDENCE OF PARTIALLY FILLED BALLOON. WAS UNABLE TO DRAW BACK ADDITIONAL STERILE WATER FROM BALLOON PORT. BALLOON PORT WAS CUT, BUT BALLOON WAS PALPATED AND NOTED TO STILL BE INFLATED. SMALL GUIDEWIRE WAS ATTEMPTED TO BE PASSED INTO CUT BALLOON PORT. IN THE PROCESS OF ATTEMPTING TO PASS WIRE, CATHETER SLIPPED OUT AND BALLOON WAS EXAMINED, AND WAS DEFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05035 | 2023-07-11 | 2023-06-16 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0119L18 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED INTO THE PATIENT AT 4:00 P.M. IN THE OPERATION ROOM. AT 6:00 P.M. SINCE THE PATIENT COMPLAINED OF LOWER ABDOMINAL PAIN, THE USER PERFORMED AN ULTRASOUND AND 600 ML OF URINE COLLECTION WAS OBSERVED. UPON CHECKING THE FOLEY CATHETER, IT SPONTANEOUSLY FELL OUT OF THE PATIENT. THE STERILE DISTILLED WATER WAS INJECTED INTO THE REMOVED CATHETER, AND WATER WAS FOUND TO BE LEAKING FROM THE BALLOON. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. BASED ON THE EVALUATION, A TEAR WAS FOUND ON THE RUBBERIZED LAYER CAUSING INTERLUMEN LEAKAGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON UNIFORM RUBBERIZED THICKNESS ETC CAUSING TORN RUBBERIZE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "SINGLE USE ONLY [WARNINGS] METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.]. DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.]. APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.]. [CONTRAINDICATIONS] METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.]. DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.]. APPLICABLE PATIENTS: DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX. [SHAPE, CONFIGURATION AND PRINCIPLES]. BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO 26 EVERY 2FR. BALLOON CATHETER. ACCESSORIES: SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) . INTENDED USE & EFFECT- EFFICACY: THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. DIRECTIONS FOR USE: METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. PRECAUTIONS FOR USE: WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. PRECAUTIONS: PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS). EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) BALLOON-RUPTURE METHOD: WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. NON-RUPTURE METHOD: ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1). IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2). IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3). IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4). B. BALLOON-RUPTURE METHOD: INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6). IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7). IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8). MALFUNCTION AND ADVERSE EVENTS: MALFUNCTION. CATHETER KINKING, DAMAGE, RUPTURE. DIFFICULTY OR FAILURE TO REMOVE THE DEVICE. OCCLUSION OF CATHETER INNER LUMENS. ENCRUSTATION. ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE. DEVICE DAMAGE DUE TO INAPPROPRIATE USE. ADVERSE EVENTS: URINARY-TRACT INFECTION. HEMORRHAGE, HEMATURIA. ALLERGY REACTION TO THE DEVICE. CALCULUS FORMATION. EDEMA. PAIN. DISCOMFORT. INJURY OF BLADDER OR URETHRAL. URETHRITIS, URINARY INCONTINENCE. RETAINED BALLOON FRAGMENTS. STORAGE METHOD AND EXPIRATION DATE: STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05032 | 2023-07-11 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® COUDE OLIVE TIP URETHRAL FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 802520 | 802520 |  | IT WAS REPORTED THAT TIPS OF THE INTERMITTENT CATHETERS WERE NOT THE SAME SIZE. SMALL AND LARGER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05034 | 2023-07-11 | 2023-06-09 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT GOT UROSCHESIS AT 16:30 ON JUNE 5, AND THEY COULD NOT URINATE ON THEIR OWN. AFTER THE DOCTOR EVALUATED THE CONDITION, THEY PLACED A URINARY CATHETER FOR THE PATIENT. DURING THE PREPARATION, THE CATHETER WAS DISINFECTED AND LUBRICATED WITH IODINE SOLUTION, AND THE CATHETER WAS INSERTED. THE INTUBATION PROCESS WENT SUCCESSFULLY. 20 ML OF STERILE WATER WAS INJECTED INTO THE BALLOON FOR FIXATION. ON THE SAME DAY, YELLOW URINE COULD BE DRAINED. AFTER TREATMENT EVALUATION, ON JUNE 9TH, AFTER CONSIDERING THAT THE SYMPTOMS OF URINARY RETENTION WERE RELIEVED, THE DOCTOR PLANNED TO REMOVE THE CATHETER, BUT THE STERILE WATER IN THE BALLOON COULD NOT BE DRAWN OUT. UPON ADJUSTING THE ANGLE OF CATHETER AND THE POSITION OF PATIENT, THE LIQUID WAS STILL NOT DRAINED OUT RESULTING IN THE FAILURE OF REMOVING THE CATHETER. ON JUNE 10TH, THE CHIEF DOCTOR TRIED TO REMOVE AGAIN BUT STILL FAILED. ON JUNE 11TH, AFTER CONSULTATION WITH THE SURGEON, THE WATER IN THE BALLOON WAS NOT DRAINED OUT EVEN AFTER REPEATED ATTEMPT. CONSIDERING THE BLOCKAGE OF THE INJECTION PORT OF THE BALLOON, THE INJECTION PORT WAS CUT TO OPEN WITH A STERILE RAZOR BLADE AND THE SYRINGE AND GUIDEWIRE WERE USED TO DRAIN THE LIQUID IN THE BALLOON BUT ALL FAILED. ON JUNE 12TH, B-ULTRASOUND EXAMINATION OF THE BLADDER SHOWED THAT THE BALLOON WAS IN THE BLADDER AREA WITHOUT TWISTING OR FOLDING. THE BALLOON WAS INFLATED AND FIXED. AFTER REPEATED ATTEMPTS BY THE EXPERIENCED DOCTOR AND THE NURSE, THE STERILE WATER IN THE AIR BAG STILL COULD NOT BE DRAWN. DURING THIS PROCEDURE, THE PATIENT URINATED MOSTLY THROUGH THE URETHRAL OPENING AND TO A LESSER EXTENT THROUGH THE URINARY CATHETER. ON JUNE 15TH, UNDER THE GUIDANCE OF SURGICAL EXPERTS AND THE DIRECTOR OF THE RESIDENT PHYSICIAN, A BLADDER PUNCTURE WAS PERFORMED UNDER B-ULTRASOUND. THE USER WANTED TO PUNCTURE THE BALLOON AND PULL OUT THE CATHETER, BUT THE BALLOON HAD GOOD ELASTICITY AND FLEXIBILITY, SEVERAL ATTEMPTS WERE PERFORMED BUT FAILED TO PUNCTURE THE BALLOON .AFTER TALKING WITH THE FAMILY MEMBERS AND INFORMING THE PATIENT OF THE CONDITION OF THE URINARY CATHETER, THE USER WENT TO THE OPERATING ROOM FOR CYSTOTOMY. THE CATHETER WAS SUCCESSFULLY REMOVED, AND A CYSTOSTOMY WAS PERFORMED CONSIDERING THE CONDITION OF THE PATIENT. THE PATIENT'S VITAL SIGNS WERE STILL STABLE AFTER THE OPERATION, AND THE FAMILY MEMBERS COULD NOT UNDERSTAND THIS PROBLEM. AFTER THE OPERATION, THE SURGEON DISSECTED THE URINARY CATHETER TO FIND OUT THE PROBLEM, AND FOUND THAT THE END OF THE BALLOON OF THE CATHETER WAS ALL ADHERED, SO THAT THE LIQUID IN THE BALLOON COULD NOT BE PUMPED OUT AT ALL. ON JUNE 16, THE PATIENT DEVELOPED FEVER. THE EVENT INCREASED THE SUFFERING OF PATIENTS AND IT WAS VERY EASILY TO CAUSE MEDICAL DISPUTES. THE PATIENT WAS STILL UNDER OBSERVATION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE, THIN RUBBERIZE WALL THAT CAUSING COLLAPSE/PINCH INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILIZATION METHOD: STERILE, STERILIZED BY G RADIATION. DO NOT USE IF PACKAGE IS OPENED OR DAMAGED. DISPOSABLE. DESTROY AFTER USE. DO NOT USE THIS PRODUCT IF YOU HAVE A LATEX ALLERGY. STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN SERIES IS COMPOSED OF TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. MATERIAL OF COMPONENTS: CATHETER BODY ¿ NATURAL LATEX. INFLATION CONICAL INTERFACE ¿ NATURAL LATEX. DEFLATION CONICAL INTERFACE ¿ NATURAL LATEX. FLUSHING CONICAL INTERFACE ¿ NATURAL LATEX. BALLOON ¿ NATURAL LATEX. VALVE ¿ POLYPROPYLENE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTIONS: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX-BASED URETHRAL CATHETERS, SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT RE-STERILIZE. FOR UROLOGICAL USE ONLY. PLEASE INSPECT VISUALLY FOR COMPLETENESS OR SURFACE WEAR OF THE PRODUCT BEFORE IT IS USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM DIRECT LIGHT, PREFERABLY STORED IN THE ORIGINAL PACKING BOX. PRECAUTIONS: FEDERAL (USA) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02858 | 2023-07-11 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | CUSTOMER REPORTED THAT AFTER THE PICC PUNCTURE WAS COMPLETED, THE JOINT WAS FOUND TO BE LOOSE, UNABLE TO BE FASTENED, EASILY CAME OFF, AND THERE WERE SAFETY RISKS, SO A NEW JOINT WAS REPLACED. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02857 | 2023-07-11 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIG JAPAN HANG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405J |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. THE PATIENT WAS TREATED AT THE PATIENT¿S HOME. ON (B)(6) 2023, WHEN THE PATIENT CAME BACK TO THE HOSPITAL, THE EXTENSION TUBE WAS FOUND TO BE BROKEN. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BROKEN CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR S/L GROSHONG CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH CATHETER DAMAGE CAUSED BY CONTACT WITH A SHARP-EDGED INSTRUMENT SUCH AS A SCALPEL. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A BREAK WAS OBSERVED APPROXIMATELY MIDWAY ALONG THE EXTENSION TUBE. MICROSCOPIC EXAMINATION OF THE CATHETER BREAK CONFIRMED IT WAS CHARACTERISTIC OF CONTACT WITH A SHARP BLADED INSTRUMENT, AND THE FEATURES OBSERVED ON THE RETURNED SAMPLE WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ THE FRACTURE SURFACES WERE REFLECTIVE IN NATURE AND CONTAINED A STRIATED PATTERN. THIS CAN OCCUR DUE TO PATTERN TRANSFER OF THE SHARPENED EDGE TYPICALLY FOUND ON A SCALPEL-TYPE INSTRUMENT. ¿ SHARPLY FORMED FRACTURE EDGES. ¿ PLANAR SHAPE TO THE FRACTURE SURFACES. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02856 | 2023-07-11 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F 55CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | CK000987 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN CUSTOMER WAS INSERTING A 4FR DUAL LUMEN PROVENA PICC AND KEPT FEELING RESISTANCE. SHE PULLED THE PICC BACK OUT AND NOTICED THAT THE WIRE WAS COMPLETELY BENT INSIDE THE PICC. SHE INSERTED A NEW PICC. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05028 | 2023-07-11 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175816 |  | IT WAS REPORTED THAT THERE WAS AN ISSUE ON AT LEAST 2 OCCASIONS WHERE THE FOLEY CATHETER BALLOON DID NOT INFLATE PROPERLY. ONE FELL OUT OF A PATIENT, THE OTHER WAS A TEST (NOT USED) WHERE THE BALLOON INFLATED HOWEVER THE BALLOON HAD A HOLE AND LEAKED. PRODUCT WAS DEFORMED, FOLEY CATHETER BALLOON FILLED UP WITH 10CC WAS LOPSIDED. PER FOLLOW UP VIA EMAIL ON 28JUN2023, IT WAS STATED THAT THE ONLY SAMPLE AVAILABLE WAS THE FOLEY THAT HAD A HOLE IN THE TUBING, NOT THE FOLEY WHERE THE BALLOON WAS LOPSIDED BECAUSE IT WAS USED, THEY COULD NOT SAVE IT. PATIENT WITH LOPSIDED FOLEY WAS OKAY, BUT WAS VERY UNCOMFORTABLE THEREFORE THEY REMOVED THE FOLEY AND THAT WAS WHEN THEY REALIZED THAT THE PRODUCT WAS DEFORMED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "BALLOON MOLDING". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED, EXCEPT FOR ANY INDIVIDUALLY PACKAGED COMPONENTS WITHIN THE TRAY WHICH ARE NOT LABELED AS STERILE. THESE COMPONENTS ARE NOT TERMINALLY STERILIZED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED. BARD® EZ-LOK® SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK (FIG. 1A) OR SLIP TIP SYRINGES (FIG. 1B). 1. OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. 2. SWAB SURFACE OF BARD® EZ-LOK® SAMPLING PORT WITH ANTISEPTIC WIPE. (FIG. 2) 3. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. (FIG. 3) 4. SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND REMOVE SYRINGE FROM SAMPLE PORT. 5. UNKINK TUBING IF NECESSARY AND TRANSFER URINE SPECIMEN INTO SPECIMEN CUP OR FOLLOW HOSPITAL PROTOCOL. DISCARD SYRINGE ACCORDING TO HOSPITAL PROTOCOL. 6. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB." THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00512 | 2023-07-11 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U3575124 | U3575124 |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE, THE RUBBER BULLET AT THE HUB OF THE BALLOON HAD ALLEGEDLY DISLODGED AFTER OPENING THE PACKAGE. THERE WAS NO PATIENT INVOLVEMENT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 10/2025). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02855 | 2023-07-11 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716000J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, ULTRASOUND EXAMINATION REVEALED THAT AIR ALLEGEDLY LEAKED. IT WAS FURTHER REPORTED THAT THE HUB OF THE INTRODUCER NEEDLE WAS SUSPECTED TO BE CRACKED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT SLIM IMPLANTABLE PORT KIT WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A CRACK WAS OBSERVED AT THE HUB OF THE INTRODUCER NEEDLE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED AIR/ GAS IN DEVICE ISSUE AND THE IDENTIFIED FRACTURE ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 11/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05047 | 2023-07-11 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE URINARY CATHETER WAS INSERTED ON (B)(6) 2023. IT WAS STATED THAT THE ON (B)(6) 2023, DURING ROUTINE NURSING CARE, THE NURSE NOTICED THAT THE URINARY CATHETER SLIPPED OUT OF THE PATIENT. IT WAS STATED THAT THE UPON INSPECTION, IT WAS OBSERVED THAT THE BALLOON APPEARED BURST AND SOME BURST FRAGMENTS WERE SEEN. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE ¿SHORT LATEX DWELL TIME" OR "LATEX DIP SPEED OUT TOO SLOW" AND MIGHT BE CONTRIBUTED BY USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE THE LATEX. VISUALLY INSPECTED THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. -USE LUER TIP SYRINGE TO INFLATE WITH STATE 10ML OF STERILE WATER OR - FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. ALLOW THE PRESSURE WITHIN TO BALLOON FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05049 | 2023-07-11 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316 | 119316 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM NEAR THE CONNECTION BETWEEN INFLATION FUNNEL OF THE FOLEY CATHETER AND SYRINGE DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05038 | 2023-07-11 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT PRIOR TO USE, THE FOLEY CATHETER BALLOON WAS FOUND BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00515 | 2023-07-11 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CRUS6A | CRUS6A |  | IT WAS REPORTED THAT DURING THE RECANALIZATION PROCEDURE, THE CATHETER TIP ALLEGEDLY BROKE WHEN THE DEVICE WAS BEING PULLED BACK OUT OF THE PATIENT. IT WAS FURTHER REPORTED THAT THE TIP OF THE CATHETER WAS UNABLE TO BE RETRIEVED AS THE VESSEL WAS COMPLETELY OCCLUDED WHERE THE TIP WAS BROKEN OFF. REPORTEDLY, THE CATHETER TIP WAS LEFT IN THE PATIENT. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 08/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02874 | 2023-07-11 | 2023-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT DURING THE DRUG DRIP, THE PATIENT WAS OBSERVED TO HAVE FLUID AT THE PICC, THE PICC WAS OPENED AND THE PATIENT'S PICC LINE WAS FOUND TO BE RUPTURED, AND THE PICC LINE WAS NOW REMOVED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02873 | 2023-07-11 | 2023-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICC LINE LEAKING JUST ABOVE WINGED PART. PICC LINE REMOVED AS PER INFUSIONAL SERVICES. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02871 | 2023-07-11 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE BECAUSE THE LUMEN OF THE CATHETER WAS LARGE. A SIMILAR EVENT OCCURRED WHEN ANOTHER NEW CATHETER CONNECTOR WAS CONNECTED TO THE CATHETER. THEREFORE, THE CATHETER WAS CUT AND RECONNECTED, THE PROCEDURE WAS COMPLETED. THERE WAS NO REPORTED PATIENT INJURY. THIS REPORT ADDRESSES THE FIRST CONNECTOR. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02865 | 2023-07-11 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY KINKED NEAR THE CUFF. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 06/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02869 | 2023-07-11 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 6F TRIPLE-LUMEN CATHETER WITH MICROEZ MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3386105 |  | IT WAS REPORTED CATHETER LEAKS IN ONE OF THE PORTS (EXTERNAL SIDE, NOT INSIDE THE PATIENT'S SKIN). MEDICAL STAFF PREFER TO CHANGE THE CATHETER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02870 | 2023-07-11 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000377 |  | IT WAS REPORTED, FRACTURED PICC, WHEN THEY PULLED THE LINE BACK 1 CM THERE WAS A SMALL HOLE, LINE REMOVED. PATIENT HAD 1 CYCLE OF FLOT, SO HAS HAD CYTOTOXIC MEDICATION THROUGH THE LINE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05074 | 2023-07-11 | 2023-06-23 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT DID NOT SPECIFY THE TIMEFRAME OF THE URINARY TRACT INFECTION. PATIENT STATED PATIENT DOES NOT OWE LMS FOR THE BILL, BECAUSE PATIENT HAD URINARY TRACT INFECTION FROM SUPPLIES AND HAD TO THROW THEM OUT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05075 | 2023-07-11 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT DURING A PRETEST, AFTER INJECTING STERILE WATER INTO THE BALLOON, THE WATER COULD NOT BE DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02872 | 2023-07-11 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS ATTEMPTED TO BE INSERTED INTO THE STEM OF THE CATHETER CONNECTOR, THE CONNECTION WAS LOOSE BECAUSE THE LUMEN OF THE CATHETER WAS LARGE. A SIMILAR EVENT OCCURRED WHEN ANOTHER NEW CATHETER CONNECTOR WAS CONNECTED TO THE CATHETER. THEREFORE, THE CATHETER WAS CUT AND RECONNECTED, THE PROCEDURE WAS COMPLETED. THERE WAS NO REPORTED PATIENT INJURY. THIS REPORT ADDRESSES THE SECOND CONNECTOR. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05069 | 2023-07-11 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM NEAR THE BIFURCATION BETWEEN INFLATION FUNNEL AND DRAINAGE FUNNEL OF THE FOLEY CATHETER DURING A PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05067 | 2023-07-11 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE CATHETER WAS DISCONTINUED FROM USE DUE TO A WART-LIKE AREA ON THE SHAFT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05068 | 2023-07-11 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY BALLOON WAS RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05060 | 2023-07-11 | 2023-06-15 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD DEVELOPED A URINARY TRACT INFECTION AND THEIR DOCTOR RECOMMENDED THAT THEY REPLACED KIT. IT WAS UNKNOWN IF THE PUREWICK FEMALE EXTERNAL CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, THE CUSTOMER'S URINE WAS DARKER THAN NORMAL, SO THEIR DOCTOR PRESCRIBED THEM A SULFUR-BASED ANTIBIOTIC FOR 7 DAYS, WHICH DID NOT CLEAR THE ISSUE. CUSTOMER HAD TEST DONE AND THEY DID NOT HAVE A UTI. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05063 | 2023-07-11 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE TWO OF THE CATHETERS WERE KINKED UP AND UNUSABLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05070 | 2023-07-11 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123612A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS NOT DRAINING WELL (EVENT DATE : (B)(6) 2023). RECENTLY THE USER HAD COUNTERED FEW CASES (LABOURING CASES ON EPIDURAL) WITH FOLEY'S CATHETER IN, WHERE THEY COULD SEE SOME URINE IN THE URINE BAG, BUT LATER ON WERE FOUND FULL BLADDER. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 03JUL2023, STATED THAT THE REPRESENTATIVE CHECKED WITH CUSTOMER AND THEY COULD NOT RECALL THE EVENT DATE OF THE PREVIOUS CASES ( EVENT DATE : UNK). | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: SALT ACCUMULATION)/BLOCK DRAINAGE LUMEN/NO DRAINAGE EYE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02883 | 2023-07-12 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER W/ SHERLOCK (TLS) STYLET (4F) (DUAL-LUMEN) (55CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9274108 |  | IT WAS REPORTED PICC BROKEN WHILE IN PATIENT. THERE IS A HOLE ON CATHETER ABOVE WINGS AT INSERTION SITE. PICC RN WAS CHANGING PICC DRESSING WHEN BLEEDING WAS NOTED DURING DRESSING CHANGE. PICC RN NOTICED HOLE IN CATHETER AT THIS POINT AND PATIENT WAS SCHEDULED FOR PICC EXCHANGE. PATIENT TO UNDERGO SECONDARY PROCEDURE FOR PICC EXCHANGE | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05091 | 2023-07-12 | 2023-06-21 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT HAS HAD 3 URINARY TRACT INFECTIONS SINCE LAST ORDER. PATIENT DID SEE DOCTOR ABOUT URINARY TRACT INFECTIONS. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, PATIENT WAS PRESCRIBED NITROFURAN AND WAS STILL CURRENTLY TAKING IT FOR AN ACTIVE URINARY TRACT INFECTION. CUSTOMER WAS STILL CURRENTLY USING THE CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
3006260740-2023-02882 | 2023-07-12 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THE PATIENT CAME TO TREATMENT AREA TO HAVE PICC CHECKED AS STATES IT WAS BLOCKED YESTERDAY WITH THE DISTRICT NURSE. PICC LINE MEASURING OK EXTERNAL LENGTH 19.5 , NO BLOOD RETURN BUT WHEN I FLUSHED IT WITH NACL 0.9%, THE FLUSH CAME OUT OF THE EXIT SITE. PICC LINE THEN REMOVED AS THERE WAS A HOLE IN THE PICC LINE. PICC LINE FLUSHED WHEN REMOVED AND HOLE IN LINE CONFIRMED. NO HARM TO PATIENT CAUSED HOWEVER PATIENTS CHEMOTHERAPY WILL NOW BE DEFERRED UNTIL A NEW PICC LINE IS PLACED AS PER CONSULTANT | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02880 | 2023-07-12 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000375 |  | IT WAS REPORTED A RISE OF INFECTION IN PATIENTS SINCE USING THIS PRODUCT. A SPECIFIC NUMBER OF AFFECTED PATIENT OR DEVICES WAS NOT PROVIDED BY THE COMPLAINANT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05095 | 2023-07-12 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT A NEW FOLEY CATHETER WAS REQUESTED TO REPLACE FOR THE PATIENT ACCORDING TO THE SURGEON'S CLINICAL NEEDS. BUT WHEN THE PACKAGE WAS OPENED, IT WAS FOUND THE PACKAGE WAS UNSEALED THUS IT WAS NOT ABLE TO BE USED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO OPERATOR'S HANDLING METHOD. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05086 | 2023-07-12 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT THE PATIENT HAD A HISTORY OF PROSTATE HYPERPLASIA FOR 10+ YEARS AND HAD LONG TERM DYSURIA, URINATION INTO A DRIP AND STRENUOUS URINATION. PATIENT'S DIAGNOSIS DURING ADMISSION WAS PROSTATIC HYPERPLASIA WITH URINARY RETENTION. ON (B)(6)2023, THE PATIENT COMPLAINED OF DYSURIA AND WAS IMMEDIATELY PROVIDED CATHETERIZATION. ON (B)(6) 2023, THE PATIENT COMPLAINED DIFFICULTY IN URINATION AND THERE WAS NO URINE DRAWN OUT OF THE CATHETER. UPON CHECKING THE CATHETER, A SMALL AMOUNT OF BLOOD WAS SEEN AT THE FRONT TIP AND IT WAS FOUND THAT THE BALLOON WAS RUPTURED THAT COULD HAVE OCCURRED BY THE PATIENT PULLING OUT THE CATHETER. AFTER THE CATHETER REMOVAL, THE PATIENT'S VITAL SIGNS WERE STABLE, NO DISCOMFORT WAS FOUND AND THE PATIENT COULD URINATE BY THEMSELVES. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. BASED ON THE ATTACHED PHOTO, IT WAS OBSERVED THAT THE BALLOON WAS BURST. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. NO FURTHER TESTING COULD BE CONDUCTED DUE TO ONLY PHOTO SAMPLE IS PROVIDED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT/ EXPOSURE TO PETROLATUM BASED PRODUCTS MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02879 | 2023-07-12 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY CUSTOMER THAT THEY HAVE BEEN HAVING MORE ISSUES WITH LEAKING THROUGH THE STOPPER ON THE T PIECE THAT THE WIRE COMES OUT OF DURING PICC INSERTIONS FOR REHQ1797. ADDITIONAL INFORMATION RECEIVED 07 JULY 2023: THE EVENT DID NOT INVOLVE AN URGENT OR LIFE THREATENING MEDICAL SITUATION. THE LEAK WAS DISCOVERED DURING INSERTION PROCEDURE. THERE WAS NOTHING INFUSING AT THE TIME, IT WAS INTERMITTENTLY FLUSHED WITH NORMAL SALINE. THERE WAS NO INTERRUPTION OR DELAY IN THE ADMINISTRATION OF MEDICATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02877 | 2023-07-12 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F MAXIMAL BARRIER TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED BY THE CUSTOMER THAT SOLO PICC SINGLE LUMEN BACK FLOWING WITH BLOOD/LEAKING BLOOD. ADDITIONAL INFORMATION RECEIVED ON 21 JUNE 2023: IT WAS STATED THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. UNABLE TO USE PICC LINE AND INFUSE NEEDED MEDICATION UNTIL REPLACED. PATIENT HAD TO TRAVEL FROM A COMMUNITY HOSPITAL TO RETURN TO FACILITY FOR CORRECTION, DELAYING TREATMENT EVEN LONGER. DEFECTIVE PICC WAS DISCARDED AFTER NEW PICC WAS PLACED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05085 | 2023-07-12 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT DURING PRETEST, IT WAS FOUND THAT THE BALLOON WAS BURST. THE WATER LEAKED FROM BALLOON. THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. THERE WAS NO PATIENT'S URINE STONE AND THE CATHETER WAS NOT PULLED. PER FOLLOW UP INFORMATION RECEIVED VIA EMAIL ON 07JUL2023, THE CATHETER WAS NOT USED ON THE PATIENT AND THE EVENT OCCURRED DURING PRETEST. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS EVALUATED AND OBSERVED THERE WAS A TEAR AT THE TOP AREA OF SAC COLLAR THAT ALLOWED WATER TO LEAK OUT. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO BUBBLE DURING DIPPING PROCESS/UNDER SIZE AND LOW SAC WEIGHT-TOO SHORT PICK-UP. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS, GLOVES, STATLOCK® FOLEY. [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1)MALFUNCTION, CATHETER KINKING, DAMAGE, RUPTURE, DIFFICULTY OR FAILURE TO REMOVE THE DEVICE, OCCLUSION OF CATHETER INNER LUMENS, ENCRUSTATION, ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE DEVICE DAMAGE DUE TO INAPPROPRIATE USE, 2) ADVERSE EVENTS: URINARY-TRACT INFECTION, HEMORRHAGE, HEMATURIA, ALLERGY REACTION TO THE DEVICE, CALCULUS FORMATION, EDEMA, PAIN, DISCOMFORT, INJURY OF BLADDER OR URETHRAL, URETHRITIS, URINARY INCONTINENCE, RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1.STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02884 | 2023-07-12 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT 3FR SL PROV ML | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | SP4153108DB |  | IT WAS REPORTED BY THE CUSTOMER THAT THE INTRODUCER GRAY SHEATH FOR THE MIDLINE WAS COMPRESSED BACK UPON INSERTION THROUGH THE SKIN LIKE AN ACCORDION. ADDITIONAL INFO RECEIVED ON 26-JUN-2023: THIS EVENT DIRECTLY INVOLVED PATIENT AND PLACER. WAS NOTICED WHEN IT WAS DIFFICULT TO INSERT. NO URGENT MEDICAL SITUATION OCCURRED AND NO ADDITIONAL HARM TO THE PATIENT. DURING ROUTINE INSERTION INTRODUCER WOULD NOT THREAD WHEN IT WAS ASSESSED AND FOUND THE ISSUE, NEW KIT OPENED TO GET INTRODUCER, LINE PLACED WITH NO ISSUES. VAT HAD AN EXTRA KIT ON THEIR CART AND WAS ABLE TO OBTAIN A NEW INTRODUCER TO COMPLETE THE PROCEDURE. ADDITIONAL COST WAS ADDED DUE TO HAVING TO OPEN ANOTHER KIT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05087 | 2023-07-12 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE ON (B)(6)2023, THE FOLEY CATHETER WAS FOUND DISLODGED DURING DIAPER CHANGE. THE MICROPORE TAPE SECURING THE TUBING WAS INTACT AND THE BALLOON WAS FULLY DEFLATED. IT WAS STATED THAT THE 10MLS OF WATER WAS INJECTED INTO THE BALLOON PORT AND WAS NOTICED TO BE LEAKING. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02881 | 2023-07-12 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA KIT SINGLE LUMEN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT "THE STOPPER ON THE STYLET LEAKS ANYTIME SALINE IS FLUSHED WITH OUR SINGLE-LUMEN PICC LINES." ADDITIONAL INFORMATION RECEIVED 26 JUNE 2023: "WHILE FLUSHING CATHETER DURING PICC INSERTION PROCEDURE, SALINE SQUIRTS OUT WHERE THE STYLET COMES THROUGH THE STOPPER. THIS HAPPENS WITH SINGLE LUMEN POWER PICC PROVENA KITS THAT HAS THIS RUBBER STOPPER." A SPECIFIC NUMBER OF AFFECTED DEVICES OR PATIENTS WAS NO PROVIDED BY THE COMPLAINANT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05093 | 2023-07-12 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165814 | 165814 |  | IT WAS REPORTED THAT AN INDWELLING 14F FOLEY CATHETER WAS REMOVED IN THE MORNING AS PER THE PHYSICIAN ORDER. 10CC SALINE WAS WITHDRAWN FROM THE INFLATED BALLOON AS PER THE PROTOCOL, BUT WHEN THEY ATTEMPTED TO REMOVE THE CATHETER FROM THE PATIENT, RESISTANCE WAS FELT. ANOTHER ATTEMPT WAS MADE TO DEFLATE THE BALLOON AND DURING THAT TIME, 1CC SALINE WAS FURTHER REMOVED. THE CATHETER WAS EASED THROUGH THE URETHRA WITH A SMALL AMOUNT OF RESISTANCE BUT THE PATIENT DID NOT STATE ANY PAIN. UPON REMOVING THE CATHETER, THE TIP WAS FOUND INTACT WITH THE BALLOON STILL INFLATED WITH APPROXIMATELY 3-5CC REMAINING. AN ATTEMPT WAS MADE TO DEFLATE THAT NOW WHICH WAS REMOVED FROM THE PATIENT, BUT THE BALLOON REMAINED INFLATED. A FAULTY CATHETER WAS SET ASIDE IN A BAG TO SHOW THE CLINIC AND ATTEMPT TO TRACE THE PRODUCT, BUT WAS ACCIDENTLY DISCARDED BY STAFF; REMOVED FROM THE FLOOR AND WAS NOT ABLE TO BE TRACED. PATIENT REPORTED TO STAFF NO DISCOMFORT OR PAIN. DISCLOSURE OCCURRED TO PATIENT AT EACH INTERVAL AND THE STAFF WERE AWARE OF THE INCIDENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02885 | 2023-07-12 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY THE CUSTOMER A LEAKING STOPPER ON A DL PICC LINE. THE PRODUCT WAS USED ON A PATIENT. THERE WAS NO HARM REPORTED. THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. THE LEAK WAS DISCOVERED WHEN CLINICIAN FLUSHED THE LUMEN WITH THE STYLET INSIDE OF IT. THE EVENT DID NOT INTERRUPT TREATMENT OR CAUSE A DELAY IN TREATMENT. THERE WERE NOT ANY SIGNS OF SYMPTOMS OR COMPLICATION THAT THE PATIENT EXPERIENCED. THE EVENT WAS NOT RELATED TO MEDICATIONS BEING INFUSED. THE CATHETER ENDED UP BEING REMOVED SO IT WAS NOT SECURED WITH ANYTHING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02922 | 2023-07-12 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F TL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1395108D3 |  | IT WAS REPORTED BY CUSTOMER THAT TRIPLE LUMEN PICC ARE HAVING ISSUES WITH THE LITTLE RUBBER PIECE WHERE GUIDEWIRE IS PULLED THROUGH AFTER WE CUT THE CATHETER, IT IS NOT ALLOWING US TO FLUSH THE LINE (NO PRESSURE SUPPORT) NOR ARE WE ABLE TO GET BLOOD RETURN FROM IT UNTIL WE PLACED A FINGER OVER THAT RUBBER HUB. THERE WAS NO SIGNIFICANT DELAY IN PROCEDURE AND NO URGENT OR LIFE THREATENING EFFECTS TO THE PATIENT. A SPECIFIC NUMBER OF AFFECTED DEVICES OR PATIENTS WAS NOT PROVIDED BY THE COMPLAINANT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02921 | 2023-07-12 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F TL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY CUSTOMER THAT "THEY HAVE BEEN HAVING ISSUES WITH LEAKING THROUGH THE STOPPER ON THE T PIECE THAT THE WIRE COMES OUT OF DURING PICC INSERTIONS FOR REHR0322." NO OTHER INFORMATION WAS PROVIDED. A SPECIFIC NUMBER OF AFFECTED DEVICES OR PATIENTS WAS NOT PROVIDED BY THE COMPLAINANT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05100 | 2023-07-12 | 2023-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PLACED AFTER THE OPERATION. WHEN THE NURSE PERFORMED PERINEAL CARE FOR THE PATIENT, THE BALLOON WAS FOUND TO BE RUPTURED AND THE FOLEY CATHETER FELL OUT. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THERE WAS NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILIZATION METHOD: STERILE, STERILIZED BY G RADIATION. DO NOT USE IF PACKAGE IS OPENED OR DAMAGED. DISPOSABLE. DESTROY AFTER USE. DO NOT USE THIS PRODUCT IF YOU HAVE A LATEX ALLERGY. STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN SERIES IS COMPOSED OF TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. MATERIAL OF COMPONENTS: CATHETER BODY ¿ NATURAL LATEX INFLATION CONICAL INTERFACE ¿ NATURAL LATEX DEFLATION CONICAL INTERFACE ¿ NATURAL LATEX FLUSHING CONICAL INTERFACE ¿ NATURAL LATEX BALLOON ¿ NATURAL LATEX VALVE ¿ POLYPROPYLENE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTIONS: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX-BASED URETHRAL CATHETERS, SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT RE-STERILIZE. FOR UROLOGICAL USE ONLY. PLEASE INSPECT VISUALLY FOR COMPLETENESS OR SURFACE WEAR OF THE PRODUCT BEFORE IT IS USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM DIRECT LIGHT, PREFERABLY STORED IN THE ORIGINAL PACKING BOX. PRECAUTIONS: FEDERAL (USA) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN." | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02891 | 2023-07-12 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT RN COMMUNICATED PATIENT COME INTO AMBULATORY CARE FROM THE COMMUNITY AS VON CALLED AND SAID PICC WAS LEAKING AROUND SITE. RN BROUGHT THEM IN TO ASSESS THE LINE AND IT SEEMED FINE EXTERNAL AND HAD A BIT OF RESISTANCE FOR FLUSH. RN COMMUNICATED THAT AS HE DID FLUSH, SALINE CAME OUT AROUND AT THE SITE SO COULD NOT LEAVE IT AS IS SO HAD TO REMOVE THE PICC. ONCE PICC WAS REMOVED, RN NOTICED A LARGE SIZED HOLE APPROXIMATELY AT THE 7CM MARK. THE LINE HAD HAD 3CM EXTERNAL LENGTH AFTER INSERTION SO THE HOLE WOULD HAVE BEEN 4CM INTERNAL, WHICH WOULD NOT HAVE BEEN AROUND ANY POINTS OF FLEXION. ADDITIONAL INFORMATION 06/27/2023 THE PATIENT WAS RECEIVING LONG TERM IV ANTIBIOTICS FOR AN INFECTED HIP POST HIP REPLACEMENT. PATIENT HAD THE PICC INSERTED ON (B)(6) 2023. WE RECEIVED A CALL ON (B)(6) 2023, REGARDING PATIENTS' PICC LINE LEAKING AROUND THE INSERTION SITE. THE PATIENT WAS BROUGHT INTO OUR CLINIC FOR ASSESSMENT. PICC INSERTION SITE LOOKED HEALTHY. PICC HAD NO BLOOD RETURN AND HAD RESISTANCE ON FLUSH WITH SALINE LEAKING AROUND SITE UPON FLUSHING. PATIENT DENIED ANY PAIN WITH FLUSH. NO SWELLING OR REDNESS NOTED ON PATIENTS' ARM. PICC WAS REMOVED ON THE SAME DATE AND A HOLE WAS NOTED AT THE 7CM MARK. THE PATIENT HAD 3CM EXTERNAL LENGTH TO HUB THEREFORE HOLE WOULD HAVE BEEN 4CM INTERNAL LENGTH. DUE TO THE PICC BEING REMOVED, PATIENT REQUIRED A PERIPHERAL IV FOR HER IV ANTIBIOTICS FOR THE WEEKEND AS WE WERE NOT ABLE TO REINSERT A PICC UNTIL MONDAY (B)(6) 2023. SHE DID NOT HAVE ANY DELAY IN HER MEDICATION, HOWEVER, DID NEED TO BE PLACED THROUGH THE REINSERTION PROCEDURE DUE TO THE PICC BEING REMOVED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR SINGLE LUMEN POWERPICC SOLO CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED NEAR THE 7 CM DEPTH MARKER. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED ¿ FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR ¿ 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER ¿ OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING) AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02901 | 2023-07-12 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT DURING IMPLANTATION OF PICC IN V.BASÍLICA, IT HAPPENED THAT AT THE END OF THE PROCEDURE, WITH THE CATHETER ALREADY POSITIONED, THERE WAS NO FLOW OR REFLUX AND THE GUIDE WIRE OF THE KIT DID NOT PASS THROUGH THE LUMEN OF THE CATHETER, REQUIRING THE EXCHANGE AND USE OF A NEW CATHETER. ADDITIONAL INFORMATION RECEIVED 06/30/2023: THERE WAS NO HARM TO PATIENT OR CAREGIVER. NO NEED FOR MEDICAL OR SURGICAL INTERVENTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A GUIDEWIRE UNABLE TO PASS THROUGH A CATHETER IS INCONCLUSIVE DUE TO THE GUIDEWIRE NOT BEING RETURNED FOR EVALUATION. ONE 5 FR D/L POWERPICC AND A T-LOCK ASSEMBLY WAS RETURNED FOR EVALUATION. FUNCTIONAL TESTS OF ATTEMPTING TO INFUSE WATER INTO THE RETURNED CATHETER WAS SUCCESSFUL, AND ATTEMPTS TO INSERT THE RETURNED STYLET INTO THE COMPLAINANT CATHETER WAS UNREMARKABLE. AN ATTEMPT TO PASS A NON-COMPLAINANT GUIDEWIRE THROUGH THE RETURNED CATHETER WAS SUCCESSFUL WITH NO OBSERVED RESISTANCE. BECAUSE THE COMPLAINANT GUIDEWIRE WAS NOT RETURNED FOR EVALUATION, THE COMPLAINT OF DIFFICULTY INSERTING A GUIDEWIRE INTO THE CATHETER IS INCONCLUSIVE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02888 | 2023-07-12 | 2023-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER W/ SHERLOCK 3CG (TPS) STYLET MAXIMAL BARRIER TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1173108D |  | IT WAS REPORTED BY CUSTOMER THAT THEY FOUND A LONG BLACK HAIR IN A KIT LAST WEEK. THEY JUST OPENED THE PACK AND IMMEDIATELY SAW IT. THEY DIDN'T USE THE KIT. ADDITIONAL INFORMATION FROM CUSTOMER RESPONSE ON 06-30-2023: NO PATIENT HARM OR URGENT SITUATION JUST USED A NEW KIT. THERE WAS A VERY BRIEF DELAY IN GETTING A NEW KIT. WE ONLY CARRY ONE OF EACH SIZE ON OUR CART SO THERE WAS A BRIEF DELAY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF FOREIGN MATERIAL IS INCONCLUSIVE DUE TO THE STATE OF THE RETURNED SAMPLE. ONE PHOTOGRAPH OF AN OPENED KIT WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED WHAT APPEARS TO BE A HAIR ON THE CSR WRAP OF THE KIT. THE OPENED STATE OF THE PACKAGING MADE IT DIFFICULT TO ASSESS WHEN AND WHERE THE FOREIGN MATERIAL ENTERED THE PACKAGE. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. MANUFACTURING CONTROLS ARE IN PLACE TO MITIGATE THE OCCURRENCE OF THIS TYPE OF FAILURE. THE COMPLAINT OF FOREIGN MATERIAL COULD NOT BE INDEPENDENTLY CONFIRMED WITHOUT EVALUATION OF A SEALED KIT; HOWEVER, THE REPORT OF FOREIGN MATERIAL HAS BEEN DOCUMENTED AND INCLUDED IN COMPLAINT TRENDING. THIS TYPE OF COMPLAINT WILL CONTINUE TO BE MONITORED AS PART OF ONGOING EFFORTS TO IDENTIFY POTENTIAL MANUFACTURING RELATED ISSUES. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02899 | 2023-07-12 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY CUSTOMER THAT STIFFENING WIRE INSIDE PICC WAS BROKEN SO WHEN WIRE WAS PULLED BACK THE END OF IT WHERE IT WAS BROKEN STAYED IN THE END OF THE PICC LINE. ADDITIONAL INFORMATION ON 06-28-2023: THIS EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. ATTEMPTED TO PULL WIRE BACK FROM PICC AND A PIECE STAYED IN THE END WHILE THE REST OF THE WIRE CAME OUT, IT ONLY BROKE IT ONE PIECE AND NO PIECES WERE RETAINED IN THE PATIENT. NO, IMAGING WAS REQUIRED AND NO PATIENT HARM REPORTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF BROKEN STYLET WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 3CG STYLET AND ONE 4FR SL POWERPICC SOLO CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND CONFIRMED TO BE BROKEN. THE FOLLOWING OBSERVATIONS WERE MADE WHICH WERE CONSISTENT WITH THIS FAILURE TYPE: ¿ MICROSCOPIC EXAMINATION REVEALED DEFORMATION OF THE STYLET TUBING AT MAGNET EDGE(S). ¿ MICROSCOPIC EXAMINATION REVEALED LOCALIZED DELAMINATION OF THE STYLET TUBING AT THE DAMAGE SITE(S), INDICATIVE OF TUBING KINKING/FOLDING. ¿ KINKING OF THE SURROUNDING STYLET TUBING, LOCATED AT MAGNET INTERFACE(S). ¿ THIS ADDITIONAL STYLET KINKING RESULTED IN A GENERAL CURL IN A SIMILAR DIRECTION. MEASUREMENTS OF THE STYLET TUBING, STYLET CORE WIRE, AND MAGNET WERE TAKEN AND CONFIRMED TO BE WITHIN SPECIFICATIONS. ALTHOUGH THE BREAK IN THE WIRE COULD BE CONFIRMED, THE ROOT CAUSE COULD NOT BE DETERMINED THROUGH EXAMINATION OF THE RETURNED SAMPLE. THE OBSERVED FEATURES WERE INDICATIVE OF CIRCUMSTANCES WHICH INCLUDED STYLET KINKING, LIKELY DURING THE INSERTION PROCESS. IT APPEARED THAT ADDITIONAL UNIDENTIFIED FACTORS MAY HAVE ALSO CONTRIBUTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02920 | 2023-07-12 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED BY THE CUSTOMER "THE PATIENT WAS ADMITTED TO THE HOSPITAL ON 2023.6.15 AFTER DIABETIC KETOSIS AS A SERIOUS CONDITION, TREATMENT: INSULIN, DOPAMINE, BUMETANIDE AND OTHER DRUGS MICRO-PUMP PUMPING, SO THE ESTABLISHMENT OF PICC INTRAVENOUS THERAPY FOR RESUSCITATION TREATMENT. TO THE PATIENT'S RIGHT UPPER ARM TO BE VENOUS ASEPTIC FORMAL OPERATION UNDER THE PLACEMENT OF THE TUBE SMOOTHLY, INSERTION DEPTH 34CM, CATHETER EXPOSED 2CM, THE TIP OF THE CATHETER IS LOCATED IN THE FOURTH INTERCOSTAL SPACE OF THE SUPERIOR VENA CAVA, SYRINGE TO DRAW BACK BLOOD WELL, SYRINGE PUSH SALINE USUALLY, BUT EXTERNAL INFUSION DEVICE GRAVITY DRIP INFUSION RATE 0 DROPS / MIN, OBSTRUCTION IS NOT SMOOTH, INFUSION PUMP PRESSURE INFUSION CATHETER USUALLY." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05098 | 2023-07-12 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123614A | 123614A |  | IT WAS REPORTED THAT THE PORT OF THE FOLEY CATHETER JUNCTION SNAPPED. PER THE FOLLOWUP INFORMATION RECEIVED ON 20JUN2023, IT WAS CONFIRMED THAT THE BASED ON THE PHOTO ATTACHED IN TRACKWISE, THE PORT OF FOLEY CATHETER JUNCTION SNAPPED WAS SILICONE CATHETER, NOT LATEX CATHETER. THE LATEX FOLEY CATHETER SEEMS LIKE THE BALLOON BURST OR RUPTURE. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. BASED ON THE ATTACHED PHOTO, IT WAS OBSERVED THAT THERE WAS IRREGULAR BURST BALLOON. HOWEVER, NO FUNCTIONAL TESTING COULD NOT BE PERFORMED SINCE THERE ARE ONLY PHOTO PROVIDED FOR INVESTIGATION. THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC AND ALSO COULD BE CONTRIBUTED BY THE USER ITSELF (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARDIA FOLEY CATHETER CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON TO BURST. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. SINGLE PATIENT USE ONLY. DO NOT REUSE. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SYRINGE. DO NOT USE NEEDLE. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. DO NOT USE IF PACKAGE IS OPENED OR DAMAGED. RECOMMENDED INFLATION CAPACITIES. 5CC BALLOON: USE 10CC STERILE WATER. 30CC BALLOON: USE 35CC STERILE WATER. DO NOT EXCEED RECOMMENDED CAPACITIES. NOTE: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LAWS AND REGULATIONS. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02910 | 2023-07-12 | 2023-06-26 | I, n, j, u, r, y | Manufacturer report | BARD TRIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. TRIALYSIS/HD CATHETER, TRIPLE LUMEN IN NECK. IT WAS REPORTED THIS OCCURRED WITH 2 DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02889 | 2023-07-12 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BEFORE INSERTION OF THE PICC LINE WHILE FILLING WITH SALINE, LEAKING OCCURRED FROM T-LOCK; LEAKAGE FROM THE YELLOW STOPPER WHEN THEY PREPARED THE PICC LINE. NO OTHER INFORMATION WAS PROVIDED. THERE WERE TWO SAMPLES RETURNED FOR INVESTIGATION, THIS FILE REPRESENTS THE SECOND RETURNED SAMPLE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02917 | 2023-07-12 | 2023-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT 4FR PEEL AWAY SHEATH UNABLE TO THREAD THROUGH TISSUE DESPITE NO ISSUE WITH INNER DILATOR. NOTED TO HAVE RIDGES ON END OF SHEATH. USED 3.5FR COOK PICC PEEL AWAY. NO PROBLEMS WITH PRODUCT. REMAINDER OF PROCEDURE COMPLETED WITHOUT INCIDENT. (B)(6)2023 ADDITIONAL INFORMATION PROVIDED: IT WAS REPORTED BY THE CUSTOMER "NO PRIOR MEDICAL HISTORY; DIAGNOSIS NEW CEREBRAL MASS. PICC NEEDED FOR LAB MONITORING AND CHEMOTHERAPY. EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING SITUATION. NO NEGATIVE OUTCOME TO PATIENT. PROCEDURE TOOK LONGER TO COMPLETE AS TWO SEPARATE KITS NEEDED TO BE OPENED IN ORDER TO GET PEEL AWAY SHEATH NEEDED FOR PROCEDURE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05114 | 2023-07-12 | 2023-05-04 | I, n, j, u, r, y | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V24S |  | IT WAS REPORTED THAT THERE WAS ANOTHER INSTANCE OF THE FOLEY CATHETER BALLOON NOT DEFLATING ON THIS PATIENT WITH A VISIT TO THE EMERGENCY ROOM AND HAD TO BREAK THE BALLOON WITH A WIRE. THE CUSTOMER WAS PART OF THE HOME CARE PROGRAM (SELF MANAGED CARE). THE CATHETER HAD BLOCKED AT CUSTOMER'S HOME AND FAMILY WAS ASSISTING IN FLUSHING. THE CATHETER HAD FAILED FROM FLUSHING AND FAMILY ATTEMPTED TO CHANGE CATHETER. THE CATHETER WAS UNABLE TO DEFLATE. THEY HAD TO GO TO EMERGENCY AT HOSPITAL AND PHYSICIAN HAD VALVE CUT BUT NO FURTHER DEFLATION OF CATHETER. THE CATHETER HAD TO BE PIERCED WITH A WIRE AND PHYSICIAN PULLED CATHETER BALLOON VIA STOMA. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02907 | 2023-07-12 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | IT WAS REPORTED THAT AFTER THE INSTALLATION OF THE EXTENSION TUBE WAS COMPLETED, LIQUIDS LEAKED FROM THE TRANSPARENT EXTENSION TUBE DURING CATHETER FLUSHING. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02892 | 2023-07-12 | 2023-06-18 | I, n, j, u, r, y | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (20G) (8CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120080T |  | IT WAS REPORTED BY CUSTOMER THAT WHILE ATTEMPTED TO PLACE MIDLINE TO LEFT CEPHALIC VESSEL. VESSEL ACCESSED WITHOUT DIFFICULTY USING 20G/8CM POWERGLIDE PRO MIDLINE, GUIDEWIRE THREADED WITHOUT DIFFICULTY BUT CATHETER MET RESISTANCE W / THREADING. WHEN MIDLINE DEVICE PULLED, NOTED CATHETER NON-INTACT AND GUIDEWIRE BENT. NON-DIAGNOSTIC ULTRASOUND OF LUE NOTED NO CATHETER SEEN IN LT CEPHALIC VESSEL, SHADOW NOTED IN SUBCUTANEOUS TISSUES NEAR LT CEPHALIC VESSEL. ADDITIONAL INFO FROM CUSTOMER: (26/6/2023). PATIENT HAS HISTORY OF MULTIPLE CO-MORBIDITIES TO INCLUDE SEPSIS, ATRIAL FIB, ENCEPHALOPATHY/DEMENTIA, ANEMIA, RESPIRATORY FAILURE, MALNUTRITION, ASPIRATION PNEUMONIA. A SAME DAY SURGERY WAS REQUIRED, SURGICAL INTERVENTION TO REMOVE THE SHEARED PIECE OF CATHETER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00516 | 2023-07-12 | 2023-05-14 | I, n, j, u, r, y | Manufacturer report | DOT PROV 4F DL MAX TLS | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S9274108D |  | IT WAS REPORTED BY CUSTOMER THAT APPROXIMATELY 3 CM OF THE TIP OF THE GUIDEWIRE BROKE OFF DURING A PICC PLACEMENT AND REQUIRED INTERVENTIONAL RADIOLOGY TO RETRIEVE THE BROKEN PORTION OF THE WIRE. ADDITIONAL INFO FROM CUSTOMER: 10-JULY-2023: CUSTOMER REPORTED "BROKEN GUIDEWIRE LEFT IN PATIENT THAT MIGRATED TO PULMONARY ARTERY. PATIENT TAKEN TO INTERVENTIONAL RADIOLOGY WHEN RADIOLOGIST ATTEMPTED TO REMOVE BROKEN WIRE. IN THE PROCESS OF ATTEMPTING RETRIEVAL, PATIENT'S AICD LEAD WAS DISLODGED CAUSING, PATIENT TO BECOME HEMODYNAMICALLY UNSTABLE. AT THIS TIME, MEDICATION WAS ADMINISTERED TO IMPROVE HEART RATE AND BLOOD PRESSURE. IN ADDITION, A TEMPORARY CARDIAC PACEMAKER HAD TO BE PLACED EMERGENTLY. A SECOND ATTEMPT WAS MADE 2 DAYS LATER AND BROKEN WIRE WAS SUCCESSFULLY RETRIEVED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02925 | 2023-07-12 | 2023-03-13 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE PATIENT 1 WAS ADMITTED TO THE TUMOR DEPARTMENT OF THIS HOSPITAL DUE TO LUNG CANCER AND HAD A CATHETER INSERTED IN (B)(6) 2022. THE CATHETER WAS MAINTAINED AT ANOTHER HOSPITAL AFTERWARDS. ON (B)(6) 2023, CHEST X-RAY SHOWED THAT THE CATHETER WAS RUPTURED AND THE RUPTURED PORTION WAS REMOVED AT THE VASCULAR SURGERY DEPARTMENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02926 | 2023-07-12 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | CUSTOMER REPORTED THAT ON (B)(6) 2022, THE PATIENT UNDERWENT PICC CATHETERIZATION FOR TREATMENT. THE PROCEDURE WAS SUCCESSFUL, AND SHE CAME TO THE HOSPITAL REPEATEDLY FOR CHEMOTHERAPY AND OTHER INFUSION TREATMENTS FROM (B)(6) 2022 TO (B)(6) 2023. DURING THE PERIOD, THE PATIENT PERFORMED CATHETER MAINTENANCE ON TIME WITHOUT DISCOMFORT. ON (B)(6) WHEN I CAME TO THE HOSPITAL ON TIME FOR CATHETER MAINTENANCE, THE SPECIALIST NURSE FOUND THAT THE PATIENT'S CATHETER WAS DAMAGED AT ABOUT 40CM AFTER INVESTIGATION AND INVESTIGATION. AFTER COMMUNICATING WITH THE DOCTOR IN CHARGE OF THE PATIENT, I EXPLAINED TO THE PATIENT THAT THE PATIENT STILL NEEDED TO UNDERGO INTRAVENOUS CHEMOTHERAPY. WITH THE CONSENT OF THE PATIENT AND HIS FAMILY, THE CATHETER WAS TRIMMED AT THE CATHETER RUPTURE AND KEPT. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02929 | 2023-07-12 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THE CATHETER WAS PLACED ON (B)(6) THE PATIENT WAS NOT AWARE OF WHEN THE FASTENERS OF THE EXTENSION TUBE CAME OFF. UPON AWARENESS, THE PATIENT IMMEDIATELY WENT BACK TO HOSPITAL FOR TREATMENT. SINCE IT HAD BEEN EXPOSED TO THE AIR, THE CLINICIAN REMOVED THE CATHETER AFTER EVALUATION AND REPLACED IT WITH A NEW ONE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02928 | 2023-07-12 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG 3F SINGLE-LUMEN PICC BASIC TRAY WITH MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7755305CE |  | IT WAS REPORTED BY THE CUSTOMER THAT A PATIENT'S MOTHER CALLS TO INFORM US THAT SHE HAS FOUND THE PICCLINE TUBING ON THE FLOOR IN THE ROOM. IN FACT, IT LOOKED AS IF IT HAD BEEN CUT OFF ABOVE THE GREEN WING. THE BANDAGE IS WELL OCCLUDED, AND THE STITCHES ARE ALSO WELL IN PLACE. WE REMOVED THE PICCLINE IMMEDIATELY TO AVOID ANY FURTHER RISK TO THE PATIENT. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05117 | 2023-07-12 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT ONLY ABOUT 5 CC OF STERILE WATER WAS RETURNED FROM THE DURING THE PRETEST, AND ANOTHER 5 CC REMAINED IN THE FOLEY CATHETER BALLOON. IT WAS STATED THAT THE ABOUT A 10 ML OF WATER WAS ASPIRATED BY SYRINGE. IT WAS NOTED THAT USUALLY, BALLOON COULD BE DEFLATED AROUND 5 SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02905 | 2023-07-12 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO HF CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1395108QD |  | IT WAS REPORTED, TRIPLE LUMEN PICC LINE WAS PLACED TODAY AND WHEN REMOVING THE STYLE IT WAS CURLED. ADDITIONAL INFORMATION 5/22/2023: PATIENT WAS NOT COMPROMISED, AND NO INJURY OCCURRED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED STYLET IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. ONE 3CG STYLET THREADED THROUGH A T-LOCK WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED A PORTION OF THE STYLET DISTAL TO THE T-LOCK WAS CURLED IN MULTIPLE LOOSE COILS. THE STYLET WAS ALSO OBSERVED TO BE BENT SLIGHTLY NEAR ITS DISTAL END. NO BREAKS WERE OBSERVED IN THE RETURNED STYLET. THE SHAPE, LOCATION, AND NATURE OF THE DAMAGE OBSERVED ON THE STYLET, AS WELL AS THE EVENT DESCRIPTION PROVIDED BY THE COMPLAINANT AND SIMILARITIES OBSERVED DURING PAST ATTEMPTS TO REPLICATE THIS FAILURE MODE, SUGGEST THE STYLET WAS PULLED AGAINST THE LUER HUB OR ANOTHER HARD SURFACE DURING REMOVAL OF THE STYLET RESULTING IN THE OBSERVED DAMAGE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02886 | 2023-07-12 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL CATHETER BASIC TRAY (W/ MICROINTRODUCER, NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | IT WAS REPORTED AIR ENTERS THROUGH THE INSERTION STYLET DURING THE PICC INSERTION WHEN ASPIRATING TO CHECK THE BLOOD RETURN, AND WHEN FLUSHING SALINE DROPS THROUGH THE INSERTION STYLET. THIS HAPPENS SINCE THE RUBBER CAP OF THE STYLET HAS CHANGED. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02909 | 2023-07-12 | 2023-06-26 | I, n, j, u, r, y | Manufacturer report | BARD TRIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY CUSTOMER THAT REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. TRIALYSIS/HD CATHETER, TRIPLE LUMEN IN NECK. IT WAS REPORTED THIS OCCURRED WITH 2 DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05094 | 2023-07-12 | 2023-06-12 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WAS HAVING A REACTION TO THE PUREWICK FEMALE EXTERNAL CATHETER AND FELT A BURNING SENSATION. THEY DID NOT TAKE THE PATIENT TO THE HOSPITAL AND STARTED BEING ABLE TO VOID AND NO LONGER NEEDS THE SYSTEM. PER FOLLOW UP VIA PHONE ON 20JUN2023, CUSTOMER STOPPED USING THE DEVICE DUE TO IT CAUSING A BURNING SENSATION AND WAS CURRENTLY BEING TREATED WITH ANTIBIOTICS BUT WAS DIAGNOSED WITH AN INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | N | LAY USER/PATIENT |  | 2023-07
3006260740-2023-02902 | 2023-07-12 | 2023-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108D |  | IT WAS REPORTED BY CUSTOMER THAT WHILE ATTEMPTING TO REMOVE THE GUIDEWIRE THROUGH THE NEEDLE, THE GUIDEWIRE BECAME LODGED WITHIN THE VESSEL. THE NURSE WAS ABLE TO REMOVE THE NEEDLE AND GUIDEWIRE ALTOGETHER, HOWEVER THE GUIDEWIRE APPEARS TO HAVE UNRAVELED APPROXIMATELY 15-20 CM FROM THE TIP. ADDITIONAL INFORMATION RECEIVED (B)(6)2023: WHILE ATTEMPTING TO PLACE A MIDLINE CATHETER, THE NURSE ATTEMPTED TO REMOVE THE GUIDEWIRE THROUGH THE NEEDLE AND THE GUIDEWIRE BECAME LODGED WITHIN THE VESSEL. THE NURSE WAS ABLE TO REMOVE THE GUIDEWIRE WITH THE NEEDLE AT THE SAME TIME, HOWEVER THE GUIDEWIRE APPEARS TO HAVE UNRAVELED APPROXIMATELY 15-20CM FROM THE TIP. THE NURSE STOPPED THE PROCEDURE IMMEDIATELY AND SEGREGATED THE DEVICE. THE PHYSICIAN WAS NOTIFIED OF THE ISSUE AND AN X-RAY WAS ORDERED TO CHECK FOR ANY RETAINED PIECES OF THE DEVICE. THE X-RAY SHOWED NO EVIDENCE OF A FOREIGN BODY, AND THE PATIENT REQUIRED NO OTHER TREATMENT RELATED TO THE EVENT. THERE WAS NO ADVERSE OUTCOME FOR THIS PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02896 | 2023-07-12 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER 5F BASIC TRAY (DUAL-LUMEN) (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6275118 |  | IT WAS REPORTED AIR ENTERS THROUGH THE INSERTION STYLET DURING THE PICC INSERTION WHEN ASPIRATING TO CHECK THE BLOOD RETURN, AND WHEN FLUSHING SALINE DROPS THROUGH THE INSERTION STYLET. THIS HAPPENS SINCE THE RUBBER CAP OF THE STYLET HAS CHANGED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02903 | 2023-07-12 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THE PATIENT UNDERWENT PERIPHERAL CENTRAL VEIN CATHETER INSERTION ON (B)(6) 2023. WHEN THE DOCTOR WITHDREW THE PUNCTURE NEEDLE AFTER SUCCESSFUL PUNCTURE, THE PUNCTURE NEEDLE WAS STUCK BY THE END OF THE GUIDE WIRE, RESULTING IN THE PUNCTURE NEEDLE FAILING TO WITHDRAW, THE PUNCTURE SHEATH FAILING TO ENTER THE BLOOD VESSEL THROUGH THE GUIDE WIRE, AND THE GUIDE WIRE BEING DRAGGED OUT OF THE ORIGINAL INSERTION LENGTH. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05106 | 2023-07-12 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316M | 119316M |  | IT WAS REPORTED THAT CUSTOMER INFLATED THE BALLOON AND SAW THAT WHEN IT WAS INFLATED WITH THE SALINE. THE FLUID LEAKED OUT OF THE BALLOON AND OUT OF THE CATHETER AND IT DEFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE, A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05110 | 2023-07-12 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT PATIENT HAD FOLEY CATHETER ISSUES. UROLOGY PHYSICIAN SAW THE PATIENT AND THEY ENDED UP HAVING TO REPLACE THE FOLEY CATHETER. THEY MENTIONED THAT THEY DISLIKED THE CLEAR CATHETERS THAT WERE COMING IN THE KITS AS THE BALLOON KEPT COLLAPSING IN ON ITSELF AND CAUSING ISSUES FOR PATIENTS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05097 | 2023-07-12 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT IMMEDIATELY AFTER CATHETER PLACEMENT AND THE BALLOON WAS FOUND BURST. THERE WAS A BALLOON RUPTURE OR TEAR. WATER LEAKED FROM THE BALLOON AND NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER. THERE WAS NO PATIENT'S URINE STONE AND THE CATHETER WAS NOT PULLED. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. VISUAL INSPECTION NOTED IRREGULAR BURST WITHOUT MISSING PIECES. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT) EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]: 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED]. (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED]. 2. APPLICABLE PATIENTS: (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED]. [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX SIZES OF CATHETERS: AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS). SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. IRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING: CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE 1. MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION CATHETER KINKING, DAMAGE, RUPTURE, DIFFICULTY OR FAILURE TO REMOVE THE DEVICE, OCCLUSION OF CATHETER INNER LUMENS, ENCRUSTATION, ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE, DEVICE DAMAGE DUE TO INAPPROPRIATE USE. 2) ADVERSE EVENTS: URINARY-TRACT INFECTION, HEMORRHAGE, HEMATURIA, ALLERGY REACTION TO THE DEVICE, CALCULUS FORMATION, EDEMA, PAIN, DISCOMFORT, INJURY OF BLADDER OR URETHRAL, URETHRITIS, URINARY INCONTINENCE, RETAINED BALLOON FRAGMENTS, [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02912 | 2023-07-12 | 2023-06-19 | I, n, j, u, r, y | Manufacturer report | FULL KIT 18G X 10 CM WITH GUARDIVA AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F118108PT |  | IT WAS REPORTED BY THE CUSTOMER THAT THE MIDLINE WAS WORKING WELL AFTER PLACEMENT ON (B)(6)2023. THE PATIENT WAS AN OUTPATIENT COMING IN FOR DAILY INFUSIONS THEREFORE DAILY ASSESSMENT WAS NOT DONE BY THE VAT TEAM BUT BY THE OUT PATIENT NURSES. THE NURSE STATED THAT LAST WEEK HE HAD A KINK IN THE CATHETER THAT PREVENTED IT FROM FLUSHING. THE NURSE BACKED OUT THE CATHETER TO STRAIGHTEN THE KINK. IT WORKED UNTIL (B)(6) 2023 WHEN THE NURSE ATTEMPTED TO FLUSH THE LINE AND THE FLUSH CAME OUT FROM UNDER THE DRESSING. THE DRESSING WAS THEN REMOVED THINKING THE CONNECTION TO THE EXTENSION WAS LOOSE. THE LINE WAS PULLED BACK AGAIN AND ONLY APPROXIMATELY 1 CM OF THE CATHETER WAS ATTACHED TO THE HUB. AN ULTRASOUND WAS USED TO VISUALIZE THE REMAINING CATHETER IN THE VEIN. A TOURNIQUET WAS PLACED TO PREVENT IT FROM DISLODGING ANY FURTHER AND THE PATIENT WAS SENT TO THE ER FOR EVALUATION. ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023: (B)(6) 2023 HE PRESENTED TO THE ED AS A TRAUMA AFTER HE INJURED HIS ELBOW WITH AN ELBOW LACERATION, SENT HOME AND RETURNED ON (B)(6) 2023 WITH INCREASED PAIN TO ELBOW. (B)(6) 2023 IRRIGATION AND DEBRIDEMENT OF LEFT MEDIAL ELBOW ABSCESS AND SEPTIC ARTHRITIS, PLACED ON IVABX. (B)(6) 2023 POWERGLIDE MIDLINE PLACED BY VAT WITHOUT COMPLICATIONS AND DISCHARGED THAT SAME DAY. THE NEXT DAY HE RETURNED TO THE ED STATING THAT HIS GIRLFRIEND HAD KICKED HIM OUT AND HE COULD NOT TAKE CARE OF HIMSELF. HE WAS GIVEN HIS IVABX DAPTOMYCIN AND DISCHARGED FROM THE ED. UNCLEAR AS TO WHETHER PATIENT RECEIVED DAILY IVABX. IT IS DOCUMENTED THAT HE RECEIVED HIS IVABX ON (B)(6)2023. PT WAS NOT SEEN AT THIS FACILITY AGAIN UNTIL (B)(6) 2023 WHEN THE MIDLINE WAS FLUSHED AND NOTED TO BE LEAKING OUT OF THE DRESSING. PER RN REPORT, THE DRESSING WAS THEN REMOVED, AND LINE WAS PULLED BACK WITH ONLY APPROXIMATELY 1CM OF CATHETER ATTACHED TO THE BASE HUB. I (VAT) WAS THEN NOTIFIED OF THE SITUATION. UPON ENTERING THE ROOM, PATIENT WAS SEATED WITH RIGHT ARM UP ON ARMREST. SMALL PUNCTURE MARK NOTED ON RIGHT ARM APPROXIMATELY 6CM ABOVE AC. ULTRASOUND WAS THEN USED TO DETERMINE IF THERE WAS ANY FOREIGN OBJECT LEFT IN VEIN. REMAINING CATHETER WAS VISUALIZED AND A TOURNIQUET WAS THEN PLACED ON PATIENTS ARM ABOVE THE ENDING OF THE VISUALIZED CATHETER. PT WAS THEN TRANSFERRED TO THE ED. PT WAS THEN TAKEN TO THE OR FOR REMOVAL OF THE FOREIGN OBJECT. OR REPORT ALL REMAINING PIECES OF CATHETER WERE RETRIEVED. PT ADMITTED FOR OBSERVATION OVERNIGHT. THIS DELAYED PATIENT FROM GETTING IVABX IN A TIMELY FASHION AS WELL ADDITIONAL EXPENSE AND HOSPITAL STAY. PRIOR TO THIS, PATIENT DENIES ANY ISSUES WITH LINE. NO URGENT MEDICAL SITUATION WAS INVOLVED. THE IVABX BEING INFUSED THROUGH THIS POWERGLIDE MIDLINE PRIOR TO THIS INCIDENT WAS DAPTOMYCIN. STAT LOCK PRO, AS WELL AS TEGADERM AND TAPE WERE USED TO SECURE THE CATHETER. PT HAS NOT BEEN BACK TO OUR FACILITY SINCE THIS INCIDENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02897 | 2023-07-12 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL CATHETER BASIC TRAY (W/ MICROINTRODUCER, NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | IT WAS REPORTED AIR ENTERS THROUGH THE INSERTION STYLET DURING THE PICC INSERTION WHEN ASPIRATING TO CHECK THE BLOOD RETURN, AND WHEN FLUSHING SALINE DROPS THROUGH THE INSERTION STYLET. THIS HAPPENS SINCE THE RUBBER CAP OF THE STYLET HAS CHANGED. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02893 | 2023-07-12 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL CATHETER BASIC TRAY (W/ MICROINTRODUCER, NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | IT WAS REPORTED THE GUIDEWIRE IS ROUGH WHEN NORMALLY HAS BEEN SMOOTH AND EASY TO INSERT. PICC WAS INSERTED ANYWAY. NO HARM FOR THE PATIENT, ONLY LONGER INSERTION TIME WITH DIFFICULT INSERTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02898 | 2023-07-12 | 2023-05-28 | I, n, j, u, r, y | Manufacturer report | POWERPICC 5F SL CATHETER BASIC TRAY (WITH 135CM NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6175335 |  | IT WAS REPORTED A PATIENT HOSPITALIZED FOR SYLVIAN INFARCTION HAD A PICC-LINE INSTALLED ON (B)(6) 2023. ON (B)(6) 2023, APPEARANCE OF A FEBRILE EPISODE AND BLOOD CULTURES ON THE PICC CAME BACK POSITIVE FOR CANDIDA ALBICANS. CONSEQUENCES: PICC REMOVED ON 05/30 AND CENTRAL VENOUS LINE INSERTED FOR ANTIFUNGAL TREATMENT. THE DEVICE WAS NOT KEPT. ADDITIONAL INFORMATION 06/23/2023: INFECTION CONTROLLED, ANTI-INFECTIVE THERAPY DISCONTINUED (B)(6) 2023. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02890 | 2023-07-12 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIG JAPAN HANG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405J |  | IT WAS REPORTED THAT THE CATHETER WAS FOUND TO BE DETACHED FROM THE CATHETER CONNECTOR. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF CATHETER DETACHMENT FROM THE CONNECTOR WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS 4FR S/L GROSHONG CATHETER. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. THE CATHETER WAS RECEIVED FULLY DETACHED FROM THE ASSEMBLED CONNECTOR. MICROSCOPIC INSPECTION OF THE PROXIMAL END OF THE CATHETER REVEALED A GLOSSY STRIATED SURFACE TYPICAL OF A SHARP INSTRUMENT CUT. INSPECTION OF THE CONNECTOR PRIOR TO AND FOLLOWING DISASSEMBLY CONFIRMED THAT NO CATHETER REMAINED WITHIN THE CONNECTOR. FOLLOWING LONGITUDINAL BISECTION OF THE DISTAL CONNECTOR SEGMENT, INSPECTION OF THE COMPRESSION SLEEVE REVEALED THAT IT HAD BEEN ADVANCED TO THE LOCKED POSITION. THE CONNECTOR BORE APPEARED WELL FORMED AND UNREMARKABLE. THE CATHETER DETACHMENT APPEARED TO BE THE RESULT OF THE CATHETER NOT BEING INSERTED THROUGH THE DISTAL CONNECTOR SEGMENT COMPRESSION SLEEVE PRIOR TO CONNECTOR ASSEMBLY. THE PRODUCT IFU PROVIDES INSTRUCTIONS FOR PROPER ASSEMBLY OF THE CATHETER/CONNECTOR SYSTEM. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02904 | 2023-07-12 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER WITH SHERLOCK 3CG TPS STYLET (3F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000821 |  | IT WAS REPORTED VAT WAS CALLED URGENTLY TO ASSESS LEAKING PICC. NOTED HOLE IN THE "PIGTAIL" PORTION OF THE PICC, CLOSE TO THE HARD PLASTIC PIECE THAT CONNECTS TO THE NEEDLELESS CONNECTOR. UNDERWENT AT LEAST 10 ADDITIONAL STICKS UNTIL NEW PICC WAS PLACED BETWEEN IV ACCESS AND LAB DRAWS. DEVICE WAS PLACED 3CM BELOW THE KNEE, AND NEONATAL GRIPLOCK WAS USED TO SECURE. PATIENT WAS BEING TREATED FOR S/P ENTEROPLASTY AND REDO ILEOCOLIC ANASTOMOSIS, AND HAS A HISTORY OF OMPHALOCEALE WITH JEJUNAL ATRESIA. INFUSING: D5 77 NS 77 ACETATE AT 5 CC HR, HEP FLUSH (B)(6) 18:57 (LAST 7 DAYS REVIEWED) CEFEPIME 40MG Q 8 LAST DOSE (B)(6) 21:58, VANCOMYCIN 160 MG Q 6 LAST DOSE (B)(6) 18:19, TPA (B)(6) 23:36 BY VAT. LINE FLUSHED BUT DID NOT HAVE BLOOD RETURN ONLY ONE DOSE NEEDED | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02908 | 2023-07-12 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 4 FR SL PROVENA MIDLINE BASIC KIT, CE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S64244118 |  | CUSTOMER REPORTED THAT ON (B)(6) , THE EIAV - IV THERAPY HOSPITAL TEAM IMPLANTS A MIDLINE IN A PATIENT ADMITTED TO GERIATRICS. THE NEXT DAY THE GERIATRICS DEPARTMENT INFORMED THE PATIENT THAT A PART HAD COME LOOSE, WITHOUT FURTHER EXPLANATION. THE PIECE IS THE DISTAL LUER OF ONE OF THE TWO LIGHTS, DETACHED FROM THE TUBE. THEY KEPT THE LINE CLAMPED AND EIAV REPLACED THE CATHETER IMMEDIATELY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02919 | 2023-07-12 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT AFTER CATHETER PLACEMENT WAS COMPLETED, WHEN FLUSHING THE CATHETER, SAND HOLES WERE FOUND AT THE CONNECTOR. LIQUID LEAKAGE OCCURRED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02906 | 2023-07-12 | 2023-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT WHILE GETTING PATIENT UP TO CHAIR VIA LIFT, ALINE BROKE AND COULD NOT BE RECONNECTED. ALINE HAD TO BE DISCONNECTED IN ORDER TO CONTROL BLEEDING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DETACHED EXTENSION SET LUER ADAPTER WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02915 | 2023-07-12 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV 4F DL MAX 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1274108D |  | IT WAS REPORTED BY CUSTOMER THAT HE CLEARED PLASTIC HUB WHICH THE INSERTION WIRE GOES THROUGH IS BOTH LEAKING AIR AND FLUID. WHEN HE DREW BACK TO CHECK FOR BLOOD FLOW, AIR GETS PULLED INTO SYRINGE. WHEN HE FLUSHED TBE SALINE SQUIRTS OUT AROUND WIRE. IT IS MORE THAN JUST A BUBBLE OR TWO OF AIR. I OBSERVED BETWEEN 2-3 CC OF AIR DRAWN UP INTO FLUSH SYRINGE. IT WAS NOT DUE TO A LOOSE HUB. UNABLE TO QUANTIFY AMOUNT OF SALINE THAT LEAKS OUT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02911 | 2023-07-12 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 3275118 |  | IT WAS REPORTED BY THE CUSTOMER, ACCORDING TO THE DOCTOR'S REPORT, THE CATHETER WAS SHOWING A FISSURE. ACCORDING TO DOCTOR THE CATHETER EXTENSION WAS PUNCTURED, IT WAS NECESSARY TO DISPENSE ANOTHER CATHETER TO CONTINUE THE PROCEDURE. NO OTHER INFORMATION PROVIDED. | H11: THE INITIAL COMPLAINT APPEARED TO BE A REPORTABLE OCCURRENCE. UPON RECEIVING ADDITIONAL INFORMATION ABOUT THE EVENT, IT WAS DETERMINED THAT A REPORTABLE EVENT HAD NOT OCCURRED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02914 | 2023-07-12 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | IT WAS REPORTED THAT UPON OPENING, LIQUIDS LEAKED FROM THE CATHETER ON THE SCALE OF ABOUT 52 CM DURING PRE-FLUSHING. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAKING CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR S/L GROSHONG CATHETER. THE SAMPLE WAS RECEIVED WITH AN INLAID STYLET. AN ATTEMPT TO INFUSE WATER THROUGH THE SAMPLE USING A 12ML SYRINGE REVEALED THE SAMPLE TO BE PATENT TO INFUSION; HOWEVER, A LEAK WAS OBSERVED AT THE SPLIT SITE. MICROSCOPIC INSPECTION OF THE LEAK SITE REVEALED A DIAGONALLY ALIGNED SPLIT BETWEEN THE 52CM AND 53CM DEPTH MARKINGS. TACTILE INSPECTION OF THE SAMPLE REVEALED TENSILE WEAKNESS IN THE VICINITY OF THE SPLIT. FOLLOWING LONGITUDINAL BISECTION OF THE CATHETER NEAR THE SPLIT SITE, MICROSCOPIC INSPECTION OF THE INSIDE SURFACE REVEALED ABRASION DAMAGE. THE DAMAGE ON THE INSIDE SURFACE WAS MORE EXTENSIVE THAN THAT OBSERVED ON THE OUTSIDE SURFACE. THE CATHETER DAMAGE WAS CONSISTENT WITH ABRASION DAMAGE CAUSED BY FRICTION BETWEEN THE CATHETER AND THE STYLET. SUCH DAMAGE CAN OCCUR IF THE STYLET IS NOT WETTED PRIOR TO MANIPULATION/REMOVAL AND IF THE CATHETER IS CONSTRAINED DURING STYLET REMOVAL. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02918 | 2023-07-12 | 2023-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 20G X 10 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F320108PT |  | IT WAS REPORTED BY THE CUSTOMER THAT CATHETER HAD HOLES WHERE THE CATHETER AND THE HUB MEET. THEY HAD TO PULL THE CATHETER AND PLACE ANOTHER ONE. IT WAS STATED IT WAS THE HOLE WAS DISCOVERED IMMEDIATELY AFTER PLACING. NO PRIMING COULD BE PERFORMED ON THIS DEVICE. CATHETER BLED AND LEAKED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02887 | 2023-07-12 | 2023-06-26 | I, n, j, u, r, y | Manufacturer report | BARD TRIPLE LUMEN POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. TRIPLE LUMEN POWERPICC UPPER ARM. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05083 | 2023-07-12 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM A PINHOLE ON THE BALLOON SECTION OF THE FOLEY CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. INSPECTED THE EXTERIOR OF THE SAMPLE AND FOUND TEAR ON THE INFLATION FUNNEL. THE TEAR MODE LOOKS LIKE ORIGINATED SHARP OBJECT FROM OUTSIDE AND STRETCHED OUT THROUGH THE LATEX LAYER UNTIL THE INNER SIDE, CREATING A CHANNEL. INFLATED BALLOON WITH WATER AND FOUND WATER LEAKING FROM THE TEAR. THE TEAR CAUSE THE WATER TO LEAK. HOWEVER THE EXACT CAUSE OF SAMPLE CONDITION COULD NOT BE DETERMINED AND COULD NOT ASSIGN A MANUFACTURING RELATED PROCESS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02878 | 2023-07-12 | 2023-06-19 | I, n, j, u, r, y | Manufacturer report | POWERPICC 3F SL SOLO MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1193108D5 |  | IT WAS REPORTED BY THE CUSTOMER, "PATIENT EVENT WITH HARM DUE TO INADVERTENT ARTERIAL PICC PLACEMENTS. SHERLOCK/3CG TECHNOLOGY WOULD SHOW AN INCREASE IN P-WAVE AMPLITUDE TO INDICATE THAT THE TIP WAS IN THE CORRECT POSITION, EVEN IF THE PICC WAS ACTUALLY IN THE ARTERIAL SYSTEM. IN THIS PATIENT EVENT WHEN 3CG TECHNOLOGY WAS USED TO CONFIRM PROPER PICC TIP PLACEMENT, THE ECG STRIPS MET CRITERIA FOR RELEASING THE PICC FOR USE. THESE PICC WAS LATER CONFIRMED BY OTHER MEASURES TO BE ARTERIAL INSTEAD OF VENOUS." PICC INSERTED (B)(6)2023 WITH ECG TIP CONFIRMATION. PICC REMOVED BY VASCULAR SURGERY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02931 | 2023-07-12 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 3F SL MBPLUS (3CG, BIOPATCH, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1173108D4 |  | IT WAS REPORTED BY CUSTOMER THAT HE CLEARED PLASTIC HUB WHICH THE INSERTION WIRE GOES THROUGH IS BOTH LEAKING AIR AND FLUID. WHEN HE DREW BACK TO CHECK FOR BLOOD FLOW, AIR GETS PULLED INTO SYRINGE. WHEN HE FLUSHED THE SALINE SQUIRTS OUT AROUND WIRE. IT IS MORE THAN JUST A BUBBLE OR TWO OF AIR. I OBSERVED BETWEEN 2-3 CC OF AIR DRAWN UP INTO FLUSH SYRINGE. IT WAS NOT DUE TO A LOOSE HUB. UNABLE TO QUANTIFY AMOUNT OF SALINE THAT LEAKS OUT. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05082 | 2023-07-12 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS BURST. THE EVENT OCCURRED AFTER 1 HOUR OF CATHETER PLACEMENT TO THE PATIENT. THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. THERE WAS NO PATIENT'S URINE STONE. THE CATHETER WAS NOT PULLED. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. EVALUATION CONFIRMED THE BALLOON WAS BURST. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS], METHOD FOR USE .DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] .DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS. PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS], METHOD FOR USE:DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS. PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX. PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD, SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 8 TO 24 EVERY 2FR. FOLEY CATHETER. ACCESSORIES, CLOSED DRAINAGE BAG, (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER, WATER SOLUBLE LUBRICANT, ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION, TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS, GLOVES, STATLOCK FOLEY, [INTENDED USE & EFFECT- EFFICACY]. THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE], METHOD OF USE. THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1, FIG.2. CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3, FIG.4. INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6. TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT: OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING: CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) PRECAUTIONS FOR USE: WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS). EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE, DIFFICULTY OR FAILURE TO REMOVE THE DEVICE, OCCLUSION OF CATHETER INNER LUMENS, ENCRUSTATION, ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE DEVICE DAMAGE DUE TO INAPPROPRIATE USE. 2) ADVERSE EVENTS: TRACT INFECTION, HEMORRHAGE, HEMATURIA, ALLERGY REACTION TO THE DEVICE, CALCULUS FORMATION, EDEMA, PAIN, DISCOMFORT, INJURY OF BLADDER OR URETHRAL URETHRITIS, URINARY INCONTINENCE RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.." H11: SECTION A THROUGH THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02927 | 2023-07-12 | 2023-03-15 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE TUMOR DEPARTMENT OF THIS HOSPITAL DUE TO CANCER AND HAD A CATHETER INSERTED IN (B)(6) 2022. THE CATHETER WAS MAINTAINED AT ANOTHER HOSPITAL AFTERWARDS. ON (B)(6) 2023, CHEST X-RAY SHOWED THAT THE CATHETER WAS IN GOOD CONDITION. BUT AFTER INFUSION OF LIQUID, LEAKAGE OCCURRED. WHEN PREPARING TO REMOVE THE CATHETER, THE CATHETER WAS FOUND RUPTURED. THE RUPTURED PORTION WAS REMOVED AT THE VASCULAR SURGERY DEPARTMENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02876 | 2023-07-12 | 2023-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED LEAKAGE FROM CATHETER DURING FLUSHING. CUSTOMER TAKES OUT THE CATHETER AND THEY SEE A HOLE 5 CM IN. THE CATHETER HAS BEEN USED FOR BLOOD TRANSFUSION AND BLOOD SAMPLES. EVENT OCCURRED DURING USE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02900 | 2023-07-12 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | CUSTOMER REPORTED THAT THE CATHETER WAS PLACED ON MAY 24, AND THE FASTENERS OF THE EXTENSION TUBE WERE CONFIRMED TO HAVE BEEN ENGAGED PROPERLY. ON JUNE 7, THE FASTENERS CAME OFF AND IT BECAME LOOSENED. THE CUSTOMER REVEALED THAT THE FASTENERS OF THE EXTENSION TUBE HAD A PROBLEM. A NEW EXTENSION TUBE WAS USED SUBSEQUENTLY. THE PATIENT WAS PACIFIED AND NO OTHER EVENT OCCURRED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. AN INVESTIGATION WAS CONDUCTED BASED ON THE KNOWN DETAIL SURROUNDING THE EVENT, THE RETURNED SAMPLE, AND THE PROVIDED PRODUCT INFORMATION. BASED ON ANALYSIS OF THE RETURNED GROSHONG CONNECTOR, THE COMPLAINT OF LOOSE CONNECTION OF THE CONNECTOR IS INCONCLUSIVE. AFTER EVALUATION OF THE RETURNED DEVICE, THE COMPLAINT COULD NOT BE CONFIRMED DUE TO THE CONDITION OF THE RETURNED SAMPLE. ATTEMPTS AT DISASSEMBLING THE DEVICE WAS UNREMARKABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02938 | 2023-07-12 | 2022-09-26 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO 4F SINGLE-LUMEN CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6194355 |  | IT WAS REPORTED, DURING A PERIOD WHERE THE USUAL PICC INSERTION SERVICE WAS UNAVAILABLE, AN ANAESTHETIC REGISTRAR INSERTED A POWERPICC SOLO, WITHOUT ANY PREVIOUS EXPERIENCE. AT THE END OF THE PROCEDURE, IT WAS NOTED THAT A GUIDEWIRE WAS UNACCOUNTED FOR, AND HAD BEEN LEFT IN THE PATIENT. PATIENT TRANSFERRED TO INTERVENTIONAL RADIOLOGY, WHERE GUIDEWIRE WAS SUCCESSFULLY RETRIEVED. INITIAL TREATMENT WAS DELAYED WHILE THE ISSUE WAS RESOLVED. SAMPLE WAS DESTROYED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02895 | 2023-07-12 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL CATHETER BASIC TRAY (W/ MICROINTRODUCER, NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | CUSTOMER REPORTED THAT AIR ENTERS THROUGH THE INSERTION STYLET DURING THE PICC INSERTION WHEN ASPIRATING TO CHECK THE BLOOD RETURN, AND WHEN FLUSHING SALINE DROPS TROUGH, THE INSERTION STYLET. DURING USE. THIS HAPPENS SINCE THE RUBBER CAP OF THE STYLET HAS CHANGED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02940 | 2023-07-12 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT W/BIOPATCH | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108BDP |  | IT WAS REPORTED BY CUSTOMER THAT WIRE SHEARED WHEN REMOVED FROM PATIENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02894 | 2023-07-12 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL CATHETER BASIC TRAY (W/ MICROINTRODUCER, NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | CUSTOMER REPORTED THAT AIR ENTERS THROUGH THE INSERTION STYLET DURING THE PICC INSERTION WHEN ASPIRATING TO CHECK THE BLOOD RETURN, AND WHEN FLUSHING SALINE DROPS TROUGH, THE INSERTION STYLET. DURING USE. THIS HAPPENS SINCE THE RUBBER CAP OF THE STYLET HAS CHANGED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02935 | 2023-07-12 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, AN EXPERIENCED CUSTOMER HAD LEAKAGE, BLEEDING, AT THE VALVE POINT AFTER THE PLACEMENT OF THE PICC. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02939 | 2023-07-12 | 2023-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV SOLO SP 3F SL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1193108PD |  | IT WAS REPORTED BY THE CUSTOMER THAT 3FR PEEL AWAY SHEATH UNABLE TO THREAD THROUGH TISSUE DESPITE NO ISSUE WITH INNER DILATOR. NOTED TO HAVE RIDGES ON END OF SHEATH. USED 3.5FR COOK PICC PEEL AWAY. NO PROBLEMS WITH PRODUCT. REMAINDER OF PROCEDURE COMPLETED WITHOUT INCIDENT. (B)(6)2023 ADDITIONAL INFORMATION PROVIDED: IT WAS REPORTED BY THE CUSTOMER "NO PRIOR MEDICAL HISTORY; DIAGNOSIS NEW CEREBRAL MASS. PICC NEEDED FOR LAB MONITORING AND CHEMOTHERAPY. EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING SITUATION. NO NEGATIVE OUTCOME TO PATIENT. PROCEDURE TOOK LONGER TO COMPLETE AS TWO SEPARATE KITS NEEDED TO BE OPENED IN ORDER TO GET PEEL AWAY SHEATH NEEDED FOR PROCEDURE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02937 | 2023-07-12 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | PATIENT REPORTED THAT THE HEATH CARE PROVIDER TRIED A PICC LINE TO GIVE MEDICINE. AFTER 5 PRICKS THEY GOT THE LINE IN, THE HCP GAVE CONTRAST AND FOUND THAT CATHETER WAS CLOTTING. TRIED A WEEK AGO AND COULD NOT GET BLOOD OUT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02930 | 2023-07-12 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) (4F) (SL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 4174118 |  | CUSTOMER REPORTED THAT THEY RECEIVED A PATIENT IN HOME TREATMENT, WITH IMPLANTED PICC CATHETER PRESENTING PARTIAL RUPTURE AND RUPTURE (BREAK) NEAR THE FINAL JUNCTION. LOSS OF MEDICATION DUE TO RUPTURE AND NECESSARY REMOVAL OF THE CATHETER. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02933 | 2023-07-12 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, AN EXPERIENCED CUSTOMER HAD MULTIPLE LEAKAGES, BLEEDING, AT THE VALVE POINT AFTER THE PLACEMENT OF THE PICC. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02932 | 2023-07-12 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, AN EXPERIENCED CUSTOMER HAD LEAKAGE, BLEEDING, AT THE VALVE POINT AFTER THE PLACEMENT OF THE PICC. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05081 | 2023-07-12 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED DURING PRETEST AND THE CATHETER WAS NOT USED ON THE PATIENT. THERE WAS NO BALLOON RUPTURE OR TEAR. THE WATER LEAKED FROM THE SHAFT AREA. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02934 | 2023-07-12 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, AN EXPERIENCED CUSTOMER HAD MULTIPLE LEAKAGES, BLEEDING, AT THE VALVE POINT AFTER THE PLACEMENT OF THE PICC. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02952 | 2023-07-13 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT SEVEN MONTHS AND ELEVEN DAYS POST A DIALYSIS CATHETER PLACEMENT, THE CLAMP ALLEGEDLY BROKE OFF. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 12/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02968 | 2023-07-13 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED THERE HAVE BEEN PATIENT EVENTS WITH HARM DUE TO INADVERTENT ARTERIAL PICC PLACEMENTS, THE SHERLOCK/3CG TECHNOLOGY WOULD SHOW AN INCREASE IN P-WAVE AMPLITUDE TO INDICATE THAT THE TIP WAS IN THE CORRECT POSITION, EVEN IF THE PICC WAS ACTUALLY IN THE ARTERIAL SYSTEM. IN THESE PATIENT EVENTS WHEN 3CG TECHNOLOGY WAS USED TO CONFIRM PROPER PICC TIP PLACEMENT, THE ECG STRIPS MET CRITERIA FOR RELEASING THE PICC FOR USE. THESE PICCS WERE LATER CONFIRMED BY OTHER MEASURES TO BE ARTERIAL INSTEAD OF VENOUS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02965 | 2023-07-13 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR PROVENA MIDLINE MAX BARRIER KIT | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108D |  | IT WAS REPORTED BY CUSTOMER THAT POWERMIDLINE IS LEAKING AT THE T CONNECTOR. THEY HAVE TO MANIPULATE TO GET IT TO WORK CORRECTLY. HAD TO SPEND EXTRA TIME WORKING AROUND THE PROBLEM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02949 | 2023-07-13 | 2023-06-19 | I, n, j, u, r, y | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, WHEN THE CATHETER WAS INSERTED THE PATIENT ALLEGEDLY EXPERIENCED VASCULAR INJURY. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02957 | 2023-07-13 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | CUSTOMER REPORTED THAT THERE WAS NO PROGRESSION OF THE CATHETER THROUGH THE INTRODUCER. ONLY WHEN THE EXCHANGE WAS PERFORMED DID THE CATHETER PROGRESS THROUGH THE PEEL WAY. THE DOCTOR REPORTS THAT WHEN TRYING TO INSERT THE CATHETER IN THE PATIENT, HE BENT THE TIP, NOT BEING POSSIBLE TO INTRODUCE IT, HE BELIEVES THAT THE TIP IS TOO SOFT OR FRAGILE. NO OTHER INFORMATION PROVIDED. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02956 | 2023-07-13 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC PROVENA DL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT THE STOPPER ON THE STYLET LEAKS ANYTIME SALINE IS FLUSHED WITH THE DUAL LUMEN POWERPICC PROVENA KITS. WHILE FLUSHING CATHETER DURING PICC INSERTION PROCEDURE, SALINE SQUIRTS OUT WHERE THE STYLET COMES THROUGH THE STOPPER. THE CUSTOMER IS UNABLE TO PROVIDE THE EXACT NUMBER OF INCIDENTS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02954 | 2023-07-13 | 2023-06-06 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED BY THE CUSTOMER "THE PATIENT WAS GIVEN A PERIPHERALLY INSERTED CENTRAL VENOUS CATHETER KIT PICC AT 10:00 A.M. ON (B)(6) 2023, AND THE PICC TUBE WAS BROKEN AND EJECTED 13 CM, BUT THE REST OF THE CATHETER WAS NOT SEEN. AT 11:29 A.M., UNDER THE GUIDANCE OF DSA, ALL THE CATHETERS WERE REMOVED, MEASURING 23 CM, AND NO STUMP REMAINED IN THE BODY. AT 11:49 A.M., THE PATIENT WAS RETURNED TO THE WARD BY THE FLAT CAR AND WAS GIVEN THREE HOURS OF LYING DOWN AS PRESCRIBED BY THE DOCTOR." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05128 | 2023-07-13 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V24S |  | IT WAS REPORTED THAT THERE WAS ANOTHER INSTANCE OF THE FOLEY CATHETER BALLOON NOT DEFLATING ON THIS PATIENT WITH A VISIT TO THE EMERGENCY ROOM AND HAD TO BREAK THE BALLOON WITH A WIRE. THE CUSTOMER WAS PART OF THE HOME CARE PROGRAM (SELF MANAGED CARE). THE CATHETER HAD BLOCKED AT CUSTOMER'S HOME AND FAMILY WAS ASSISTING IN FLUSHING. THE CATHETER HAD FAILED FROM FLUSHING AND FAMILY ATTEMPTED TO CHANGE CATHETER. THE CATHETER WAS UNABLE TO DEFLATE. THEY HAD TO GO TO EMERGENCY AT HOSPITAL AND PHYSICIAN HAD VALVE CUT BUT NO FURTHER DEFLATION OF CATHETER. THE CATHETER HAD TO BE PIERCED WITH A WIRE AND PHYSICIAN PULLED CATHETER BALLOON VIA STOMA. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02958 | 2023-07-13 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT ON (B)(6) 2023, THE PATIENT WAS REQUIRED TO HAVE AN INDWELLING PICC FOR TREATMENT, AND THE TIP OF THE PUNCTURE SHEATH WAS FOUND TO BE UNEVEN DURING THE PUNCTURE PROCESS. AFTER THE PUNCTURE SHEATH WAS REPLACED, THE PICC WAS SUCCESSFULLY INDWELLING WITHOUT AFFECTING THE PATIENT'S CONDITION. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05123 | 2023-07-13 | 2023-06-23 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT CUSTOMER STATED THAT THE PATIENT DEVELOPED A RASH. REPRESENTATIVE ADVISED TO STOP USE. NO MEDICAL INTERVENTION WAS PROVIDED. PER FOLLOW UP VIA PHONE ON 28JUN2023, IT WAS REPORTED THAT THEY RECEIVED PRESCRIBED CREAM TO TREAT THEIR RASH THEY GOT FROM USING THE PUREWICK. THE RASH HAD GOTTEN BETTER AND THE PATIENT HAD COMPLETELY STOPPED USING THE DEVICE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05125 | 2023-07-13 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 | FOL0102 | FOL0102 |  | IT WAS REPORTED THAT THE STICKY SUBSTANCE OF THE STATLOCK CAME OFF EASILY. THEY WERE TRANSFERRED TO PATIENT SERVICES PER REQUEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-07
3006260740-2023-02943 | 2023-07-13 | 2023-06-19 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT PROCEDURE, THE RESIDUAL GUIDEWIRE FRAGMENTS WERE ALLEGEDLY FOUND INSIDE THE PULMONARY VEIN. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 04/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00517 | 2023-07-13 | 2023-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQF75104 | CQF75104 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A LEFT FISTULA, THE PTA BALLOON ALLEGEDLY HAD HOLES IN THE UPPER ONE THIRD AND LOWER TWO THIRD AT FIRST INFLATION. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 04/2025), G3. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02942 | 2023-07-13 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT TWO WEEKS AND THREE DAYS POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER TUBE ALLEGEDLY BROKE. IT WAS FURTHER REPORTED THAT BLOOD WAS ALLEGEDLY COMING OUT OF THE TUBE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 11/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02941 | 2023-07-13 | 2023-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT SIX YEARS SEVEN MONTHS AND SEVENTEEN DAYS POST PORT PLACEMENT PROCEDURE, THE X-RAY EXAMINATION REVEALED THAT THE CATHETER WAS DAMAGED NEAR THE SUBCUTANEOUS TUNNEL AND THERE WAS NO LEAKAGE OF MEDICINE BECAUSE ONLY ROUTINE PORT FLUSHING HAS BEEN PERFORMED. IT WAS FURTHER REPORTED THAT PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02948 | 2023-07-13 | 2023-06-20 | I, n, j, u, r, y | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 7707540 | 7707540 |  | IT WAS REPORTED THAT ONE YEAR FOUR MONTHS AND NINE DAYS POST A PORT PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY GOT BROKEN AND GOT SEPARATED. IT WAS FURTHER REPORTED THAT THE BROKEN CATHETER ALLEGEDLY DISLODGED IN THE HEART. REPORTEDLY THE PORT WAS REMOVED. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 01/2024) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05119 | 2023-07-13 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT STERILE WATER WAS DIFFICULT TO INJECT INTO THE FOLEY CATHETER AND THE CATHETER WAS DISCONTINUED TO USE. IT WAS STATED THAT THE AFTER SEVERAL ATTEMPTS OF WATER INJECTION, THE FOLEY CATHETER BALLOON INFLATED BUT WAS SEVERELY ASYMMETRICAL.. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05122 | 2023-07-13 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED ON (B)(6) 2023. IT WAS STATED THAT ON (B)(6) 2023, THE PATIENT'S URINARY CATHETER WAS OBSERVED TO BE INTACT AND SECURED DURING ROUTINE NURSING CARE. IT WAS STATED THAT THE NURSE NOTICED THAT THE PATIENT'S BLADDER WAS DISLODGED WHEN THE BALLOON BURST AND TORN INTO TWO PIECES. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 27JUN2023, NURSE NOTICED THAT THE PATIENT'S URINARY WAS DISLODGED WITH THE BALLOON BURST AND TORN INTO 2 PIECES AND NO INFORMATION ON THE MISSING PIECES IN THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO THIN SAC CAUSING BY SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02951 | 2023-07-13 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT SOMETIME POST A DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY KINKED RIGHT AFTER THE DEPLOYMENT. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 06/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02944 | 2023-07-13 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | A710962 | A710962 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT WAS ABLE TO BE FLUSHED AND OBTAIN A BLOOD RETURN, BUT IT WAS SLUGGISH. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05142 | 2023-07-14 | 2023-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. CATHETER PLACEMENT PERIOD IS 1 DAY. THE CATHETER WAS SPONTANEOUSLY FELL OFF DUE TO RAPID DECREASE IN STERILE WATER. AFTER SPONTANEOUS REMOVAL, 10 CC OF WATER WAS INJECTED TO THE CATHETER AGAIN, AND WHEN CHECKED, ABOUT 6 CC OF WATER HAD DRAWN AFTER ONE DAY. THE SAME EVENT OCCURRED FOR THREE CONSECUTIVE DAYS IN THE SAME PATIENT (EVENT DATES (B)(6), 2023). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F, THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05139 | 2023-07-14 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT THERE WAS NO URINE NOTED IN THE FOLEY CATHETER/BAG. DECISION WAS MADE TO REMOVE. BLOOD WAS NOTED ON TIP OF FOLEY AND TIP ALSO APPEARED TO BE MISSING. DOCTOR PREFORMED A FLEXIBLE CYSTOSCOPY AND THE FOLEY PLACED WITHOUT DIFFICULTY. NO FOREIGN BODY NOTED ON VISUAL CYSTOSCOPY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05150 | 2023-07-14 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT PLACED AN INDWELLING FOLEY CATHETER AT 11:00 ON (B)(6) 2023 DUE TO POSTOPERATIVE ACUTE URINARY RETENTION TO DRAIN THE URINE. WHEN THE CATHETER WAS REMOVED AT 9:30 ON (B)(6) 2023, THE FLUID IN THE BALLOON COULD NOT BE DRAWN OUT, AND THE CATHETER COULD NOT BE REMOVED. THEN, THE CATHETER WAS CUT OFF IN THE MIDDLE, AND THE CATHETER WAS PULLED OUTWARD SLOWLY. AFTER THE LIQUID IN THE BALLOON FLOWED OUT SLOWLY, THE CATHETER WAS PULLED OUT WITHOUT CAUSING OBVIOUS ADVERSE EFFECTS ON THE PATIENT. BUT IF THE CATHETER WAS NOT REMOVED, IT WOULD BE PROBABLY REMOVED UNDER CYSTOTOMY, CAUSING SERIOUS DAMAGE TO THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05149 | 2023-07-14 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT CUSTOMER WERE CURRENTLY EXPERIENCING AN ISSUE WITH A TEMPERATURE SENSING FOLEY NOT READING IN 6112. THE FOLEY WAS WORKING INITIALLY AND THEN SUDDENLY QUIT WORKING. THE CABLES WERE ALL EXCHANGED TO RULE THAT OUT. PER FOLLOW UP VIA EMAIL ON 27JUN2023, IT WAS REPORTED THAT THEY HAD SAVED THE TEMPERATURE SENSING FOLEY THAT FAILED ON THEIR PATIENT IN 6118 OVER THE WEEKEND. ALSO STATED THAT THE PATIENT WAS STILL ADMITTED AND THIS WAS THE SECOND BARD TEMP SENSING FOLEY THAT THEY HAD RETURNED TO THE COMPANY WITH THE SAME UNRESOLVED ISSUE. PER FOLLOW UP VIA PHONE 06JULY2023, CUSTOMER HAVE BEEN WORKING WITH THEIR LOCAL REPRESENTATIVE AS THIS ISSUE WITH THE TEMPERATURE SENSING FOLEYS HAS BEEN ONGOING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05147 | 2023-07-14 | 2023-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THE PATIENT COULD NOT URINATE BY THEMSELVES, SO THEY BEGAN TO USE A DISPOSABLE STERILE CATHETER FOR URINARY CATHETERIZATION ON MAY 28, 2023. THE CATHETER WAS LUBRICATED WITH IODINE, AND INJECTED 20ML OF STERILE WATER FOR INJECTION INTO THE BALLOON, WITH A VOLUME OF 30ML. ON MARCH 3, IT WAS FOUND THAT THE FOLEY CATHETER FELL OUT BY ITSELF, THE BALLOON SHRANK, AND ABOUT 5 ML OF STERILE WATER FOR INJECTION WAS WITHDRAWN. THE PATIENT COULD NOT URINATE BY THEMSELVES. THEN A NEW CATHETER WAS REPLACED IMMEDIATELY, AND THE URINE WAS DRAINED NORMALLY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE POOR SAMPLE CONDITION. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. BASED ON THE ATTACHED PHOTO, OBSERVED CATHETER BALLOON WAS SLIGHTLY DEFLATED. HOWEVER, THE REPORTED EVENT COULD NOT BE CONFIRMED DUE TO POOR SAMPLE CONDITION. FURTHER FUNCTIONAL TESTING COULD NOT BE PERFORMED IF ONLY BASED ON THE ATTACHED PHOTOS. THEREFORE, THE REPORTED EVENT IS INCONCLUSIVE DUE TO POOR SAMPLE CONDITION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO INSUFFICIENT LATEX STRENGTH, LOW/NON UNIFORM RUBBERIZED THICKNESS ETC CAUSING TORN RUBBERIZE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05145 | 2023-07-14 | 2023-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED FOR THE PATIENT AS REQUESTED. DURING THE PROCESS, IT WAS FOUND THAT THE FOLEY CATHETER BALLOON WAS BURST, AND A NEW CATHETER WAS REPLACED IN TIME FOR USE WITHOUT ADVERSE EFFECTS. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 03JUL2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05146 | 2023-07-14 | 2023-06-14 | I, n, j, u, r, y | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT HAD AN INDWELLING FOLEY CATHETER AFTER OPERATION. AFTER USE, IT WAS FOUND THAT THE BALLOON RUPTURED. THE PATIENT'S URINE OVERFLOWED FROM THE WOUND CAUSING URINARY FISTULA. STATED THAT ANOTHER CATHETER WAS INDWELLED. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 03JUL2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05144 | 2023-07-14 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05143 | 2023-07-14 | 2023-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED ,THAT WHEN STERILE WATER WAS INJECTED INTO THE FOLEY CATHETER IN THE SLEEVE, DURING THE PRE-TEST BEFORE INSERTION, WATER LEAKED FROM THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02916 | 2023-07-14 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 3F SL MBPLUS (3CG, BIOPATCH, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1173108D4 |  | IT WAS REPORTED BY CUSTOMER THAT HE CLEARED PLASTIC HUB WHICH THE INSERTION WIRE GOES THROUGH IS BOTH LEAKING AIR AND FLUID. WHEN HE DREW BACK TO CHECK FOR BLOOD FLOW, AIR GETS PULLED INTO SYRINGE. WHEN HE FLUSHED TUBE SALINE SQUIRTS OUT AROUND WIRE. IT IS MORE THAN JUST A BUBBLE OR TWO OF AIR. I OBSERVED BETWEEN 2-3 CC OF AIR DRAWN UP INTO FLUSH SYRINGE. IT WAS NOT DUE TO A LOOSE HUB. UNABLE TO QUANTIFY AMOUNT OF SALINE THAT LEAKS OUT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05148 | 2023-07-14 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT DURING A PRETEST, FIXATION WATER LEAKED FROM THE EDGE AREA OF THE CAP OF THE FOLEY CATHETER INFLATION VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05151 | 2023-07-14 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THE NURSE TESTED THE FOLEY CATHETER BEFORE USING THE PRODUCT ON A PATIENT, AND THE FIRST CATHETER WOULD NOT DEFLATE. SO THEY PULLED ANOTHER CATHETER WITH THE SAME LOT NUMBER AND GOT THE SAME RESULT, CATHETER WOULD NOT DEFLATE. THEY REMOVED/QUARANTINED ALL PRODUCTS WITH SAME LOT NUMBER AS A SAFETY PRECAUTION AS THEY WERE CLINICALLY UNACCEPTABLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05156 | 2023-07-14 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THE PACKETS WERE DIFFICULT TO BREAK, PATIENT HAD TO STAND ON THEM TO BREAK. PATIENT SAID ABOUT 6-8 IN A BOX.. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "SEAL STRENGTH TOO HIGH". THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: INSTRUCTIONS FOR USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05160 | 2023-07-14 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE URINE LEAKED FROM FOLEY CATHETER BUT THE LEAKAGE LOCATION WAS UNCLEAR. AS PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 03JUL2023, EVEN THOUGH CHECKING THE RETURNED SAMPLE, REPRESENTATIVE COULD NOT VISUALLY CONFIRM THE LEAK LOCATION. AN INVESTIGATION FOR LEAK SHOULD BE PERFORMED ON BOTH THE CATHETER AND THE DRAIN BAG. AS PER INVESTIGATOR'S NOTIFICATION RECEIVED ON 05DEC2023, THE TEMPERATURE WIRE WAS PROTRUDING FROM THE SHAFT, THE BALLOON WAS MUSHROOMED, SAMPLE PORT DISCONNECTING FROM TUBING AND THE SCRATCHES/ABRASIONS FOUND ON THE FOLEY CATHETER AS PER INVESTIGATOR'S NOTIFICATION RECEIVED ON 20DEC2023, STATED THAT THE TEMPERATURE WIRE ON THE CATHETER WAS PROTRUDING OUT FROM THE SHAFT. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED 1 OPENED URINARY DRAINAGE BAG ATTACHED TO A 3-WAY TEMP SENSING CATHETER. VISUAL INSPECTION NOTED THAT THE TEMPERATURE WIRE ON THE CATHETER WAS PROTRUDING OUT FROM THE SHAFT AND THE BALLOON WAS MUSHROOMED. FURTHER INSPECTION NOTED THAT THE SAMPLE PORT CAME LOOSE FROM THE TUBING OF THE DRAINAGE BAG VERY EASILY AND WOULD NOT LOCK BACK IN PLACE WITH THE TUBING. IT APPEARED AS IF THERE WAS SCRATCHES OR ABRASIONS AROUND THE CONNECTION OF THE SAMPLE PORT CONNECTOR AND DRAINAGE BAG TUBING. INFLATED BALLOON OF THE CATHETER WITH 10 ML OF SOLUTION AND THE CATHETER RESTED WITH NO LEAKS NOTED. DEFLATED BALLOON PASSIVELY WITH NO CUFFING OR LEAKS NOTED. FLUSHED THE SAMPLE PORT AND FILLED THE DRAINAGE BAG WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THERE WAS A LEAK AROUND THE SAMPLE PORT CONNECTOR. FIRST PHOTO SAMPLE SHOWS OVERVIEW OF DRAIN BAG WITH TUBING ATTACHED. SECOND PHOTO SAMPLE ZOOMS IN ON COLLECTION CONTAINER AND THE SAMPLE PORT CONNECTOR SECURELY CLOSED AND SEATED. THIRD PHOTO SAMPLE SHOWS ON TOP VIEW OF BACK OF URINE COLLECTION CATHETER. FOURTH PHOTO SAMPLE SHOWS TOP OVERVIEW OF CATHETER. FIFTH PHOTO SAMPLE SHOWS TEMPERATURE WIRE THERMISTOR PROTRUDING OUT. BASED ON PHYSICAL SAMPLE RECEIVED PRODUCT DOES NOT MEET SPECIFICATIONS WHICH STATES "CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE." ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE IS INAPPROPRIATE PACKAGE DESIGN. THE DHR REVIEW DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "BARD® EZ-LOK® SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK OR SLIP TIP SYRINGES. 1. OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. 2. SWAB SURFACE OF BARD® EZ-LOK® SAMPLING PORT WITH ANTISEPTIC WIPE. 3. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. 4. SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND REMOVE SYRINGE FROM SAMPLE PORT. 5. UNKINK TUBING IF NECESSARY AND TRANSFER URINE SPECIMEN INTO SPECIMEN CUP OR FOLLOW HOSPITAL PROTOCOL. DISCARD SYRINGE ACCORDING TO HOSPITAL PROTOCOL. 6. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB. PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED" CORRECTION- D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02985 | 2023-07-14 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR (RADEL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770352 |  | IT WAS REPORTED THE LOLLIPOP GOES MISSING AND IT'S NOT IN THE CORRECT PLACEMENT WHEN DOING A PICC LINE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF THE LOLLIPOP GOES MISSING WAS UNCONFIRMED. THE STYLET ICON WAS VISIBLE AND DID NOT DISAPPEAR. THE LOLLIPOP IS NOT IN THE CORRECT PLACEMENT WHEN DOING A PICC LINE WAS UNCONFIRMED. THE MAGNETIC TRACKING WAS TESTED AND FUNCTIONED PROPERLY, DISPLAYING THE STYLET ICON IN THE CORRECT LOCATIONS. THE ROOT CAUSE OF THE REPORTED FAILURE IS INCONCLUSIVE AS THE REPORTED ISSUE COULD NOT BE REPRODUCED DURING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05161 | 2023-07-14 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACK INFECTION AND THEY WAS BLAMED ON THE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION. PER CUSTOMER CONTACT (WIFE - (B)(6) VIA PHONE ON (B)(6) 2023, THE PATIENT HAS BEEN USING THE MAGIC3GO FOR YEARS. BEGINNING ON (B)(6) 2023, HE STARTED GETTING UTIS. THERE HAS BEEN 4 UTIS TOTAL OF VARIOUS DEGREES, ONE OF WHICH REQUIRED HOSPITALIZATION. HE WAS PRESCRIBED MEDICATION BUT DID NOT HAVE THE NAME HANDY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05164 | 2023-07-14 | 2023-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT FOLEY CATHETER RUPTURED WHILE INDWELLING IN PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00523 | 2023-07-14 | 2023-06-16 | I, n, j, u, r, y | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80144 | ATG80144 |  | IT WAS REPORTED THAT, DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY BURST AND SEPARATED FROM THE CATHETER AND LEFT INSIDE THE PATIENT'S BODY. IT WAS FURTHER REPORTED THAT THE DETACHED BALLOON WAS TRIED TO REMOVE USING A BASKET, BUT IT WAS HARD TO SEE UNDER THE ANGIOGRAM IMAGING DUE TO THE BALLOON MATERIAL. IT WAS ALSO REPORTED THAT A STENT WAS USED TO ATTACH IT TO THE VESSEL WALL IN PATIENT'S BODY SO THAT THE BALLOON COULD NO LONGER BE MOVED. REPORTEDLY THERE WAS A FOREIGN SUBSTANCE WAS ALLEGEDLY FOUND ON THE END OF THE CATHETER UPON REMOVAL. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS AND IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 03/2026). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00524 | 2023-07-14 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8094 | VA8094 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, TIP OF THE PTA BALLOON CATHETER ALLEGEDLY FOUND TO BE BROKEN. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 01/2026) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05167 | 2023-07-14 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS LEAKING. PER EMAIL ON 06JULY2023, CONTACT STATED THAT THIS ISSUE HAPPENED TWICE ON THE SAME PATIENT, HOWEVER THEY DO NOT HAVE THE LOT/BATCH NUMBERS. THIS WAS THE ONLY INCIDENCE OF THIS OCCURRING. THERE WAS NO HARM TO THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE IS AVAILABLE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "BAD FIT WITH CONNECTOR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿PROPER TECHNIQUES FOR URINARY CATHETER INSERTION. ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION. ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE. ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED. ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT. ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE. ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED. DIRECTIONS FOR USE. WASH HANDS AND DON CLEAN GLOVES. EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT. 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERIURETHRAL AREA. 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL 5. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP. 6. DON STERILE GLOVES. 7. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN. NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE. 8. POSITION FENESTRATED DRAPE ON PATIENT. 9. SATURATE 3 FOAM SWAB STICKS IN POVIDONE IODINE. 10.ATTACH THE WATER FILLED SYRINGE TO THE INFLATION PORT. NOTE: IT IS NOT NECESSARY TO PRE-TEST THE FOLEY CATHETER BALLOON. 11. REMOVE FOLEY CATHETER FROM WRAP AND LUBRICATE CATHETER. 12. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS. NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA INORA. MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD. 13. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND. A. WHEN CATHETER TIP HAS NTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE. B. INSERT CATHETER TWO MORE INCHES AND INFLATE CATHETER BALLOON. 14. INFLATE CATHETER BALLOON USING ENTIRE 10CC OF STERILE WATER PROVIDED IN THE PREFILLED SYRINGE. NOTE: USE OF LESS THAN 10CC CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON. 15. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK. 16. SECURE THE FOLEY CATHETER TO THE PATIENT USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED (SEE STATLOCK® FOLEY STABILIZATION DEVICE IFU). NOTE: PLEASE MAKE SURE PATIENT IS APPROPRIATE FOR USE OF STATLOCK® STABILIZATION DEVICE 17. POSITION HANGER ON BED RAIL AT THE FOOT OF THE BED. NOTE: EXERCISE CARE TO KEEP BAG OFF THE FLOOR. 18. USE GREEN SHEETING CLIP TO SECURE DRAINAGE TUBE TO THE SHEET. MAKE SURE TUBE IS NOT KINKED. 19. INDICATE TIME AND DATE OF CATHETER INSERTION ON PROVIDED LABELS. PLACE DESIGNATED LABELS ON PATIENT CHART AND DRAINAGE SYSTEM. 20. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL. FOLEY CATHETER REMOVAL. 1. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. 2. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. 3. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. 4. IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. 5. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT¿. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02986 | 2023-07-14 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER "AFTER PICC PLACEMENT ON (B)(6) 2023, THE NURSE FOUND FLUID LEAKAGE FROM THE EXPOSED END OF THE PICC CATHETER WHILE INFUSING HIM ON (B)(6) 2023, AND THEN STOPPED THE INFUSION AND FOUND ANOTHER INTRAVENOUS ACCESS. THE INFUSION WAS STOPPED, REPOSITIONED, THE CATHETER WAS ADJUSTED, AND THE BROKEN SKIN WAS REMOVED AND REFITTED FOR FIXATION. EXTENDED INFUSION TIME". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02979 | 2023-07-14 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, CATHETER CUFF APPEARING PAPERY WITH NO FIBROSIS. IT WAS FURTHER REPORTED THAT THE CATHETER SPONTANEOUSLY CAME OUT FROM PATIENT BODY. REPORTEDLY THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2024) H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05136 | 2023-07-14 | 2023-06-21 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAS HAD A LOT OF MEDICAL BILLS FROM GETTING URINARY TRACT INFECTION FROM THE MAGIC 3 GO INTERMITTENT CATHETER. PER FOLLOW UP VIA PHONE ON 26JUNE2023, THEY WERE GETTING URINARY TRACT INFECTION (UTIS) ALL LAST YEAR EVERY TIME THEY USED THE CATHETER. CUSTOMER STATED THEY HAD TO PAY A LOT IN MEDICAL BILLS AND PAYING FOR THE ANTIBIOTICS. CUSTOMER STATED THEY WOULD CALL BACK TO LEAVE THE NAME OF THE ANTIBIOTIC THEY WERE PRESCRIBED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INDICATIONS FOR USE: THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS, NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND / OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS, URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA, REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
3006260740-2023-02978 | 2023-07-14 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE IN THE RIGHT INTERNAL JUGULAR VEIN, CATHETER ALLEGEDLY DETACHED DURING ITS DELIVERY THROUGH THE SUBCUTANEOUS TUNNEL. IT WAS FURTHER REPORTED THAT THE CATHETER WAS REMOVED AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE UNSEALED 19CM HEMOSTAR D/L CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS OBSERVED APPROXIMATELY 5.0CM FROM THE DISTAL END OF THE Y-BODY. THE DISTAL CATHETER SEGMENT WAS RETURNED AND MEASURED APPROXIMATELY 19.0CM. THE CIRCUMFERENTIAL BREAK OF THE DISTAL CATHETER WAS OBSERVED TO HAVE UNEVEN EDGES AND SURFACE AS OBSERVED TO BE GLOSSY/FINELY GRANULAR. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS ONE UNSEALED SAMPLE TRAY CONTAINING A 19CM HEMOSTAR CATHETER AND THE CATHETER WAS NOTED TO BE BROKEN INTO TWO SEGMENTS NEAR THE CUFF AREA. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND MATERIAL SEPARATION ISSUES. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DETACHMENT ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 01/2024), G3, H6 (DEVICE). H11: H6 (METHOD, RESULT, CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00522 | 2023-07-14 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | U4150625RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE GUIDEWIRE WAS UNABLE TO BE INSERTED THROUGH THE CATHETER, AND THE CATHETER WAS ALLEGEDLY FOUND TO BE KINKED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 07/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02982 | 2023-07-14 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD TRIPLE LUMEN PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT THE STOPPER ON THE STYLET LEAKS ANYTIME SALINE IS FLUSHED WITH OUR POWERPICC PROVENA TRIPLE LUMEN KITS. WHILE FLUSHING CATHETER DURING PICC INSERTION PROCEDURE, SALINE SQUIRTS OUT WHERE THE STYLET COMES THROUGH THE STOPPER. NO OTHER INFORMATION PROVIDED. CUSTOMER IS UNABLE TO PROVIDE THE NUMBER OF INCIDENTS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05141 | 2023-07-14 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. CATHETER PLACEMENT PERIOD IS 1 DAY. THE CATHETER WAS SPONTANEOUSLY FELL OFF DUE TO RAPID DECREASE IN STERILE WATER. AFTER SPONTANEOUS REMOVAL, 10 CC OF WATER WAS INJECTED TO THE CATHETER AGAIN, AND WHEN CHECKED, ABOUT 6 CC OF WATER HAD DRAWN AFTER ONE DAY. THE SAME EVENT OCCURRED FOR THREE CONSECUTIVE DAYS IN THE SAME PATIENT (EVENT DATES (B)(6)2023, (B)(6) 2023 AND (B)(6) 2023). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02988 | 2023-07-14 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED BY THE CUSTOMER THAT TWO DIFFERENT INCIDENCES WITH SAME PRODUCT ON THE SAME DAY. BLOOD LEAKING FROM THE PICC. UNABLE TO USE PICC LINES AND INFUSE NEEDED MEDICATION UNTIL REPLACED. NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED (B)(6)2023: THE EVENT DID NOT INVOLVE AN URGENT OR LIFE THREATENING SITUATION. ONE PATIENT HAD TO TRAVEL FROM A COMMUNITY HOSPITAL TO RETURN TO FACILITY FOR CORRECTION, DELAYING TREATMENT EVEN LONGER. THIS REPORT ADDRESSES THE SECOND DEVICE AND PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02987 | 2023-07-14 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED BY THE CUSTOMER THAT TWO DIFFERENT INCIDENCES WITH SAME PRODUCT ON THE SAME DAY. BLOOD LEAKING FROM THE PICC. UNABLE TO USE PICC LINES AND INFUSE NEEDED MEDICATION UNTIL REPLACED. NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED 21 JUNE 2023: THE EVENT DID NOT INVOLVE AN URGENT OR LIFE THREATENING SITUATION. ONE PATIENT HAD TO TRAVEL FROM A COMMUNITY HOSPITAL TO RETURN TO FACILITY FOR CORRECTION, DELAYING TREATMENT EVEN LONGER. THIS REPORT ADDRESSES THE FIRST DEVICE AND PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02973 | 2023-07-14 | 2023-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716070J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, WHEN THE NON-CORING NEEDLE WAS PUNCTURED INTO THE PORT BODY AFTER THE PORT WAS IMPLANTED AND HEPARIN SALINE WAS ATTEMPTED TO BE INFUSED, INFUSION COULD NOT BE PERFORMED. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORD IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F, THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02974 | 2023-07-14 | 2023-06-20 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT ELEVEN MONTHS AND NINETEEN DAYS POST A PORT PLACEMENT, AN X-RAY EXAMINATION WAS PERFORMED AND REVEALED THAT THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THORACOTOMY WAS PERFORMED TO REMOVE THE DISTAL CATHETER SEGMENT. REPORTEDLY, THE PORT SYSTEM WAS REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02975 | 2023-07-14 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1678300 | 1678300 |  | IT WAS REPORTED THAT DURING PREPARATION OF A PORT PLACEMENT PROCEDURE, THERE WAS A GAP WHILE ATTACHING THE CATHETER TO THE HUB USING THE CATHETER LOCK, THE CATHETER ALLEGEDLY DIDN¿T GO ALL THE WAY DOWN TO THE END OF THE HUB OF THE PORT. IT WAS FURTHER REPORTED THAT THE TOP PIECE OF THE PORT THAT HOOKS INTO THE CATHETER ALLEGEDLY BROKE AND WAS DISLODGED IN THE CATHETER. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 08/2024) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
MW5119509 | 2023-07-14 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARD MAGIC 3 GO COUDE | CATHETER, STRAIGHT | BARD / C.R. BARD, INC. | 50814G | 50814G |  | I JUST LEARNED BARD(THE MANUFACTURER C.R. BARD INC., A WHOLLY OWNED SUBSIDIARY OF BD (BECTON, DICKINSON, & COMPANY), HAS A RECALL ON CERTAIN CATHETERS. BARD MAGIC 3 GO 50814G CATHETERS HAD A PROBLEM 2 1/2 YEARS AGO AND THEY DID NOT WISH TO ANNOUNCE A RECALL WAS NECESSARY. I RETURNED THE BAD CATHETERS TO BARD AND 6 MONTHS LATER ALL THEY REPORTED WAS "POSSIBLE INTERNAL OR EXTERNAL DAMAGE". THE CATHETERS ACTUALLY HAD A "BURR" ON THE TIP THAT CAUSED BLEEDING WHEN INSERTED. CASE NUMBERS FROM BARD WERE (B)(4). THEIR QUALITY CONTROL IS ACTUALLY LIBERATOR WHO SELLS PRODUCT, NOT MANUFACTURER. IN AUGUST/SEPTEMBER OF 2021, SEVERAL EMAILS WERE SENT TO (B)(4) (BARD MANAGER IN MARKETING) AND TO ONE OF THEIR EXECUTIVE VP'S ABOUT A MONTH AGO. NEVER RECEIVED A RESPONSE FROM EITHER. I EVEN REFERENCED THE "FILE OR CASE" NUMBERS RELATED TO MY FAULTY CATHETERS. I WOULD SAY THEY HAVE ABOUT 10 MANAGERS & VP'S FOR EVERY 10 WORKERS. REFERENCE REPORTS MW5119510, MW5119511, MW5119512. |  | N | Y |  |  | 2023-07
MW5119510 | 2023-07-14 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARD MAGIC 3 GO COUDE | CATHETER, STRAIGHT | BARD / C.R. BARD, INC. |  |  |  | I JUST LEARNED BARD(THE MANUFACTURER C.R. BARD INC., A WHOLLY OWNED SUBSIDIARY OF BD (BECTON, DICKINSON, & COMPANY), HAS A RECALL ON CERTAIN CATHETERS. BARD MAGIC 3 GO 50814G CATHETERS HAD A PROBLEM 2 1/2 YEARS AGO AND THEY DID NOT WISH TO ANNOUNCE A RECALL WAS NECESSARY. I RETURNED THE BAD CATHETERS TO BARD AND 6 MONTHS LATER ALL THEY REPORTED WAS "POSSIBLE INTERNAL OR EXTERNAL DAMAGE". THE CATHETERS ACTUALLY HAD A "BURR" ON THE TIP THAT CAUSED BLEEDING WHEN INSERTED. CASE NUMBERS FROM BARD WERE (B)(4). THEIR QUALITY CONTROL IS ACTUALLY LIBERATOR WHO SELLS PRODUCT, NOT MANUFACTURER. IN AUGUST/SEPTEMBER OF 2021, SEVERAL EMAILS WERE SENT TO (B)(4) (BARD MANAGER IN MARKETING) AND TO ONE OF THEIR EXECUTIVE VP'S ABOUT A MONTH AGO. NEVER RECEIVED A RESPONSE FROM EITHER. I EVEN REFERENCED THE "FILE OR CASE" NUMBERS RELATED TO MY FAULTY CATHETERS. I WOULD SAY THEY HAVE ABOUT 10 MANAGERS & VP'S FOR EVERY 10 WORKERS. REFERENCE REPORTS MW5119509, MW5119511, MW5119512. |  | N | Y |  |  | 2023-07
MW5119511 | 2023-07-14 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARD MAGIC 3 GO COUDE | CATHETER, STRAIGHT | BARD / C.R. BARD, INC. | 50814G | 50814G |  | I JUST LEARNED BARD(THE MANUFACTURER C.R. BARD INC., A WHOLLY OWNED SUBSIDIARY OF BD (BECTON, DICKINSON, & COMPANY), HAS A RECALL ON CERTAIN CATHETERS. BARD MAGIC 3 GO 50814G CATHETERS HAD A PROBLEM 2 1/2 YEARS AGO AND THEY DID NOT WISH TO ANNOUNCE A RECALL WAS NECESSARY. I RETURNED THE BAD CATHETERS TO BARD AND 6 MONTHS LATER ALL THEY REPORTED WAS "POSSIBLE INTERNAL OR EXTERNAL DAMAGE". THE CATHETERS ACTUALLY HAD A "BURR" ON THE TIP THAT CAUSED BLEEDING WHEN INSERTED. CASE NUMBERS FROM BARD WERE (B)(4). THEIR QUALITY CONTROL IS ACTUALLY LIBERATOR WHO SELLS PRODUCT, NOT MANUFACTURER. IN AUGUST/SEPTEMBER OF 2021, SEVERAL EMAILS WERE SENT TO (B)(4) (BARD MANAGER IN MARKETING) AND TO ONE OF THEIR EXECUTIVE VP'S ABOUT A MONTH AGO. NEVER RECEIVED A RESPONSE FROM EITHER. I EVEN REFERENCED THE "FILE OR CASE" NUMBERS RELATED TO MY FAULTY CATHETERS. I WOULD SAY THEY HAVE ABOUT 10 MANAGERS & VP'S FOR EVERY 10 WORKERS. REFERENCE REPORTS MW5119509, MW5119510, MW5119512. |  | N | Y |  |  | 2023-07
MW5119512 | 2023-07-14 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARD MAGIC 3 GO COUDE | CATHETER, STRAIGHT | BARD / C.R. BARD, INC. | 50814G | 50814G |  | I JUST LEARNED BARD(THE MANUFACTURER C.R. BARD INC., A WHOLLY OWNED SUBSIDIARY OF BD (BECTON, DICKINSON, & COMPANY), HAS A RECALL ON CERTAIN CATHETERS. BARD MAGIC 3 GO 50814G CATHETERS HAD A PROBLEM 2 1/2 YEARS AGO AND THEY DID NOT WISH TO ANNOUNCE A RECALL WAS NECESSARY. I RETURNED THE BAD CATHETERS TO BARD AND 6 MONTHS LATER ALL THEY REPORTED WAS "POSSIBLE INTERNAL OR EXTERNAL DAMAGE". THE CATHETERS ACTUALLY HAD A "BURR" ON THE TIP THAT CAUSED BLEEDING WHEN INSERTED. CASE NUMBERS FROM BARD WERE (B)(4). THEIR QUALITY CONTROL IS ACTUALLY LIBERATOR WHO SELLS PRODUCT, NOT MANUFACTURER. IN AUGUST/SEPTEMBER OF 2021, SEVERAL EMAILS WERE SENT TO (B)(6) (BARD MANAGER IN MARKETING) AND TO ONE OF THEIR EXECUTIVE VP'S ABOUT A MONTH AGO. NEVER RECEIVED A RESPONSE FROM EITHER. I EVEN REFERENCED THE "FILE OR CASE" NUMBERS RELATED TO MY FAULTY CATHETERS. I WOULD SAY THEY HAVE ABOUT 10 MANAGERS & VP'S FOR EVERY 10 WORKERS. REFERENCE REPORTS MW5119509, MW5119510, MW5119511. |  | N | Y |  |  | 2023-07
1018233-2023-05173 | 2023-07-15 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THERE WAS ANOTHER BURST BALLOON IN URINARY CATHETER. THE SUPRAPUBIC CATHETER WAS INSERTED TO REPLACE THE LAST PROBLEM REPORTED TO BARD. THIS TIME, THE PATIENT WAS AT HOME AND HAD KEPT THE CATHETER AFTER IT FELL OUT. THERE HAD BEEN NO TENSION ON THE INDWELLING CATHETER WHILE INDWELLING AND THIS HAD NOT HAPPENED WITH OTHER DEVICES THIS PATIENT HAS HAD INSERTED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 05JUL2023, STATED THAT THE BURST BALLOON WAS WHAT CAUSED THE CATHETER TO FALL OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05171 | 2023-07-15 | 2023-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. CATHETER PLACEMENT PERIOD IS 1 DAY. THE CATHETER WAS SPONTANEOUSLY FELL OFF DUE TO RAPID DECREASE IN STERILE WATER. AFTER SPONTANEOUS REMOVAL, 10 CC OF WATER WAS INJECTED TO THE CATHETER AGAIN, AND WHEN CHECKED, ABOUT 6 CC OF WATER HAD DRAWN AFTER ONE DAY. THE SAME EVENT OCCURRED FOR THREE CONSECUTIVE DAYS IN THE SAME PATIENT (EVENT DATE# (B)(4).). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05170 | 2023-07-15 | 2023-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED ON (B)(6) 2023, AND WAS WELL SECURED WITH MICROPORE TAPE. IT WAS STATED THAT ON (B)(6) 2023, DURING ROUTINE NURSING CARE, THE NURSE NOTICED THAT THE URINARY CATHETER SLIPPED OUT OF THE PATIENT. IT WAS STATED THAT UPON INSPECTION, IT WAS OBSERVED THAT THE BALLOON WAS DEFLATED AND THAT THERE WAS A TEAR ON THE BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE LOW LATEX VISCOSITY, SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. THEREFORE, NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02990 | 2023-07-15 | 2023-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT THREE YEARS FIVE MONTHS ELEVEN DAYS POST PLACEMENT PROCEDURE, THE PORT ALLEGEDLY SEVERAL CRACKS AT THE DISTAL END OF CATHETER. IT WAS FURTHER REPORTED THAT THE DISTAL END OF THE CATHETER ALLEGEDLY LEAKED. IT WAS ALSO REPORTED THAT THE PATIENT HAD SEVERAL PROBLEMS OVER THREE YEARS INCLUDING PAIN WITH INJECTION BUT THERE IS NO INFORMATION REGARDING THE TREATMENT. REPORTEDLY, THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 11/2022) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02991 | 2023-07-15 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616070 | 1616070 |  | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT PROCEDURE, THE WIRES WAS ALLEGEDLY STICKING OUT OF THE SKIN. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 11/2023) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02993 | 2023-07-16 | 2023-07-05 | I, n, j, u, r, y | Manufacturer report | POWERPICC 3F SL SOLO MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1193108D5 |  | CUSTOMER REPORTED "THERE HAS BEEN A PATIENT EVENT WITH HARM DUE TO INADVERTENT ARTERIAL PICC PLACEMENT. WE WOULD LIKE BD TO ADDRESS WHY THE SHERLOCK/3CG TECHNOLOGY WOULD SHOW AN INCREASE IN P-WAVE AMPLITUDE TO INDICATE THAT THE TIP WAS IN THE CORRECT POSITION, EVEN IF THE PICC WAS ACTUALLY IN THE ARTERIAL SYSTEM. IN THE PATIENT EVENT WHEN 3CG TECHNOLOGY WAS USED TO CONFIRM PROPER PICC TIP PLACEMENT, THE ECG STRIPS MET CRITERIA FOR RELEASING THE PICC FOR USE. THE PICC WAS LATER CONFIRMED BY OTHER MEASURES TO BE ARTERIAL INSTEAD OF VENOUS. THE SITUATION HAS BEEN EXTENSIVELY REVIEWED VIA OUR INSTITUTIONAL QUALITY AND PATIENT SAFETY PROCESS. PICC RNS HAVE RECEIVED REFRESHER EDUCATION AND OPPORTUNITIES TO ENHANCE PATIENT SAFETY ARE BEING EXPLORED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02992 | 2023-07-16 | 2023-07-05 | I, n, j, u, r, y | Manufacturer report | DOT PROVPPSOLO 4FDL MBP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1294108D5 |  | CUSTOMER REPORTED "THERE HAS BEEN A PATIENT EVENT WITH HARM DUE TO INADVERTENT ARTERIAL PICC PLACEMENT. WE WOULD LIKE BD TO ADDRESS WHY THE SHERLOCK/3CG TECHNOLOGY WOULD SHOW AN INCREASE IN P-WAVE AMPLITUDE TO INDICATE THAT THE TIP WAS IN THE CORRECT POSITION, EVEN IF THE PICC WAS ACTUALLY IN THE ARTERIAL SYSTEM. IN THE PATIENT EVENT WHEN 3CG TECHNOLOGY WAS USED TO CONFIRM PROPER PICC TIP PLACEMENT, THE ECG STRIPS MET CRITERIA FOR RELEASING THE PICC FOR USE. THE PICC WAS LATER CONFIRMED BY OTHER MEASURES TO BE ARTERIAL INSTEAD OF VENOUS. THE SITUATION HAS BEEN EXTENSIVELY REVIEWED VIA OUR INSTITUTIONAL QUALITY AND PATIENT SAFETY PROCESS. PICC RNS HAVE RECEIVED REFRESHER EDUCATION AND OPPORTUNITIES TO ENHANCE PATIENT SAFETY ARE BEING EXPLORED." PICC REMOVED BY VASCULAR SURGEON. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03024 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE FOURTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05209 | 2023-07-17 | 2023-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166V14S |  | IT WAS REPORTED THAT ON (B)(6) 2023, DURING PREOPERATIVE CATHETERIZATION THE NURSE INFLATED THE WATER SAC PRIOR USE TO TEST IF THE WATER SAC WOULD BREAK OR COULD BE DEFLATED SMOOTHLY. IT WAS FOUND THAT THE LIQUID IN THE BALLOON COULD NOT BE DRAINED EVEN IF THE SYRINGE WAS FORCIBLY CONNECTED INTO THE INJECTION PORT AND THE VALVE WAS EMBEDDED INTO THE LUMEN BUT STILL IT FAILED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05207 | 2023-07-17 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION (UTI) FROM FALL OF 2022 TO (B(6) 2023 BY USING THE MAGIC3 GO INTERMITTENT CATHETER OF DIFFERENT LOT NUMBERS LOT# JUFV0450, LOT# JUFZ0455, LOT# JUFZ0782, LOT# JUFZ0926, LOT# JUGP9222, LOT# JUGQ3272, LOT# JUGR0160, LOT# JUGR1384, LOT# JUGU1966. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05206 | 2023-07-17 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE FIVE OF THESE MAGIC 3 COUDE INTERMITTENT CATHETERS ONLY HAD AIR INSTEAD OF SOLUTION IN THE PACKET. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05203 | 2023-07-17 | 2023-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05204 | 2023-07-17 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS STATED THAT THE CUSTOMER STARTED USING THE FOLEY CATHETER ON (B)(6). IT WAS STATED THAT THE CATHETER WAS FOUND TO HAVE BEEN WITHDRAWN SPONTANEOUSLY FROM THE PATIENT ON (B)(6). IT WAS STATED THAT THE BALLOON HAD DEFLATED. IT WAS NOTED THAT THE EVENT OCCURRED 6 DAYS AFTER THE PLACEMENT. IT WAS STATED THAT THE THERE WAS NO BALLOON RUPTURE OR TEAR. NO MATERIALS WERE POSSIBLY THERE IN CONTACT WITH THE CATHETER AND THE CATHETER WAS INSERTED BY HAND. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. SAMPLE WAS EVALUATED AND FOUND TEAR ON SHAFT OF CATHETER THAT ALLOW WATER TO LEAK OUT. THE TEAR LOOKS LIKE IT HAD BEEN EXPOSED TO SHARP OBJECT. THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE (EXAMPLE: USER RELATED/OPERATOR ERROR/ROUGH HANDLING OR EXPOSE TO SHARP OBJECT/MECHANICAL FORCE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05201 | 2023-07-17 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | INTP14 | INTP14 |  | IT WAS REPORTED THAT THE NURSE WAS IN THE PROCESS OF OBTAINING URINE FROM A CLOSED SYSTEM URINARY CATHETER, WHEN THEY ACCESSED THE NEEDLELESS URINE SPECIMEN PORT, THE BLUE PART GOT DETACHED AND BROKE. THIS IS THE SECOND TIME IT HAPPENED IN THIS UNIT IN THE LAST 2 WEEKS. PATIENT WAS ALREADY SCHEDULED FOR ORGAN PROCUREMENT AND WAS REMOVED IN OPERATING ROOM. ALSO STATED THAT IT WAS NOT ABLE TO RETRIEVE THE DEFECTIVE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05200 | 2023-07-17 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX®LUBRICATH® FOLEY CATHETER ALCOCK MODEL | FOLEY CATHETER ALCOCK MODEL | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0134L22 |  | IT WAS REPORTED THAT FOUR DAYS AFTER PLACEMENT, A POPPING SOUND WAS HEARD WHILE THE FOLEY CATHETER WAS IN PLACE, AND WHEN CHECKED, THE BALLOON HAD BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05198 | 2023-07-17 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2023 FOR DYSURIA OVER A MONTH. THE DIAGNOSIS WERE PROSTATIC HYPERPLASIA AND CYSTS OF BOTH KIDNEYS. THE PATIENT HAD DIFFICULTY IN URINATION AND SO CATHETERIZATION WAS GIVEN ON (B)(6) 2023. ON (B)(6) 2023, THE CATHETER FELL OUT AND WAS FOUND RUPTURED WHICH WAS REPLACED WITH A NEW DISPOSABLE LATEX CATHETER. THE OPERATION WAS SUCCESSFUL WITH THE PATIENT'S VITAL SIGNS STABLE AND NO DISCOMFORT OBSERVED. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. BASED ON THE ATTACHED PHOTO, IT WAS OBSERVED THAT THE BALLOON WAS BURST. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. NO FURTHER TESTING COULD BE CONDUCTED DUE TO ONLY PHOTO SAMPLE IS PROVIDED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT/ EXPOSURE TO PETROLATUM BASED PRODUCTS MECHANICAL FAILURE/OPERATOR ERROR)/ SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03006 | 2023-07-17 | 2023-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT, "ON (B)(6) 2023, WHEN THE NURSE FLUSHED THE TUBE BEFORE GIVING THE PATIENT CHEMOTHERAPY, THE PICC EXTENSION TUBE HAD A BREAK, AND THE LIQUID WAS EXTRAVASATED, AND THE INFUSION WAS SUSPENDED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05195 | 2023-07-17 | 2023-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | INTP14 | INTP14 |  | IT WAS REPORTED THAT ATTEMPTING TO STRAIGHT CATHETER PATIENT. NURSE TECH OPENED INTERMITTENT STERILE KIT AND STARTED UNFOLDING PAPER OR WRAPPING WHEN IT WAS NOTED THAT CATHETER WAS ON THE OUTSIDE OF THE PAPER OR WRAPPING. COULD NOT OPEN LAST FOLD OF PAPER WITHOUT CONTAMINATING CATHETER OR BREAKING STERILITY. SECOND KIT WAS OBTAINED. SECOND KIT ALSO HAD PART OF CATHETER ON TOP OF PAPER OR NOT IN TRAY. THE NURSE TECH WAS ABLE TO CAREFULLY, AND SLOWLY, SLIDE THE PAPER OR WRAPPING OPEN WITHOUT CONTAMINATING CATHETER. CONCERNED THAT THERE WERE MORE KITS PACKAGED THIS WAY AND THAT THERE WOULD BE WASTE AND OR PATIENTS STRAIGHT CATHED WITH CONTAMINATED CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05194 | 2023-07-17 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | A119108 | 119108 |  | IT WAS REPORTED THAT THERE HAVE BEEN SOME PATIENTS THAT HAD A 10F OR 8FR INDWELLING CATHETER AND THE NURSE HAD DIFFICULTY REMOVING IT. APPARENTLY, THE BALLOON CAUSED A RIDGE AFTER BEING INFLATED AND MAKES IT UNREMOVABLE. THEY HAVE REINFLATED AND DELATED AND THIS SOMETIMES WORKS. THE CATHETERS CITED WERE A119108 AND A119110. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05191 | 2023-07-17 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 | FOL0102 | FOL0102 |  | IT WAS REPORTED THAT THE PATIENT HAD TELEFLEX SILICONE FOLEY IN 14FR AND IT WOULD NOT FIT THE CLAMP OF FOL0102 AND ASKED IF IT WOULD FIT FOL0105. REPRESENTATIVE EXPLAINED THAT SOME MANUFACTURER'S MIGHT HAVE A WALL THICKNESS THE CLAMP COULD NOT ACCOMMODATE AND STATED THAT THE CURE BRAND FOLEY WORKED AND THAT THEY COULD NOT GUARANTEE THAT THE FOL0105 WILL ACCOMMODATE. CUSTOMER WOULD ALSO LIKE TO MAKE A PRODUCT COMPLAINT ABOUT THE BARDIA FOLEY INSERTION TRAYS (802010 - MULTIPLE LOTS) AND QUESTIONED THE STERILITY OF THE PRODUCT AND STATED THAT THE CONTENTS WERE HAPHAZARDLY THROWN INTO THE KIT. PER FOLLOW UP VIA PHONE ON 07JUL2023, IT WAS REPORTED THAT FOR THE LAST YEAR AND A HALF THE PATIENT WAS TEMPORARILY CATHARIZED. UROLOGIST RECOMMENDED A 48 HOUR USAGE. THEY WERE CURRENTLY GETTING THEIR PRODUCTS VIA INSURANCE THROUGH THE SUPPLIER, PERSONABLY DELIVERED. THE RECENT ORDERS OF BARDIA TRAYS HAD ARRIVED WITH A STICKER DETAILING THAT IT SUPERSEDED THE STERILITY STATEMENT. THEY WERE DEEPLY CONCERNED WITH THE RESPONSES THEY HAD RECEIVED REGARDING THE STERILITY OF THE PRODUCTS. STERILITY DATE: 08-20-18 / PK764766 / LOT# NGGW2844. THIS REPORTED EVENT WAS AN INQUIRY REGARDING THE TELEFLEX 14FR DID NOT HAVE A COMPATIBLE 2-WAY STATLOCK AND THE CUSTOMER WAS LOOKING INTO A 3-WAY STATLOCK THAT WAS DESIGNED THICKER FOR PATIENT'S ANATOMY. THEY WANTED TO KNOW WHAT VARIETIES WERE AVAILABLE TO BEST FIT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INCORRECT ORIENTATION OF CATHETER IN CLAMP". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "APPLICATION TECHNIQUE PREP 1. PLACE FOLEY CATHETER INTO RETAINER. DIRECTIONAL ARROW SHOULD POINT TOWARDS CATHETER TIP, AND BALLOON INFLATION ARM SHOULD BE NEXT TO THE CLAMP HINGE. 2. CLOSE LID, BEING CAREFUL TO AVOID PINCHING THE CATHETER. 3. IDENTIFY SECUREMENT SITE BY LAYING THE DEVICE RETAINER ON THE FRONT OF THE THIGH, LEAVING 1 INCH OF CATHETER SLACK BETWEEN INSERTION SITE AND THE STATLOCK® DEVICE RETAINER. 4. AFTER PLACING THE STATLOCK® STABILIZATION DEVICE OFF TO THE SIDE, CLEANSE AND DEGREASE THE SECUREMENT SITE WITH ALCOHOL PER HOSPITAL POLICY. LET SKIN DRY. 5. APPLY SKIN PROTECTANT, IN DIRECTION OF HAIR GROWTH, TO AREA LARGER THAN SECUREMENT SITE. ALLOW TO DRY COMPLETELY (10-15 SECONDS). 6. USING PERMANENT MARKER, WRITE INITIALS AND DATE OF APPLICATION ON THE STATLOCK® DEVICE ANCHOR PAD. NOTE: ALWAYS SECURE CATHETER INTO THE STATLOCK® DEVICE RETAINER BEFORE APPLYING ADHESIVE PAD ON SKIN. PLACE AND PEEL 7. ALIGN THE STATLOCK® STABILIZATION DEVICE OVER SECUREMENT SITE LEAVING 1 INCH OF CATHETER SLACK. MAKE SURE LEG IS FULLY EXTENDED. 8. WHILE HOLDING THE RETAINER TO KEEP THE PAD IN PLACE, PEEL AWAY PAPER BACKING, ONE SIDE AT A TIME AND PLACE TENSION-FREE ON SKIN. REMOVAL TECHNIQUE DISENGAGE 1. OPEN RETAINER BY PRESSING RELEASE BUTTON WITH THUMB, THEN LIFT TO OPEN. 2. REMOVE FOLEY CATHETER FROM THE STATLOCK® DEVICE. DISSOLVE 3. WIPE THE EDGE OF THE PAD USING AT LEAST 5-6 ALCOHOL PADS UNTIL A CORNER LIFTS. THEN CONTINUE TO STROKE UNDERSURFACE OF PAD WITH ALCOHOL TO DISSOLVE ADHESIVE PAD AWAY FROM SKIN. DO NOT PULL OR FORCE PAD TO REMOVE." SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05187 | 2023-07-17 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT DURING A PRETEST, THE FOLEY CATHETER BALLOON ASYMMETRICALLY INFLATED AND THE WATER WAS UNABLE TO BE COMPLETELY REMOVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05193 | 2023-07-17 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165814 | 165814 |  | IT WAS REPORTED THAT FOLEY CATHETER WAS NOT DRAINING AND WAS LEAKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05182 | 2023-07-17 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166V12S |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS NOT DRAINING WELL (EVENT DATE : (B)(6) 2023). RECENTLY THE USER HAD COUNTERED FEW CASES (LABOURING CASES ON EPIDURAL) WITH FOLEY'S CATHETER IN, WHERE THEY COULD SEE SOME URINE IN THE URINE BAG, BUT LATER ON WERE FOUND FULL BLADDER. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 03JUL2023, STATED THAT THE REPRESENTATIVE CHECKED WITH CUSTOMER AND THEY COULD NOT RECALL THE EVENT DATE OF THE PREVIOUS CASES (EVENT DATE : UNK). | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: SALT ACCUMULATION)/BLOCK DRAINAGE LUMEN/NO DRAINAGE EYE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
2020394-2023-00521 | 2023-07-17 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DORADO | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | DR4084 | DR4084 |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE, THE RE-SHEATHER WAS ALLEGEDLY FUSED TO THE BALLOON AND PULLED THE BALLOON OFF THE SHAFT. IT WAS FURTHER REPORTED THAT THE WHEN THE COVER WAS TRIED TO REMOVE, IT ALLEGEDLY GOT STUCK ON THE BALLOON AND THE CATHETER ALLEGEDLY SNAPPED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2026) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05199 | 2023-07-17 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON BURST. IT WAS STATED THAT THE CATHETER WAS FOUND FELL OUT OF THE PATIENT IN THE WARD ON THE NEXT DAY AFTER PLACEMENT. IT WAS NOTED THAT THE BALLOON AREA WAS TORN. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 30JUN2023, STATED THAT THERE WERE NO MISSING PIECES. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. EVALUATION CONFIRMED THE BALLOON WAS BURST. DISSECTED CATHETER AND FOUND THERE WAS LIKE RESIDUE OF JELLY INSIDE THE SAC. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS: 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [HE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS: (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS: (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX. SHAPE, CONFIGURATION AND PRINCIPLES: BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX. SIZES OF CATHETERS: AVAILABLE IN SIZES 8 TO 24 EVERY 2FR. 1. FOLEY CATHETER. 2. ACCESSORIES CLOSED DRAINAGE BAG. (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER. WATER SOLUBLE LUBRICANT. ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION. TWEEZERS. GAUZE PADS. WATERPROOF SHEET. COTTON BALLS. GLOVES. STATLOCK® FOLEY. INTENDED USE & EFFECT- EFFICACY: THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. DIRECTIONS FOR USE: 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1). (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2). FIG.1. FIG.2. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3). (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4). FIG.3. FIG.4. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5). FIG.5. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7). FIG.6. FIG.7. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT: 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8). FIG. 8. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING: CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9). FIG. 9. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS). (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD: WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10). FIG. 10. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11). FIG. 11. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12). FIG. 12. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13). FIG. 13. B. BALLOON-RUPTURE METHOD: 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14). FIG. 14. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15). FIG. 15. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16). FIG. 16. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17). FIG. 17. 3. MALFUNCTION AND ADVERSE EVENTS. 1) MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE. DIFFICULTY OR FAILURE TO REMOVE THE DEVICE OCCLUSION OF CATHETER INNER LUMENS. ENCRUSTATION. ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE. DEVICE DAMAGE DUE TO INAPPROPRIATE USE. 2) ADVERSE EVENTS: URINARY-TRACT INFECTION. HEMORRHAGE, HEMATURIA. ALLERGY REACTION TO THE DEVICE. CALCULUS FORMATION. EDEMA. PAIN. DISCOMFORT. INJURY OF BLADDER OR URETHRAL. URETHRITIS, URINARY INCONTINENCE. RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03004 | 2023-07-17 | 2023-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THE PATIENT'S PICC CATHETER SLIPPED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03009 | 2023-07-17 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393230 | 5393230 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE IN THE KIDNEY, THE CATHETER WAS ALLEGEDLY KINKED DUE TO THE PHYSICIAN BEING BENT THE CATHETER AT THE CUFF UPON INSERTING AND WAS UNABLE TO BE INSERTED. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURES AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. HOWEVER THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED IMPROPER PROCEDURE ISSUE AS THE INSERTION PROCEDURE WAS PERFORMED AGAINST THE INSTRUCTION FOR USE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INSTRUCTIONS FOR USE WAS REVIEWED: THE INSTRUCTIONS FOR USE STATES "DO NOT APPLY PRESSURE CAUSING THE DISTAL TIP OF THE CATHETER TO FOLD OVER THE STEPPED SECTION". H10: B5, D4 (EXPIRY DATE: 06/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03010 | 2023-07-17 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5743690 | 5743690 |  | IT WAS REPORTED THAT PRIOR TO A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE WHITE RUBBER TUBE OF THE TUNNELLER INSIDE THE DIALYSIS CATHETER WAS ALLEGEDLY FOUND TO BE RUPTURED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS THE TUNNELER ALONG WITH THE WHITE RUBBER TUBE. THE TUBE PART HAS BEEN BROKEN. BASED ON THE PHOTO REVIEW, THE REPORTED FRACTURE CAN BE CONFIRMED. THEREFORE, THE REPORTED DEVICE DAMAGED PRIOR TO USE HAS BEEN KEPT INCONCLUSIVE AS THE EXACT TIME OF THE EVENT OCCURRED IS UNKNOWN. A DEFINITIVE ROOT CAUSE FOR THE REPORTED EVENT COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 05/2024), G3, H6 (DEVICE, METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03023 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03025 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE FIFTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03035 | 2023-07-17 | 2023-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 20G X 8 CM WITH PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F220088PT |  | IT WAS REPORTED BY CUSTOMER THAT THE CATHETER APPEARED TO BE "CLIPPED" CLOSE TO CATHETER HUB/PROXIMAL LOCATION. WHEN NEEDLE & GUIDEWIRE WAS REMOVED, SCRUB TECH NOTICED IT LOOKED CLIPPED. CATHETER KINKED WHEN HE TRIED TO ADVANCE AND IT WAS ASPIRATING AIR. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03030 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE TENTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03028 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE EIGHTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03029 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE NINTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03026 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE SIXTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03027 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE SEVENTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03021 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03020 | 2023-07-17 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2174108 |  | CUSTOMER REPORTED THAT A POWERPICC BROKE. DURING THE INFUSION OF DRUGS WAS FOUND THE LEAKAGE OF SUBSTANCE INFUSED FROM THE POINT OF EMERGENCE OF THE CATHETER. IMMEDIATELY IT IS UNCORKED THE INFUSION AND REMOVED THE PICC. ONCE REMOVED THERE WAS A PARTIAL BREAKAGE OF THE DEVICE ABOUT 8-10 CM FROM THE ANCHOR FINS. IT WAS NECESSARY TO REPOSITION ANOTHER DEVICE TO COMPLETE THERAPY. THE OPERATOR DOES NOT REPORT ANY HARM TO THE PATIENT. ADDITIONAL INFORMATION RECEIVED ON 13 JULY THE PATIENT WAS BEING ADMINISTERED SALINE SOLUTION WITH AN ANTIHISTAMINE IN IT. AFTER A FEW MINUTES THE PATIENT BEGAN TO REPORT BURNING AND FLUID LEAKING FROM THE EXIT SITE. SUBSEQUENTLY THE INFUSION WAS STOPPED AND THE PICC WAS REMOVED, WHICH WAS THEN PUNCTURED AT A DISTANCE OF 6 CM FROM THE EXIT SITE. THERE WAS NO OVERFLOW OF CHEMOTHERAPY AND THE OPERATOR WAS EQUIPPED WITH ALL PERSONAL PROTECTIVE EQUIPMENT. PREVIOUSLY THE PATIENT HAD REGULARLY PERFORMED 5 CYCLES OF CETUXIMAB/FOLFOX FROM THE PICC WITHOUT ANY PROBLEMS . | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR SINGLE LUMEN POWERPICC. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED NEAR THE 6 CM DEPTH MARKER. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED. 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER. OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING). AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03019 | 2023-07-17 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT THE CATHETER WAS FOUND RUPTURED AT THE CONNECTOR DURING MAINTENANCE. THE CONNECTOR WAS SEPARATED FROM THE CATHETER, MAKING IT IMPOSSIBLE TO CONTINUE USING THE PRODUCT. CATHETER HAD TO BE REMOVED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A SPLIT IN THE CATHETER EXTENSION TUBING IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. ONE 5 FR D/L POWERPICC WAS RETURNED FOR EVALUATION. A MICROSCOPIC OBSERVATION OF THE EXTENSION TUBING AND THE RETURNED PURPLE LUER REVEALED A COMPLETE BREAK IN THE EXTENSION TUBING WITH SOME EXTENSION TUBING STILL VISIBLE INSIDE THE DISTAL SECTION OF THE PURPLE LUER. BEACH MARKS WERE OBSERVED IN THE EXTENSION TUBING FRACTURE SURFACES, AND THE FRACTURE EDGES OF THE EXTENSION TUBING WERE UNEVEN AS SEEN WITH MATERIAL FATIGUE OF THE EXTENSION TUBING. BASED ON THE OBSERVATIONS OF THE BREAK IN THE EXTENSION TUBING OF THE RETURNED CATHETER, THE COMPLIANT OF A BREAK IN THE EXTENSION TUBING IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03018 | 2023-07-17 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG 4F SINGLE-LUMEN PICC (60CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7717405 |  | IT WAS REPORTED THAT WHEN PLACING THE CATHETER IN THIS PATIENT ON (B)(6) 2023, THE OPERATOR PRE-FLUSHED THE CATHETER IN ADVANCE TO INSPECT THE CATHETER FOR INTEGRITY, FINDING EVIDENT LIQUID LEAKAGE ABOUT 16 CM FROM THE CATHETER TIP. ONE TINY SPLIT WAS OBSERVED AT THE SITE. THE CATHETER WAS REPLACED WITH A NEW ONE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03034 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE ELEVENTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03032 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE THIRTEENTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03031 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE FOURTEENTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03033 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE TWELFTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03014 | 2023-07-17 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "PICC LEAKS AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT IS ON THE NEW ONE THAT HAS A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. NO PATIENT IMPACT REPORTED." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03015 | 2023-07-17 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8275118 |  | IT WAS REPORTED, "PATIENT WITH INDICATION FOR PASSAGE OF PICC, SINCE HE HAS RESTRICTED VENOUS ACCESS AT THE MOMENT AND HAS ORDERED VANCOMYCIN, MEROPENEM, VORICONAZOLE PLUS BASIC MEDICATIONS. HAND WASHING IS PERFORMED, THE PUNCTURE SITE IS VERIFIED, A PATIENT WITH VERY DIFFICULT VASCULAR ACCESS, FOR WHICH AN ULTRASOUND-GUIDED PROCEDURE IS PERFORMED, THE BASILIC VEIN IS ANATOMICALLY LOCATED WITH A LINEAR TRANSDUCER, A MEASUREMENT OF THE VASCULAR DIAMETER IS PERFORMED ACCORDING TO THE PASSAGE OF THE CATHETER 5FR, TRAJECTORY IS VERIFIED WITH DOPPLER, VEIN IN OPTIMAL CONDITIONS, A SINGLE PUNCTURE IS PERFORMED WITH GUIDE NEEDLE, VENOUS RETURN IS OBTAINED, METALLIC GUIDE WIRE IS INSERTED, VEIN GUIDANCE IS CONFIRMED WITH ULTRASOUND, VENOUS DILATOR IS INSERTED WITH COMPLICATIONS. A DERMOTOMY OF APPROXIMATELY 3MM IS PERFORMED, HOWEVER, WHEN TRYING TO ADVANCE THE DILATOR, DAMAGE TO IT IS OBSERVED. ITS DISTAL END LOSES ITS NORMAL CONTOUR AND FLOURISHES, WHICH IS WHY WHEN TRYING TO ADVANCE IT, VENOUS ACCESS IS LOST, ADDED TO THIS, THE METAL GUIDE BENDS WITH VERY EASILY. I CLARIFY THAT DESPITE THE AFOREMENTIONED, NO GREATER FORCE WAS EXERTED DURING THE PASSAGE OF THE DILATOR. IT IS EXPLAINED TO THE PATIENT, DESPITE THE EASE OF PUNCTURE OF THE VENOUS ACCESS, IT WAS NOT POSSIBLE FOR ME TO PASS THE DILATOR. GAUZE IS LEFT WITH A TRANSPARENT DRESSING APPLYING PRESSURE. SERVICE WHERE THE EVENT OCCURRED: HOSPITALIZATION." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03012 | 2023-07-17 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT ONE WEEK AFTER INSERTED IN THIS PATIENT, LEAKAGE OCCURRED AT THE INSERTION SITE DURING MAINTENANCE. THE CATHETER WAS PULLED OUT BY ABOUT 3CM AND LIQUIDS LEAKED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF FLUID LEAKAGE FROM THE CATHETER WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. A SINGLE VIDEO OF THE IMPLICATED 4FR S/L GROSHONG NXT CATHETER WAS RETURNED FOR EVALUATION. A FUNCTIONAL TEST OF THE INSERTED CATHETER WAS CONDUCTED; AS THE PICC WAS FLUSHED, A SPRAYING LEAK OF A CLEAR FLUID WAS SEEN EMANATING FROM THE TUBING. THE LEAK WAS LOCATED IN THE CATHETER SHAFT TUBING, BETWEEN THE TWO-PIECE CONNECTOR AND THE INSERTION SITE. IT WAS REPORTED THAT THE CATHETER LEAK HAD OCCURRED AT THE INSERTION SITE AND THE CATHETER HAD BEEN RETRACTED IN ORDER TO EXPOSE THE LEAK. THERE WERE NO VISIBLE DISTINGUISHING FEATURES WHICH COULD AID IN IDENTIFICATION OF A SPECIFIC ROOT CAUSE. IF THE PHYSICAL SAMPLE IS RECEIVED AT A LATER DATE, THIS INVESTIGATION WILL BE UPDATED WITH THE RELEVANT INFORMATION. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE IMPLICATED PRODUCT BATCH INDICATED NO ISSUES POTENTIALLY RELATED TO PRODUCT MANUFACTURE, ASSEMBLY, AND PACKAGING. BD APPRECIATES YOUR COMMUNICATING THIS CIRCUMSTANCE AND WILL RECORD THIS INFORMATION AS VALUABLE FEEDBACK INTO COMPLAINT TRENDING AND ONGOING QUALITY EFFORTS. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11 | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03022 | 2023-07-17 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED PICCLINES CAN LEAK AIR WHEN ASPIRATING AND CHECKING FOR BACKFLOW. IT WAS NOTICED IT ON SEVERAL OCCASIONS, BUT IT IS ONLY ON THE NEW ONES THAT HAVE A SMALL YELLOW "HAT" WHERE LEADERS IN THE PICC LINE ENTER. THE OLD ONES HAD A SLIGHTLY LARGER YELLOW "HAT". NO PATIENT IMPACT REPORTED. IT WAS REPORTED THIS OCCURRED WITH FOURTEEN DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03013 | 2023-07-17 | 2023-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK PICC PLUS SLIDING POST, CRESCENT PAD, INCLUDES SKIN PREP | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | VPPCSPCE |  | IT WAS REPORTED THAT WHEN THE NURSE CHECKED THE INSERTION SITE, THE PLASTIC PART OF THE STATLOCK WAS FOUND TO BE PEELED OFF. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03011 | 2023-07-17 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE DISTAL END OF THE GUIDE WIRE WAS NOT SMOOTH ENOUGH TO MAKE IT POSSIBLE TO REMOVE THE INTRODUCER NEEDLE. THE INTRODUCER NEEDLE WAS HARDLY REMOVED AFTER ITS DISTAL END WAS CUT OFF. THE INTRODUCER SHEATH WAS UNABLE TO BE INSERTED OVER THE GUIDE WIRE. IT TOOK A LONG TIME TO INSERT IT, LEADING TO A COMPLAINT ABOUT THE CLINICIAN BY THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-02998 | 2023-07-17 | 2023-06-22 | D, e, a, t, h | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA RIGHT INTERNAL JUGULAR VEIN, THE PATIENT REPORTEDLY EXPIRED DURING CARDIAC ARREST DUE TO AN AIR EMBOLISM. IT WAS FURTHER REPORTED THAT BETWEEN THE TIME REMOVING DILATOR AND INSERTING THE CATHETER, AIR WAS DRAWN INTO THE VESSEL, RESULTING IN THE PATIENT DEVELOPED AN AIR EMBOLISM. THE CATHETER SHOULD HAVE BEEN INSERTED IMMEDIATELY AFTER REMOVING DILATOR; HOWEVER, THE PHYSICIAN HAD NOT PREPARED THE CATHETER IN ADVANCE, AND TIP OF SHEATH REMAINED OPEN FOR SOME TIME UNTIL INSERTING CATHETER. REPORTEDLY, A HEAD CT SCAN CONFIRMED AIR ACCUMULATION IN THE BRAIN. AFTER THIRTY-FOURTY MINUTES OF LIFE SUPPORT, THE PATIENT EXPIRED. NO AUTOPSY WAS PERFORMED, AND THE CAUSE OF DEATH IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 04/2025). H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03050 | 2023-07-18 | 2023-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN THE NURSE PLACED THE PICC CATHETER IN THIS PATIENT, RESISTANCE WAS FELT WHEN THE SELDINGER WAS BEING INTRODUCED IN THE PROCESS OF PUNCTURE AND IT WAS REBOUNDED. UPON REMOVAL, BIFURCATIONS WERE FOUND AT THE PROXIMAL END OF THE SELDINGER. A NEW SELDINGER WAS USED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED SHEATH IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE 4.5 FR MICROEZ MICROINTRODUCER WAS PROVIDED FOR EVALUATION. THREE PHOTO SAMPLES OF A MICROINTRODUCER WERE ALSO PROVIDED AND CORRESPONDED WITH THE CONDITION OF THE RETURNED PHYSICAL SAMPLE. THE DILATOR WAS NOT RETURNED AND THE SHEATH HAS NOT BEEN SPLIT. DAMAGE AT THE DISTAL END OF THE SHEATH WAS NOTED. MICROSCOPIC INSPECTION OF THE GREY SHEATH REVEALED INWARD BUNCHED DEFORMATION. A LONGITUDINAL SPLIT WAS ALSO OBSERVED EXTENDING FROM THE DISTAL END NEAR THE BUNCHED REGION. BASED ON THE CONDITION OF THE RETURNED SAMPLE, POSSIBLE CONTRIBUTING FACTORS INCLUDE ADVANCEMENT AGAINST RESISTANCE, STEEP INSERTION ANGLE, AND INADEQUATE SKIN NICK, SINCE BUNCHING AND SPLIT DAMAGE WAS FOUND TO BE PRESENT ON THE SHEATH, THE COMPLAINT IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03045 | 2023-07-18 | 2023-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE VIA JUGULAR VEIN, INTRODUCER NEEDLE ALLEGEDLY LEAKED AIR. IT WAS FURTHER REPORTED THAT THE INTRODUCER NEEDLE WAS FOUND RUPTURED WHEN DRAWING SALINE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. ONE ELECTRONIC VIDEO WAS PROVIDED FOR REVIEW. THE VIDEO SHOWS THE CLINICIAN WAS TRYING TO ASPIRATE SALINE THROUGH THE INTRODUCER NEEDLE AND IT WAS UNSUCCESSFUL. THEN DURING FLUSHING THE SALINE AND A LEAK WAS OBSERVED THROUGH THE HUB OF THE INTRODUCER NEEDLE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND AIR/GAS IN DEVICE ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 03/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05220 | 2023-07-18 | 2023-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER HAD NO ENOUGH LUBE AND PATIENT GOT A YEAST INFECTION. PATIENT WAS ON MEDS NOW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05219 | 2023-07-18 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | A119110 | 119110 |  | IT WAS REPORTED THAT THERE HAVE BEEN SOME PATIENTS THAT HAD A 10F OR 8FR INDWELLING CATHETER AND THE NURSE HAD DIFFICULTY REMOVING IT. APPARENTLY, THE BALLOON CAUSED A RIDGE AFTER BEING INFLATED AND MAKES IT UNREMOVABLE. THEY HAVE REINFLATED AND DELATED AND THIS SOMETIMES WORKS. THE CATHETERS CITED WERE A119108 AND A119110. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT BALLOON DESIGN (BALLOON WALL THICKNESS EXCESSIVE)". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION. ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED. ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT. ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE. ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED. DIRECTIONS FOR USE: 1. WASH HANDS AND DON CLEAN GLOVES. 2. EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT. 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERIURETHRAL AREA. 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL. 5. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP. 6. DON STERILE GLOVES. 7. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE 8. POSITION FENESTRATED DRAPE ON PATIENT. 9. SATURATE 3 FOAM SWAB STICKS IN POVIDONE IODINE. 10. ATTACH THE WATER FILLED SYRINGE TO THE INFLATION PORT NOTE: IT IS NOT NECESSARY TO PRE-TEST THE FOLEY CATHETER BALLOON 11. REMOVE FOLEY CATHETER FROM WRAP AND LUBRICATE CATHETER. 12. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY. FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA MINORA. MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD. 13. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND A. WHEN CATHETER TIP HAS ENTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE B. INSERT CATHETER TWO MORE INCHES AND INFLATE CATHETER BALLOON. 14. INFLATE THE BALLOON WITH STERILE WATER TO THE VOLUME STATED ON THE INFLATION LUMEN OF THE CATHETER BY USING A WATER-FILLED SYRINGE. DO NOT EXCEED RECOMMENDED CAPACITIES AS STATED ON THE INFLATION LUMEN OF THE CATHETER. NOTE: USE OF LESS THAN THE RECOMMENDED VOLUME CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON. 15. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK. 16. SECURE THE FOLEY CATHETER TO THE PATIENT. USE THE STATLOCK ® FOLEY STABILIZATION DEVICE IF PROVIDED (SEE STATLOCK ® FOLEY STABILIZATION DEVICE INSTRUCTIONS FOR USE) NOTE: PLEASE MAKE SURE PATIENT IS APPROPRIATE FOR USE OF STATLOCK® STABILIZATION DEVICE. 17. POSITION HANGER ON BED RAIL AT THE FOOT OF THE BED NOTE: EXERCISE CARE TO KEEP BAG OFF THE FLOOR. 18. USE GREEN SHEETING CLIP TO SECURE DRAINAGE TUBE TO THE SHEET. MAKE SURE TUBE IS NOT KINKED. 19. INDICATE TIME AND DATE OF CATHETER INSERTION ON PROVIDED LABELS. PLACE DESIGNATED LABELS ON PATIENT CHART AND DRAINAGE SYSTEM. 20. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL. 21. FOLEY CATHETER REMOVAL. 1. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. 2. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. 3. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. 4. IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. 5. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05222 | 2023-07-18 | 2023-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166V14S |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE NURSE PERFORMED PREOPERATIVE CATHETERIZATION AND INSERTED THE CATHETER INTO THE PATIENT'S URETHRA. WHEN INJECTED WATER INTO THE AIR BAG AND REMOVED THE SYRINGE, IT WAS FOUND THAT THE INJECTION VALVE (AIRBAG PLUG) FELL OFF. THE AIR PLUG WAS NOT FIRMLY CONNECTED TO THE URINARY CATHETER AS PER THE DEVICE FAILURE PERFORMANCE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE CAP NOT FULLY FITTED ON TO FUNNEL. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05223 | 2023-07-18 | 2023-06-24 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER HAD NOT ENOUGH LUBE AND PATIENT GOT A YEAST INFECTION. PATIENT WAS ON MEDS NOW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05227 | 2023-07-18 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123518A |  | IT WAS REPORTED THAT A HOMECARE PATIENT WITH SUPRAPUBIC CATHETER WHEN COMPLETED FLUSHING, THE CATHETER BROKE AND THE BALLOON DID NOT DEFLATE. THE PATIENT HAD TO VISIT ER TO HAVE CENTRAL LINE WIRE TO BREAK BALLOON THROUGH PORT AND REMOVE NUB AND THEN WAS CHANGED A NEW CATHETER. THEY HAVE ALSO HAD THE EXPERIENCE WITH THE CATHETER CRACKING AGAIN IN (B)(6) 2023 WHEN OTHER CATHETER CHANGES COMPLETED AS WELL. FIRST DEVICE WAS A 16F AND THE SECOND DEVICE WAS 18F. AS PER THE EMAIL RECEIVED ON 28JUN2023, THERE WAS AN ADDITIONAL SAMPLE RETURNED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05232 | 2023-07-18 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT THE CUSTOMER RECEIVED SOME OF THE INTERMITTENT CATHETER THAT WERE EATEN AT THE TOP IN THE LAST 2 ORDERS. STATED THAT THE CATHETER WAS ¿SPLIT- CHEWED¿ IN THE END (LOT# JUGW0773, LOT# JUGT1003). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05231 | 2023-07-18 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI18 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON DID NOT INFLATE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "LOW LATEX MODULUS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿ VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. USING PROPER ASEPTIC METHODS, REMOVE CATHETER FROM PACKAGE. PREPARE PATIENT PER HOSPITAL/NURSING RECOMMENDED PROCEDURE. PROCEED WITH CATHETERIZATION USING STANDARD TECHNIQUES. INFLATE THE 10ML BALLOON WITH A MAXIMUM OF 10ML STERILE WATER BY USING A WATER-FILLED LUER-TIP SYRINGE. DO NOT USE A NEEDLE TIP SYRINGE TO INFLATE BALLOON. CONNECT CATHETER TO COLLECTION CONTAINER. TO DEFLATE THE 10ML BALLOON, GENTLY REINSERT THE LUER TIP SYRINGE INTO THE VALVE AND ASPIRATE. RECOMMENDED INFLATION CAPACITIES 5ML BALLOON: USE 5ML STERILE WATER. 10ML BALLOON: USE 10ML STERILE WATER. 30ML BALLOON: USE 35ML STERILE WATER. DO NOT EXCEED RECOMMENDED CAPACITIES.¿ H11: SECTION A THROUGH F THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05233 | 2023-07-18 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER WAS DIFFICULT TO REMOVE FROM THE FOLEY CATHETER DURING THE PRETEST. WHEN ATTEMPTING TO DRAIN WATER AFTER INJECTING STERILE DISTILLED WATER DURING A PRETEST, THE WATER COULD NOT BE REMOVED WELL, AND THE BALLOON BECAME ASYMMETRICAL SHAPE. THERE WERE MANY PROBLEMS AT THE SAME FACILITY, AND IT IS WAS A PROBLEM. PER MAIL RECEIVED FROM IBC ON 04JUL2023, STATED THAT WATER COULD NOT BE REMOVED WELL. THE BALLOON BECAME ASYMMETRICAL SHAPE BECAUSE THE WATER REMAINED IN THE MIDDLE. PER MAIL RECEIVED FROM IBC ON 05JUL2023, STATED THAT THE MANY COMPLAINTS AT THE SAME FACILITY REFERRED TO ALL THE OTHER COMPLAINTS THAT HAD BEEN ALREADY RECORDED IN THE PAST. IT DID NOT MEANT THAT THERE WERE NEW COMPLAINTS WITH UNKNOWN DETAILS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05235 | 2023-07-18 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT URINE LEAKED FROM THE BLADDER SIDE OF THE FOLEY CATHETER BIFURCATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05230 | 2023-07-18 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER CREEVY MODEL | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 01267520 | 01267520 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON COULD NOT DEFLATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05244 | 2023-07-18 | 2023-06-26 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT THE PATIENT ASKED THE FOLEY CATHETER WAS MADE WITH CHEMICALS. THE INSIDE OF THE TUBING WAS TURNING A BRIGHT PURPLE AND PATIENT WAS HAVING URINARY TRACT INFECTION SYMPTOMS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE POSSIBILITY DUE TO USER MEDICATION/DIET. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSERT FOLEY CATHETERS ONLY FOR APPROPRIATE INDICATIONS AND LEAVE IN PLACE ONLY AS LONG AS NEEDED. CDC GUIDELINES FOR APPROPRIATE INDICATIONS FOR INDWELLING URETHRAL CATHETER USE: PATIENT HAS ACUTE URINARY RETENTION OR BLADDER OUTLET OBSTRUCTION. NEED FOR ACCURATE URINE OUTPUT MEASUREMENTS. USE FOR SELECTED SURGICAL PROCEDURES. TO ASSIST IN HEALING OF OPEN SACRAL OR PERINEAL WOUNDS. PATIENT REQUIRES PROLONGED IMMOBILIZATION. TO IMPROVE COMFORT FOR END OF LIFE CARE. PROPER TECHNIQUES FOR URINARY CATHETER INSERTION: PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION. INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN THE PATIENT RECORD. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: SECURE THE FOLEY CATHETER. USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED. MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. EMPTY THE COLLECTION BAG REGULARLY USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT. GENERALLY, DRAINAGE IS ACCOMPLISHED BY INSERTING THE CATHETER THROUGH THE URETHRA AND INTO THE BLADDER. HOWEVER, DRAINAGE IS SOMETIMES ACCOMPLISHED BY SUPRAPUBIC OR OTHER PLACEMENT OF THE CATHETER, SUCH AS NEPHROSTOMY TRACT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05238 | 2023-07-18 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT THE CUSTOMER RECEIVED SOME OF THE INTERMITTENT CATHETER THAT WERE EATEN AT THE TOP IN THE LAST 2 ORDERS. STATED THAT THE CATHETER WAS ¿SPLIT- CHEWED¿ IN THE END (LOT# JUGW0773, LOT# JUGT1003). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05241 | 2023-07-18 | 2023-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DUE TO STERILE WATER LEAK. THERE WAS NO BALLOON RUPTURE OR TEAR. THE LOCATION OF LEAKAGE WAS UNKNOWN. THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05248 | 2023-07-18 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165L16 | 0165L16 |  | IT WAS REPORTED THAT FOLEY CATHETER DEFLATED ON ITS OWN WITH ONE OF EACH OF THE PRODUCTS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE 1) SECURE THE FOLEY CATHETER. USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED. 2) MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. 3) MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. 4) KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. 5) EMPTY THE COLLECTION BAG REGULARLY USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH. FOLEY CATHETER REMOVAL 1) TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TOMAKE THE SYRINGE ¿STICK¿ IN THE VALVE 2) ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. 3) USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. 4) IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. 5) SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05245 | 2023-07-18 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE NURSE CAME ROUND TO CHANGE CATHETER DUE TO PATIENT BLOCKING (MATERIAL # Z01). THE PATIENT HAD ONE CATHETER AT HOME, NURSE OPENED AND STATED THAT IT WAS FAULTY. THEY DID NOT EXPLAIN WHY IT WAS FAULTY (MATERIAL # D226914). PER ADDITIONAL INFORMATION RECEIVED FROM IBC ON 30JUN2023, STATED THAT THEY WERE NOT SURE ABOUT THE CATHETER THAT WAS IN THE PATIENT BEING FROM THE SAME LOT. BLOCKING WAS CAUSED BY A SEDIMENT BUILD UP IN THE PIPE OF THE CATHETER AND WAS VERY COMMON, IT WAS NOT A FAULT OF THE CATHETER THEY NEEDED CHANGING. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 10JUL2023, STATED THAT NO REASON WAS GIVEN REGARDING THE FAULTY CATHETER AND PATIENT DOES NOT KNOW. BLOCKING WAS NORMAL, WAS SEDIMENT BUILD UP WHICH WAS A NATURAL EVENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO DRAINAGE EYE OCCLUSION BLOCKED DRAINAGE LUMEN CAUSED BY PLC FAILURE AND ALSO MIGHT BE CONTRIBUTE BY NO DRAINAGE EYE OR USER RELATED (SALT ACCUMULATION). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05251 | 2023-07-18 | 2023-06-29 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION FROM THE INTERMITTENT CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
3006260740-2023-03051 | 2023-07-18 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THERE WAS LEAKAGE BY THE WINGS OF THE CATHETER. WHEN FLUSHING THE CATHETER THERE IS LEAKAGE BY THE WINGS DURING USE. ADDITIONAL INFORMATION 07/10/2023: THE PATIENT GOT HER TREATMENT IN A PIVC INSTEAD. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAK IN THE CATHETER IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. ONE 4FR SL POWERPICC SOLO CATHETER WAS RETURNED FOR EVALUATION. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A CIRCUMFERENTIAL SPLIT WAS OBSERVED IN THE CATHETER BODY ADJACENT TO THE MOLDED SUTURE WING. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED. ¿ FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR. ¿ 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER. ¿ OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING). AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03049 | 2023-07-18 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000952 |  | IT WAS REPORTED BY CUSTOMER THAT DURING INSERTION ON "HEALTHY" 28 YR OLD PT AFTER SKIN NICK UNABLE TO INSERT INTRODUCER/SHEATH. WHEN REMOVED IT WAS NOTED THAT THE INTEGRITY OF THE OUTER PORTION OF THE INTRODUCER/SHEATH WAS COMPROMISED. IT WAS NECESSARY TO ADD ANOTHER INTRODUCER/SHEATH TO THE STERILE FIELD IN ORDER TO COMPLETE PICC INSERTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED MICROINTRODUCER SHEATH IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE PHOTOGRAPH OF A MICROINTRODUCER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED THE DISTAL PORTION OF THE MICROINTRODUCER SHEATH WAS DAMAGED AND PLASTICALLY DEFORMED. BLOOD RESIDUE WAS OBSERVED ON THE DISTAL TIP OF THE MICROINTRODUCER IN THE RETURNED PHOTOGRAPH. NO OTHER DETAILS OF THE DAMAGE COULD BE DISCERNED IN THE RETURNED PHOTOGRAPH. WHILE THE EXACT CAUSE OF THE OBSERVED PLASTIC DEFORMATION COULD NOT BE DETERMINED FROM THE RETURNED PHOTOGRAPH, POSSIBLE CAUSES OF THE PLASTIC DEFORMATION INCLUDE EXCESSIVE INSERTION FORCE, INADEQUATE SKIN NICK, AND DAMAGE PRIOR TO ATTEMPTED USE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05260 | 2023-07-18 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 129416M | 129416M |  | IT WAS REPORTED THAT THE CUSTOMER HAD A TEMPERATURE SENSING FOLEY AND WOULD LIKE TO SEND FOR ANALYSIS. ALSO STATED THAT PER THE NURSES ON 6 NORTH THE TEMPERATURE PROBE WAS NOT WORKING, DESPITE CHANGING OUT CORDS. | THE REPORTED EVENT IS CONFIRMED, CAUSE UNKNOWN. THE USED THREE-WAY EGGSHELL COLORED LATEX FOLEY CATHETER WAS RETURNED WITHOUT THE ORIGINAL PACKAGING. NO KINKS OR VISUAL DEFORMITIES WERE NOTED ON THE SURFACE OF THE CATHETER, HOWEVER, SOME RESIDUE WAS PRESENT WITHIN THE EYELET. THE RESISTANCE READING WAS TAKEN AT ROOM TEMPERATURE ROOM TEMPERATURE AND FOUND TO BE 2.567 K OHMS. THE CATHETER WAS SUBMERGED IN A WATER BATH AND THE RESISTANCE READINGS WERE TAKEN AT 3 TEMPERATURES; VALUES BELOW. CONTINUITY TESTS WERE PERFORMED; 25 ºCELSIUS: 2.265 K OHMS RESISTANCE, DISPLAYED 25.1 ºCELSIUS; 37 ºCELSIUS: 1.356 K OHMS, DISPLAYED 37 ºCELSIUS; 41 º CELSIUS: 1.159 K OHMS, DISPLAYED 41.3 ºCELSIUS. THE TEMPERATURE BASED ON THE RESISTANCE READINGS WERE THEN CALCULATED; THE TEMPERATURE CALCULATED FROM THE RESISTANCE READING TAKEN AT 25º AND 41 ºCELSIUS WERE FOUND TO BE OUT OF SPECIFICATIONS. THOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE SENSOR WIRE TOO WEAK OR THERMISTOR WIRE TOO WEAK. THE DEVICE WAS USED FOR TREATMENT PURPOSES. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. NOTE: COMPATIBLE WITH APPROPRIATE 400-SERIES TEMPERATURE MONITORS. INTERCHANGEABILITY + 0.2ºC AT 37ºC. AS WITH ALL TEMPERATURE PROBES, IN THE PRESENCE OF RF ENERGY SOURCES, LOCAL HEATING, TEMPERATURE ERRORS, AND PROBE DAMAGE MAY OCCUR. IN MEDICAL USE, UNPLUG THE TEMPERATURE-SENSING CATHETER AT THE TEMPERATURE MONITOR BEFORE ACTIVATING ELECTROSURGICAL OR OTHER TYPES OF DIRECT COUPLED RF ENERGY SOURCES." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03053 | 2023-07-18 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR (RADEL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770352 |  | IT WAS REPORTED THE SENSOR LOSES TIP LOCATION/GIVING INACCURATE READINGS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF SENSOR LOSES TIP LOCATION/GIVING INACCURATE READINGS WAS UNCONFIRMED. THE SENSOR¿S MAGNET TRACKING AND ECG WERE EVALUATED AND FUNCTIONED PROPERLY. THE ROOT CAUSE OF THE REPORTED FAILURE IS INCONCLUSIVE AS THE REPORTED ISSUE COULD NOT BE REPRODUCED DURING EVALUATION. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03058 | 2023-07-18 | 2023-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT W/BIOPATCH AND PROBE COVER | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108BDP |  | IT WAS REPORTED VIA MEDWATCH THAT DURING INSERTION OF A 3FR POLYMIDLINE, ONE HALF OF THE HUB OF THE PEEL-AWAY SHEATH POPPED OFF OF THE ALREADY SPLIT BODY OF THE SHEATH WHILE IT WAS STILL IN THE PATIENT. THE CLINICIAN WAS ABLE TO GRASP THE END WITH MY FINGERS AND REMOVE IT ENTIRELY FROM THE PATIENT'S ARM. THE HUB AND THE BODY ARE NOT SUPPOSED TO SEPARATE FROM EACH OTHER AT ANY TIME DURING THE PROCEDURE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05266 | 2023-07-18 | 2023-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175814 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED ON (B)(6) 2023. IT WAS STATED THAT THE BALLOON IN THE CATHETER DEFLATED OR BROKE WHIST INSERTED IN THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "HIGH MODULUS SILICONE". A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿BARD® UROLOGICAL CATHETER: WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE SILICONE. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LAWS AND REGULATIONS. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LAWS AND REGULATIONS.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05265 | 2023-07-18 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L20 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER LEAKED FROM THE TIP OF THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03037 | 2023-07-18 | 2023-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6CF INT W SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806060 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA THE RIGHT INTERNAL JUGULAR VEIN, THE CATHETER WAS ADVANCED THROUGH HALF OF THE CONNECTING HANDLE BUT NOT YET THROUGH THE BUMP. IT WAS FURTHER REPORTED THAT WHEN THE FASTENERS WERE ENGAGED, NO "CLICK" WAS HEARD, AND THE CATHETER WAS NOT PULLED BY THE FASTENERS TO COVER THE ENTIRE CONNECTING HANDLE. REPORTEDLY, THE FASTENERS WERE EXTREMELY LOOSE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A VIDEO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 09/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03041 | 2023-07-18 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808050 | 1808050 |  | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED SORENESS. IT WAS FURTHER REPORTED THAT THERE IS NO INFORMATION REGARDING THE ADDITIONAL INTERVENTION PERFORMED OR MEDICATION PRESCRIBED TO TREAT THE SORENESS. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03046 | 2023-07-18 | 2023-04-26 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6CF INT W SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806060 |  | IT WAS REPORTED THAT FOUR MONTHS AND TEN DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, THE CATHETER WAS ALLEGEDLY FOUND TO BE RUPTURED. IT WAS FURTHER REPORTED THAT A CRACK WAS CONFIRMED ON THE CATHETER ONE CENTIMETER AWAY FROM THE PROXIMAL PORT HUB. IT WAS ALSO REPORTED THAT UPON INFUSION, THE SALINE ALLEGEDLY LEAKED. REPORTEDLY, THE CATHETER WAS ALLEGEDLY MIGRATED AND EMBOLIZED INSIDE THE PATIENT, AND THE PATIENT EXPERIENCED CHEST PAIN. HOWEVER, IT WAS CONFIRMED THAT CHEST PAIN IMPROVED AFTER THE REMOVAL OF THE PORT. REPORTEDLY, A NEW PORT WAS IMPLANTED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: (EXPIRATION DATE: 04/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03048 | 2023-07-18 | 2022-12-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT MRI ISP, 8FR GROSHONG, INT WO SP, ATTSL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808561 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, DEVICE ALLEGEDLY DISENGAGED. IT WAS FURTHER REPORTED THAT THE LOCK WAS ALLEGEDLY COULD NOT BE LOCKED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A VIDEO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 04/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03047 | 2023-07-18 | 2022-12-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT MRI ISP, 8FR GROSHONG, INT WO SP, ATTSL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808561 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, DEVICE ALLEGEDLY DISENGAGED. IT WAS FURTHER REPORTED THAT THE LOCK ALLEGEDLY COULD NOT BE LOCKED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A VIDEO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 04/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05303 | 2023-07-19 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WOULD NOT INFLATE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINCH LUMEN/COLLAPSE LUMEN/THICK SAC THICKNESS/VALVE DAMAGE/SHORT INFLATION LUMEN. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05305 | 2023-07-19 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION (UTI) FROM FALL OF 2022 TO JAN 2023 BY USING THE MAGIC3 GO INTERMITTENT CATHETER OF DIFFERENT LOT NUMBERS LOT# JUFV0450, LOT# JUFZ0455, LOT# JUFZ0782, LOT# JUFZ0926, LOT# JUGP9222, LOT# JUGQ3272, LOT# JUGR0160, LOT# JUGR1384, LOT# JUGU1966. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INDICATIONS FOR USE THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05267 | 2023-07-19 | 2023-05-26 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT KEPT GETTING URINARY TRACT INFECTIONS, BUT PATIENT WAS NOT SURE IF IT WAS FROM PURE WICK FEMALE EXTERNAL CATHETER. PATIENT WAS GOING TO TALK TO DOCTOR. NO MEDICAL INTERVENTION WAS REPORTED. PER CUSTOMER VIA PHONE ON (B)(6) 2023, CUSTOMER STATES FECAL MATTER GETS INTO THE WICK CAUSING AN URINARY TRACT INFECTION. CUSTOMER STATES THAT THE RUBBER NEEDS TO BE EXTENDED TO PREVENT THIS. CUSTOMER STATES THAT SHE DOES NOT KNOW IF THE WICK WAS CAUSING THE INFECTIONS. CUSTOMER WAS TREATED WITH ANTIBIOTICS. CUSTOMER ALSO STATES THAT THERE WAS BROWN STAIN THAT IS ON THE WHITE PART OF THE WICK THAT DOES NOT COME FROM HER. PER CUSTOMER VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT CUSTOMER SPOKE WITH HER UROLOGIST, AND HE INFORMED HER THAT SHE HAS A URINARY TRACT INFECTION. THE DOCTOR PRESCRIBED HER ANTIBIOTICS (MONURAL). CUSTOMER ALSO STATED THAT THE DOCTOR COULD NOT DETERMINE IF THE UTI CAME FROM THE PUREWICK CATHETER OR IF IT CAME FROM THE CUSTOMER SITTING IN URINE IN HER PADDING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05302 | 2023-07-19 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS WITHDRAWN SPONTANEOUSLY AND THE HOSPITAL CONFIRMED THAT THE BALLOON WAS INTACT UPON ANALYSIS. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINHOLE IN BALLOON OR CUFF. PHOTO: THERE ARE 2 PHOTO SAMPLES ATTACHED TO THIS INVESTIGATION FROM THE CUSTOMER. THE FIRST PHOTO IS AN OVERVIEW OF THE CATHETER. THE SECOND PHOTOS SHOW THE BURST FOUND ON THE BALLOON. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY SILICONE FOLEY CATHETER. VISUAL INSPECTION NOTED THE CATHETER BALLOON HAD BURST MEASURING 0.1795". FURTHER INSPECTION NOTED THE CATHETER BALLOON HAD NO MISSING PIECES. THIS IS OUT OF SPECIFICATION WHICH STATES, "BALLOON MUST NOT BE TORN". A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05306 | 2023-07-19 | 2023-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT FOLLOWED THE DOCTOR'S ADVICE AND PLACED AN INDWELLING CATHETER UNDER THE LUBRICATION OF NORMAL SALINE AT 15:15 PM ON (B)(6) 2023, DUE TO THE PATIENT'S DIFFICULTY IN URINATING. AFTER INSERTION, 0.9% OF NORMAL SALINE 20ML WAS INJECTED INTO THE BALLOON TO FIX IT. THE VOLUME OF THE BALLOON WAS ABOUT THE SIZE OF A TABLE TENNIS BALL, THE PROCESS WENT SMOOTHLY. AT 17:53 AFTER THE URINARY CATHETER WAS INSERTED, THE NURSE IN CHARGE OF THE BED FOUND THAT THE PATIENT'S URINARY CATHETER HAD PROLAPSED BY ITSELF. AFTER INSPECTION, IT WAS FOUND THAT THE CAUSE WAS RUPTURE OF THE URINARY CATHETER BALLOON. AFTER REPORTING TO THE DOCTOR, THEY CONTINUED TO OBSERVE. TWO HOURS LATER, THE PATIENT WAS FOUND TO HAVE HEMATURIA, AND THE PHYSICAL EXAMINATION REVEALED BLOOD CLOTS IN THE PERINEUM. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY." | N | Y | HEALTH PROFESSIONAL |  | 2023-07
17349502 | 2023-07-19 | 2023-05-21 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP FOLEY TRAY SYSTEM BARDEX IC COMPLETE CARE | CATHETER, UROLOGICAL (ANTIMICROBIAL) AND ACCESSORIES | C. R. BARD, INC. | A303316A | A303316A |  | FOLEY BALLOON BURST INSIDE THE PATIENT'S BLADDER DURING INFLATION WITH 10ML NORMAL SALINE. NO IMPACT NOTED TO PATIENT. FOLEY BALLOON IMMEDIATELY REMOVED WHEN BURST WAS HEARD. THE BALLOON WAS INSPECTED WITH CHARGE RN (REGISTERED NURSE) AND A TEAR WAS NOTED ON THE OUTER LAYER OF THE FOLEY TIP, BELOW THE DRAINAGE HOLE. A NEW FOLEY WAS PLACED. MANUFACTURER RESPONSE FOR FOLEY CATHETER, BARD (PER SITE REPORTER). EMAIL SENT TO BD SALES REP AND FILE PRODUCT COMPLAINT. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05275 | 2023-07-19 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS INSERTED IN THE UROLOGY OUTPATIENT DEPARTMENT, IT WAS WITHDRAWN SPONTANEOUSLY A FEW HOURS LATER IN THE HOSPITAL WARD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05274 | 2023-07-19 | 2023-06-22 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT PATIENT DID NOT WANT TO USE THE PUREWICK FEMALE EXTERNAL CATHETER DUE TO REOCCURRING INFECTIONS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR INFECTIONS. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, CUSTOMER DOCTOR ADVISED THEM TO STOP USE OF THE PUREWICK DUE TO EXPERIENCING FREQUENT URINARY TRACT INFECTION (UTIS). CUSTOMER WAS PRESCRIBED AN UNKNOWN ANTIBIOTIC FOR THE UTIS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05276 | 2023-07-19 | 2023-06-29 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123518A |  | IT WAS REPORTED THAT THE PATIENT WAS GETTING INFECTIONS AND THE DOCTORS SAID THAT WAS DUE TO KIDNEY STONES. THE PATIENT THOUGHT THE REASON FOR STONE FORMATION COULD BE BECAUSE OF CATHETER. IT WAS UNCLEAR IF THE LOT NUMBER WAS REQUESTED. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05269 | 2023-07-19 | 2023-06-26 | I, n, j, u, r, y | Manufacturer report | BARD® FLIP-FLO¿ CATHETER VALVE | FLIP-FLO CATHETER VALVE | C.R. BARD, INC. (COVINGTON) -1018233 | BFF20 | BFF20 |  | IT WAS REPORTED THAT IN AN ONLINE SURVEY A PHYSICIAN STATED THAT THE PATIENT EXPERIENCED UTI (URINARY TRACT INFECTION) PRIOR TO THE DEVICE PLACEMENT AND 15 DAYS AFTER THE DEVICE PLACEMENT NON-SYMPTOMATIC UTI COMPLICATIONS WERE DIRECTLY ATTRIBUTABLE TO THE DEVICE, AND PLACEMENT OF WIRED CATHETER WAS NEEDED IN RELATION TO BFF20 ¿ FLIP-FLO CATHETER VALVE WITH 180° LEVER TAP. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "WARNING: THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/ OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05279 | 2023-07-19 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THERE WAS LEAKAGE/SEEPAGE FROM CATHETER INSERTION POINT. ALSO STATED THAT THE BAG WAS IN USE FOR 22 DAYS AND THE LEAKAGE BEGAN OCCURRING ON (B)(6) 2023. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05285 | 2023-07-19 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT HAD 3 CATHETERS WITH BENT TIPS AND ONE THAT DID NOT HAVE ANY FLUID IN IT. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR." IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05291 | 2023-07-19 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WOULD NOT DEFLATE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INCORRECT BALLOON DESIGN (BALLOON WALL THICKNESS EXCESSIVE)". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿DIRECTIONS FOR USE: PROPER TECHNIQUES FOR URINARY CATHETER INSERTION: PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION. INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT. ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE. LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED. STERILE UNLESS PACKAGE IS OPENED OR DAMAGED, EXCEPT FOR ANY INDIVIDUALLY PACKAGED COMPONENTS WITHIN THE TRAY WHICH ARE NOT LABELED AS STERILE. THESE COMPONENTS ARE NOT TERMINALLY STERILIZED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. BARD® EZ-LOK® SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK (FIG. 1A) OR SLIP TIP SYRINGES (FIG. 1B). 1. OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. 2. SWAB SURFACE OF BARD® EZ-LOK® SAMPLING PORT WITH ANTISEPTIC WIPE. (FIG. 2) 3. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. (FIG. 3) 4. SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND REMOVE SYRINGE FROM SAMPLE PORT. 5. UNKINK TUBING IF NECESSARY AND TRANSFER URINE SPECIMEN INTO SPECIMEN CUP OR FOLLOW HOSPITAL PROTOCOL. DISCARD SYRINGE ACCORDING TO HOSPITAL PROTOCOL. 6. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05290 | 2023-07-19 | 2023-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER DID NOT WORK RIGHT AND PATIENT HAD TO FORCE THE CATHETER IN BECAUSE THEY WOULD NOT INSERT. IT IS UNCLEAR IF THE LOT NUMBER WAS REQUESTED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "CATHETER TOO STIFF FOR INSERTION, ROUGH OR IRREGULAR SHAPE PROTRUSIONS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-07
1018233-2023-05297 | 2023-07-19 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS LEAKING. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05296 | 2023-07-19 | 2023-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE ON (B)(6)2023, A THREE-LUMEN FOLEY CATHETER WAS INDWELLED FOR THE PATIENT. AFTER THE INDWELLING, 15ML OF NORMAL SALINE WAS INJECTED INTO THE BALLOON, DRAINAGE AND FIXATION WAS GOOD. IT WAS STATED THAT THE NO ABNORMALITY WAS FOUND. AT 9:20 ON THE 18TH, THE FOLEY CATHETER FELL OUT OF THE URETHRA BY ITSELF. THE INTEGRITY OF THE URINARY CATHETER WAS CHECKED, NO PIECES MISSING. NORMAL SALINE WAS INJECTED TO CHECK THAT THE BALLOON WAS NOT BROKEN. THEN REPLACED A NEW URINARY CATHETER FOR THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05298 | 2023-07-19 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT A FOLEY CATHETER WAS IMPLANTED FOR THE PATIENT. IT WAS STATED THAT AFTER 26 HOURS, THE URINARY CATHETER WAS FOUND TO HAVE FALLEN OUT BY ITSELF AND THE WATER BAG WAS FOUND TO HAVE RUPTURED, SO A NEW CATHETER WAS RE-INDWELLED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05299 | 2023-07-19 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT AFTER THE OPERATION ON (B)(6) 2023, THE DOCTOR GAVE A NO. 18 THREE-LUMEN CATHETER INDWELLING AND AT THE SAME TIME, THE BLADDER IRRIGATION WAS GIVEN, BUT THE IRRIGATION LUMEN OF THE CATHETER WAS FOUND TO BE NOT SMOOTH, SO THE CATHETER WAS REPLACED IMMEDIATELY. THE PATIENT HAS A POTENTIAL BLEEDING RISK, AND THE INDWELLING CATHETER WAS INVASIVE. THE OPERATION HAD POTENTIAL RISK TO BRING SECONDARY INJURY TO THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. HOWEVER, THE POTENTIAL ROOT CAUSE COULD BE MANUFACTURING RELATED DUE TO NO IRRIGATION EYE OR POOR IRRIGATION EYE CREATED DURING EYE BURNING PROCESS. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05301 | 2023-07-19 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123622A |  | THE NURSE CHANGED THE CATHETER FOR THE PATIENT, DISINFECTED IT ACCORDING TO THE REQUIREMENTS, INSERTED THE CATHETER, SAW THE URINE FLOW OUT, THEN INSERTED ANOTHER 2 CM, AND PREPARED TO INJECT STERILE WATER INTO THE BALLOON, BUT IT COULD NOT BE INJECTED, AND THEN INSERTED ANOTHER 1 CM, STILL WAS UNABLE TO INJECT. THE USER ATTEMPTED TO PULL OUT A SECTION OF THE URINARY CATHETER AND TRY TO INJECT IT, BUT IT STILL COULD NOT BE INJECTED. THE USER REPLACED THE CATHETER WITH A NEW ONE AND INSERTED IT AGAIN, AND INJECTED STERILE WATER FOR INJECTION FROM THE AIR BAG, AND FIXED IT PROPERLY. THE EVENT INCREASED PATIENT SUFFERING AND HOSPITAL CONSUMABLE COSTS. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW-UP INFORMATION RECEIVED FROM (B)(6) ON 06JUL2023, STATED THAT THE SUFFERING THE PATIENT HAD WAS PAIN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINCH LUMEN/COLLAPSE LUMEN/THICK SAC THICKNESS/VALVE DAMAGE/SHORT INFLATION LUMEN. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05277 | 2023-07-19 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT UPON ATTEMPTING TO INFLATE THE BALLOON, THE WATER LEAKED OUT. THE LEAK SEEMED TO BE CAUSED BY THE EXTRA SPACE AT THE CONNECTION BETWEEN THE INFLATION VALVE AND THE SYRINGE. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. THE DHR REVIEW IS NOT REQUIRED AND THE RISK REVIEW AND LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05300 | 2023-07-19 | 2023-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT TEMPERATURE DISPLAY FAILURE (POOR TEMPERATURE DISPLAY) OCCURRED DURING SURGERY. | THE REPORTED EVENT WAS NOT CONFIRMED BECAUSE THE AFFECTED PRODUCT MEETS SPECIFICATIONS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS NOT CONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING: CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03076 | 2023-07-20 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT WHEN PLACING THE CATHETER, THE ICU CUSTOMER PRE-FLUSHED THE CATHETER BEFORE TRIMMING IT. WHEN PULLING OUT THE STYLET, THE CUSTOMER FOUND WHITE FLOCCULE FOREIGN MATTERS ON THE WIRE. DUE TO THE CONCERN ABOUT THE PRESENCE OF THE FOREIGN MATTER IN THE CATHETER, THE CUSTOMER DARED NOT TO USE IT AND FILED THE COMPLAINT. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03078 | 2023-07-20 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT WHEN PLACING THE CATHETER, THE ICU CUSTOMER PRE-FLUSHED THE CATHETER BEFORE TRIMMING IT. WHEN PULLING OUT THE STYLET, THE CUSTOMER FOUND WHITE FLOCCULE FOREIGN MATTERS ON THE WIRE. DUE TO THE CONCERN ABOUT THE PRESENCE OF THE FOREIGN MATTER IN THE CATHETER, THE CUSTOMER DARED NOT TO USE IT AND FILED THE COMPLAINT. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF APPLICABLE INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF WHITE FOREIGN MATTER ON STYLETS IS CONFIRMED BUT THE EXACT CAUSE IS UNKNOWN. THREE 5 FR D/L POWERPICC CATHETERS WITH THEIR T-LOCK ASSEMBLIES WERE RETURNED OPENED IN THEIR ORIGINAL PACKAGING FOR EVALUATION. A MICROSCOPIC OBSERVATION REVEALED WHITE RESIDUES THROUGHOUT EACH RETURNED ENTIRE STYLET. THE MATERIAL APPEARED BUNCHED AND PEELING. THE MATERIAL APPEARED TO PARTIALLY COAT THE STYLET WIRE. IT APPEARED THAT THE HYDROPHILIC COATING APPLIED TO THE STYLET WIRE DURING PRODUCT MANUFACTURE WAS BUNCHING AND PEELING. IT COULD NOT BE DETERMINED WHAT CAUSED THE DELAMINATION OF THE COATING. CONSEQUENTLY, THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE EXPOSURE TO INCOMPATIBLE CHEMICALS, UNIDENTIFIED ENVIRONMENTAL FACTORS AFFECTING THE PROPERTIES OF THE COATING, WITHDRAWAL OF THE STYLET THROUGH A DRY T-LOCK PRIOR TO FLUSHING THE SYSTEM, AND WITHDRAWAL OF THE STYLET ACROSS THE EDGE OF THE T-LOCK BODY. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03077 | 2023-07-20 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT WHEN PLACING THE CATHETER, THE ICU CUSTOMER PRE-FLUSHED THE CATHETER BEFORE TRIMMING IT. WHEN PULLING OUT THE STYLET, THE CUSTOMER FOUND WHITE FLOCCULE FOREIGN MATTERS ON THE WIRE. DUE TO THE CONCERN ABOUT THE PRESENCE OF THE FOREIGN MATTER IN THE CATHETER, THE CUSTOMER DARED NOT TO USE IT AND FILED THE COMPLAINT. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05354 | 2023-07-20 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION (UTI) FROM FALL OF 2022 TO ON (B)(6) 2023 BY USING THE MAGIC3 GO INTERMITTENT CATHETER OF DIFFERENT LOT NUMBERS: JUFV0450, LOT#: JUFZ0455, LOT#: JUFZ0782, LOT#: JUFZ0926, LOT#: JUGP9222, LOT#: JUGQ3272, LOT#: JUGR0160, LOT#: JUGR1384, LOT#: JUGU1966. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE": REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05344 | 2023-07-20 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998006 | 998006 |  | IT WAS REPORTED THAT THE DOCTOR FOUND THAT THE BALLOON OF THE CATHETER LEAKED WATER UPON OPENING THE PACKAGE AND DID NOT USE IT ON THE PATIENT. IT WAS STATED THAT THE KIT WAS REPLACED WITH A NEW KIT OF THE X-FORCE U30 URETEROSCOPIC BALLOON DILATION CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05352 | 2023-07-20 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION (UTI) FROM FALL OF 2022 TO JAN 2023 BY USING THE MAGIC3 GO INTERMITTENT CATHETER OF DIFFERENT LOT NUMBERS LOT# JUFV0450, LOT# JUFZ0455, LOT# JUFZ0782, LOT# JUFZ0926, LOT# JUGP9222, LOT# JUGQ3272, LOT# JUGR0160, LOT# JUGR1384, LOT# JUGU1966. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05356 | 2023-07-20 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION (UTI) FROM (B)(6) BY USING THE MAGIC3 GO INTERMITTENT CATHETER OF DIFFERENT LOT NUMBERS LOT# JUFV0450, LOT# JUFZ0455, LOT# JUFZ0782, LOT# JUFZ0926, LOT# JUGP9222, LOT# JUGQ3272, LOT# JUGR0160, LOT# JUGR1384, LOT# JUGU1966. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS: URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
3006260740-2023-03067 | 2023-07-20 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER "THE METAL GUIDE HAS FRAYED INSIDE THE NEEDLE. TO INTERVENE AND RECOVER THE ENTIRE GUIDE." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03075 | 2023-07-20 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D5 |  | IT WAS REPORTED BY CUSTOMER THAT PICC HAS A HOLE IN IT ON A PATIENT. ADDITIONAL INFO FROM CUSTOMER: (06-JULY-2023). NO HARM TO THE PATIENT, REMOVED THE PICC AND PLACED A NEW ONE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED CATHETER IS CONFIRMED. ONE 4 FR SINGLE LUMEN POWERPICC WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLE. A LARGE SPLIT WAS OBSERVED IN THE EXTENSION LEG TUBING, ABOUT 1 CM PROXIMAL TO THE MOLDED JOINT. A MICROSCOPIC OBSERVATION REVEALED THE EDGES OF THE SPILT WERE SOMEWHAT UNEVEN BUT SHARP AND DISTINCT. STRIATIONS WERE OBSERVED ON THE FRACTURE SURFACES, WHICH WERE MOSTLY FLAT AND LUSTROUS. THE CHARACTERISTICS OF THE SPLIT OBSERVED IN THE RETURNED SAMPLE SUGGEST THE CATHETER WAS DAMAGED BY A SHARP INSTRUMENT SUCH AS A SCALPEL. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03083 | 2023-07-20 | 2023-06-20 | I, n, j, u, r, y | Manufacturer report | FULL KIT 18G X 10 CM WITH GUARDIVA AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F118108PT |  | IT WAS REPORTED VIA MEDWATCH "THE PATIENT HAD AN 18G RIGHT UPPER ARM MIDLINE (BASILIC VEIN) CATHETER PLACED ON (B)(6) 2023. ON (B)(6) 2023 WHILE AT OUTPATIENT THERAPY TO GET HIS IV ANTIBIOTICS THE LINE WAS NOTED TO BE LEAKING. THE RN PULLED DOWN THE DRESSING AND NOTICED THE CATHETER HUB AND PART OF THE CATHETER HAD BROKEN OFF, LEAVING A PART OF THE CATHETER IN THE PATIENT'S ARM. THE PATIENT WAS TAKEN TO THE ED(EMERGENCY DEPT) AND THEN WENT TO THE OPERATING ROOM FOR A VENOTOMY WHERE THE BROKEN PIECE IN THE PATIENT'S ARM WAS SUCCESSFULLY REMOVED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05329 | 2023-07-20 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L22 |  | IT WAS REPORTED THAT WATER LEAKAGE OCCURRED FROM THE FOLEY CATHETER BALLOON SECTION (LOT# MYFU1698). IT WAS STATED THAT THE USED PRODUCTS HAD BEEN DISCARDED. THE CUSTOMER REQUESTED THAT ALL OTHER LOTS WITH THE SAME CATALOG NUMBER BE INVESTIGATED, SO THEY WERE COLLECTED. IT WAS STATED THAT DUE TO THE FREQUENT OCCURRENCE OF DEFECTS (OTHER EXTERNAL REFERENCE # (B)(4)AND OTHER EXTERNAL REFERENCE #(B)(4)) THIS YEAR, REPRESENTATIVES RECEIVED A REQUEST TO CHECK ALL OTHER LOTS OF THE SAME PART NUMBER(LOT# MYFY3438,LOT# MYEW1667, LOT# MYFP2501), SO THEY COLLECTED THEM. IT WAS STATED THAT THERE WERE MANY PROBLEMS WITH THIS PRODUCT. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 10JUL2023, IT WAS STATED THAT NO OTHER MALFUCTION FOUND IN THE DEVICE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 18JUL2023, IT WAS CONFIRMED THAT THE DEVICE WAS USED ON THE PATIENT. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTIONS: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05328 | 2023-07-20 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L20 |  | IT WAS REPORTED THAT A FEW HOURS AFTER SURGERY, THE CATHETER BALLOON SPONTANEOUSLY FELL OUT OF THE PATIENT. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05327 | 2023-07-20 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT. UPON CHECKING THE BALLOON, IT REVEALED TWO DAMAGES ON THE BALLOON. AS PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 11JUL2023, THE BALLOON WAS BURST. THE EXACT LOCATION OF DAMAGE WAS PROVIDED AS ONLY ON BALLOON. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. THERE ARE 3 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE FIRST PHOTO ATTACHED IS AN OVERVIEW OF A 3-WAY TEMPERATURE SENSING CATHETER. THE SECOND PHOTO SHOW THE RUPTURED CATHETER BALLOON AND TIP. THE THIRD PHOTO SHOWS A WHITE HANDWRITTEN DOCUMENT. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A -WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED THE CATHETER BALLOON HAD RUPTURED MEASURING 0.2295". NO MISSING PIECES NOTED. THIS IS OUT OF SPECIFICATION WHICH STATES, "BALLOON MUST NOT BE TORN". A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINHOLE IN BALLOON OR CUFF. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05324 | 2023-07-20 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE BALLOON WAS RUPTURED AND DID NOT INFLATE AFTER INJECTING STERILE WATER. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 11JUL2023, STATED THAT WHEN THE SAMPLE WAS CHECKED, IT WAS FOUND THAT THERE WAS A PINHOLE IN THE BALLOON, NOT A BALLOON RUPTURE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05323 | 2023-07-20 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT AFTER INSERTION OF THE CATHETER, ATTEMPTED TO INJECT WATER WITH THE SYRINGE IN THE KIT, BUT WAS UNABLE TO DO SO SUCCESSFULLY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05321 | 2023-07-20 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION (UTI) FROM FALL OF 2022 TO JAN 2023 BY USING THE MAGIC3 GO INTERMITTENT CATHETER OF DIFFERENT LOT NUMBERS LOT# JUFV0450, LOT# JUFZ0455, LOT# JUFZ0782, LOT# JUFZ0926, LOT# JUGP9222, LOT# JUGQ3272, LOT# JUGR0160, LOT# JUGR1384, LOT# JUGU1966. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INDICATIONS FOR USE THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS: URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05320 | 2023-07-20 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION (UTI) FROM FALL OF 2022 TO (B)(6) 2023 BY USING THE MAGIC3 GO INTERMITTENT CATHETER OF DIFFERENT LOT NUMBERS LOT# JUFV0450, LOT# JUFZ0455, LOT# JUFZ0782, LOT# JUFZ0926, LOT# JUGP9222, LOT# JUGQ3272, LOT# JUGR0160, LOT# JUGR1384, LOT# JUGU1966. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INDICATIONS FOR USE" THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA AND IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-07
3006260740-2023-03073 | 2023-07-20 | 2023-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5733690 | 5733690 |  | IT WAS REPORTED THAT PRIOR TO A DIALYSIS CATHETER PLACEMENT PROCEDURE IN THE RIGHT INTERNAL JUGULAR VEIN, THE HEALTHCARE PERSONNEL FOUND THAT THE SHEATH HAD ALLEGEDLY FRACTURED BEFORE BEING USED. IT WAS FURTHER REPORTED THAT THE GUIDEWIRE HAD ALLEGEDLY BECOME TORTUOUS. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 19CM HEMOSPLIT D/L CATHETER KIT WAS RECEIVED FOR EVALUATION AND THREE PHOTOS WERE PROVIDED FOR REVIEW. VISUAL AND MICROSCOPIC EVALUATIONS WERE PERFORMED. THE T-HANDLE VALVE WAS NOTED TO BE DAMAGED. MATERIAL DEFORMATION TO DISTAL TIP OF THE DILATOR. THE PHOTO SHOWS THE INTRODUCER SHEATH IN WHICH THE VALVE CAP WAS NOTED TO BE SEPARATED AND THE SHEATH WAS PARTIALLY PEELED. THE GUIDEWIRE WAS NOTED TO BE KINKED AND TWISTED IN THE PHOTOS. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED GUIDEWIRE TWIST AND THE IDENTIFIED DEFORMATION DUE TO COMPRESSIVE STRESS, SEPARATION FAILURE AND MATERIAL SEPARATION ISSUES. HOWEVER, THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED FRACTURE ISSUE AS NO SHEATH FRACTURE WAS NOTED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 06/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05313 | 2023-07-20 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE PATIENT SAID THAT THEY DID NOT SEE WHICH WAY THE MAGIC 3 COUDE INTERMITTENT CATHETER NEED TO BE INSERTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
3006260740-2023-03071 | 2023-07-20 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | CUSTOMER REPORTED CATHETER THAT LEAKED FROM THE RUBBER-STOPPER WHEN PICC LINE WAS BEING INSERTED. PICC LINE AVAILABLE FOR EVALUATION. ADDITIONAL INFORMATION PROVIDED JULY 5: THE CATHETER DID NOT BREAK. EVENT DID NOT INTERRUPT TREATMENT OR CAUSE CLINICALLY SIGNIFICANT DELAY TO TREATMENT. THERE WERE TWO SAMPLES RETURNED FOR INVESTIGATION, THIS FILE REPRESENTS THE SECOND RETURNED SAMPLE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED, AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05316 | 2023-07-20 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE STERILE WATER COULD NOT BE INJECTED. WHEN THE FOLEY CATHETER WAS INSERTED AND STERILE WATER WAS INJECTED INTO THE CATHETER DURING A ROUTINE CATHETER REPLACEMENT TO THE PATIENT, IT WAS QUITE DIFFICULT TO INJECT THE WATER. WHEN THE SYRINGE PLUNGER WAS PUSHED A LITTLE MORE, THE STERILE WATER LEAKED FROM THE SYRINGE AND THE CATHETER VALVE. THEREFORE THE CATHETER WAS PULLED OUT AND ANOTHER CATHETER WAS INSERTED. THE CATHETER COULD BE PLACED AS USUAL BUT THE CATHETER THAT COULD NOT BE INJECTED WITH STERILE WATER WAS VERY HARD AND DIFFICULT TO INJECT. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 11JUL2023, THE CATHETER WAS NOT HARD, BUT THE STERILE WATER INFUSION WAS HARD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05307 | 2023-07-20 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | SURESTEP INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | INTP14 | INTP14 |  | IT WAS REPORTED THAT THE NURSE WAS IN THE PROCESS OF OBTAINING URINE FROM A CLOSED SYSTEM URINARY CATHETER, WHEN THEY ACCESSED THE NEEDLELESS URINE SPECIMEN PORT, THE BLUE PART GOT DETACHED AND BROKE. THIS IS THE SECOND TIME IT HAPPENED IN THIS UNIT IN THE LAST 2 WEEKS. PATIENT WAS ALREADY SCHEDULED FOR ORGAN PROCUREMENT AND WAS REMOVED IN OR. ALSO STATED THAT IT WAS NOT ABLE TO RETRIEVE THE DEFECTIVE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03065 | 2023-07-20 | 2023-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7957505J |  | IT WAS REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. MORPHINE HYDROCHLORIDE WAS INFUSED USING A SYRINGE PUMP. GRADUALLY AN ALARM TO NOTIFY CATHETER OCCLUSION SOUNDED. ON (B)(6) 2023, CHANGED THE ROUTE THROUGH THE LUMEN OF THE RED LUER CONNECTOR SIDE AND THE CATHETER WAS LOCKED WITH LOW DOSE HEPARIN SALINE. ON (B)(6) 2023, WHEN THE CATHETER WAS FLUSHED WITH HEPARIN SALINE, LEAKAGE OCCURRED AND A CRACK WAS FOUND. THE CATHETER MAY HAVE BEEN KINKED. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 5FR D/L GROSHONG CATHETER. THE DAMAGE OBSERVED IN THE RETURNED SAMPLE WAS CHARACTERISTIC OF OVER-PRESSURIZATION (BURST) DAMAGE. THIS CAN OCCUR THROUGH THE USE OF SYRINGES SMALLER THAN 10ML, BY FLUSHING AGAINST AN OCCLUSION, OR DUE TO EXCESSIVE FORCE APPLIED DURING INFUSION PROCEDURES. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED IN THE DISTAL LUMEN EXTENSION TUBE, JUST DISTAL OF THE WHITE SLEEVE. THIS CATHETER DAMAGE WAS TYPICAL OF A BURST, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: 'C' SHAPED SPLIT; TENSILE WEAKNESS AT THE FRACTURE SITE (DUE TO MATERIAL BALLOONING PRIOR TO BURST); GRANULAR FRACTURE SURFACE TEXTURE (TYPICAL OF TEARING FAILURE MODES); THE CATHETER MATERIAL WAS VISIBLY LIGHTER IN COLOR AT THE FRACTURE SITE. AN OCCLUSION WAS OBSERVED DISTAL TO THE BURST SITE, AND THIS MAY HAVE CONTRIBUTED TO THE FAILURE. FORCEFUL FLUSHING AGAINST AN OCCLUSION CAN CAUSE THIS TYPE OF FAILURE. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03072 | 2023-07-20 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | CUSTOMER REPORTED CATHETER THAT LEAKED FROM THE RUBBER-STOPPER WHEN PICC LINE WAS BEING INSERTED. ADDITIONAL INFORMATION PROVIDED JULY 5: THE CATHETER DID NOT BREAK. EVENT DID NOT INTERRUPT TREATMENT OR CAUSE CLINICALLY SIGNIFICANT DELAY TO TREATMENT. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03066 | 2023-07-20 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER, 4F DL PICC KINKED IN ARM COUPLE DAYS AFTER BEING INSERTED. ADDITIONAL INFO FROM CUSTOMER: ((B)(6) 2023): IT WAS REPORTED "SOFT TISSUE SACROCOCCYGEAL MASS. THIS WAS AN INPT. LEFT BASILIC VEIN WAS USED. PICC WAS REMOVED BY THE PROVIDER. DELAY IN IV MEDICATIONS. PICC HAD TO BE REPLACED AND THE LEFT ARM USED. DELAY IN CARE. PATIENT HAD TO ENDURE ANOTHER PROCEDURE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A KINK IN THE CATHETER WAS INCONCLUSIVE DUE TO THE SAMPLE CONDITION. A SINGLE RADIOGRAPHIC IMAGE WAS RETURNED FOR EVALUATION OF THIS COMPLAINT. THE IMAGE WAS OF THE ARM AND CHEST WITH AN ANTECUBITAL PICC. ONLY THE PROXIMAL END OF THE PICC WAS VISIBLE IN THE IMAGE. FROM THIS VIEW, IT COULD NOT CONCLUSIVELY BE DETERMINED IF A KINK WAS PRESENT IN THE CATHETER. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05312 | 2023-07-20 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 265716 | 265716 |  | IT WAS REPORTED THAT THE PATIENT HAD TO CHANGE THE FOLEY AFTER 2 DAYS OF LEAKAGE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO WRONG PRODUCT SIZE SELECTED. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: 1) SECURE THE FOLEY CATHETER. USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED. 2) MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. 3) MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. 4) KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. 5) EMPTY THE COLLECTION BAG REGULARLY USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH. FOLEY CATHETER REMOVAL: 1) TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE 2) ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY 3) USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION 4) IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL 5) SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-07
3006260740-2023-03064 | 2023-07-20 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED VIA THE LEFT BASILIC VEIN ON (B)(6) 2023. THE CATHETER WAS IN USE AND ADMINISTERED WITHOUT PROBLEMS UNTIL (B)(6) 2023. DURING THAT TIME, THE DRESSING WAS CHANGED EVERY WEEK. ON (B)(6) 2023, THE CATHETER WAS FOUND TO BE BROKEN INSIDE THE OVERSLEEVE. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR S/L GROSHONG CATHETER. USAGE RESIDUES WERE ABUNDANT THROUGHOUT THE SAMPLE. THE CATHETER WAS ASSEMBLED WITH A TWO-PIECE CONNECTOR. THE SAMPLE WAS RECEIVED IN TWO SEGMENTS WHICH APPEARED TO SHARE A COMPLETE BREAK WITHIN THE DISTAL CONNECTOR SEGMENT. CURVED SHAPE MEMORY WAS OBSERVED NEAR THE BREAK SITE. MICROSCOPIC INSPECTION OF THE BREAK REVEALED A SHARPLY DEFINED, GRANULAR FRACTURE SURFACE. TACTILE INSPECTION OF THE SAMPLE REVEALED SEVERE TENSILE WEAKNESS, MATERIAL NECKING AND DISCOLORATION IN THE VICINITY OF THE BREAK. THE BREAK FEATURES, SHAPE MEMORY AND TENSILE WEAKNESS WERE CONSISTENT WITH MATERIAL FAILURE DUE TO TENSILE (PULLING) STRESS. THE LOCATION OF THE DAMAGE SUGGESTED THAT CATHETER PLACEMENT, SECUREMENT AND MAINTENANCE CONDITIONS MAY HAVE CONTRIBUTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03063 | 2023-07-20 | 2023-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THERE WAS A HOLE IN THE PICC LINE. THE INCIDENT OCCURRED DURING USE. PATIENT CAME IN FOR CARE AND MAINTENANCE ON 27/6. WHEN FLUSHING THE PICC LINE, IT LEAKS WITH SALINE AND A LITTLE BLOOD COMES FROM THE INJECTION SITE. A DECISION IS MADE TO WITHDRAW THE PICC LINE. A CLEAR HOLE IS SEEN 5CM FROM THE 0-MARKING. THE PATIENT IS DOING WELL BUT HAS HAD SEVERAL PICC-LINE INSERTIONS AND IS NOW GETTING ANOTHER ONE. THE PATIENT HAS HAD TWO TREATMENTS (CYTOTOXIC) WITH A PUMP IN THE PICC LINE THAT HE RECEIVED (B)(6) 2023. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A TEAR IN THE CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR SL POWERPICC SOLO CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED AT THE 6CM MARKER. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED ¿ FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05330 | 2023-07-20 | 2023-06-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION (UTI) FROM FALL OF 2022 TO ON (B)(6) 2023 BY USING THE MAGIC3 GO INTERMITTENT CATHETER OF DIFFERENT LOT NUMBERS: JUFV0450, LOT#: JUFZ0455, LOT#: JUFZ0782, LOT#: JUFZ0926, LOT#: JUGP9222, LOT#: JUGQ3272, LOT#: JUGR0160, LOT#: JUGR1384, LOT#: JUGU1966. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND / OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND / OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND / OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-07
1018233-2023-05332 | 2023-07-20 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT THERE WAS A PUNCTURE NEEDLE DAMAGE IN THE FOLEY CATHETER. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED IT WAS UNKNOWN WHAT THE CUSTOMER MEANT BY PUNCTURE NEEDLE DAMAGE. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ MECHANICAL FAILURE/OPERATOR ERROR. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNING: THIS IS A SINGLE USE DEVICE. DO NO RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH LEAD TO INJURY, ILLNESS, OR DEATH OF THE PATIENT. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE THE BALLOON TO BURST. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05333 | 2023-07-20 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THERE WAS A PUNCTURE NEEDLE DAMAGE IN THE FOLEY CATHETER. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 14JUL2023, STATED IT WAS UNKNOWN WHAT THE CUSTOMER MEANT BY PUNCTURE NEEDLE DAMAGE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX.. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05334 | 2023-07-20 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175816 |  | IT WAS REPORTED THAT FOLEY TRAY 10CC SYRINGE OF STERILE WATER TIP BROKE OFF IN THE INFLATION PORT. PER FOLLOW UP VIA EMAIL ON 28JUN2023, IT WAS STATED THAT THE ONLY SAMPLE AVAILABLE WAS THE FOLEY THAT HAD A HOLE IN THE TUBING NOT THE FOLEY WHERE THE BALLOON WAS LOPSIDED BECAUSE IT WAS USED. THEY COULD NOT SAVE IT. PATIENT WITH LOPSIDED FOLEY WAS OKAY, BUT WAS VERY UNCOMFORTABLE THEREFORE THEY REMOVED THE FOLEY AND THAT WAS WHEN THEY REALIZED THAT THE PRODUCT WAS DEFORMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05336 | 2023-07-20 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT AND UPON CHECKING THE BALLOON, IT REVEALED TWO DAMAGES. AS PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 11JUL2023, THE BALLOON WAS BURST. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. THERE ARE 3 PHOTOS THAT WERE SENT BY THE CUSTOMER. THE FIRST PHOTO ATTACHED IS AN OVERVIEW OF A 3-WAY TEMPERATURE SENSING CATHETER AND SYRINGE. THE SECOND PHOTO SHOW THE RUPTURED CATHETER BALLOON AND TIP. THE THIRD PHOTO SHOWS A WHITE HANDWRITTEN DOCUMENT. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A -WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED THE CATHETER BALLOON HAD RUPTURED MEASURING 0.2635". NO MISSING PIECES NOTED. THIS IS OUT OF SPECIFICATION WHICH STATES, "BALLOON MUST NOT BE TORN". A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINHOLE IN BALLOON OR CUFF. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05339 | 2023-07-20 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165818 | 165818 |  | IT WAS REPORTED THAT THERE WAS A RIDGING ON THE FOLEY CATHETER BALLOON WHEN REMOVING. AS PER FOLLOW UP VIA EMAIL ON (B)(6) 2023, IT WAS STATED THAT THE CUSTOMER DID NOT HAVE THE ORIGINAL PACKAGING, SO NO LOT OR BATCH NUMBER. THE CATHETER WAS INSERTED IN A MALE PATIENT WITH NO CONCERNS IDENTIFIED AND THEN REMOVED 28 DAYS LATER FOR A ROUTINE CHANGE. THERE WAS NO RESIDUAL FLUID LEFT IN THE BALLOON, BUT A MARKED RIDGING PRESENT. THIS CAUSED THE PATIENT PAIN AND BLEEDING AT THE TIME, BUT FORTUNATELY THIS SETTLED DOWN WITHIN A FEW HOURS, AND OUTPUT REPORTED AS CLEAR OR AMBER URINE SUBSEQUENTLY. NURSES DID NOT PRETEST BALLOONS, AND IT WAS FILLED WITH 10ML AS PER PRODUCT SPECIFICATIONS. PATIENT WAS A YOUNG MALE POST TRAUMATIC INJURY, AND WAS UNABLE TO VERBALIZE THE EXPERIENCE, BUT THERE WAS RESISTANCE UPON REMOVAL AND NONVERBAL EXPRESSION OF PAIN. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05343 | 2023-07-20 | 2023-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT DURING A PRETEST, THE BALLOON DID NOT INFLATE DUE TO A HOLE IN THE BALLOON. WHEN ATTEMPTING TO INJECT STERILE WATER INTO THE BALLOON, THE BALLOON DID NOT INFLATE, AND UPON CHECKING IT, THERE WAS A HOLE FOUND IN THE BALLOON. AS PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON (B)(6)2023, THE WATER LEAKED FROM THE BALLOON AND SO THE CATHETER COULD NOT BE INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05341 | 2023-07-20 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT CATHETER VALVE DISCONNECTED DURING PRETEST. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. SAMPLE HAS BEEN EVALUATED AND INSPECTED THE EXTERIOR OF THE SAMPLE AND FOUND THE CAP WAS DETACHED FROM INFLATION FUNNEL. DISSECTED THE CATHETER AND THE EDGES OF THE INFLATION FUNNEL WAS STILL MET THE MANUFACTURING SPECIFICATION AND THE ACCEPTANCE CRITERIA. TRIED TO ASSEMBLE THE CAP AND IT WAS FUNCTIONAL AND SAC OF CATHETER CAN BE INFLATED AND DEFLATED WITHOUT ANY DIFFICULTIES. HOWEVER, THE EXACT CAUSE ON HOW AND WHEN THE EVENT OCCURRED COULD NOT BE DETERMINE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE, ¿ROUGH SHAFT/INFLATION FUNNEL WEIGHT OUT OF SPECIFICATION/MISHANDLING CATHETER. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05342 | 2023-07-20 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L22 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE BALLOON. ON (B)(6) 2023, THE CATHETER WAS INSERTED IN THE OPERATION ROOM AND WHILE RETURNING TO THE PATIENT'S ROOM, THE CATHETER DID NOT WITHDRAW WHEN CHECKED. AT 17:00 THE PATIENT REPORTED THAT THE PUBIC AREA WAS WET AND THEN THE CATHETER WAS REMOVED. UPON TESTING THE REMOVED BALLOON BY INJECTING STERILE WATER, IT LEAKED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05361 | 2023-07-20 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT AFTER INSERTING THE CATHETER INTO THE PATIENT, ATTEMPTED TO INJECT STERILE WATER WITH THE SYRINGE IN THE SET, BUT THE TIP OF THE SYRINGE FELL OFF. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 13JUL2023, IT WAS NOT A BROKEN SYRINGE, BUT A DISCONNECTED INFLATION VALVE OF THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. SAMPLE HAS BEEN EVALUATED AND INSPECTED THE EXTERIOR OF THE SAMPLE AND FOUND THE CAP WAS DETACHED FROM INFLATION FUNNEL. DISSECTED THE CATHETER AND THE EDGES OF THE INFLATION FUNNEL WAS STILL MET THE MANUFACTURING SPECIFICATION AND THE ACCEPTANCE CRITERIA. TRIED TO ASSEMBLE THE CAP AND IT WAS FUNCTIONAL AND SAC OF CATHETER CAN BE INFLATED AND DEFLATED WITHOUT ANY DIFFICULTIES. HOWEVER, THE EXACT CAUSE ON HOW AND WHEN THE EVENT OCCURRED COULD NOT BE DETERMINE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE, ¿ROUGH SHAFT/INFLATION FUNNEL WEIGHT OUT OF SPECIFICATION/MISHANDLING CATHETER. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05362 | 2023-07-20 | 2023-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HEMATURIA FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 2551H22 |  | IT WAS REPORTED THAT DURING A PRETEST, A DAMAGE WAS FOUND IN THE BALLOON AREA OF THE FOLEY CATHETER. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 19JUL2023, DID NOT CLARIFY IF IT WAS A PINHOLE DAMAGE OR BALLOON BURST AS THE SAMPLE WASN'T REACHED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR PATIENT TREATMENT. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), USED FOLEY CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED THE CATHETER BALLOON WAS INFLATED WITH 30 ML OF DI WATER AND NO BALLOON BURST WAS OBSERVED. WITH THE SYRINGE ATTACHED THE BALLOON DEFLATED PASSIVELY WITH NO ISSUES. FLUSH THE DRAINAGE LUMEN AND NO ISSUES OBSERVED. GROSS VISUAL EVALUATION 2: TWO PHOTO SAMPLES WERE PROVIDED COULD NOT BE EVALUATED FOR THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE INVESTIGATION WAS UNCONFIRMED. HOWEVER, A DHR WAS COMPLETED BEFORE UNCONFIRMING THE EVENT. THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05363 | 2023-07-20 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRETEST, THE STERILE WATER COULD NOT BE REMOVED AND THE BALLOON REMAINED INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03082 | 2023-07-20 | 2023-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD NIAGARA DIALYSIS CATHETER | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | CUSTOMER REPORTED THAT WHEN THE NURSE MEASURED THE BLOOD GLUCOSE OF THE PATIENT, THERE WAS NO ANOMALY IN THE TRANSFUSION AND THE INDWELLING NEEDLE WAS IN GOOD CONDITION. WHEN THE NURSE OF THE RESPONSIBILITY GROUP TOOK OVER THE AFTERNOON AND MEASURED THE PATIENT'S BLOOD PRESSURE AT 14:10, IT WAS FOUND THAT THE HEPARIN CAP OF THE PATIENT'S INDWELLING NEEDLE HAD BEEN SEPARATED FROM THE INDWELLING NEEDLE, THE TWO LIQUID NEEDLES FOR INFUSION AT THE END OF THE HEPARIN CAP HAD FALLEN OFF, AND BLOOD WAS SEEN FLOWING FROM THE BED SHEET AND THE ARM OF THE PUNCTURE SIDE WITH THE INFUSION LIQUID. THE DEPARTMENT REPORTED THAT THE HEPARIN CAP INTERFACE OF THIS INDWELLING NEEDLE WAS GENERALLY LOOSE. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03079 | 2023-07-20 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE SHEATH OF SELDINGER WAS RUPTURED DURING CATHETER PLACEMENT. THE PROBLEM MADE IT IMPOSSIBLE TO PROCEED TO PUNCTURE. THE OPERATOR UNPACKED A NEW SELDINGER FOR RE-PUNCTURE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAK IS INCONCLUSIVE DUE TO THE STATE OF THE RETURNED SAMPLE. ONE VIDEO OF A MICROINTRODUCER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE VIDEO SHOWED A MICROINTRODUCER BEING FLUSHED; HOWEVER, THE ALLEGED LEAK CANNOT BE SEEN IN THE RETURNED VIDEO. NO USE RESIDUE OR DAMAGE COULD BE DISCERNED IN THE RETURNED VIDEO. THEREFORE, THIS COMPLAINT IS INCONCLUSIVE AT THIS TIME. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03080 | 2023-07-20 | 2023-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY CUSTOMER THAT HAIR FOUND INSIDE THE STERILE PICC CATHETER INSERTION KIT. HAIR APPEARED TO BE STUCK IN THE NEEDLE CAP. ADDITIONAL INFORMATION RECEIVED (B)(6)2023: THE EVENT WAS DISCOVERED DURING SET UP FOR A PICC INSERTION. THE EVENT DID NOT INVOLVE AN URGENT OR LIFE-THREATENING MEDICAL SITUATION. THERE WAS NO PATIENT HARM, INJURY, COMPLICATION OR NEGATIVE OUTCOME THAT OCCURRED AS A RESULT OF THE EVENT. THE EVENT DID NOT INTERRUPT THE ADMINISTRATION OF ANY MEDICATION, OR CAUSE A CLINICALLY SIGNIFICANT DELAY IN MEDICATION, THAT NEGATIVELY IMPACTED THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03091 | 2023-07-20 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REPLACEMENT CONNECTOR FOR 4F GROSHONG NXT CLEARVUE PICC WITH EXTENSION LEG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7812400 |  | IT WAS REPORTED BY THE CUSTOMER THAT ONE CASE OF BARB BREAKAGE ON ONE SIDE OF THE WING PART OF THE CONNECTOR OCCURRED AT THE PICC PLACEMENT CENTER OF THE HOSPITAL. THIS HAPPENED WHEN ENGAGING THE BARBS ON THE WING PART OF THE CONNECTOR WITH THE GROOVE OF THE OVERSLEEVE PORTION AFTER THE TWO PARTS WERE ALIGNED DURING CONNECTOR INSTALLATION. SINCE THERE WAS NO ACCESSORY FOR REPLACEMENT AT THAT TIME, IT WAS RETAINED FOR ONE WEEK AND REPLACED DURING MAINTENANCE AT WEEK 2. NO OTHER INFORMATION PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED CONNECTOR WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR S/L GROSHONG CONNECTOR. THE SAMPLE WAS RECEIVED ASSEMBLED. NO CATHETER WAS VISIBLE WITHIN THE CONNECTOR. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. ONE LOCKING ARM WAS COMPLETELY DETACHED FROM THE CONNECTOR AND WAS NOT RETURNED. A CRACK WAS OBSERVED AT THE BASE OF THE REMAINING LOCKING ARM. MICROSCOPIC INSPECTION OF THE FRACTURE SITES REVEALED DULLY GRANULAR FRACTURE SURFACES. NO MECHANICAL DAMAGE WAS OBSERVED IN THE VICINITIES OF THE FRACTURES AND NO MATERIAL DEFORMATION WAS OBSERVED AT THE FRACTURE SITES. AN ATTEMPT TO REPLICATE THE FAILURE BY MANIPULATING A NON-COMPLAINANT SAMPLE RESULTED IN EXTENSIVE MATERIAL DEFORMATION AND DISCOLORATION. THE FRACTURE FEATURES WERE CONSISTENT WITH BRITTLE MATERIAL FAILURE. ATTEMPTS TO REPLICATE THE FRACTURE FEATURES WERE UNSUCCESSFUL, SUGGESTING THAT AN UNIDENTIFIED ENVIRONMENTAL FACTOR(S) CONTRIBUTED TO MATERIAL EMBRITTLEMENT. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05348 | 2023-07-20 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER INSERTION TRAY | FOLEY CATHETER INSERTION TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | 802010 | 802010 |  | IT WAS REPORTED THAT PATIENT STATED THAT ONE OF THE ITEMS WAS PUNCTURED. THE HEAD HAD A PUNCTURE THROUGH THE TAP. PER FOLLOW UP VIA EMAIL RECEIVED ON 11JUL2023, CUSTOMER STATED THAT NO ISSUE WITH A PATENT. THE PRODUCT ARRIVED WITH A PUNCTURE IN THE PACKAGE. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. 1 SAMPLE WERE CONFIRMED TO EXHIBIT THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED PHOTO SAMPLE NOTED ONE UNOPENED (WITH ORIGINAL PACKAGING), BARDIA INSERTION TRAY. VISUAL INSPECTION OF THE PHOTO SAMPLE NOTED A HOLE PUNCTURE IN THE PACKAGE. THIS DOES NOT MEET SPECIFICATION WHICH STATES "NO DAMAGE OR HOLES ON TRAYS OR TYVEK ARE ALLOWED". A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE ¿MATERIAL HANDLING." THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
MW5119831 | 2023-07-20 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | HUBLESS SILICONE FLAT DRAIN 10MM | CATHETER, IRRIGATION | C.R. BARD, INC. | 0070440 |  |  | A 10MM FLAT DRAIN WAS REMOVED FROM A PATIENT, WHO WAS ONE WEEK POST OPERATION. PER THE RESIDENT WHO REMOVED THE DRAIN, THE DRAIN WAS REMOVED WITHOUT INCIDENT AND WITHOUT ANY UNUSUAL RESISTANCE. ONCE REMOVED IT WAS NOTED BY THE RESIDENT A PORTION OF THE DRAIN APPEARED TO BE MISSING POSSIBLY CUT/TORN. THE RESIDENT ALSO NOTED A CUT/TEAR PROXIMALLY NEAR THE MISSING PORTION OF DRAIN. A RETAINED PORTION OF THE DRAIN WAS CONFIRMED VIA CT (COMPUTERIZED TOMOGRAPHY) SCAN. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-07
MW5119822 | 2023-07-20 | 2023-07-12 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARD POWERGLIDE PRO MIDLINE CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | F218108PT |  |  | POWERGLIDE MIDLINE CATHETER GUIDEWIRE FAILED TO DEPLOY FROM CASING AND CURLED UP IN THE CASING. NO APPARENT HARM TO THE PATIENT. NECESSITATED REMOVAL AND A NEW INSERTION. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05347 | 2023-07-20 | 2023-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 33416 | 33416 |  | IT WAS REPORTED THAT THE CUSTOMER WAS SLEEPING IN BED AND UPON AWAKENING HAD FOUND THE FOLEY CATHETER LYING BESIDE THEM. THE CUSTOMER EXAMINED THE CATHETER AND FOUND THE BALLOON HAD RUPTURED. THE CUSTOMER THEN PHONED THE HOME CARE OFFICE FIRST THING IN THE MORNING, AND REQUESTED TO COME IN AND HAD SAME REINSERTED, AS WAS UNABLE TO VOID. THE OLD CATHETER WAS EXAMINED, AND THE RUPTURED BALLOON APPEARS TO BE ONE UNIT, NO MISSING PIECES NOTED. NO APPARENT HARM REACHED THE PATIENT. THE IMPACT OF THE INCIDENT WAS PATIENT WAS UNABLE TO VOID. THE PATIENT REQUIRED UNEXPECTED OR PROLONGED CARE. | THE REPORTED EVENT WAS CONFIRMED AND 1 SAMPLE WAS CONFIRMED TO EXHIBIT THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE "HIGH MODULUS LATEX". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03096 | 2023-07-20 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS FOUND THAT THE END OF THE GUIDE WIRE WAS BENT AND DISINTEGRATED WHEN THE GUIDE WIRE WAS THREADED THROUGH THE PUNCTURE NEEDLE, RESULTING IN FAILURE TO REMOVE THE PUNCTURE NEEDLE SMOOTHLY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BENT GUIDEWIRE WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. TWO PHOTOGRAPHS OF THE IMPLICATED NITINOL GUIDEWIRE WERE RETURNED FOR EVALUATION. THE DISTAL COILED END OF THE GUIDEWIRE CONTAINED A SLIGHTLY BENT SHAPE MEMORY. IT IS UNKNOWN IF THE BEND IN THE WIRE WAS SEVERE ENOUGH TO CAUSE THE REPORTED DIFFICULTY THREADING THE GUIDEWIRE THROUGH THE NEEDLE AS IT COULD NOT BE TESTED WITHOUT THE PHYSICAL SAMPLE. IT COULD NOT BE DETERMINED FROM THE PHOTOGRAPHS IF THE GUIDEWIRE CONTAINED ANY ADDITIONAL DAMAGE OR BREAKS NEAR THE DISTAL END; EVALUATION OF THE PHYSICAL DEVICE WOULD NEED TO BE COMPLETED TO CONCLUSIVELY DETERMINE IF ADDITIONAL DAMAGE WAS PRESENT ON THE GUIDEWIRE. THERE WERE NO DISTINGUISHING FEATURES IN THE RETURNED PHOTOGRAPHS WHICH COULD AID IN IDENTIFICATION OF A SPECIFIC ROOT CAUSE FOR THE BEND IN THE WIRE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03093 | 2023-07-20 | 2023-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE CATHETER LEAKED WATER DURING CATHETER PLACEMENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAK IS CONFIRMED, AND THE ROOT CAUSE COULD NOT BE DETERMINED. ONE VIDEO OF A LEAKING GROSHONG NXT CLEARVUE WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE VIDEO SHOWED THE CATHETER BEING FLUSHED WITH SALINE. A STEADY STREAM OF FLUID WAS OBSERVED COMING OUT OF THE CATHETER NEAR THE DISTAL END, PROXIMAL TO THE VALVE. NO OTHER DETAILS OF THE APPARENT DAMAGE COULD BE SEEN IN THE RETURNED VIDEO; THEREFORE, THE CAUSE OF THE LEAK IS UNKNOWN. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03086 | 2023-07-20 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/SHERLOCK 3CG TPS STYLET 4F MAXIMAL BARRIER TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED BY CUSTOMER THAT PATIENT CAME IN WITH A PICC THAT WAS LEAKING. PATIENT DENIED HHC NURSE USING ANY SCISSORS NEAR WHEN DOING DRESSING CHANGE. FRACTURE AT THE HUB. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03088 | 2023-07-20 | 2023-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 6F TRIPLE-LUMEN CATHETER W/ MICROEZ MICROINTRODUCER NURSE BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3386105 |  | CUSTOMER REPORTED THE PATIENT HAD A NEW CATHETER EXCHANGE ON A 6FR TL POWERPICC ON (B)(6) AT NIGHT. ADDITIONAL INFORMATION RECEIVED 7/13/2023: CATHETER WAS REPLACED DUE TO A LEAK. THE DOCTOR PASSED THE CATHETER, THERE WAS NO FLOW AND REFLUX, AN IMAGE EXAM WAS PERFORMED AND DETECTED THE NEED FOR THIS EXCHANGE, DR. CAN COMPLEMENT AS NECESSARY. THE CATHETER WAS EXCHANGED IN A SURGICAL CENTER BY THE SAME DOCTOR WHO DID THE PREVIOUS ONE. PATIENT RECEIVED SEDATION, AND THE PATIENT AND FAMILY WERE DISSATISFIED, BUT NO INFECTION, FOR THE TIME BEING IDENTIFIED. DR. FOLLOWED THE CASE CLOSELY. TREATMENT CONTINUED WITH THE NEW CATHETER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03092 | 2023-07-20 | 2023-07-03 | I, n, j, u, r, y | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (22G) (8CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F122080T |  | IT WAS REPORTED BY THE CUSTOMER THAT CATHETER SHEARED WHEN PULLED BACK ONTO DEVICE AFTER ADVANCING CATHETER WINGS. PATIENT UNDERWENT SURGERY TO REMOVE SHEARED CATHETER. ADDITIONAL INFORMATION RECEIVED 7/14/2023: CUSTOMER REPORTS THAT THE BREAK WAS DISCOVERED WHEN SCANNING THE PATIENT'S ARM. X-RAY AND ULTRASOUND WERE USED TO CONFIRM CATHETER PIECE REMAINED IN THE PATIENT. PATIENT COMPLAINED OF PAIN AT THE SITE. ANOTHER POWERGLIDE MIDLINE WAS PLACED IN THE OTHER ARM, AND THEY HAD TO STOP OTHER MEDS TO GIVE THE MED THEY NEEDED GIVE. CATHETER PIECE WAS SUCCESSFULLY REMOVED VIA SURGERY AT BEDSIDE. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED MIDLINE CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS 22GA X 8CM POWERGLIDE PRO MIDLINE CATHETER. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. THE CATHETER TERMINATED APPROXIMATELY 3.5CM FROM THE MOLDED JOINT. MICROSCOPIC INSPECTION OF THE BREAK SURFACE REVEALED A PARTIALLY GLOSSY AND PARTIALLY GRANULAR FRACTURE SURFACE. A LONGITUDINALLY ALIGNED SCORING MARK WAS OBSERVED ON THE INSIDE SURFACE OF THE CATHETER SHAFT LEADING INTO THE FRACTURE SITE. THE FRACTURE FEATURES WERE CONSISTENT WITH DAMAGE CAUSED BY CONTACT BETWEEN THE CATHETER SHAFT AND THE TIP OF THE INTRODUCER NEEDLE. SUCH DAMAGE CAN OCCUR IF THE CATHETER IS WITHDRAWN ONTO THE NEEDLE SHAFT AND IF THE NEEDLE IS RE-INSERTED INTO THE CATHETER FOLLOWING CATHETER ADVANCEMENT. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03089 | 2023-07-20 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 20 GAUGE POWERGLIDE PRO MIDLINE CATHETER 10CM INTERMEDIATE KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6F120101 |  | CUSTOMER REPORTED THAT A POWERGLIDE PRO CATHETER PLACEMENT WAS ATTEMPTED IN THE BASILIC VEIN. THE PATIENT COMPLAINED ABOUT SEVERE PAIN DURING THE INSERTION, WHICH WAS A SIGN FOR US TO CHECK THE CATHETER ITSELF. WHEN VISUALLY EXAMINING THE TIP OF THE CATHETER, WE SEE THAT THE A PIECE OF THE GUIDEWIRE IS STUCK (0.2-0.3MM) OUT OF THE CATHETER. THE GUIDEWIRE WAS NOT ADVANCED BEFORE THE INSERTION. NO OTHER INFORMATION PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF DIFFICULTY ADVANCING THE GUIDEWIRE OF A POWERGLIDE PRO IS CONFIRMED AND WAS DETERMINED TO BE USE RELATED. ONE 20 GA POWERGLIDE PRO WAS RETURNED FOR EVALUATION. THE GUIDEWIRE IN THE RETURNED POWERGLIDE PRO WAS OBSERVED TO HAVE SOME MISALIGNED COILS AT THE DISTAL END OF THE GUIDEWIRE. BASED ON THE MISALIGNED COILS OBSERVED ALONG THE RETURNED GUIDEWIRE OF THE POWERGLIDE PRO, THE COMPLAINT OF DIFFICULT DEVICE REMOVAL IS CONFIRMED. INSERTION INTO RESISTANCE SUCH AS TISSUE OR STEEP INSERTION ANGLES CAN CAUSE DAMAGE TO THE GUIDEWIRE THUS CAUSING DIFFICULTY DURING REMOVAL OF THE GUIDEWIRE. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03081 | 2023-07-20 | 2023-06-26 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405J |  | IT WAS REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. THE PHYSICIAN ACCIDENTALLY CUT ALSO THE STYLET UPON ADJUSTING THE CATHETER AND A FEW DAYS LATER IT WAS UNABLE TO DRAW BLOOD. ON (B)(6) 2023, THE STYLET WAS CONFIRMED REMNANT IN THE VESSEL WHEN THE CATHETER WAS REMOVED. REPORTEDLY, THE REMNANT STYLET WAS REMOVED FROM THE PATIENT BODY. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03090 | 2023-07-20 | 2023-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F)(SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED BY THE CUSTOMER THAT AFTER PRE-FLUSHING AND BEFORE PUNCTURE, IT WAS FOUND THAT THE EXTENSION TUBING CONNECTION LEAKED LIQUIDS. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING INVESTIGATION ELEMENTS WERE DETERMINED TO BE APPLICABLE AND WERE COMPLETED AS PART OF THIS INVESTIGATION: COMPLAINT HISTORY REVIEW, DHR RECORDS, SAMPLE ANALYSIS, REVIEW OF RISK DOCUMENTATION. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAK IS INCONCLUSIVE. ONE PHOTOGRAPH OF A 4FR POWERPICC WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED THE MOLDED JOINT AND A PORTION OF THE EXTENSION TUBING. NO DAMAGE WAS OBSERVED IN THE RETURNED PHOTOGRAPH. INSPECTION OF THE RETURNED PHOTOGRAPH WAS INSUFFICIENT TO IDENTIFY THE ALLEGED ISSUE OR A ROOT CAUSE OF THE REPORTED ISSUE; THEREFORE, THIS COMPLAINT IS INCONCLUSIVE AT THIS TIME. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03084 | 2023-07-20 | 2023-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH 15 CM PRE-CURVE KIT | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5554150 |  | CUSTOMER REPORTED THAT AFTER THE JUGULAR VEIN PUNCTURE WAS SUCCESSFUL, IT WAS DIFFICULT TO PLACE THE GUIDE WIRE, AND THERE WAS RESISTANCE TO PULLING BACK THE GUIDE WIRE, LEADING TO PARTIAL DISCONNECTION OF THE GUIDE WIRE AND BENDING OF THE HEAD SEGMENT. THE PUNCTURE NEEDLE AND GUIDE WIRE WERE REMOVED, AND THE ARROW TUBE PACK FOR THORACENTESIS WAS REMOVED AGAIN. THE ARROW TUBE PACK GUIDE WIRE WAS USED TO GUIDE THE PUNCTURE SUCCESSFULLY. NO OTHER INFORMATION PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
3006260740-2023-03087 | 2023-07-20 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG 4F SINGLE-LUMEN PICC PERIPHERALLY INSERTED CVC BASIC TRAY | CATHETER, PERCUTANEOUS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7715405 |  | IT WAS REPORTED BY THE CUSTOMER THAT DURING INJECTING SALINE PRIOR TO PLACEMENT, THE CATHETER WAS FOUND DAMAGED AND SEVERELY LEAKED LIQUIDS. IT WAS REPORTED THIS OCCURRED WITH 2 DEVICES THAT WERE RETURNED FOR INVESTIGATION. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF LEAKING CATHETERS WAS UNCONFIRMED BECAUSE THE PROBLEM COULD NOT BE REPRODUCED. THE PRODUCT RETURNED FOR EVALUATION WAS TWO 4FR S/L GROSHONG CATHETERS. BOTH SAMPLES WERE RECEIVED WITH INLAID STYLETS. THE FIRST SAMPLE (SAMPLE 1) EXHIBITED BENDS ALONG THE LENGTH OF THE STYLET NEAR THE 20CM, 26CM, AND 30CM DEPTH MARKINGS. THE SECOND SAMPLE (SAMPLE 2) EXHIBITED A BEND IN THE STYLET JUST DISTAL OF THE 5CM MARKING. ATTEMPTS TO INFUSE WATER THROUGH BOTH SAMPLES USING A 12ML SYRINGE REVEALED BOTH SAMPLES TO BE PATENT TO INFUSION AND ASPIRATION WITH NO OBSERVED LEAKS. NO LEAKS WERE OBSERVED DURING PRESSURIZATION OF THE SAMPLES. MICROSCOPIC INSPECTION OF THE SAMPLES DID NOT REVEAL ANY DAMAGE OR EVIDENCE OF LEAKAGE. NO DEFICIENCIES WERE DISCOVERED DURING EVALUATION OF THE RETURNED SAMPLES. CONSEQUENTLY THIS COMPLAINT IS UNCONFIRMED AT THIS TIME. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05368 | 2023-07-20 | 2023-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119108 | 119108 |  | IT WAS REPORTED THAT NO OUTFLOW OF URINE AFTER CATHETER INSERTION. WHEN THE CATHETER WAS REMOVED AND WATER WAS INJECTED, IT COULD NOT BE DRAINED. POSSIBLE BLOCKAGE AT THE TIP OF THE CATHETER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. IT IS UNKNOWN IF THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITH ORIGINAL PACKAGING), ALL-SILICONE TEMP SENSING FOLEY CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 3.5ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 2 MINUTES WITH NO CUFFING OR LEAKS NOTED, RETURNING 3.5ML OF SOLUTION. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE LABELING REVIEW WAS NOT REQUIRED AS THE REPORTED EVENT WAS UNCONFIRMED. CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07
1018233-2023-05370 | 2023-07-20 | 2023-07-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT THE CUSTOMER EXPERIENCED A PRODUCT COMPLAINT WITH PRODUCT NUMBER A942216. THE FOLEY CATHETER FELL OUT OF THE PATIENT AND THE NURSE RECOGNIZED THAT THE TIP WAS MISSING (NGHQ2185). THE CUSTOMER SAVED THE CATHETER, AND IT WAS AVAILABLE TO BE RETURNED. NO PATIENT COMPLICATION WAS REPORTED. THE CUSTOMER WAS CONCERNED ABOUT THE FOLEY TRAYS AND OPEN LOT NUMBER, THEY NOTICED OTHER ISSUES WITH FOLEY CATHETERS, KINK TUBING (NGHG1002) AND CUT FOLEY CATHETER (NGGZ1611). THESE FOLEY CATHETERS WERE NOT INSERTED IN PATIENTS. THEY HAD SAMPLES OF THESE TRAYS TOO. THEY QUARANTINED LOT NUMBER ALL THE THREE LOTS WHICH HAD BEEN REMOVED FROM CIRCULATION. AFTER SPEAKING WITH CUSTOMER, THEY PULLED ABOUT 3 CASES AND INDIVIDUAL FOLEY TRAYS FROM THE FLOOR. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT CAUSED THE REPORTED FAILURE. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED A 2-WAY CATHETER. VISUAL INSPECTION NOTED THAT THE TIP OF THE CATHETER WAS BROKEN OFF. THIS IS OUT OF SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INADEQUATE MATERIAL SELECTION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-07