## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name | year_month
3006260740-2023-00448 | 2023-02-15 | 2022-12-06 | I, n, j, u, r, y | Manufacturer report | POWERPORT MRI ISP 9.6 FR. SILILCONE SINGLE LUMEN CATHETER | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1859660 | 1859660 |  | IT WAS REPORTED THAT FOUR MONTHS AND TWO DAYS POST PORT PLACEMENT, THE DEVICE ALLEGEDLY HAD A LEAK. IT WAS FURTHER REPORTED THAT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN AND VISION CHANGES AND PAIN MEDICATION AND INFLAMMATION MEDICATION WAS PRESCRIBED TO THE PATIENT. REPORTEDLY, THE DEVICE WAS REMOVED AND REPLACED. THE CURRENT PATIENT STATUS WAS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT MRI IMPLANTABLE PORT ATTACHED TO A CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE, AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED FLUID LEAK ISSUE, AS NO KINKS OR SPLITS WERE NOTED THROUGHOUT THE CATHETER AND NO LEAKS WERE OBSERVED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00454 | 2023-02-15 | 2023-01-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D1 |  | IT WAS REPORTED BY BD SALES REPRESENTATIVE ¿DURING CLINICAL VISIT, STAFF REPORTED THAT AIR IS BEING INTRODUCED INTO THE WIRE CONNECTOR ADAPTER OF THE PICC LINE. A 4 FRENCH SINGLE LUMEN PICC KIT WAS OPENED FOR DEMONSTRATION. THE LINE WAS PRIMED WITH SALINE AND AIR IS ENTERING EITHER FROM THE HOLE THE WIRE PASSES THROUGH OR AT THE CONNECTION TO THE PLASTIC (THE RUBBERIZED PIECE THE WIRE STYLET PASSES THROUGH). | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGW2246 SHOWED FIVE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00451 | 2023-02-15 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT A PICC LINE LEAKED SALINE FROM THE RUBBER STOPPER WHERE THE INSERTION WIRE TRAVELS THROUGH THEM. SALINE LEAKED WHEN THE CATHETER WAS FLUSHED, AND ALSO DREW AIR BUBBLES IN THE SALINE SYRINGE WHEN ASPIRATING BACK FOR BLOOD RETURN DURING THE INSERTION PROCEDURE. CUSTOMER STATES THE RUBBER STOPPER "VISUALLY LOOKS DIFFERENT THAN OUR OTHER 5F DOUBLE LUMENS (WITH DIFFERENT LOT NUMBERS)." THERE WAS NO PATIENT HARM REPORTED. PATIENT DOES NOT HAVE AN INFECTIOUS DISEASE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, TREND ANALYSIS, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 3CG TLS STYLET. THE WIRE WAS INSERTED THROUGH THE T-LOCK ASSEMBLY SEPTUM. THE WHITE WIRE WAS CUT DISTAL OF THE SPIKED TETHER ASSEMBLY, WHICH WAS NOT RETURNED FOR EVALUATION. AN ATTEMPT TO INFUSE WATER THROUGH T-LOCK USING A 12ML SYRINGE REVEALED THE SAMPLE TO BE PATENT TO INFUSION; HOWEVER, A LEAK WAS OBSERVED IN THE SEPTUM WHERE THE WIRE PASSED THROUGH. MICROSCOPIC INSPECTION OF THE SEPTUM REVEALED GAPS ON TWO SIDES WHERE THE WIRE PASSED THROUGH AN INCISION. LEAKING WATER WAS OBSERVED AT THOSE GAPS. THE GAPS AT THE WIRE INSERTION SITE RESULTED IN THE OBSERVED LEAKAGE AND MAY HAVE OCCURRED AS PART OF DEVICE ASSEMBLY. THE DEVICE IS A SUPPLIED COMPONENT AND BD IS WORKING CLOSELY WITH THE SUPPLIER TO ADDRESS THIS TYPE OF EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00449 | 2023-02-15 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT 2 PICCS WERE INSERTED AND BOTH HAD ISSUES WITH THE RUBBER STOPPERS WHERE THE INSERTION WIRE TRAVELS THROUGH THEM. THEY LEAKED SALINE WHEN FLUSHING WHERE THE WIRE INSERTS INTO THE STOPPER. CUSTOMER STATES THAT THE RUBBER STOPPER "VISUALLY LOOKS DIFFERENT THAN OUR OTHER 5F DOUBLE LUMENS (WITH DIFFERENT LOT NUMBERS)." THIS REPORT ADDRESSES THE FIRST DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGZ0461 SHOWED TWO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. THE COMPLAINTS FOR THIS BATCH NUMBER HAVE BEEN REPORTED FROM THE SAME FACILITY IN CANADA. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00450 | 2023-02-15 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT 2 PICCS WERE INSERTED AND BOTH HAD ISSUES WITH THE RUBBER STOPPERS WHERE THE INSERTION WIRE TRAVELS THROUGH THEM. THEY LEAKED SALINE WHEN FLUSHING WHERE THE WIRE INSERTS INTO THE STOPPER. CUSTOMER STATES THAT THE RUBBER STOPPER "VISUALLY LOOKS DIFFERENT THAN OUR OTHER 5F DOUBLE LUMENS (WITH DIFFERENT LOT NUMBERS)." THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGZ0461 SHOWED TWO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. THE COMPLAINTS FOR THIS BATCH NUMBER HAVE BEEN REPORTED FROM THE SAME FACILITY IN CANADA. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00920 | 2023-02-15 | 2023-01-12 | I, n, j, u, r, y | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 | FOL0102 | FOL0102 |  | IT WAS REPORTED THAT THE FOLEY STATLOCK IRRITATED THE PATIENTS SKIN AND ENDED UP BREAKING OUT IN LOCATION. REPRESENTATIVE ADVISED TO CHANGE THE STATLOCK TO ANOTHER ONE, BUT THE PATIENT DECLINED. NO MEDICAL INTERVENTION WAS REPORTED. PER ADDITIONAL INFORMATION RECEIVED, VIA PHONE ON 27JAN2023, STATED THAT THE PATIENT STOPPED USING THE PRODUCT DUE TO SKIN IRRITATION AND SEEKS MEDICAL ATTENTION AND WAS PRESCRIBED A CREAM. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCOMPLETE BIOCOMPATIBILITY ASSESSMENT OF ADHESIVE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INDICATIONS FOR USE: THE STATLOCK® DEVICE IS A STABILIZATION DEVICE FOR COMPATIBLE CATHETERS. CONTRAINDICATIONS: KNOWN TAPE OR ADHESIVE ALLERGIES. WARNINGS AND PRECAUTIONS: 1. DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, DIAPHORETIC OR NON-ADHERENT SKIN, OR WHEN THE ACCESS DEVICE IS NOT MONITORED DAILY. 2. OBSERVE UNIVERSAL BLOOD AND BODY FLUID PRECAUTIONS AND INFECTION CONTROL PROCEDURES, DURING APPLICATION AND REMOVAL OF THE STATLOCK® DEVICE. 3. MINIMIZE CATHETER MANIPULATION DURING APPLICATION AND REMOVAL OF THE STATLOCK® DEVICE. 4. DAILY MAINTENANCE: A. THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY INDICATED, AT LEAST EVERY SEVEN DAYS. B. IF PAD BECOMES SOILED, WASH WITH SOAP/WATER, SALINE OR HYDROGEN PEROXIDE. DO NOT USE ALCOHOL OR PREPACKAGED BATHING SYSTEMS, WHICH COULD LEAD TO EARLY LIFTING. C. IF SHOWERING/BATHING, COVER WITH PLASTIC WRAP OR WATERPROOF DRESSING. D. CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. E. USE CLINICAL JUDGMENT ON THE REMOVAL OF THE STATLOCK® STABILIZATION DEVICE IF THE PATIENT EXPERIENCES ANY FLUID SHIFTS THAT MAY INTERFERE WITH SKIN INTEGRITY. APPLICATION TECHNIQUE PREP: 1. PLACE FOLEY CATHETER INTO RETAINER. DIRECTIONAL ARROW SHOULD POINT TOWARDS CATHETER TIP, AND BALLOON INFLATION ARM SHOULD BE NEXT TO THE CLAMP HINGE. 2. CLOSE LID, BEING CAREFUL TO AVOID PINCHING THE CATHETER. 3. IDENTIFY SECUREMENT SITE BY LAYING THE DEVICE RETAINER ON THE FRONT OF THE THIGH, LEAVING 1 INCH OF CATHETER SLACK BETWEEN INSERTION SITE AND THE STATLOCK® DEVICE RETAINER. 4. AFTER PLACING THE STATLOCK® STABILIZATION DEVICE OFF TO THE SIDE, CLEANSE AND DEGREASE THE SECUREMENT SITE WITH ALCOHOL PER HOSPITAL POLICY. LET SKIN DRY. 5. APPLY SKIN PROTECTANT, IN DIRECTION OF HAIR GROWTH, TO AREA LARGER THAN SECUREMENT SITE. ALLOW TO DRY COMPLETELY (10-15 SECONDS). 6. USING PERMANENT MARKER, WRITE INITIALS AND DATE OF APPLICATION ON THE STATLOCK® DEVICE ANCHOR PAD. NOTE: ALWAYS SECURE CATHETER INTO THE STATLOCK® DEVICE RETAINER BEFORE APPLYING ADHESIVE PAD ON SKIN. PLACE AND PEEL: 7. ALIGN THE STATLOCK® STABILIZATION DEVICE OVER SECUREMENT SITE LEAVING 1 INCH OF CATHETER SLACK. MAKE SURE LEG IS FULLY EXTENDED. 8. WHILE HOLDING THE RETAINER TO KEEP THE PAD IN PLACE, PEEL AWAY PAPER BACKING, ONE SIDE AT A TIME AND PLACE TENSION-FREE ON SKIN. REMOVAL TECHNIQUE DISENGAGE: 1. OPEN RETAINER BY PRESSING RELEASE BUTTON WITH THUMB, THEN LIFT TO OPEN. 2. REMOVE FOLEY CATHETER FROM THE STATLOCK® DEVICE. DISSOLVE: 3. WIPE THE EDGE OF THE PAD USING AT LEAST 5-6 ALCOHOL PADS UNTIL A CORNER LIFTS. THEN CONTINUE TO STROKE UNDERSURFACE OF PAD WITH ALCOHOL TO DISSOLVE ADHESIVE PAD AWAY FROM SKIN. DO NOT PULL OR FORCE PAD TO REMOVE. THE STATLOCK® STABILIZATION DEVICE SHOULD BE MONITORED DAILY AND REPLACED WHEN CLINICALLY INDICATED, AT LEAST EVERY 7 DAYS. THE CATHETER INSERTION SITE SHOULD BE TREATED PER LOCAL POLICY RECOMMENDATIONS. THE STATLOCK® DEVICE IS CONTRAINDICATED ON PATIENTS WITH KNOWN TAPE AND ADHESIVE ALLERGIES. ALCOHOL AND ACETONE MAY WEAKEN THE ADHESIVE BOND BETWEEN THE STATLOCK® DEVICE PAD AND THE SKIN. CARE SHOULD BE TAKEN WHEN USING THESE SOLUTIONS WHILE PERFORMING CATHETER SITE CARE. PLEASE CONSULT PRODUCT INSERTS AND LABELS FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. FOR ADDITIONAL INFORMATION ON THE STATLOCK® FOLEY STABILIZATION DEVICE, PLEASE CALL YOUR LOCAL BARD REPRESENTATIVE.¿ | Y | N | LAY USER/PATIENT |  | 2023-02
1018233-2023-00918 | 2023-02-15 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | MAGIC3 INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT HAD TO USE 3 MAGIC 3 GO MALE HYDROPHILIC INTERMITTENT CATHETERS IN ONE SITTING DUE TO THE CATHETER WAS HAVING THE TIP CUT OFF AND CAUSED BLEEDING. IT WAS A MANUFACTURING DEFECT AND THE PATIENT HAD 5 OUT OF 90 CATHETERS THAT DID NOT WORK PROPERLY. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "MATERIAL WEAKENS OVER TIME". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-00919 | 2023-02-15 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0102L20 | 0102L20 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS LEAKING AND IT CAUSED THEIR PENIS TO BE RED AND IRRITATED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿BAD FIT WITH CONNECTOR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-00916 | 2023-02-15 | 2023-01-26 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION WHILE USING THE PUREWICK FEMALE EXTERNAL CATHETER AND THE DOCTOR SAID IT WAS CHRONIC. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. PER CUSTOMER CONTACT, VIA PHONE ON (B)(6) 2023, CUSTOMERS AUTHORIZED THE CAREGIVER STATED THAT THE PATIENT HAD CHRONIC URINARY TRACK INFECTION SO THE DOCTOR JUST RECOMMENDED STOPPING USING OF THE PUREWICK FOR NOW TO SEE HOW THINGS GO. THE CAREGIVER WAS PRESCRIBED CIPROFLOXACINE FOR THE PATIENT LAST URINARY TRACK INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | N | LAY USER/PATIENT |  | 2023-02
1018233-2023-00915 | 2023-02-15 | 2022-10-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT MENTIONED THAT THE INSTRUCTIONS FOR USE (IFU) DOES NOT PROVIDE SUFFICIENT INFORMATION ABOUT THE PRODUCT AND ITS MEDICAL APPLICATION AND THE CUSTOMER DID NOT AGREE WITH THE STATEMENT: ¿THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD / BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S).¿ IN RELATION TO BARDEX® I.C. AQUAFIL FOR THE EUROPEAN COMMUNITY (FRENCH SIZE 12,14,16,18,20,22) (PCN D236512S, PCN D236514S, PCN D236516S, PCN D236518S, PCN D236520S AND PCN D236522S)BARDEX® I.C. 2-WAY COUNCIL TIP FOLEY CATHETER (STRIP PACK) (18 FRENCH) AND BARDEX® I.C. 2-WAY COUNCIL TIP FOLEY CATHETER TRAYS. ALSO STATED NO, THE BENEFITS OF USING THIS PRODUCT DO NOT OUTWEIGH THE POTENTIAL RISKS WHEN ASKED IF THEY BELIEVE THAT THE BENEFITS OF USING THE BARD/BECTON DICKINSON I.C. FOLEY CATHETER AND/OR COMPREHENSIVE CARE TRAY OUTWEIGH THE POTENTIAL RISKS REFERRING TO THE PROS OUTWEIGH THE CONS IN RELATION TO BARDEX I.C. FEMALE LENGTH, 2-WAY FOLEY CATHETER (14 FRENCH) AND BARDEX® I.C. AQUAFIL FOR THE EUROPEAN COMMUNITY (24 FRENCH). PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON (B)(6)2022, PATIENT STATED THAT A LEAFLET TO EXPLAIN HOW IT WAS USED WOULD BE HELPFUL. BOOKLET DID NOT PROVIDE THAT INFORMATION FOR NEW STARTERS OR STUDENTS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ILLEGIBLE PRINT & MISSING PRINT. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE PRODUCT CATALOG NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00917 | 2023-02-15 | 2023-01-26 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION BUT WAS UNSURE IF IT WAS FROM THE PUREWICK FEMALE EXTERNAL CATHETER USE OR DUE TO HAVING URINARY RETENTION AND LOOSE BOWL MOVEMENTS. PATIENT HAD BEEN USING THE PUREWICK PRODUCT FOR MORE THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. PER CUSTOMER CONTACT VIA PHONE ON 09FEB2023, CUSTOMER STATED THAT THEY DID NOT THINK THE URINARY TRACT INFECTION WAS DUE TO PURE WICK USE. THE PATIENT WAS ON HOSPICE AND HAD A LOT OF SEDIMENT IN THE URINE. PATIENT WAS PRESCRIBED AN ANTIBIOTIC FOR THE URINARY TRACT INFECTION BUT THEY DID NOT KNOW THE NAME OF IT. THE ANTIBIOTIC LED TO BOWEL INCONTINENCE, SO THEY HAD TO STOP PURE WICK USE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-02
1018233-2023-00914 | 2023-02-15 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT THE CUSTOMER USED BARD INDWELLING URINARY CATHETERS AT THE HOSPITAL. THE CUSTOMER WOULD LIKE TO KNOW HOW LONG THEY WERE RECOMMENDED TO STAY IN SITU IN A CLIENT. THE CUSTOMER ASKED WHETHER THERE WAS AN EXPIATORY DATE AFTER INSERTION AND WHEN IT SHOULD BE CHANGED. A FEW STAFF HAD NOTICED THE BREAKDOWN OF THE PRODUCT IF LEFT IN FOR A WHILE. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 13FEB2023, STATED THAT THIS WAS SECOND HAND INFORMATION THE CUSTOMER WAS GIVEN ABOUT A PREVIOUS PATIENT UPON DISCHARGE AND THE DISCHARGE NURSE WAS ASKING FOR MORE INFORMATION OF THIS PRODUCT AFTER THE FACT. THE CUSTOMER DO NOT HAVE ANY OTHER INFORMATION BECAUSE THEY RECEIVED THIS QUESTION POSED TO THEM AFTER THE FACT. THE CUSTOMER STATED THAT THEY WOULD ENSURE TO PROVIDE WITH THIS INFORMATION IF THIS HAPPENED AGAIN. PER ADDITIONAL INFORMATION RECEIVED FROM CUSTOMER ON 15FEB2023, STATED THAT THIS WAS SECOND HAND INFORMATION THE CUSTOMER WAS GIVEN ABOUT A PREVIOUS PATIENT UPON DISCHARGE AND SOMETHING A NURSE HAD NOTED AND THEN ASKED THIS QUESTION. THE CUSTOMER DO NOT HAVE ANY OTHER INFORMATION AND WAS ONLY INQUIRING ABOUT INFORMATION FOR THEIR OWN KNOWLEDGE NOT PUTTING IN A COMPLAINT. THE CUSTOMER STATED THAT THEY WOULD ENSURE TO PROVIDE WITH THIS INFORMATION IF THIS HAPPENED AGAIN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE, ¿OPERATOR ERROR OR MACHINE MALFUNCTION". THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE PRODUCT CATALOG NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00912 | 2023-02-15 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123518A |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON ASYMMETRICALLY INFLATED. THERE WERE MANY SIMILAR CASES, BUT WHEN THE BALLOON WAS INFLATED, THE BALLOON PART WAS DEFORMED. THE BULGE WAS BIASED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 02FEB2023, CLARIFIED THAT THE BALLOON ASYMMETRICALLY INFLATED. THE REPORTER OF THIS COMPLAINT MADE VAGUE REFERENCES TO SIMILAR CASES IN THE PAST AND DID NOT POINT TO A SPECIFIC CASE. THUS, DETAILS FOR SIMILAR CASES WERE NOT AVAILABLE. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE NON-UNIFORM FINISH DIP COVERAGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE(S) THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® SILICONE-COATED FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILICONE COATING <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 30 EVERY 2FR 1. BALLOON CATHETER [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00913 | 2023-02-15 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THERE WAS A BLACK FOREIGN SUBSTANCE MIXED IN THE FOLEY CATHETER. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 14FEB2023, STATED THAT THE DETAILS REGARDING THE FOREIGN MATERIAL LOCATION WAS NOT PROVIDED FROM CUSTOMER. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DEFECT CONTRIBUTED BY VENDOR OR CUSTOMER/ IMPROPER HANDLING/ UNCLEAN MACHINE /WORKING AREA. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. (THE URETHRA MAY BE INJURED) (2) DO NOT PULL THE CATHETER HARD. (THE BLADDER/URETHRA MAYBE INJURED) 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER (WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED). CONTRAINDICATIONS 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE ON PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00905 | 2023-02-15 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT THE 18 FR CATHETER CAME OFF THE BAG AND CAUSED THE PATIENT TO LEAK URINE. THE PATIENT NEVER HAD A PROBLEM USING 16 FR CATHETER, BUT IN TWO WEEKS OF USING 18 FR CATHETER, BAGS WORN OFF. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO DRAINAGE FUNNEL TOO SHORT/ IMPROPER INSERTION BY OPERATOR/ USER RELATED (PULLING BY USER/TAMPER EVIDENT SEAL REMOVED BY USER). THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE PRODUCT CATALOG NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-00908 | 2023-02-15 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURE STEP¿ FOLEY TRAY SYSTEM BARDEX® LUBRICATH® FOLEY CATHETER TRAY | FOLEY CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 |  | 909116M |  | IT WAS REPORTED THAT WHEN THE BALLOON WAS CHECKED FOR INFLATION BEFORE USE, IT COULD INFLATE, BUT FIXATION WATER LEAKED FROM THE SYRINGE SIDE OF THE SYRINGE CONNECTION PORT. DUE TO THIS WATER LEAKAGE, THE PRODUCT WAS JUDGED TO BE DEFECTIVE, AND THE DEVICE WAS NOT USED. SUBSEQUENTLY, ANOTHER PRODUCT WITH THE SAME LOT NUMBER WAS USED, AND NO DEFECTS WERE FOUND IN THAT PRODUCT. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED. EVALUATION CONFIRMED THAT THE RETURNED SYRINGE HAD CAUSED THE LEAKAGE AT INFLATION VALVE DURING INFLATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO DEFECTIVE COMPONENT FROM SUPPLIER. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER SHAPE, CONFIGURATION AND PRINCIPLES: BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES INTENDED USE & EFFECT- EFFICACY: THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. DIRECTIONS FOR USE: 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDO-ELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG.THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR. 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00907 | 2023-02-15 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER CHANGED FROM PREVIOUS ONE AND THEY WERE HAVING A HARD TIME GETTING THE NEW ONE DRAINED. STATED THAT THE STOPPER DID NOT OPEN UP ENOUGH, RESTRICTING THE FLOW, AND ALSO IT DID NOT OPEN ALL THE WAY UP. STATED THAT IT WAS HARD TO GET IT TO OPEN. CUSTOMER HAD CARPER TUNNEL, SO IT WAS DIFFICULT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: SALT ACCUMULATION)/BLOCK DRAINAGE LUMEN/NO DRAINAGE EYE.) THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-00899 | 2023-02-15 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165PL08 |  | IT WAS REPORTED THAT FOLEY CATHETER TIP SIZE WAS FOUND TO BE INCONSISTENT WITH OTHER CATHETERS OF THE SAME BATCH. THE SAME ISSUE WAS REPORTED IN ANOTHER RECORD. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE PRODUCTS MEET SPECIFICATION. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT RESTERILE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00902 | 2023-02-15 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT THE TEMPERATURE SENSING FOLEY CATHETER BALLOON WAS DEFLATING INTO A RING AROUND THE CATHETER THAT WAS CAUSING IT TO GET STUCK INSIDE THE PATIENT. THIS OCCURRED ON 2 SEPARATE UNITS ON 2 SEPARATE PATIENTS AND ONE OF THE PATIENTS NEEDED TO HAVE THE BALLOON REMOVED BY A UROLOGIST. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON 26JAN2023, IT WAS STATED THAT THE CUSTOMER HAD TWO MORE INSTANCES IN THE LAST MONTH WHERE STAFF HAD REPORTED THAT THEY HAD DIFFICULT DEFLATING THE FOLEY BALLOON. PER FOLLOW UP VIA EMAIL ON 15FEB2023, THERE WAS DELAY IN PATIENT CARE AND BEING ABLE TO REMOVE THE PRODUCTS, THERE WAS NO PATIENT INJURY OR IMPACT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE PRODUCT CATALOG NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00074 | 2023-02-15 | 2023-01-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U41502H4RX | U41502H4RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED UNDER NOMINAL PRESSURE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 01/2024. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00075 | 2023-02-15 | 2023-01-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | U4150630RX |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON CATHETER WAS ALLEGEDLY FOUND TO BE INCORRECTLY DESCRIBED IN THE LABEL. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 07/2025. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00080 | 2023-02-15 | 2022-11-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 014 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | US14150610 | US14150610 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE SUPERFICIAL FEMORAL ARTERY VIA A CONTRALATERAL APPROACH, THE GUIDEWIRE ALLEGEDLY GOT STUCK AFTER INFLATING. IT WAS FURTHER REPORTED THAT THE BALLOON AND GUIDEWIRE WERE REMOVED TOGETHER AS ONE UNIT. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ULTRASCORE 014 PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, THE STRAIN RELIEF WAS NOTED TO BE DISLODGED FROM THE Y-BODY. NO KINKS WERE NOTED TO THE CATHETER SHAFT AND NO ANOMALIES TO THE LUERS/Y-BODY. DURING FUNCTIONAL EVALUATION, THE SAMPLE GUIDEWIRE UNABLE TO BE FLUSHED. THE PATENCY OF THE GUIDE-WIRE LUMEN WAS TESTED USING A RETURNED GUIDEWIRE AND WAS UNABLE TO BE INSERTED ALL THE WAY THROUGH. THE CATHETER WAS STRIPPED AND UNDER MICROSCOPIC EXAMINATION, DRY BLOOD WAS NOTED WITHIN THE INNER GUIDEWIRE LUMEN. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED GUIDEWIRE REMOVAL ISSUE AS FUNCTIONAL TESTING COULD NOT BE PERFORMED DUE TO DRIED BLOOD WITHIN THE GUIDEWIRE LUMEN. HOWEVER, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED DISLODGED STRAIN RELIEF AS THE STRAIN RELIEF WAS NOTED TO BE DISLODGED FROM THE Y-BODY DURING VISUAL EVALUATION. A DEFINITIVE ROOT CAUSE FOR THE REPORTED GUIDEWIRE REMOVAL ISSUE AND IDENTIFIED DISLODGED STRAIN RELIEF COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRY DATE: 08/2024) . | N | Y | HEALTH PROFESSIONAL |  | 2023-02
MW5114992 | 2023-02-15 | 2023-02-02 | I, n, j, u, r, y | Voluntary report | BARD SURESTEP FOLEY TRAY SYSTEM 16 FR TEMPERATURE-SENSING | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | A119216M | A119216M |  | URETHRAL CATHETER REMOVED DUE TO COMPLAINTS OF BLADDER RETENTION. BULB DEFLATED HAVING ONLY 3 CC OF A CLEAR YELLOW COLOR FLUID. CATHETER BULB RECHECK AND WAS RE-INFLATED TO RULE OUT MALFUNCTION AND IT WAS NOTED THE BULB DID HAVE A LEAK. PATIENT WAS ABLE TO URINATE ON HER OWN AND WAS NOT RE-CATHETERIZED. NO INJURY TO PATIENT WAS NOTED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00893 | 2023-02-15 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L18 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON DID NOT INFLATE DURING THE PRETEST. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE CONTRIBUTED BY TORN RUBBERIZE ON RUBBERIZE LAYER THAT MAY LEAD TO LUMEN-TO-LUMEN LEAK. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE(S) THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00891 | 2023-02-15 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NATURAL¿ NON-ADHESIVE MALE EXTERNAL CATHETER | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 38303 | 38303 |  | IT WAS REPORTED THAT THE PATIENT HAD TO PINCH AND PULL THE MALE EXTERNAL CATHETER OFF AND WHICH NOW HAD CAUSED A GROWTH ON THE END OF HIS PENIS. NO MEDICAL ATTENTION REPORTED. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS INITIAL MDR WAS REPORTED IN ERROR. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-00897 | 2023-02-15 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT THE NURSING STAFF WAS ENCOUNTERING RESISTANCE WITH THE REMOVAL OF A FOLEY CATHETER ON THE PATIENT. THE BALLOON WAS TOTALLY DEFLATED AND THE FOLEY CATHETER WAS STUCK APPROXIMATELY 2 INCHES UP THE SHAFT FROM THE END OF THE PENIS. WARM BLANKET HAD BEEN TRIED AND THE PHYSICIAN ASSISTANT AND PHYSICIAN WERE CONTACTED. PHYSICIAN ASSISTANT WAS ABLE TO REMOVE WITH SOME ADDITIONAL FORCE REQUIRED. UPON INSPECTION OF THE CATHETER IT WAS NOTED THAT THE BALLOON WAS TOTALLY DEFLATED BUT LEFT A RIMMED EDGE TO THE FOLEY CATHETER WHICH WAS NOT NORMAL. NO BLOOD NOTED ON THE CATHETER OR EVIDENCE OF ANY TRAUMA. NO SERIOUS HARM WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00886 | 2023-02-15 | 2023-01-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED THAT THE TUBING OF SOME OF THE INTERMITTENT CATHETERS LOOKED THICKER THAN OTHERS. NOT ALL CATHETERS WERE THICKER. THE REPRESENTATIVE ASKED IF THIS WAS CAUSING THE CUSTOMER ANY ISSUES AND THE CUSTOMER STATED, NO, THEY JUST THOUGHT OF REPORTING THESE AS SOME TUBES LOOKED THICKER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿OPERATOR ERROR.¿ THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿FOR UROLOGICAL USE ONLY. INTENDED FOR USE BY PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICES ARE PASSED TO THE URINARY BLADDER VIA THE URETHRA. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. 1. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. 2. PRIOR TO OPENING THE SEALED CATHETER POUCH, APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. 3. WET THE CATHETER BY HOLDING THE PACKAGE WITH THE PRINTED SIDE UP AND TIP THE PACKAGE END-TO-END THREE TO SIX TIMES TO WET THE CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. 4. PEEL OPEN THE PACK AT THE FUNNEL/HANDLE END JUST ENOUGH TO EXPOSE 5 CM (2¿) OF THE CATHETER, OR FOR PRODUCTS WITH AN INSERTION SLEEVE, EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 5. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. 6. WASH YOUR HANDS AGAIN. 7. USING THE CATHETER WITH INSERTION SLEEVE (GRIPPER): A. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. B. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. C. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. WITHOUT INSERTION SLEEVE (GRIPPER): A. HOLD THE HANDLE/FUNNEL END AND REMOVE THE CATHETER FROM THE PACKAGING. B. GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA AND ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. 8. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. 9. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER.¿ | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-00936 | 2023-02-16 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON ASYMMETRICALLY INFLATED. THERE WERE MANY SIMILAR CASES, BUT WHEN THE BALLOON WAS INFLATED, THE BALLOON PART WAS DEFORMED. THE BULGE WAS BIASED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, CLARIFIED THAT THE BALLOON ASYMMETRICALLY INFLATED. THE REPORTER OF THIS COMPLAINT MADE VAGUE REFERENCES TO SIMILAR CASES IN THE PAST AND DID NOT POINT TO A SPECIFIC CASE. THUS, DETAILS FOR SIMILAR CASES WERE NOT AVAILABLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00935 | 2023-02-16 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165814 | 165814 |  | IT WAS REPORTED THAT THE DESCRIPTION ON THE PACKAGING DIFFERS TO THE DESCRIPTION ON THE VALVE OF THE FOLEY CATHETER. THE CUSTOMER CONTINUED TO USE THE SAME CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00456 | 2023-02-16 | 2023-01-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT ONE YEAR ELEVEN MONTHS AND FIFTEEN DAYS POST PORT PLACEMENT, THE DEVICE ALLEGEDLY HAD DIFFICULTY IN FLUSHING. IT WAS FURTHER REPORTED THAT THE SEPTUM OF THE PORT WAS ALLEGEDLY MOVING UP AND DOWN WHEN ATTEMPTING TO FLUSH USING NORMAL PRESSURE. REPORTEDLY, THE PORT WAS REMOVED. THE PROCEDURE WAS COMPLETED REPLACING ANOTHER DEVICE SUCCESSFULLY. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00084 | 2023-02-16 | 2023-01-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 018 PTA DILATATION CATHETER | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U813054 | U813054 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE SHAFT OF THE PTA BALLOON WAS ALLEGEDLY BROKEN. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ULTRAVERSE 018 PTA DILATATION CATHETER WAS RETURNED FOR EVALUATION. NO EVIDENCE OF A BREAK NOR ANY DEFORMITY WAS NOTED DURING THE VISUAL EVALUATION. A FUNCTIONAL TEST WAS NOT PERFORMED DUE TO THE NATURE OF THE COMPLAINT. THEREFORE, THE INVESTIGATION WAS UNCONFIRMED FOR THE REPORTED BREAK AS THERE WAS NO EVIDENCE OF A BREAK OR ANY DEFORMITY NOTED DURING THE VISUAL EVALUATION OF THE RETURNED DEVICE. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BREAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 11/2024), G3 . H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00459 | 2023-02-16 | 2023-01-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN,WHITE, 4.2F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 0601610 | 0601610 |  | IT WAS REPORTED THAT DURING PREPARATION OF A CATHETER REPAIRING PROCEDURE, THE ADHESIVE ON THE DEVICE WAS ALLEGEDLY SOLID AND WOULDN'T COME OUT OF TUBE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: A VOLUNTARY RECALL HAS BEEN INITIATED FOR CATHETER REPAIR KITS WHICH ARE PRODUCT CATALOG/LOT NUMBER SPECIFIC. REPORTEDLY THE CATHETER REPAIR KITS HAVE HARDENED/COAGULATED ADHESIVE. HARDENED OR COAGULATED ADHESIVE HAS THE POTENTIAL TO CAUSE DELAYS WHILST AN ALTERNATIVE REPAIR KIT, ADHESIVE OR REPLACEMENT CATHETER IS OBTAINED; THEREFORE, POTENTIALLY PROLONGING SURGERY OR NECESSITATING EXCHANGE. TO DATE THERE HAVE BEEN NO ADVERSE EVENTS WORLDWIDE REPORTED RELATED TO THIS ISSUE. AS RESULT OF THE FIELD ACTION, THIS EVENT IS BEING REPORTED AS A MALFUNCTION REPORTABLE EVENT. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 09/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00951 | 2023-02-16 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165PL08 |  | IT WAS REPORTED THAT FOLEY CATHETER TIP SIZE WAS FOUND TO BE INCONSISTENT WITH OTHER CATHETERS OF THE SAME BATCH. THE SAME ISSUE WAS REPORTED IN ANOTHER RECORD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00946 | 2023-02-16 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTERMITTENT ROCHESTER CATHETERS | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER NOTICED A CHANGE IN THE SIZE AND AMOUNT OF WATER LUBE IN SOME OF THE MAGIC3 CATHETERS THEY RECEIVED. ALSO STATED THAT THE CATHETERS CAUSED SOME IRRITATION DUE TO LACK OF THE LUBE PROVIDED. PER FOLLOW UP VIA EMAIL RECEIVED ON 09FEB2023, THEY DID NOT HAVE THE LOT NUMBERS AND NOTICED THAT THE SIZE OF THE PACKETS CHANGED ONE SOME OF THEM AND THE AMOUNT CHANGED ALSO. THERE WAS NO PATIENT INJURY BUT JUST LITTLE CHUNKS OF BLOOD THAT THE DOCTOR DESCRIBED AT TRAUMA. THEY WERE STILL ABLE TO VOID THE INDIANA POUCH BUT COULD VISIBLY SEE THE CATHETER CATCHING ONTO THE STOMA WHEN PULLED OUT. THEY ALSO STATED THAT THEY LIKED THE 50816G THAN THE 50616. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-00945 | 2023-02-16 | 2023-01-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE INFLATION VALVE OF THE FOLEY CATHETER AND THE BALLOON DID NOT INFLATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00462 | 2023-02-16 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D5 |  | IT WAS REPORTED BY THE CUSTOMER THAT "WHEN ASPIRATING FROM THE SIDE PORT OF THE PICC STYLET HE NOTICES AIR COMING OUT WHERE THE RUBBER STOPPER AND STYLE JOIN/MEET. WHEN FLUSHING HE NOTICES FLUID COMING OUT OF THE SAME AREA." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGW0935 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00461 | 2023-02-16 | 2023-01-19 | D, e, a, t, h | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TIP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED THAT DURING POST MORTEM THE REMAINS OF A PICC WERE FOUND IN THE PERSONS ATRIUM. PICC INSERTED (B)(6) 2022, X RAY ON ADMISSION TO HOSPITAL ON (B)(6) 2022 REPORTED THAT THE DEVICE WAS IN A CORRECT POSITION. THE PICC WAS REMOVED AT SOMETIME AFTER THIS ON PATIENTS ADMISSION TO HOSPITAL BUT THERE IS NO DOCUMENTATION AS TO WHEN THIS HAPPENED. THE PATIENT WAS UNDERGOING CHEMOTHERAPY AND IT IS UNKNOWN AT THIS TIME AS TO THE CAUSE OF DEATH, AND THE FINDINGS OF THE REMAINS OF THE PICC WAS ONLY AFTER THE PATIENTS DEATH ON POSTMORTEM. ADDITIONAL INFORMATION RECEIVED ON 2/2/2023: CUSTOMER REPORTS THAT THE CATHETER FRAGMENTS FOUND DURING POST-MORTEM CARE ARE BELIEVED TO BE FROM THE PICC. THE CAUSE OF DEATH IS BEING INVESTIGATED BY THE CORONER AND IS CURRENTLY UNASCERTAINED. THE PATIENT EXPIRED (B)(6) 2022. ADDITIONAL INFORMATION RECEIVED: DEVICE BEING USED FOR PARENTERAL NUTRITION. INSERTED IN THE RIGHT BASILIC VEIN. LIDOCAINE USED AT INSERTION SITE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00929 | 2023-02-16 | 2023-01-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316 | 119316 |  | IT WAS REPORTED THAT STERILIZED WATER DID NOT COME OUT DURING THE PRETEST OF THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00473 | 2023-02-16 | 2022-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE DRESSING IN PICC KIT HAS LOW ADHERENCE TO SKIN, EVEN WITH EXTRA GLUE IT'S REQUIRED TO REQUEST FOR ANOTHER DEVICE TO AVOID THE EXTERIORIZATION RISK. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFS4713 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00082 | 2023-02-16 | 2022-12-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT AFTER A RECANALIZATION PROCEDURE OF CHRONIC TOTAL OCCLUSION IN THE SUPERFICIAL FEMORAL ARTERY VIA A CONTRALATERAL APPROACH, THE DEVICE WAS ALLEGEDLY FOUND TO BE FLUSHED FROM THE CENTRAL PART OF THE SHAFT. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN D2 AND G4. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 10/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00937 | 2023-02-16 | 2023-01-27 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT URINARY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE URINARY TRACT INFECTIONS. THEY STATED THAT THEY HAD STOPPED USING INTERMITTENT CATHETERS AND HAD A DOCTORS APPOINTMENT IN THREE MONTHS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: MAGIC3"INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN WARNINGS. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-02
1018233-2023-00925 | 2023-02-16 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT WHEN THE BALLOON WAS INFLATED IN THE PRETEST, THERE WAS A HOLE AND WATER LEAKED FROM THE FOLEY CATHETER. | THE REPORTED EVENT WAS CONFIRMED-MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS INFLATED WITH WATER AND OBSERVED WATER LEAKING THRU DRAINAGE EYE. DISSECTED THE CATHETER AND OBSERVED TEAR CREATED ON RUBBERIZE LAYER OF CATHETER. TEAR WAS EXAMINED UNDER MICROSCOPE AND NOTED TO BE ROUNDED AND NOT SMOOTH AND IDENTIFIED THIS IS RELATED TO MANUFACTURING. A POTENTIAL ROOT CAUSE FOR THIS POTENTIAL FAILURE MODE COULD BE WIRE PRESSING ERROR. A REVIEW OF THE MANUFACTURING PROCESS INDICATES THAT THE PROCESS IS CAPABLE OF PRODUCING OF THIS TYPE OF DEFECT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00926 | 2023-02-16 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PLACED ON (B)(6), SPONTANEOUS REMOVAL WAS CONFIRMED ON THE (B)(6). THE PRODUCT WAS REPLACED AND PLACED AGAIN, BUT SPONTANEOUS WITHDRAWAL WAS CONFIRMED ON THE (B)(6), ONE DAY LATER. WHEN WATER WAS REINJECTED OUTSIDE THE BODY AND THE BALLOON PART WAS OBSERVED, IT WAS CONFIRMED THAT THE WATER SEEPED OUT, ALTHOUGH NO SCRATCHES WERE VISIBLE WITH THE NAKED EYE. SINCE NO PRE-TEST WAS DONE, IT WAS NOT POSSIBLE TO CONFIRM WHETHER THE SCRATCHES WERE PRESENT BEFORE USE. IN ADDITION, WHEN WE ASKED THE OPINION OF DR. UROLOGY, ALSO COMMENTED ON THE POSSIBILITY OF DAMAGE TO THE BALLOON PART (FROM THE PATIENT) DUE TO FOREIGN SUBSTANCES IN THE URINE. ALSO STATED THAT IF THEY COULD INCLUDE THE FACTORS THAT WERE CONSIDERED TO BE THE CAUSE OF THE BALLOON RUPTURE WHEN THEY PREPARE THE REPORT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
16385098 | 2023-02-16 | 2022-10-05 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDIA® FOLEY CATHETER SILICONE COATED | CATHETER, COUDE | C. R. BARD, INC. | 123520A | 123520A |  | FOLEY REMOVED DUE TO PATIENT REPORTING IT WAS LEAKING, PER CHART REVIEW FOLEY DUE TO BE CHANGED, THEREFORE PATIENT REQUESTING NEW CATHETER TO BE INSERTED. UPON REMOVAL OF OLD CATHETER, A SUBSTANCE OF PALE YELLOW LANDED ON PATIENT'S VULVA. UPON INSPECTION THE MATERIAL WAS HARD AND NON-PALPABLE, WHEN SQUEEZED IT BROKE INTO 3 PIECES. CLOSE EXAMINATION OF THE CATHETER TIP/DEFLATED BALLOON. WITH VISUAL AND TACTILE INSPECTION, IT WAS NOTED THAT WHEN THE TIP OF THE CATHETER WAS SQUEEZED/LIGHTLY SCRAPED WITH A GLOVED FINGERNAIL, PARTICLES WERE DISLODGING FROM THE CATHETER AND APPEARED TO BE CATHETER MATERIAL DISLODGING FROM THE CATHETER. |  | N | Y |  |  | 2023-02
3006260740-2023-00457 | 2023-02-16 | 2023-01-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602833 | 0602833 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, AIR WAS ALLEGEDLY ASPIRATED IN DEVICE DURING ASPIRATION BETWEEN THE SYRINGE AND THE INTRODUCER NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 09/2025. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00081 | 2023-02-16 | 2023-01-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80144 | ATG80144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON WAS ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT THE PTA BALLOON WAS ALLEGEDLY DETACHED FROM THE SHAFT OF THE CATHETER. IT WAS ALSO REPORTED THAT INITIALLY THE PROXIMAL END OF THE BALLOON DETACHED AND THE BALLOON PROLAPSED DISTALLY, BUT THE DISTAL END OF THE BALLOON REMAINED WEAKLY ATTACHED. FURTHERMORE, THE CATHETER WAS REMOVED, HOWEVER DISTAL END OF THE BALLOON CAME INTO THE TIP OF THE SHEATH. THE SHEATH WAS THEN REMOVED FROM THE PATIENT LEAVING THE WIRE IN PLACE. REPORTEDLY, AT THAT POINT, THE PTA BALLOON WAS COMPLETELY DETACHED FROM THE SHAFT, BUT A TINY BIT OF IT BARELY PROTRUDED OUT OF THE SKIN. THE PROTRUDING BIT OF BALLOON WAS GRASPED WITH A CLAMP AND REMOVED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00455 | 2023-02-16 | 2023-01-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT ONE DAY POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED ITCHING AROUND THE INCISION SITE. HOWEVER, THERE IS NO FURTHER INFORMATION REGARDING ADDITIONAL INTERVENTION OR ADDITIONAL MEDICATION PRESCRIBED TO TREAT ITCHING. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3. H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
MW5115011 | 2023-02-16 | 2023-01-11 | I, n, j, u, r, y | Voluntary report | POWERPORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. | 1859600 | 1859600 |  | POWERPORT REMOVAL BY SURGEON. ONLY A PORTION OF THE PORT CAME OUT. A REMAINING REMNANT OF THE CATHETER WAS SEVERED INTERNALLY AND LEFT IN THE PATIENT. THE PATIENT THEN HAD TO BE SCHEDULED FOR SURGERY FOR A FOREIGN BODY REMOVAL. |  | Y | N | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00458 | 2023-02-16 | 2022-11-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER, SINGLE-LUMEN, 4.2F | CHRONIC CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 0600520 | 0600520 |  | IT WAS REPORTED THAT TWENTY EIGHT DAYS POST CHRONIC CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY SPONTANEOUSLY PRESENTED BUBBLES, WHILE THE PATIENT HAD ALREADY TURNED OFF PARENTERAL NUTRITION AND WASHED THE CATHETER ACCORDING TO THE PROTOCOL. IT WAS FURTHER REPORTED THAT IT REMAINED WITH EBB AND FLOW, BUT WHEN INFUSED AGAIN IT ALLEGEDLY BROKE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRY DATE: 01/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00968 | 2023-02-17 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT A DYSFUNCTION WAS NOTED IN FOLEY CATHETERS ON 24-JAN-2023. THE USER REQUIRED THEM DURING TRANSURETHRAL RESECTION OF THE PROSTATE SURGERY. WHEN TESTING TWO OF THE FOLEY CATHETERS, BEFORE INTRODUCING THEM TO THE PATIENT, THE USER NOTICED THE DYSFUNCTION OF THE FIXATION BALLOON FILLING AND EMPTYING SYSTEM. DURING THE SAME MANEUVER WITH A THIRD PROBE, THERE SEEMED TO BE NO FAILURE, BUT WHEN TRYING TO DEFLATE THE FIXATION BALLOON WITH THE PROBE ALREADY INSERTED IN THE PATIENT, IT WAS NOT POSSIBLE AFTER TRYING FOR APPROXIMATELY 10 MINUTES. THE USER DID NOT SUCCEED, SO THEY CUT THE FILLING AND EMPTYING SYSTEM, MANAGING TO DEFLATE THE BALLOON AND REMOVE THE FOLEY CATHETER. THE USER SENT PHOTOGRAPHIC EVIDENCE OF THE TWO FOLEY CATHETERS WITH THE INFLATED BALLOON THAT FAILED TO DEFLATED DURING TEST PRIOR PLACEMENT AND THE THIRD DEVICE WHICH WAS CUT AFTER BEING PLACED INTO PATIENT. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 01-FEB-2023, STATED THAT THERE WAS NO HARM TO THE PATIENT. THE DEVICE WAS REMOVED BY CUTTING THE INFUSE (INFLATING) CHANNEL, AND BY REPLACING IT BY ANOTHER DEVICE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE POOR SAMPLE CONDITION. SAMPLE HAS BEEN EVALUATED AND OBSERVED NO ABNORMALITIES. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE THIN RUBBERIZE WALL THAT CAUSING COLLAPSE/PINCH INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE OVER ASPIRATED, INCORRECT SYRINGE)/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." CORRECTION: D,H. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00969 | 2023-02-17 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT A DYSFUNCTION WAS NOTED IN FOLEY CATHETERS ON (B)(6) 2023. THE USER REQUIRED THEM DURING TRANSURETHRAL RESECTION OF THE PROSTATE SURGERY. WHEN TESTING TWO OF THE FOLEY CATHETERS, BEFORE INTRODUCING THEM TO THE PATIENT, THE USER NOTICED THE DYSFUNCTION OF THE FIXATION BALLOON FILLING AND EMPTYING SYSTEM. DURING THE SAME MANEUVER WITH A THIRD PROBE, THERE SEEMED TO BE NO FAILURE, BUT WHEN TRYING TO DEFLATE THE FIXATION BALLOON WITH THE PROBE ALREADY INSERTED IN THE PATIENT, IT WAS NOT POSSIBLE AFTER TRYING FOR APPROXIMATELY 10 MINUTES. THE USER DID NOT SUCCEED, SO THEY CUT THE FILLING AND EMPTYING SYSTEM, MANAGING TO DEFLATE THE BALLOON AND REMOVE THE FOLEY CATHETER. THE USER SENT THE PHOTOGRAPHIC EVIDENCE OF THE TWO FOLEY CATHETERS WITH THE INFLATED BALLOON THAT WAS TESTED PRIOR PLACEMENT AND THE THIRD DEVICE WHICH WAS CUT AFTER BEING PLACED INTO PATIENT. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 01-FEB-2023, STATED THAT THERE WAS NO HARM TO THE PATIENT. THE DEVICE WAS REMOVED BY CUTTING THE INFUSE (INFLATING) CHANNEL, AND BY REPLACING IT BY ANOTHER DEVICE. | THE REPORTED EVENT IS INCONCLUSIVE-POOR SAMPLE CONDITION DUE TO 1 OF RETURNED FOLEY CATHETER HAS BEEN CUT AT THE INFLATION FUNNEL AND FULL INVESTIGATION COULD NOT BE PERFORMED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE THIN RUBBERIZE WALL THAT CAUSING COLLAPSE/PINCH INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT. SAMPLE HAS BEEN EVALUATED AND OBSERVED NO ABNORMALITIES. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00979 | 2023-02-17 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT STATED ONE PACKAGE WAS EMPTY. NO ADDITIONAL INFORMATION PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿FAILURE OF CATHETER DETECTION SENSOR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS; URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
3006260740-2023-00482 | 2023-02-17 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV SP 4F DL MAX | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108PD |  | IT WAS REPORTED BY THE CUSTOMER, 'WIRE IN THE KIT IS NOT SMOOTH ON INSERTION'. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00964 | 2023-02-17 | 2023-01-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THEY WERE TREATING A PATIENT ON ARCTIC SUN DEVICE. THE DEVICE STOPPED READING THE PATIENT TEMPERATURE. NURSE ADVANCED THE FOLEY AND MADE SURE IT WAS FIRMLY CONNECTED TO THE TEMPERATURE IN CABLE. MIS CONFIRMED THE CABLE WAS CONNECTED TO TEMPERATURE PORT 1. THEY HAVE ANOTHER PROBE GOING TO THE BEDSIDE MONITOR, BUT THE CONNECTOR WAS NOT THE SAME AS THE ONE ON THE FOLEY. THE TEMPERATURE IN CABLE WAS ZIP-TIED AND COULD NOT BE REMOVED BY THE NURSES. MIS SUGGESTED TRYING ANOTHER FOLEY OR CHECKING IF BIOMED COULD REPLACE THE CABLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00483 | 2023-02-17 | 2023-01-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602830 | 0602830 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY BENT WHILE IMPLANTED. IT WAS FURTHER REPORTED THAT THE CATHETER WAS UNABLE TO PUT IT IN. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 02/2025. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00987 | 2023-02-17 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BIPOLAR ELECTRODE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT BALLOON OF THE TEMPORARY PACING ELECTRODE CATHETER HAD BURST. PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON 09FEB2023, STATED THAT THERE WAS NO MISSING PIECES OF A BALLOON. | THE REPORTED EVENT IS UNCONFIRMED AS THE REPORTED PROBLEM COULD NOT BE REPRODUCED. THE USED TEMPORARY PACING ELECTRODE WITH ATTACHED LUER SYRINGE WAS RETURNED WITH THE ORIGINAL PACKAGING. THE DEVICE WAS USED FOR TREATMENT PURPOSES. THE DEVICE HAD MET RELEVANT SPECIFICATIONS. NO ROOT CAUSE WAS REQUIRED AS THE REPORTED EVENT IS UNCONFIRMED. AS THE REPORTED EVENT IS UNCONFIRMED, A DHR REVIEW IS NOT REQUIRED. AS THE REPORTED EVENT IS UNCONFIRMED, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00977 | 2023-02-17 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTERMITTENT ROCHESTER CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER HAD URINARY PROBLEMS AND WAS NOT ABLE TO USE THE CATHETER RIGHT. ALSO STATED THAT THEY COULD NOT GET IT TO STAY ON AND WANTED INSTRUCTION ON HOW TO USE PRODUCT AND ASKED IF THERE WAS ANOTHER, EASIER TO USE. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-00963 | 2023-02-17 | 2023-01-28 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TARCT INFECTION FROM USING THE PUREWICK FEMALE EXTERNAL CATHETER. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS: DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS." "PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE". " RECOMMENDATIONS: ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-02
1018233-2023-00982 | 2023-02-17 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER WAS NOT ABLE TO BE INJECTED INTO THE FOLEY CATHETER AND LEAKED FROM SIDE OF THE INFLATION VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00085 | 2023-02-17 | 2022-12-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQ7586 | CQ7586 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY FAILED TO INFLATE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE CONQUEST PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, NO KINKS WERE NOTED TO THE CATHETER SHAFT AND NO ANOMALIES TO THE LUERS AND BIFURCATE. DURING FUNCTIONAL TESTING, AN IN-HOUSE PRESTO INFLATION DEVICE WAS USED TO INFLATE THE BALLOON AND WATER WAS NOTED STREAMING FROM THE BALLOON. THE BALLOON FIBERS WERE STRIPPED AND UNDER MICROSCOPIC EXAMINATION, A PINHOLE RUPTURE WAS NOTED TO THE BALLOON. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE INFLATION ISSUE AND IDENTIFIED PINHOLE RUPTURE AS WATER WAS NOTED STREAMING FROM A PINHOLE RUPTURE ON THE BALLOON UNDER MICROSCOPIC EXAMINATION. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. EXPIRY DATE: 04/2025. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00481 | 2023-02-17 | 2023-01-13 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808060 | 1808060 |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS POST PORT PLACEMENT PROCEDURE, THE PORT ALLEGEDLY HAD SUCTION ISSUE. THE PATIENT IS BEING TREATED FOR COLON, RECTAL CANCER WHICH METASTASIZED INTO BONE CANCER. IT WAS FURTHER REPORTED THAT THE PORT WAS IDENTIFIED WITH FIBRIN SHEATH AND THE HEALTH CARE PROVIDER ORDERED A DE-CLOTTING AGENT TO RESOLVE THIS ISSUE. THE PATIENT IS REPORTED TO BE STABLE. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRY DATE: 01/2020. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00484 | 2023-02-17 | 2023-01-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY ONE MONTH POST PORT PLACEMENT, BLOOD RETURN COULD NOT BE ALLEGEDLY CONFIRMED. IT WAS FURTHER REPORTED THAT THE PORT ALLEGEDLY LEAKED NEAR THE PORT STEM AND THE VERTEX OF THE VASCULAR INSERTION SITE. REPORTEDLY, THE CATHETER ALLEGEDLY HAD A CRACK ON THE CATHETER AND THE PORT STEM WAS KINKED. THE DEVICE WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE POWERPORT IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. GROSS, MICROSCOPIC VISUAL, TACTILE EVALUATION AND FUNCTIONAL TESTING WERE PERFORMED. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, TWO CHEST RADIOGRAPHS AND ONE FLUOROSCOPIC VIEW OF THE CHEST WERE PROVIDED FOR REVIEW. A KINK WAS NOTED ON THE ATTACHED CATHETER, PROXIMAL TO THE DISTAL END OF THE CATH-LOCK. THE IMAGE SHOWS THE DISRUPTION OF CATHETER INTEGRITY CAUSING FLUID LEAKAGE. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED CATHETER DAMAGE AND FLUID LEAK ISSUES. HOWEVER THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED DEFORMATION ISSUE AS NO DAMAGE WAS NOTED ON THE PORT STEM. ALSO THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED SUCTION ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
16395871 | 2023-02-17 | 2023-01-20 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARD URINE METER FOLEY TRAY WITH LUBRI-SIL ALL-SILICONE FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | 942216 | 942216 |  | PATIENT EXPERIENCING AN UNUSUAL AMOUNT OF DISCOMFORT WITH THE INSERTED CATHETER AND ONGOING URGENCY TO URINATE. |  | N | Y |  |  | 2023-02
3006260740-2023-00497 | 2023-02-20 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 7F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0603870CE |  | IT WAS REPORTED THAT DURING PORT DEVICE IMPLANT PROCEDURE, THE PATIENT ALLEGEDLY EXPERIENCED ALLERGY. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 7F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 7F PRODUCTS IS IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00999 | 2023-02-20 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE-SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER IN THE FOLEY CATHETER BALLOON WAS UNABLE TO BE REMOVED DURING A PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00994 | 2023-02-20 | 2023-01-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | URETHRAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE TOUCHLESS PLUS CATHETER INSIDE DETACHED AND CAUSED A MESS. PATIENT SAID THE DESIGN OF THIS CATHETER IS TO BE A CLOSED SYSTEM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01007 | 2023-02-20 | 2023-01-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123614A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS USED FOR URINE DRAINAGE, THEN IT WAS FOUND THAT THE CATHETER BALLOON WAS DAMAGED, AND THE CATHETER WAS REMOVED. PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON 09FEB2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATER ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-00996 | 2023-02-20 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316M | 119316 |  | IT WAS REPORTED THAT CUSTOMER HAVING ISSUES WITH THE FOLEY CATHETER BALLOONS IN THEIR KITS DEFLATING DURING CASES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01000 | 2023-02-20 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILIZED WATER WAS UNABLE TO BE REMOVED FROM THE BALLOON AFTER THE FOLEY CATHETER PLACEMENT TO THE PATIENT. ABOUT 10 ML OF WATER WAS ASPIRATED BY SYRINGE. IT WAS STATED THAT THE USER TRIED TO DEFLATE USING SYRINGE AND THE BALLOON WAS DEFLATED COMPLETELY. USUALLY, BALLOON CAN BE DEFLATED WITHIN 20-40 SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00090 | 2023-02-20 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE OF A CALCIFIED TARGET LESION IN THE SUPERFICIAL FEMORAL ARTERY VIA AN IPSILATERAL ANTEGRADE APPROACH, THE TIP OF THE CATHETER ALLEGEDLY DETACHED. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY VIBRATED WEAKLY. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED . AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 10/2023). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01004 | 2023-02-20 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT AFTER TURNING PATIENT IN BED, PATIENTS FOLEY CATHETER WAS NOTED TO BE OUT FROM THE FOLEY STAT LOCK. WHEN STAFF WAS ABOUT TO REATTACH THE FOLEY BALLOON PORT OR URINE DRAINAGE PORT TO THE STATLOCK NOTICED THAT THE BALLOON PORT WAS SEVERED AND SOON RIGHT AFTER THE CATHETER THEN SLIPPED OUT OF PATIENT¿S PENIS AND BALLOON WAS ALREADY DEFLATED. PATIENT WAS NOT INJURED WHILE THIS INCIDENT OCCURRED. STATED THAT THE STAFF HEARD IN TIER 4 THIS MORNING THAT THERE WERE ALSO REPORTS OF THE DRAINAGE BAG EMPTYING VERY SLOWLY. THIS SOUNDS LIKE AN AIR FILTER PROBLEM TO THEM. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "LATEX MATERIAL TOO WEAK TO WITHSTAND NORMAL FORCES APPLIED DURING USE". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00488 | 2023-02-20 | 2023-01-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT IT WAS DIFFICULT TO INSERT THE PROXIMAL END OF THE INTRODUCER SHEATH IN SELDINGER INTO THE VESSEL. CRACKS WERE FOUND WITH THE GREY SLEEVE, AFFECTING ITS USE. VIGOROUS FORCES WERE NOT APPLIED IN THE PROCESS OF INSERTION. INSTEAD, GENTLE FORCES WERE APPLIED AND CEASED IMMEDIATELY ONCE RESISTANCE WAS MET. THE DEVICE WAS REMOVED. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REFY3459 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00490 | 2023-02-20 | 2022-12-28 | I, n, j, u, r, y | Manufacturer report | GROSHONG 4F SL PICC (60 CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7717405 |  | IT WAS REPORTED THAT ON (B)(6) 2022, UNDER THE GUIDANCE OF B-ULTRASOUND, THE CENTRAL VENOUS CATHETER KIT AND ACCESSORIES 7717405 (THREE-WAY VALVE PICC KIT) WERE INSERTED FROM THE BASILIC VEIN OF THE LEFT UPPER EXTREMITY INTO THE SUPERIOR VENA CAVA FOR INTRAVENOUS INFUSION OF DRUGS. PREVENT EXTRAVASATION OF CHEMOTHERAPY DRUGS. AFTER THE CATHETER IS INSERTED, THE CATHETER WILL BE DISINFECTED AND CHANGED EVERY WEEK, AND THE CATHETER WILL BE FLUSHED AND OTHER MAINTENANCE. FOR CHEMOTHERAPY: (B)(6) 2022, (B)(6) 2022, 2 CYCLES OF PACLITAXEL 130MG/M2 + CISPLATIN 80MG/M2 CHEMOTHERAPY, (B)(6) 2022, (B)(6) 2022, (B)(6) 2022, (B)(6) 2022, 4 CYCLES OF DOCETAXEL + CAMRELIZUMAB TREATMENT. (B)(6) 2022, (B)(6) 2022,(B)(6) 2022, (B)(6) 2022, (B)(6) 2022, 5 CYCLES OF CAPECITABINE 1000MG/M2/BID/D1-14+CARRELIZUMAB MONOCLONAL ANTIBODY 200MG TREATMENT. FOR THE TREATMENT OF ESOPHAGEAL CANCER. (B)(6) 2022, B-ULTRASOUND OF LEFT UPPER EXTREMITY VEIN SHOWED LEFT BASILIC VEIN, AXILLARY VEIN, AND SUBCLAVIAN VEIN WITH LOCAL THROMBUS IN LEFT BASILIC VEIN, AXILLARY VEIN, AND SUBCLAVIAN VEIN. FORM. ON (B)(6) 2022, 4000IU OF LOW MOLECULAR WEIGHT HEPARIN CALCIUM FOR INJECTION WAS STARTED, ONCE EVERY 12 HOURS, AND ANTICOAGULANT THERAPY WAS INJECTED SUBCUTANEOUSLY. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. DE | Y | N | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00494 | 2023-02-20 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT POST PORT PLACEMENT AFTER BEING DIAGNOSED WITH CANCER ON RIGHT SIDE VIA JUGULAR. THE PATIENT EXPERIENCED SEVERE PAIN AND SHORTNESS OF BREATH. HOWEVER, THERE WAS NO INFORMATION REGARDING TREATMENT OR MEDICATION PROVIDED FOR PAIN. THE CURRENT STATUS IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00496 | 2023-02-20 | 2022-05-09 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT APPROXIMATELY ONE MONTH AND TWENTY-TWO DAYS POST PORT PLACEMENT AFTER BEING DIAGNOSED WITH CANCER IN RIGHT BREAST. IT WAS REPORTED THAT THE DEVICE ALLEGEDLY HAD A LEAK DUE TO THE SILICONE CRACKING. IT WAS FURTHER REPORTED THAT THE PATIENT ALLEGEDLY EXPERIENCED PAIN, SWELLING AND REDNESS THAT BEGAN TO TRAVEL FROM FRONT RIGHT SHOULDER UP TO THE NECK, AND THE PORT WAS REMOVED DUE TO THESE COMPLICATIONS. REPORTEDLY THE PORT SITE WAS RUPTURED FROM CONSTANT INFECTION AND THE PATIENT WAS TESTED POSITIVE FOR METHICILLIN-SUSCEPTIBLE STAPHYLOCOCCUS AUREUS BLOOD INFECTION. INFECTION MEDICATIONS WERE PROVIDED FOR TREATMENT. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 03/2023. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00485 | 2023-02-20 | 2023-01-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7274118 |  | IT WAS REPORTED BY THE CUSTOMER "LEAKING CATHETER IN INVERTED CONE. PATIENT IS WITHOUT CENTRAL ACCESS AND WITH MULTIPLE PUNCTURES." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFT3533 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01040 | 2023-02-21 | 2022-11-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414 |  | IT WAS REPORTED THAT IT WAS DIFFICULT TO DEFLATE IN THE FOLEY CATHETER DURING PRETEST. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER, WATER SOLUBLE LUBRICANT, ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION, TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS GLOVES. [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01042 | 2023-02-21 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT THE BALLOON WAS DEFORMED. DURING A PRETEST PRIOR TO USE, THE SHAPE OF THE CATHETER BALLOON WAS DEFORMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01043 | 2023-02-21 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT THE BALLOON WAS DEFORMED. DURING A PRETEST, THE SHAPE OF THE CATHETER BALLOON WAS OBSERVED TO BE DEFORMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01041 | 2023-02-21 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT LEAD WIRE PENETRATED THE FOLEY CATHETER MATERIAL. THE USER ASKED WHETHER IT WAS POSSIBLE FOR SUCH AN EVENT TO OCCUR AND WHETHER THERE WERE ANY MEASURES TO PREVENT THIS FROM HAPPENING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01044 | 2023-02-21 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE TEMP SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119214M | 119214M |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON FOLDED BACK ON ITSELF WHEN IT WAS REMOVED RESULTING IN DIFFICULT REMOVAL. IT WAS CREATED A RING THAT ADDED TO THE CATHETER CIRCUMFERENCE. THIS WAS THE SECOND TIME THIS HAS HAPPENED ON THIS UNIT . THE FIRST TIME WAS ALMOST COULD NOT GET THE CATHETER OUT OF THE LITTLE BOY . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01047 | 2023-02-21 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE PATIENT STATED THE MAGIC 3 GO INTERMITTENT CATHETER PLASTIC INFLATED AND THERE WAS CONDENSATION ON THE INSIDE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01057 | 2023-02-21 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE BALLOON WAS BROKEN DURING USE IN THE EMERGENCY DEPARTMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01055 | 2023-02-21 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119214M | 119214M |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON FOLDED BACK ON ITSELF WHEN IT WAS REMOVED RESULTING IN DIFFICULT REMOVAL. IT WAS CREATED A RING THAT ADDED TO THE CATHETER CIRCUMFERENCE. THIS WAS THE SECOND TIME THIS HAS HAPPENED ON THIS UNIT. THE FIRST TIME WAS ALMOST COULD NOT GET THE CATHETER OUT OF THE LITTLE BOY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01050 | 2023-02-21 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS CAUSING ISSUES WHEN REMOVING IT FROM MEN, AND THERE WAS SOME SORT OF LIP ON THE END WHEN THEY DEFLATE THE BALLOON, AND IT GOT CAUGHT DURING REMOVAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01037 | 2023-02-21 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT A FOLEY CATHETER WAS INSERTED IN THE SURGERY, AND THE WATER IN THE BALLOON DID NOT COME OUT AT ALL WHEN IT WAS REMOVED IN THE WARD THE NEXT DAY; THERE WAS NO LEAK AND THE BALLOON WAS PARTIALLY DEFLATED. BUT THE CATHETER COULD BE REMOVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
16410205 | 2023-02-21 | 2022-09-28 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARD® URETHRAL CATHETERIZATION TRAY WITH VINYL CATHETER | CATHETER, UROLOGICAL | C. R. BARD, INC. | 772417 | 772417 |  | RN CARING FOR PATIENT ATTEMPTED TO PERFORM A STRAIGHT CATHETERIZATION TO OBTAIN URINE SPECIMEN. THE CATHETER FOLDED OVER ON ITSELF. WRITER TOOK A NEW, CLEAN, CATHETER AND TRIED TO PERFORM ON MANIKIN, AND SAME THING HAPPENED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01025 | 2023-02-21 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE PATIENT HAD TWO MAGIC 3 GO MALE COUDE INTERMITTENT CATHETERS WHERE THE BURST PACKET WAS EMPTY AND THE PATIENT COULD NOT USE THEM AND HAD ONE WHERE THE TIP WENT TO ONE SIDE AND TWISTED. THE PATIENT TRIED TO INSERT THE TWISTED CATHETER AND IT CAUSED A LITTLE TRAUMA AND A BIT OF BLEEDING. NO ADDITIONAL INFORMATION PROVIDED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." | N | Y | LAY USER/PATIENT |  | 2023-02
2020394-2023-00096 | 2023-02-21 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE TIP OF THE CATHETER ALLEGEDLY DETACHED. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED MATERIAL SEPARATION ISSUE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 10/2024) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-02
16408509 | 2023-02-21 | 2022-12-06 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP¿ FOLEY TRAY SYSTEM WITH BARD® LUBRI-SIL® TEMP SENSING FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | A119216M | A119216M |  | FOLEY WAS PLACED BY OPERATING ROOM RN AROUND 1115, AROUND 1255 IT WAS DISCOVERED THAT THE FOLEY CATHETER WAS NO LONGER INSIDE THE BLADDER. BALLOON WAS FILLED WITH 10CC OF AIR TO CHECK FOR HOLE. NO OBVIOUS SIGNS OF HOLES. AFTER SYRINGE WAS REMOVED, THE CATHETER BALLOON SLOWLY DEFLATED. |  | N | Y |  |  | 2023-02
3006260740-2023-00498 | 2023-02-21 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1618000 | 1618000 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN. IT WAS FURTHER REPORTED THAT, THE PORT WAS REMOVED. HOWEVER, THERE IS NO FURTHER INFORMATION REGARDING ADDITIONAL INTERVENTION OR PRESCRIBED ADDITIONAL MEDICATION TO TREAT PAIN. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00500 | 2023-02-21 | 2023-01-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT, THE CATHETER ALLEGEDLY CRACKED AND BROKEN. IT WAS FURTHER REPORTED THAT, THE DEVICE WAS REMOVED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00499 | 2023-02-21 | 2023-01-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PEEL APART INTRODUCER KIT, 10F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 |  | 0601110CE |  | IT WAS REPORTED THAT DURING A CHRONIC CATHETER PLACEMENT PROCEDURE, THE WIRE ALLEGEDLY BENT WHEN ATTEMPTED TO PULL IT OUT. IT WAS FURTHER REPORTED THAT THE WIRE ALLEGEDLY BROKE OFF AT ONE END. REPORTEDLY, THE DEVICE WAS REMOVED FROM THE PATIENT. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE PEEL APART INTRODUCER KIT, 10F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE PEEL APART INTRODUCER KIT, 10F ARE IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 02/2026. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00097 | 2023-02-21 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U4150410RX | U4150410RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF CHRONIC TOTAL OCCLUSION IN THE SUPERFICIAL FEMORAL ARTERY VIA AN IPSILATERAL ANTEGRADE APPROACH, THE PTA BALLOON ALLEGEDLY RUPTURED AT 3 ATM. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 08/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01015 | 2023-02-21 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT STERILIZED WATER LEAKED INTO THE DRAINAGE LUMEN WHEN IT WAS CONFIRMED THAT IT WAS NATURALLY REMOVED AFTER PLACEMENT. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO TORN RUBBERIZE ON RUBBERIZE LAYER THAT MAY LEAD TO LUMEN-TO-LUMEN LEAK. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE(S) THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01009 | 2023-02-21 | 2022-12-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI12 | 0165SI12 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON HAD BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01010 | 2023-02-21 | 2022-12-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILIZED WATER WAS DIFFICULT TO ENTER IN THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
16408508 | 2023-02-21 | 2022-12-04 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDEX® LUBRICATH® FOLEY CATHETER COUDE TIP | CATHETER, COUDE | C. R. BARD, INC. | 0168L18 | 0168L18 |  | AFTER INSERTING CATHETER AND GETTING URINE RETURN, THE BALLOON ON THE CATHETER WOULD NOT INFLATE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01014 | 2023-02-21 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE PATIENT HAD TWO MAGIC 3 GO MALE COUDE INTERMITTENT CATHETERS WHERE THE BURST PACKET WAS EMPTY AND THE PATIENT COULD NOT USE THEM AND HAD ONE WHERE THE TIP WENT TO ONE SIDE AND TWISTED. THE PATIENT TRIED TO INSERT THE TWISTED CATHETER AND IT CAUSED A LITTLE TRAUMA AND A BIT OF BLEEDING. NO ADDITIONAL INFORMATION PROVIDED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01097 | 2023-02-22 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE NURSE STATED ARCTIC SUN DEVICE STOPPED EARLIER ALERTING ABOUT SOMETHING WITH THE PROBE. THEY TURNED OFF THE DEVICE AND BACK ON. MACHINE SAID IT NEEDED WATER, SO THEY ADDED WATER AND RESTARTED. PATIENT WAS MOVING AND SHIVERING. MEDICATIONS WERE ADMINISTERED. FOLEY PROBE WAS IN USE. TARGET TEMPERATURE WAS 37C, PATIENT TEMPERATURE WAS 35C. TWO NURSES TALKING AT SAME TIME. MIS WALKED THROUGH ACCESSING EVENT LOG WHICH SHOWS ALARM 15 UNABLE TO OBTAIN A STABLE PATIENT TEMPERATURE X 4, ALERT 50 PATIENT TEMPERATURE 1 ERRATIC AND ALARM 5 WATER RESERVOIR EMPTY. MIS EXPLAINED THAT THEY RECEIVED ALARM 5 BECAUSE THEY POWERED OFF DEVICE BEFORE DRAINING THE PADS. WATER WAS NOT ACTUALLY NEEDED. MIS EXPLAINED WHEN ALARM 15 AND ALERT 50 OCCUR TOGETHER AND REPEATEDLY IT WAS LIKELY A PROBE OR TEMPERATURE CABLE ISSUE. MIS HAD NURSE FLEX CABLE, BUT FIRST NURSE DENIED ANY CHANGE. MIS ADVISED IF EITHER ALERT 50 OR ALARM 15 RETURN THEY NEED TO CONTACT BIOMED ABOUT REPLACING CABLE OR UTILIZE CABLE FROM ANOTHER DEVICE AVAILABLE ON THE FLOOR. MIS EXPLAINED IMPORTANCE OF ACCURATE PATIENT TEMPERATURE INPUT AS IT DRIVES THERAPY THAT WAS WATER TEMPERATURE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01108 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L18 | 0119L18 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT MENTIONED THAT THE PATIENT EXPERIENCED URINARY TRACT INFECTION, HEMATURIA LASTING GREATER THAN 24 HOURS, BLADDER FISTULA AND BLADDER PERFORATION PRIOR TO THE DEVICE PLACEMENT, HEMATURIA LASTING GREATER THAN 24 HOURS, BLADDER FISTULA, BLADDER TRAUMA , BLADDER PERFORATION) AND CALCIFICATION AFTER THE DEVICE PLACEMENT AND HEMATURIA LASTING GREATER THAN 24 HOURS COMPLICATIONS WERE DIRECTLY ATTRIBUTABLE TO THE DEVICE AND FISTULA WAS NEEDED IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 18. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01103 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L18 | 0119L18 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THEY WERE ABLE TO SUCCESSFULLY IRRIGATE WHILE USING THIS PRODUCT BECAUSE GOT CLAMPED OFF IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 18. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01101 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L20 | 0119L20 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF UNCLEAR INFORMATION , MISSING INFORMATION , LACK OF MAINTENANCE DOCUMENTATION OR GUIDELINES, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 20. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01102 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L20 | 0119L20 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THEY WERE ABLE TO SUCCESSFULLY PLACE THE FOLEY CATHER IN A PATIENT BECAUSE IT WAS TO BIG IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 20. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01106 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L20 | 0119L20 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF UNCLEAR INFORMATION, MISSING INFORMATION, LACK OF MAINTENANCE DOCUMENTATION OR GUIDELINES, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 20. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01099 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | COUDE INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0103L14 | 0103L14 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF EXPIRATION DATE ERROR , ILLEGIBLE INFORMATION , INACCURATE INFORMATION , UNCLEAR INFORMATION AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO TIEMANN MODEL, LUBRICATH®, 2-WAY, 30CC BALLOON, FRENCH SIZE 14. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00503 | 2023-02-22 | 2023-02-02 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D5 |  | IT WAS REPORTED BY THE CUSTOMER, "RECEIVED A CALL FROM OUR ONCOLOGY DEPT. REGARDING EXTRAVASATION OF A PICC LINE THAT WAS INSERTED ON (B)(6) 2023. THE PATIENT RECEIVED 600ML OF CHEMOTHERAPEUTIC AGENT AND NOTED MODERATED SWELLING ON RIGHT FOREARM AT LEAST 1-2CM BELOW PUNCTURE SITE. RN PULLED THE PICC AFTER DISCUSSION WITH ONCOLOGY MD. THE PICC WAS 1 CM OUT FROM THE PUNCTURE SITE AND THE AREA WHERE IT WAS LEAKING WAS AT 2.5CM MARK". | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGU1678 SHOWED NO OTHER SIMILAR PRODUCT OTHCOMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01111 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0103L14 | 0103L14 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF EXPIRATION DATE ERROR, ILLEGIBLE INFORMATION, INACCURATE INFORMATION, UNCLEAR INFORMATION AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO TIEMANN MODEL, LUBRICATH®, 2-WAY, 30CC BALLOON, FRENCH SIZE 14. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01112 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0103L14 | 0103L14 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF EXPIRATION DATE ERROR, ILLEGIBLE INFORMATION, INACCURATE INFORMATION, UNCLEAR INFORMATION AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO TIEMANN MODEL, LUBRICATH®, 2-WAY, 30CC BALLOON, FRENCH SIZE 14. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01100 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L16 | 0119L16 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF UNCLEAR INFORMATION, MISSING INFORMATION, LACK OF MAINTENANCE DOCUMENTATION OR GUIDELINES, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 16. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00504 | 2023-02-22 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D2 |  | IT WAS REPORTED BY THE CUSTOMER THAT "HAIR WAS LOCATED IN PICC TRAY OUTSIDE OF WRAP. TRAY WAS NOT USED." NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGU1681 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00517 | 2023-02-22 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD2 |  | IT WAS REPORTED, CRACK IN CLEAR T-PORT WHICH IS WHAT THE STYLET GOES THROUGH, PICC HAD TO BE EXCHANGED FOR A NEW ONE. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGT3425 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00506 | 2023-02-22 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REFURB, SHERLOCK 3CG TCS | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770131R |  | IT WAS REPORTED THE USER IS HAVING ACCURACY ISSUES WITH THE SENSOR; STATING THAT IT WILL SHOW THE PICC IS IN PLACE BUT IT'S OFF CONSIDERABLY. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01113 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0103L14 | 0103L14 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF EXPIRATION DATE ERROR, ILLEGIBLE INFORMATION , INACCURATE INFORMATION, UNCLEAR INFORMATION AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO TIEMANN MODEL, LUBRICATH®, 2-WAY, 30CC BALLOON, FRENCH SIZE 14. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01093 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® COUDE OLIVE TIP URETHRAL FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 802516 | 802516 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE DULL AND HARD TO USE, DID NOT WORK VERY WELL. AS PER ADDITIONAL INFORMATION RECEIVED VIA SAMPLE FORM ON (B)(6) 2023, IT WAS STATED THAT THE CUSTOMER HAS SENT EXAMPLE OF A GOOD ONE WITH A SLICK OUTER SURFACE AND A BAD ONE THAT WAS DULL AND DID NOT PASS SMOOTH. CUSTOMER HAVE RECEIVED WHOLE BOXES OF EACH KIND AND NOT SURE WHY IT VARIES FROM BOX TO BOX. THEY COULD SEE ON THE OUTER SURFACE THAT THE GOOD ONE WAS SHINY WITH A SLICK OUTER SURFACE. THE BAD ONE WAS DULL WITH NO OUTER COATING. HAVING TO USE THE DULL ONE CAUSES MUCH DISCOMFORT BECAUSE IT DID NOT PASS THROUGH THE SKIN. NO MEDICAL INTERVENTION WAS REPORTED (LOT# NGGW0481). AS PER SAMPLE RECEIVED ON (B)(6) 2023, IT WAS NOTED THAT THE INTERMITTENT CATHETERS WITH LOT# NGGS3185 WAS RETURNED FOR EVALUATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01109 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L18 | 0119L18 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT MENTIONED THAT THE PATIENT HAD URINARY TRACT INFECTION, HEMATURIA LASTING GREATER THAN 24 HOURS, BLADDER FISTULA AND BLADDER PERFORATION PRIOR TO THE DEVICE PLACEMENT, HEMATURIA LASTING GREATER THAN 24 HOURS, CALCIFICATION, BLADDER FISTULA, BLADDER TRAUMA AND BLADDER PERFORATION AFTER THE DEVICE PLACEMENT, AND BLADDER TRAUMA COMPLICATIONS WERE DIRECTLY ATTRIBUTABLE TO THE FOLEY CATHETER AND NO MEDICAL OR SURGICAL INTERVENTION WAS NEEDED IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 18 . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01090 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE FOLEY CATHETER BALLOON AND THE BALLOON DID NOT INFLATE DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01092 | 2023-02-22 | 2023-01-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT A FOLEY CATHETER WAS INDWELLED FOR THE PATIENT ON JANUARY 5. ON JANUARY 18, URINE LEAKAGE WAS FOUND, AND THE DOCTOR ORDERED TO STOP CATHETERIZATION. WHILE REMOVING, THE WATER IN THE BALLOON COULD NOT BE DRAWN OUT WITH REPEATED ATTEMPT. THEN PART OF THE CATHETER WAS CUT OFF, AND TRIED SEVERAL TIMES TO DRAW THE WATER FROM THE DRAINAGE HOLE, BUT IT ALSO FAILED. THE USER ASKED THE UROLOGIST FOR EMERGENCY CONSULTATION. THE EXTRACTION WITH THE SYRINGE FAILED AGAIN. THE GASTRIC TUBE AND THE GUIDE WIRE OF THE CENTRAL VENOUS CATHETER WERE USED TO TRY TO PUNCTURE THE AIR BAG, BUT ALL FAILED. FINALLY, THE URINARY CATHETER WAS PUNCTURED AND REMOVED UNDER THE GUIDANCE OF ULTRASOUND B. THE URINARY CATHETER WAS REMOVED. THE PATIENT COULD URINATE SPONTANEOUSLY WITHOUT HEMATURIA. AS PER DESCRIPTION OF EVENT CAUSE ANALYSIS STATED LONG RETENTION TIME, OPERATION REASONS, LUBRICANTS ,PRODUCT QUALITY PROBLEMS COULD NOT BE RULED OUT AND THAT THERE WAS NO SPECIFIC RETENTION TIME IN THE INSTRUCTION MANUAL. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 10FEB2023, STATED THAT THERE WAS NO DESCRIPTION ABOUT RETENTION TIME IN CHINESE IFU OF FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00502 | 2023-02-22 | 2023-01-26 | I, n, j, u, r, y | Manufacturer report | POWERPORT DUO M.R.I. IMPLANTABLE PORT, CHRONOFLEX DUAL-LUMEN, 9.5F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1829570 | 1829570 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT WAS ALLEGEDLY FOUND LEAKING. IT WAS FURTHER REPORTED THAT WHEN CHEMO WAS INJECTED INTO THE PORT, THE PATIENT ALLEGEDLY EXPERIENCED COMPLETE SKIN BURN. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 12/2023). | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00501 | 2023-02-22 | 2023-01-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER, SINGLE-LUMEN, 4.2F | CHRONIC CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 0600520 | 0600520 |  | IT WAS REPORTED THAT SOME TIME POST CATHETER PLACEMENT PROCEDURE, THE INFUSION SYSTEM THAT WAS CONNECTED WITH THE CHRONIC CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE CATHETER HAS TO BE CUT ALLEGEDLY AND REPAIRED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE BROVIAC S/L REPAIR KIT CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED TO THE DISTAL END OF THE CATHETER AND THE EDGES OF THE COMPLETE CIRCUMFERENTIAL BREAK ON THE DISTAL END OF THE CATHETER WERE NOTED TO BE UNEVEN. THE SURFACE WAS NOTED TO BE GLOSSY WITH STRIATIONS THROUGHOUT. THE PHOTO SHOWS ONE BROKEN CATHETER SEGMENT AND BROKEN HUB PLACED INSIDE A TRANSPARENT BAG. THEREFORE THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE TO DISCONNECT ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN AND CANNOT VERIFIED FROM THE PROVIDED PHOTO. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 09/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01062 | 2023-02-22 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE PACKAGING WAS STERILE. BUT THE PLASTIC IN THE PACKAGING OF FOLEY CATHETER WAS TOO WEAK AND BROKE. IN ORDER TO BE ABLE TO OPEN THE PACKAGE AND POUR STERILT, THE PLASTIC NEEDED TO BE STRONGER AND THE PERFORATIONS TO BE EASY TO TEAR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
16416367 | 2023-02-22 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | PERIPHERAL CATHETER INSERTION KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. |  |  |  | A NURSE WAS ATTEMPTING TO INSERT A MIDLINE CATHETER INTO A PATIENT'S ARM. SHE WAS HAVING TROUBLE ADVANCING THE CATHETER AND SOMETHING DIDN'T FEEL RIGHT. SHE AND ANOTHER NURSE WHO WAS STANDING NEARBY HEARD AN AUDIBLE UNUSUAL CLICK. SHE DECIDED TO REMOVE THE CATHETER AND MAKE ANOTHER FRESH ATTEMPT. WHEN SHE WAS WITHDRAWING THE CATHETER SHE FELT RESISTANCE TRYING TO REMOVE IT FROM THE PATIENT'S SKIN. A STRONG PULL REMOVED IT AND IT WAS NOTED THAT THE GUIDEWIRE HAD EXITED THE SIDE OF THE CATHETER ABOUT ONE INCH FROM THE TIP WAS COILED INTO A SMALL KNOT. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01073 | 2023-02-22 | 2023-02-01 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L18 | 0119L18 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT MENTIONED THAT THE PATIENT EXPERIENCED URINARY TRACT INFECTION, HEMATURIA LASTING GREATER THAN 24 HOURS, BLADDER FISTULA AND BLADDER PERFORATION PRIOR TO THE DEVICE PLACEMENT, HEMATURIA LASTING GREATER THAN 24 HOURS, BLADDER FISTULA, BLADDER TRAUMA , BLADDER PERFORATION AND CALCIFICATION AFTER THE DEVICE PLACEMENT AND HEMATURIA LASTING GREATER THAN 24 HOURS COMPLICATIONS WERE DIRECTLY ATTRIBUTABLE TO THE DEVICE AND FISTULA WAS NEEDED IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 18. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE, -STIFFNESS OF THE PRODUCT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." CORRECTIONS: B. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01069 | 2023-02-22 | 2023-02-01 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L18 | 0119L18 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT MENTIONED THAT THE PATIENT EXPERIENCED URINARY TRACT INFECTION, HEMATURIA LASTING GREATER THAN 24 HOURS, BLADDER FISTULA AND BLADDER PERFORATION PRIOR TO THE DEVICE PLACEMENT, HEMATURIA LASTING GREATER THAN 24 HOURS, BLADDER FISTULA, BLADDER TRAUMA, BLADDER PERFORATION AND CALCIFICATION AFTER THE DEVICE PLACEMENT AND BLADDER TRAUMA COMPLICATIONS WERE DIRECTLY ATTRIBUTABLE TO THE DEVICE AND NO MEDICAL OR SURGICAL INTERVENTION WAS NEEDED IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 18. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00101 | 2023-02-22 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE OF A CALCIFIED TARGET LEGION IN THE SUPERFICIAL FEMORAL ARTERY VIA AN IPSILATERAL ANTEGRADE APPROACH, THE TIP OF THE CATHETER ALLEGEDLY DETACHED AFTER USING ABOUT 25 SECONDS. THE PROCEDURE WAS COMPLETED USING ANOTHER CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN COMMON DEVICE NAME AND PMA/510K. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRY DATE: 10/2024. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01075 | 2023-02-22 | 2023-02-01 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THERE WAS SUCTION ISSUES WITH THE PUREWICK URINE COLLECTION SYSTEM SO THE PATIENT WAS WAKING UP WET. THE PATIENT ALSO HAD A URINARY TRACT INFECTIONS. THE PATIENT HAD BEEN USING THIS PRODUCT FOR MORE THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION WAS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-02
1018233-2023-01078 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L16 | 0119L16 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF UNCLEAR INFORMATION, MISSING INFORMATION, LACK OF MAINTENANCE DOCUMENTATION OR GUIDELINES, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 16. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01077 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L16 | 0119L16 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF UNCLEAR INFORMATION, MISSING INFORMATION, LACK OF MAINTENANCE DOCUMENTATION OR GUIDELINES, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 16. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01076 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0103L14 | 0103L14 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF EXPIRATION DATE ERROR, ILLEGIBLE INFORMATION, INACCURATE INFORMATION, UNCLEAR INFORMATION AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO TIEMANN MODEL, LUBRICATH®, 2-WAY, 30CC BALLOON, FRENCH SIZE 14. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01083 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TIEMANN MODEL UROLOGICAL CATHETER COUDE TIP | UROLOGICAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 010116 | 010116 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE NOT ALIGNED, THEREFORE CAUSING AN ISSUE WITH INSERTION. PER ADDITIONAL INFORMATION ON 01FEB2023, THE BUMP AT THE END OF THE CATHETER DID NOT PROPERLY LINE UP WITH THE CURVE IN THE COUDE TIP. IT WAS OFF CENTERED AND WAS CAUSING THE PATIENT TO HAVE A FALSE IDEA ON WHERE THE TIP OF THE CATHETER REALLY WAS. THIS CAUSED THE PATIENT A LOT OF PAIN. NO MEDICAL INTERVENTION WAS REPORTED. PER CUSTOMER VIA PHONE ON 21FEB2023, IT WAS REPORTED THAT CUSTOMER'S ISSUE WAS RESOLVED. PER SAMPLE FORM RECEIVED ON 01MAR2023, IT WAS STATED THAT THE CUSTOMER WAS TOLD BY DR.OFFICE ABOUT THE ALIGNMENT. ALSO STATED THAT WHILE INSERTING CATHETER, THEY FELT SHARP PAIN UPON ENTERING THE BLADDER AND AGAIN AFTER RETRIEVING IT, A SCRAPING SENSATION. | THE REPORTED EVENT IS UNCONFIRMED AS THE PROBLEM COULD NOT BE REPRODUCED. TWO UNOPENED AND ONE USED CATHETER WITH THE ORIGINAL PACKAGING WERE RETURNED. THE CATHETERS WERE LABELED 1-3 FOR IDENTIFICATION WITH 3 BEING THE USED CATHETER. THE CATHETERS WERE SLIGHTLY BENT FROM THE WAY THEY WERE STORED FOR EVALUATION, NO ADDITIONAL COMPLAINTS WILL BE CREATED. THE CATHETER WAS REMOVED FROM THE PACKAGING AND LAID FLAT, THE COUDE NOTCH APPEARED TO LINE UP WITH THE CURVE. NO ABNORMALITIES WERE NOTED ON THE RETURNED CATHETER AND IT MET SPECIFICATIONS "ACCEPT ANY CURVE UNLESS CATHETER IS COMPLETELY STRAIGHT." THE CATHETER WAS THEN TAPED DOWN IN ORDER TO REMAIN FLAT/STRAIGHT AND THE COUDE INDICATOR NOTCH WAS STILL ALIGNED WITH THE CURVE FOR ALL THREE CATHETERS. THE DEVICE WAS USED FOR TREATMENT PURPOSES. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED, A LABEL/PACKAGING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01088 | 2023-02-22 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE CUSTOMER OPENED ANOTHER BOX AND 2 OF THE CATHETERS WERE DAMAGED, 1 HAD A BURST PACKET, STILL APPEARED TO BE FILLED WITH FLUID, WATER SHOWING INSIDE THE CATHETER AND 1 CATHETER HAD SIGNIFICANT KINKS, ONE END WAS 3 INCHES OFF THE TABLE WHEN LAID FLAT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01072 | 2023-02-22 | 2023-02-01 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L18 | 0119L18 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT MENTIONED THAT THE PATIENT HAD URINARY TRACT INFECTION, HEMATURIA LASTING GREATER THAN 24 HOURS, BLADDER FISTULA AND BLADDER PERFORATION PRIOR TO THE DEVICE PLACEMENT, HEMATURIA LASTING GREATER THAN 24 HOURS, CALCIFICATION, BLADDER FISTULA, BLADDER TRAUMA AND BLADDER PERFORATION AFTER THE DEVICE PLACEMENT, AND BLADDER TRAUMA COMPLICATIONS WERE DIRECTLY ATTRIBUTABLE TO THE FOLEY CATHETER AND NO MEDICAL OR SURGICAL INTERVENTION WAS NEEDED IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 18. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01086 | 2023-02-22 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L20 | 0119L20 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE OF UNCLEAR INFORMATION, MISSING INFORMATION, LACK OF MAINTENANCE DOCUMENTATION OR GUIDELINES, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 20. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01130 | 2023-02-23 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER(SILICONE) | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175816 |  | IT WAS REPORTED THAT PATIENT STATED THAT THE FOLEY CATHETER CAME OUT UPON HEARING A POPPING SOUND. THE CATHETER WAS OUT, AND THE BALLOON WAS FULLY DEFLATED. THIS HAPPENED TWICE BACK-TO-BACK WITH EXACTLY THE SAME CIRCUMSTANCES. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "TOOLING DAMAGED (CORE PINS)". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. FOR UROLOGICAL USE ONLY. STERILE UNLESS PACKAGE IS OPENED OR DAMAGED, EXCEPT FOR ANY INDIVIDUALLY PACKAGED COMPONENTS WITHIN THE TRAY WHICH ARE NOT LABELED AS STERILE. THESE COMPONENTS ARE NOT TERMINALLY STERILIZED. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LAWS, REGULATIONS, AND POLICIES. KEEP AWAY FROM SUNLIGHT. DO NOT STORE AT FREEZING TEMPERATURES, KEEP AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN ORIGINAL BOX. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILISE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. CONTAINS OR PRESENCE OF PHTHALATES: DI (2-ETHYLHEXYL) PHTHALATE (DEHP) IS A PLASTICIZER USED IN SOME POLYVINYL CHLORIDE MEDICAL DEVICES. DEHP HAS BEEN SHOWN TO PRODUCE A RANGE OF ADVERSE EFFECTS IN EXPERIMENTAL ANIMALS, NOTABLY LIVER TOXICITY AND TESTICULAR ATROPHY. ALTHOUGH THE TOXIC AND CARCINOGENIC EFFECTS OF DEHP HAVE BEEN WELL ESTABLISHED IN EXPERIMENTAL ANIMALS, THE ABILITY OF THIS COMPOUND TO PRODUCE ADVERSE EFFECTS IN HUMANS IS CONTROVERSIAL. THERE IS NO EVIDENCE THAT NEONATES, INFANTS, PREGNANT AND BREAST FEEDING WOMEN EXPOSED TO DEHP EXPERIENCE ANY RELATED ADVERSE EFFECTS. HOWEVER, A LACK OF EVIDENCE OF CAUSATION BETWEEN DEHP-PVC AND ANY DISEASE OR ADVERSE EFFECT DOES NOT MEAN THAT THERE ARE NO RISKS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01138 | 2023-02-23 | 2020-06-05 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TEMPORARY PACING ELECTRODE CATHETER WAS KINKED. THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01136 | 2023-02-23 | 2020-05-07 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TEMPORARY PACING ELECTRODE CATHETER WAS KINKED. THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01139 | 2023-02-23 | 2020-09-29 | I, n, j, u, r, y | Manufacturer report | TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TEMPORARY PACING ELECTRODE CATHETER WAS KINKED. THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01141 | 2023-02-23 | 2021-06-17 | I, n, j, u, r, y | Manufacturer report | TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TEMPORARY PACING ELECTRODE CATHETER WAS KINKED. THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01104 | 2023-02-23 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L18 | 0119L18 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT MENTIONED THAT THE PATIENT EXPERIENCED URINARY TRACT INFECTION, HEMATURIA LASTING GREATER THAN 24 HOURS, BLADDER FISTULA AND BLADDER PERFORATION PRIOR TO THE DEVICE PLACEMENT, HEMATURIA LASTING GREATER THAN 24 HOURS, BLADDER FISTULA, BLADDER TRAUMA, BLADDER PERFORATION AND CALCIFICATION AFTER THE DEVICE PLACEMENT AND BLADDER TRAUMA COMPLICATIONS WERE DIRECTLY ATTRIBUTABLE TO THE DEVICE AND NO MEDICAL OR SURGICAL INTERVENTION WAS NEEDED IN RELATION TO LUBRICATH®, 3-WAY, 5CC BALLOON, FRENCH SIZE 18. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01146 | 2023-02-23 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0103L24 | 0103L24 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, PATIENT DID NOT AGREE THAT THE BENEFITS OF USING THE BARD/BECTON DICKINSON BIOCATH PRODUCT OUTWEIGH THE POTENTIAL RISKS BECAUSE IT MIGHT GET IN A KNOT IN RELATION TO TIEMANN MODEL, LUBRICATH, 2-WAY, 30CC BALLOON, FRENCH SIZE 24. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01147 | 2023-02-23 | 2022-08-15 | I, n, j, u, r, y | Manufacturer report | TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE FLOATING BALLOON CANNOT BE INFLATED. ALSO STATED THAT THE PATIENT HAS BEEN EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 09FEB2023, USER EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01148 | 2023-02-23 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT ABOUT 30 MINUTES AFTER THE FOLEY CATHETER WAS PLACED, IT CAME OUT, SO WHEN IT WAS CHECKED, THE BALLOON HAD BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01152 | 2023-02-23 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THESE INTERMITTENT CATHETERS WERE UNUSABLE BECAUSE THEY WERE NOT LUBRICATED ENOUGH. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01120 | 2023-02-23 | 2019-07-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TIP OF BALLOON LEAKED AIR SO THAT IT WAS NOT ABLE TO BE USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01140 | 2023-02-23 | 2021-05-08 | I, n, j, u, r, y | Manufacturer report | TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TEMPORARY PACING ELECTRODE CATHETER WAS KINKED. THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01121 | 2023-02-23 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0103L20 | 0103L20 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, PATIENT DID NOT AGREE THAT THE BENEFITS OF USING THE BARD/BECTON DICKINSON BIOCATH PRODUCT OUTWEIGH THE POTENTIAL RISKS BECAUSE IT MIGHT GET IN A KNOT IN RELATION TO TIEMANN MODEL, LUBRICATH, 2-WAY, 30CC BALLOON, FRENCH SIZE 20. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01114 | 2023-02-23 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE CUSTOMER OPENED ANOTHER BOX AND 2 OF THE CATHETERS WERE DAMAGED , 1 HAD A BURST PACKET, STILL APPEARED TO BE FILLED WITH FLUID , WATER SHOWING INSIDE THE CATHETER AND 1 CATHETER HAD SIGNIFICANT KINKS , ONE END WAS 3 INCHES OFF THE TABLE WHEN LAID FLAT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01123 | 2023-02-23 | 2023-01-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE STERILIZED WATER FILLED WITH 10 CC DECREASED TO ABOUT 2 CC AFTER ONE WEEK OF INDWELLING FOLEY CATHETER. ALSO STATED THAT THE URINE LEAKAGE INCREASED FROM THE SIDE OF THE FOLEY CATHETER. OSMOTIC PRESSURE AND SEMIPERMEABLE MEMBRANE ACTION WAS EXPLAINED. PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON 17FEB2023, STATED THAT OSMOTIC PRESSURE AND SEMIPERMEABLE MEMBRANE ACTION MEANS THAT THE WATER IN THE BALLOON MIGHT BE REDUCED DUE TO OSMOTIC PRESSURE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01119 | 2023-02-23 | 2019-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE FLOATING BALLOON ON THE TEMPORARY PACING ELECTRODE COULD NOT BE INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00523 | 2023-02-23 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS INSERTED, BLOOD RETURN WAS CONFIRMED WITHOUT APPLYING NEGATIVE PRESSURE. WHEN THE CATHETER WAS CHECKED, LEAKAGE WAS FOUND APPROXIMATELY 1 CM PROXIMAL TO THE TIP OF THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFY3463 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01126 | 2023-02-23 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS REMOVED SPONTANEOUSLY. A PRE-TEST HAS BEEN CARRIED OUT, AND THERE WAS NO PROBLEM AT THAT TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01118 | 2023-02-23 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0103L18 | 0103L18 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, PATIENT DID NOT AGREE THAT THE BENEFITS OF USING THE BARD/BECTON DICKINSON BIOCATH PRODUCT OUTWEIGH THE POTENTIAL RISKS BECAUSE IT MIGHT GET IN A KNOT IN RELATION TO TIEMANN MODEL, LUBRICATH, 2-WAY, 30CC BALLOON, FRENCH SIZE 18. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01117 | 2023-02-23 | 2023-01-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT A FOLEY CATHETER WAS INDWELLED FOR THE PATIENT ON (B)(6). ON (B)(6), URINE LEAKAGE WAS FOUND, AND THE DOCTOR ORDERED TO STOP CATHETERIZATION. WHILE REMOVING, THE WATER IN THE BALLOON COULD NOT BE DRAWN OUT WITH REPEATED ATTEMPT. THEN PART OF THE CATHETER WAS CUT OFF, AND TRIED SEVERAL TIMES TO DRAW THE WATER FROM THE DRAINAGE HOLE, BUT IT ALSO FAILED. THE USER ASKED THE UROLOGIST FOR EMERGENCY CONSULTATION. THE EXTRACTION WITH THE SYRINGE FAILED AGAIN. THE GASTRIC TUBE AND THE GUIDE WIRE OF THE CENTRAL VENOUS CATHETER WERE USED TO TRY TO PUNCTURE THE AIR BAG, BUT ALL FAILED. FINALLY, THE URINARY CATHETER WAS PUNCTURED AND REMOVED UNDER THE GUIDANCE OF ULTRASOUND B . THE URINARY CATHETER WAS REMOVED. THE PATIENT COULD URINATE SPONTANEOUSLY WITHOUT HEMATURIA. AS PER DESCRIPTION OF EVENT CAUSE ANALYSIS, STATED LONG RETENTION TIME, OPERATION REASONS, LUBRICANTS AND PRODUCT QUALITY PROBLEMS COULD NOT BE RULED OUT. ALSO STATED THAT THERE WAS NO SPECIFIC RETENTION TIME IN THE INSTRUCTION MANUAL. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 10FEB2023, IT WAS STATED THAT THERE WAS NO DESCRIPTION ABOUT RETENTION TIME IN CHINESE IFU OF FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01115 | 2023-02-23 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50000000 | UNK |  | IT WAS REPORTED THAT THE ARCTIC SUN DEVICE STOPPED EARLIER ALERTING ABOUT SOMETHING WITH THE PROBE. THEY TURNED OFF THE DEVICE AND BACK ON. MACHINE SAID IT NEEDED WATER, SO THEY ADD WATER AND RESTARTED. PATIENT WAS MOVING AND SHIVERING. MEDICATIONS WERE ADMINISTERED. FOLEY PROBE IN USE. THE TARGET TEMPERATURE WAS 37C AND PATIENT TEMPERATURE WAS 35C. TWO NURSES TALKING AT SAME TIME. MIS WALKED NURSE THROUGH ACCESSING EVENT LOG WHICH SHOWED ALARM 15 (UNABLE TO OBTAIN A STABLE PATIENT TEMPERATURE) X 4, ALERT 50 (PATIENT TEMPERATURE 1 ERRATIC) AND ALARM 5 (WATER RESERVOIR EMPTY). MIS EXPLAINED THAT THEY RECEIVED ALARM 5 (WATER RESERVOIR EMPTY) BECAUSE THEY POWERED OFF DEVICE BEFORE DRAINING THE PADS . WATER WAS NOT ACTUALLY NEEDED. MIS EXPLAINED WHEN ALARM 15 (UNABLE TO OBTAIN A STABLE PATIENT TEMPERATURE) AND ALERT 50 (PATIENT TEMPERATURE 1 ERRATIC) OCCUR TOGETHER AND REPEATEDLY IT WAS LIKELY A PROBE OR TEMPERATURE CABLE ISSUE . MIS HAD NURSE TO FLEX CABLE, BUT FIRST NURSE DENIED ANY CHANGE. MIS ADVISED THAT IF EITHER ALERT 50 (PATIENT TEMPERATURE 1 ERRATIC) OR ALARM 15 (UNABLE TO OBTAIN A STABLE PATIENT TEMPERATURE) RETURN THEY NEED TO CONTACT BIOMED ABOUT REPLACING CABLE OR UTILIZE CABLE FROM ANOTHER DEVICE AVAILABLE ON THE FLOOR. MIS EXPLAINED IMPORTANCE OF ACCURATE PATIENT TEMPERATURE INPUT AS IT DRIVES THERAPY WHICH IS THE WATER TEMPERATURE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00105 | 2023-02-23 | 2022-11-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQF7594 | CQF7594 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED WITHIN NOMINAL BURST PRESSURE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE CONQUEST 40 PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. A IMAGE WAS PROVIDED FOR EVALUATION. DURING VISUAL EVALUATION, FRAYED FIBERS WERE NOTED TO THE BALLOON. NO KINKS WERE NOTED TO THE CATHETER SHAFT AND NO ANOMALIES TO THE LUERS, BIFURCATE OR GLUE FILLETS. DURING FUNCTIONAL EVALUATION, AN IN-HOUSE PRESTO INFLATION DEVICE WAS USED TO INFLATE THE BALLOON AND WATER WAS NOTED LEAKING FROM THE BALLOON. THE BALLOON FIBERS WERE STRIPPED AND UNDER MICROSCOPIC EXAMINATION, A PINHOLE RUPTURE WAS NOTED TO THE BALLOON. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. ONE IMAGE WAS PROVIDED AND REVIEWED. A ZOOMED IN IMAGE AT AN UNKNOWN LOCATION SHOWED A WIRE AND CATHETER WITHIN A VESSEL. AN ANGIOPLASTY BALLOON APPEARED AT LEAST PARTIALLY INFLATED. NUMEROUS ARROWS FROM IMAGING SOFTWARE WERE POINTING TOWARDS AN AREA THAT MIGHT HAVE REPRESENTED CONTRAST EXTRAVASATION OUTSIDE THE BALLOON BUT INSIDE VESSEL, POSSIBLY REPRESENTING BALLOON RUPTURE. THE PROVIDED MEDIA WAS NOT OF SUFFICIENT QUALITY TO DEFINITELY CONFIRM A RUPTURE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED BALLOON RUPTURE AS A PINHOLE RUPTURE WAS NOTED TO THE BALLOON UNDER MICROSCOPIC EXAMINATION OF THE RETURNED DEVICE. THE INVESTIGATION IS ALSO CONFIRMED FOR THE IDENTIFIED FRAYED FIBERS AS FRAYED FIBERS WERE NOTED ON THE BALLOON DURING VISUAL EVALUATION OF THE RETURNED DEVICE. A DEFINITIVE ROOT CAUSE FOR THE REPORTED PINHOLE BALLOON RUPTURE AND IDENTIFIED FRAYED FIBERS COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. EXPIRY DATE: 05/2025. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00522 | 2023-02-23 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602680 | 0602680 |  | IT WAS REPORTED THAT DURING PREPARATION OF A PORT PLACEMENT PROCEDURE, THE HAIR WAS ALLEGEDLY FOUND INSIDE THE PACKAGE. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 06/2026). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00521 | 2023-02-23 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN HEMODIALYSIS/APHERESIS CATHETER, DUAL-LUMEN,13.5F | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 0600690 | 0600690 |  | IT WAS REPORTED THAT DURING A DRAINAGE CATHETER PLACEMENT PROCEDURE, THE CONNECTION BETWEEN THE CATHETER AND TUNNELER ALLEGEDLY LOOSENED. FURTHER, IT WAS UNABLE TO BE USED AS EACH TIME DOCTOR TRIED TO ¿PULL¿ THE CATHETER THROUGH, IT WOULD DETACH FROM THE TUNNELER UNDER PATIENT¿S SKIN. THE PROCEDURE WAS COMPLETED USING THE SAME DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRY DATE: 12/2023). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00518 | 2023-02-23 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT, WHEN THE PORT SYSTEM WAS REMOVED, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN. IT WAS FURTHER REPORTED THAT THE DISTAL CATHETER SEGMENT AND THE PORT BODY WERE REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00104 | 2023-02-23 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U41502H30RX | U41502H30RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED AT NOMINAL PRESSURE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 05/2025. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01153 | 2023-02-23 | 2023-02-02 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETER WAS CAUSING TOO MANY URINARY TRACT INFECTIONS. PATIENT HAD BEEN USING THE PUREWICK PRODUCT FOR MORE THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTIONS. PER CUSTOMER VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE CUSTOMER WAS HOSPITALIZED DUE TO URINARY TRACT INFECTION AND PRESCRIBED ANTIBIOTICS. THE DOCTOR DID NOT CONFIRM THAT THE CATHETERS CAUSED THE URINARY TRACT INFECTION, HOWEVER URGED THE PATIENT TO STOP USE IMMEDIATELY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | N | LAY USER/PATIENT |  | 2023-02
1018233-2023-01127 | 2023-02-23 | 2021-04-26 | I, n, j, u, r, y | Manufacturer report | TEMPERATURE PACING ELECTRODE CATHETER | TEMPERATURE PACING ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE FLOATING BALLOON CANNOT BE INFLATED. ALSO STATED THAT THE PATIENT HAS BEEN EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 09FEB2023, USER EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00529 | 2023-02-23 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "WHEN PLACING THE PROBE COVER ON THE PROBE, THERE IS A SMALL HOLE THAT GEL COMES THROUGH WHEN PLACING THE RUBBER BANDS AROUND THE PROBE. THIS HAS HAPPENED FOUR TIMES NOW. HAD TO REPLACE PROVE COVER AND STERILE GLOVES BEFORE MOVING FORWARD WITH PICC" THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGW0892 SHOWED THREE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGW0892) HAVE BEEN REPORTED FROM THE SAME FACILITY. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00527 | 2023-02-23 | 2023-02-06 | I, n, j, u, r, y | Manufacturer report | FULL KIT 20G X 10 CM WITH PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F220108PT |  | IT WAS REPORTED BY THE CUSTOMER, "WE HAD A 20/10 PG CATHETER TIP SHEAR AND DISLODGE INTO A VEIN TODAY. PATIENT WILL HAVE TO HAVE IT REMOVED BY VASCULAR SURGERY TEAM. WIRE ADVANCED, CATHETER DID NOT, ENTIRE UNIT REMOVED FROM ARM AND A PIECE OF CATHETER WAS MISSING. VERY LITTLE TO NO RESISTANCE PULLING PG OUT." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE ANALYSIS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED POWERGLIDE PRO CATHETER WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE PHOTOGRAPH WHICH DEPICTED A 20GA X 10CM POWERGLIDE PRO MIDLINE CATHETER ASSEMBLY. THE DEVICE WAS ASSEMBLED, WITH THE CATHETER OVERLAYING THE NEEDLE SHAFT. BLOOD RESIDUE WAS VISIBLE ON THE CATHETER SHAFT. THE CATHETER TERMINATED APPROXIMATELY 1CM FROM THE TIP OF THE NEEDLE. THE CATHETER WAS OBSERVED TO BE FRACTURED; HOWEVER, INSPECTION OF THE PHOTOGRAPH WAS INSUFFICIENT TO IDENTIFY THE CAUSE OF THE FRACTURE. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE SHEARING THE CATHETER AGAINST THE NEEDLE TIP DURING ATTEMPTED DEVICE PLACEMENT AND CONTACT WITH A SHARP INSTRUMENT. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00531 | 2023-02-23 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "WHEN PLACING THE PROBE COVER ON THE PROBE, THERE IS A SMALL HOLE THAT GEL COMES THROUGH WHEN PLACING THE RUBBER BANDS AROUND THE PROBE. THIS HAS HAPPENED FOUR TIMES NOW. HAD TO REPLACE PROVE COVER AND STERILE GLOVES BEFORE MOVING FORWARD WITH PICC" THIS REPORT ADDRESSES THE FOURTH DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGW0892 SHOWED THREE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGW0892) HAVE BEEN REPORTED FROM THE SAME FACILITY. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00530 | 2023-02-23 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "WHEN PLACING THE PROBE COVER ON THE PROBE, THERE IS A SMALL HOLE THAT GEL COMES THROUGH WHEN PLACING THE RUBBER BANDS AROUND THE PROBE. THIS HAS HAPPENED FOUR TIMES NOW. HAD TO REPLACE PROVE COVER AND STERILE GLOVES BEFORE MOVING FORWARD WITH PICC" THIS REPORT ADDRESSES THE THIRD DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGW0892 SHOWED THREE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGW0892) HAVE BEEN REPORTED FROM THE SAME FACILITY. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00528 | 2023-02-23 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "WHEN PLACING THE PROBE COVER ON THE PROBE, THERE IS A SMALL HOLE THAT GEL COMES THROUGH WHEN PLACING THE RUBBER BANDS AROUND THE PROBE. THIS HAS HAPPENED FOUR TIMES NOW. HAD TO REPLACE PROVE COVER AND STERILE GLOVES BEFORE MOVING FORWARD WITH PICC" THIS REPORT ADDRESSES THE FIRST DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGW0892 SHOWED THREE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGW0892) HAVE BEEN REPORTED FROM THE SAME FACILITY. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00520 | 2023-02-23 | 2022-05-09 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT ONE MONTH AND TWENTY-TWO DAYS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED SWELLING, REDNESS AND PAIN AROUND THE PORT SITE AND FELT FEVERISH. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING TREATMENT OR MEDICATION FOR THESE COMPLICATIONS. IT WAS FURTHER REPORTED THAT THE PORT ALLEGEDLY HAD A RUPTURE ON ITS OWN AND WAS DRAINING PUSS AND BLOOD. REPORTEDLY, THE PORT WAS REMOVED. FURTHERMORE, THE PATIENT EXPERIENCED INFECTION AND WAS TESTED POSITIVE FOR METHICILLIN-SUSCEPTIBLE STAPH AUREUS BACTERIAL INFECTION, AND ANTIBIOTICS WERE ADMINISTERED FOR THE INFECTION. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 03/2023). THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
MW5115164 | 2023-02-23 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARD POWERGLIDE PRO MIDLINE CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. | F220088PT | F220088PT |  | CHECKED DEVICE BEFORE BREAKING SKIN AND FOUND THAT THE GUIDEWIRE WAS MISSING FROM THE DEVICE. DEVICE UNUSABLE WITHOUT GUIDEWIRE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00519 | 2023-02-23 | 2023-01-27 | I, n, j, u, r, y | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808001 | 1808001 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE DEVICE ALLEGEDLY MIGRATED. REPORTEDLY, THE PATIENT EXPERIENCED RASH, ITCHING, BURNING, REDNESS AND SWELLING AROUND THE PORT ACCESS AREA. THE CURRENT PATIENT STATUS WAS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 07/2023. DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01202 | 2023-02-24 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L22 |  | IT WAS REPORTED THAT WATER IN THE FOLEY CATHETER BALLOON LEAKED AND SLIPPED OFF. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01203 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO ¿THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) DOES NOT PROVIDE SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS OR THEIR MEDICAL APPLICATION(S)" BECAUSE "ILLEGIBLE INFORMATION, MISSING INFORMATION, LACK OF MAINTENANCE DOCUMENTATION OR GUIDELINES, AND INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONALS" IN RELATION TO THE BIOCATH® FOLEY CATHETER, 3-WAY CATHETER, 30 ML BALLOON, NO FRENCH SIZE SPECIFIED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00486 | 2023-02-24 | 2023-01-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER ¿THE PATIENT CAME FOR TREATMENT, EPIRUBICIN AND CYCLOPHOSPHAMIDE. NO BACKFLOW, EASY TO FLUSH SO THE TREATMENT IS STARTED. AFTER THE DROP GOES IN, IT STARTS TO ITCH UNDER THE STING, NO REDNESS OR PAIN. NO BACKFLOW NOW, INCREASED SWELLING ABOUT 5 CM WHICH INCREASES TO 9 CM. PICC LINE DRAWN, NO OBVIOUS HOLE IN PICC LINE. THE PATIENT IS PERFECTLY FINE, HAS HAD SEVERAL FOLLOW-UPS OVER TIME. THE ITCH CAME FIRST AND AFTER THAT ON INSPECTION THE SWELLING WAS SEEN. NO VISIBLE LEAKAGE OF EPIRUBICIN, INTERPRETED AS THE SWELLING WAS THE IRRIGATION DROP WITH NACL. PLASTIC SURGEONS WERE CONSULTED, AND A DECISION WAS MADE NOT TO GIVE SAVENE. PATIENT WAS MONITORED AT THE RECEPTION FOR A FEW HOURS AND WANTED TO GO HOME BEFORE THE AGREED TIME TO STAY; STATED THEY ¿FELT GOOD¿. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGQ4610 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00543 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 5F SINGLE-LUMEN CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6195355 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN MAKING THE PICC DRESSING AND RINSING THE VALVE, THE SALINE DRIPPED FROM THE PICC LINE AS WELL AS A FEW DROPS OF BLOOD WHILE I WAS PUTTING A CAP ON. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFZ2091 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00546 | 2023-02-24 | 2022-12-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER PTFE PEEL-APART SHEATH 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT THERE WERE BURRS AT THE PROXIMAL END OF THE GUIDE WIRE IN THE SELDINGER. THIS MADE THE GUIDE WIRE UNUSABLE FOR PUNCTURE. THE SELDINGER KIT WAS REPLACED WITH A NEW ONE FOR PUNCTURE. THE CUSTOMER REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. 2/6/2023 - THE RETURNED GUIDEWIRE EXHIBITED A KINK. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GUIDEWIRE IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. TWO 0.018 IN. GUIDEWIRES IN PLASTIC HOOPS WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLES. THE DISTAL END OF ONE SAMPLE WAS OBSERVED TO BE BENT AND KINKED. A MICROSCOPIC OBSERVATION REVEALED SEVERAL OF THE MOST PROXIMAL COILS OF THE GUIDEWIRE WITH A DAMAGED DISTAL END WERE OBSERVED TO BE DISJOINTED AND MISALIGNED BUT CONNECTED TO THE WELD TIP. THIS GUIDEWIRE WAS ALSO OBSERVED TO BE BENT AND KINKED IN SEVERAL LOCATIONS NEAR ITS DISTAL END. THE MOST PROXIMAL COIL OF THE OTHER RETURNED GUIDEWIRE WAS ALSO FOUND TO BE DISJOINTED AND MISALIGNED BUT CONNECTED TO THE WELD TIP. TWO PHOTOGRAPHS WERE ALSO RETURNED FOR EVALUATION, AND WHILE ONE APPEARS TO BE OF ONE OF THE RETURNED SAMPLES, THE OTHER PHOTOGRAPH APPEARS TO BE OF A SAMPLE THAT WAS NOT RETURNED. THE COILED WIRE OF THE GUIDEWIRE IN THIS RETURNED PHOTOGRAPH WAS OBSERVED TO BE DETACHED FROM THE WELD TIP. WHILE THE EXACT CAUSE OF THE DAMAGE OBSERVED IN THE RETURNED GUIDEWIRES, AS WELL AS THE RETURNED PHOTOGRAPH, IS UNKNOWN, POSSIBLE CAUSES INCLUDE DAMAGE DURING MANUFACTURING, PACKAGING, HANDLING, OR USE. A BATCH HISTORY REVIEW (BHR) OF REFV0807 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00548 | 2023-02-24 | 2022-12-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER PTFE PEEL-APART SHEATH 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT THERE WERE BURRS AT THE PROXIMAL END OF THE GUIDE WIRE IN THE SELDINGER. THIS MADE THE GUIDE WIRE UNUSABLE FOR PUNCTURE. THE SELDINGER KIT WAS REPLACED WITH A NEW ONE FOR PUNCTURE. THE CUSTOMER REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. ON 2/6/2023 - THE RETURNED GUIDEWIRE EXHIBITED A KINK. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GUIDEWIRE IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. TWO 0.018 IN. GUIDEWIRES IN PLASTIC HOOPS WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLES. THE DISTAL END OF ONE SAMPLE WAS OBSERVED TO BE BENT AND KINKED. A MICROSCOPIC OBSERVATION REVEALED SEVERAL OF THE MOST PROXIMAL COILS OF THE GUIDEWIRE WITH A DAMAGED DISTAL END WERE OBSERVED TO BE DISJOINTED AND MISALIGNED BUT CONNECTED TO THE WELD TIP. THIS GUIDEWIRE WAS ALSO OBSERVED TO BE BENT AND KINKED IN SEVERAL LOCATIONS NEAR ITS DISTAL END. THE MOST PROXIMAL COIL OF THE OTHER RETURNED GUIDEWIRE WAS ALSO FOUND TO BE DISJOINTED AND MISALIGNED BUT CONNECTED TO THE WELD TIP. TWO PHOTOGRAPHS WERE ALSO RETURNED FOR EVALUATION, AND WHILE ONE APPEARS TO BE OF ONE OF THE RETURNED SAMPLES, THE OTHER PHOTOGRAPH APPEARS TO BE OF A SAMPLE THAT WAS NOT RETURNED. THE COILED WIRE OF THE GUIDEWIRE IN THIS RETURNED PHOTOGRAPH WAS OBSERVED TO BE DETACHED FROM THE WELD TIP. WHILE THE EXACT CAUSE OF THE DAMAGE OBSERVED IN THE RETURNED GUIDEWIRES, AS WELL AS THE RETURNED PHOTOGRAPH, IS UNKNOWN, POSSIBLE CAUSES INCLUDE DAMAGE DURING MANUFACTURING, PACKAGING, HANDLING, OR USE. A BATCH HISTORY REVIEW (BHR) OF REFV0807 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01200 | 2023-02-24 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS URETHRAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE CUSTOMER HAD LEAKING PROBLEMS AND HAD A TOUCHLESS INTERMITTENT CATHETER THAT WAS ATTACHED TO A SELF CONTAINED BAG AND STATED THIS DEVICE LEAKS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
3006260740-2023-00550 | 2023-02-24 | 2023-01-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F DL FT MAX BARRIER PLUS KIT W/GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1295108FD2 |  | IT WAS REPORTED BY THE CUSTOMER THAT "RN STATES SHE GOT CALLED TO ASSESS A PICC THAT WAS LEAKING. WHEN RN WENT TO ASSESS THE PATIENT THE PICC WAS AROUND THE 6CM MARK. STATES THE PREVIOUS NURSE CHARTED THAT WHEN SHE PLACED THE PICC IT WAS AT THE 0CM MARK. WHEN RN WENT TO FLUSH THE LINE, SHE NOTED LEAKING AROUND THE 6CM MARK. NO FURTHER INFORMATION PROVIDED ON HOW PLACEMENT WAS. ONLY INFORMATION THAT THE ACCOUNT WAS ABLE TO GIVE WAS WHEN THEY WENT TO ASSESS THE PICC. STAT LOCK USED WAS THE BD STAT LOCK USED IN THE PICC KITS. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGV2140 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00551 | 2023-02-24 | 2022-12-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER PTFE PEEL-APART SHEATH 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT THERE WERE BURRS AT THE PROXIMAL END OF THE GUIDE WIRE IN THE SELDINGER. THIS MADE THE GUIDE WIRE UNUSABLE FOR PUNCTURE. THE SELDINGER KIT WAS REPLACED WITH A NEW ONE FOR PUNCTURE. THE CUSTOMER REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. (B)(6) 2023 - THE RETURNED GUIDEWIRE EXHIBITED A KINK. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GUIDEWIRE IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. TWO 0.018 IN. GUIDEWIRES IN PLASTIC HOOPS WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLES. THE DISTAL END OF ONE SAMPLE WAS OBSERVED TO BE BENT AND KINKED. A MICROSCOPIC OBSERVATION REVEALED SEVERAL OF THE MOST PROXIMAL COILS OF THE GUIDEWIRE WITH A DAMAGED DISTAL END WERE OBSERVED TO BE DISJOINTED AND MISALIGNED BUT CONNECTED TO THE WELD TIP. THIS GUIDEWIRE WAS ALSO OBSERVED TO BE BENT AND KINKED IN SEVERAL LOCATIONS NEAR ITS DISTAL END. THE MOST PROXIMAL COIL OF THE OTHER RETURNED GUIDEWIRE WAS ALSO FOUND TO BE DISJOINTED AND MISALIGNED BUT CONNECTED TO THE WELD TIP. TWO PHOTOGRAPHS WERE ALSO RETURNED FOR EVALUATION, AND WHILE ONE APPEARS TO BE OF ONE OF THE RETURNED SAMPLES, THE OTHER PHOTOGRAPH APPEARS TO BE OF A SAMPLE THAT WAS NOT RETURNED. THE COILED WIRE OF THE GUIDEWIRE IN THIS RETURNED PHOTOGRAPH WAS OBSERVED TO BE DETACHED FROM THE WELD TIP. WHILE THE EXACT CAUSE OF THE DAMAGE OBSERVED IN THE RETURNED GUIDEWIRES, AS WELL AS THE RETURNED PHOTOGRAPH, IS UNKNOWN, POSSIBLE CAUSES INCLUDE DAMAGE DURING MANUFACTURING, PACKAGING, HANDLING, OR USE. A BATCH HISTORY REVIEW (BHR) OF REFV0807 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00554 | 2023-02-24 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 22 GAUGE POWERGLIDE PRO MIDLINE CATHETER 8CM BASIC KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6F122080 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE CATHETER WAS BLOCKED WITH THE GUIDEWIRE, WHILE IN THE PATIENT'S VEIN. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REFY1175 SHOWED THREE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. THE COMPLAINTS FOR THIS BATCH NUMBER HAVE BEEN REPORTED FROM THE SAME FACILITY IN BELGIUM. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00532 | 2023-02-24 | 2023-01-30 | I, n, j, u, r, y | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT SOME TIME POST PORT PLACEMENT, THE PORT WAS ALLEGEDLY BULGED OUT OF PATIENT CHEST. IT WAS FURTHER REPORTED THAT, THE PATIENT ALLEGEDLY EXPERIENCED SKIN EROSION AND SOME DISCOLORATION AROUND THE PORT INCISION SITE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00535 | 2023-02-24 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616070 | 1616070 |  | IT WAS REPORTED THAT PRIOR TO PORT DEVICE IMPLANT PROCEDURE, THE HUB OF THE PORT ALLEGEDLY BROKE OFF INSIDE THE CATHETER WHILE CONNECTING THEM TOGETHER. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 04/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00534 | 2023-02-24 | 2022-09-08 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT HAS BEEN REPORTED THAT THREE MONTHS AND NINETEEN DAYS POST PORT PLACEMENT, THE CATHETER TIP ALLEGEDLY DISLOCATED FROM VENA CAVA TO VENA JUGULARIS. IT WAS FURTHER REPORTED THAT THE PORT WAS REMOVED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 02/2023). | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00547 | 2023-02-24 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000821 |  | IT WAS REPORTED BY THE CUSTOMER THAT, 3FR SV REPORTED TO HAVE HOLE IN IT. PICC FOUND LEAKING WHERE THE PURPLE HUB ATTACHES TO THE EXTERNAL LUMEN. SECURACATH WAS USED AS THE CATHETER SECUREMENT. PATIENT NEEDED ANOTHER PICC LINE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGN3878 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01197 | 2023-02-24 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123520A | 123520A |  | IT WAS REPORTED THAT WATER IN THE FOLEY CATHETER BALLOON COULD NOT ESCAPE. THE BALLOON WAS CUT OFF. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON BURST. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE PATIENT USE ONLY. DO NOT REUSE AND RESTERILIZE. FOR UROLOGICAL USE ONLY. USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PROTOCOL, THE MAY BE CUT OFF. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT". THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00545 | 2023-02-24 | 2023-02-05 | I, n, j, u, r, y | Manufacturer report | POWERPICC HF CATHETER WITH SHERLOCK 3CG TIP TPS STYLET FULL TRAY (5F) TL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1385108QNS |  | IT WAS REPORTED THE PATIENT IS EXPERIENCING THE CATHETER KINKING UNDER THE INSERTION SITE SEVERAL DAYS POST-INSERTION. X-RAYS HAVE CONFIRMED THE KINK. ON 2/14/2023: ADDITIONAL INFORMATION: LINE HAD TO BE REMOVED, KINK WOULD NOT PULL STRAIGHT BUT KEPT RE-KINKING UNDER THE SKIN AT THE 9CM MARK, JUST BELOW THE AXILLA BETWEEN THE INSERTION SITE AND THE ARMPIT. LINE HAD TO BE REPLACED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01189 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0174L20 |  | IT WAS REPORTED THAT THE PHYSICIAN IN AN ONLINE SURVEY STATED THAT NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE INACCURATE INFORMATION, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO BARD 3-WAY 5CC BIOCATH HYDROGEL COATED LATEX FOLEY CATHETER, FRENCH SIZE 20. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01193 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE DOCTOR WAS NOT ABLE TO INFLATE THE BALLOON WHILE INSTALLING THE FOLEY CATHETER. THE DOCTOR REPORTED THAT IT LOOKED LIKE THERE WAS A BLOCKAGE IN IT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01173 | 2023-02-24 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 0165L14 |  | IT WAS REPORTED THAT URINE FLOW WAS CHECKED WHEN THE CATHETER WAS PLACED TO THE PATIENT, BUT NO URINE FLOWED. THE USER INSERTED AND CHECKED WITH ANOTHER CATHETER, THE URINE FLOW WAS CONFIRMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01166 | 2023-02-24 | 2019-03-05 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TEMPORARY PACING ELECTRODE CATHETER WAS BROKEN. IT WAS STATED THAT THE USER WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. AS PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON 10FEB2023, IT WAS STATED THAT THE DEVICE WAS USED ON THE PATIENT AND THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01167 | 2023-02-24 | 2019-08-06 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE FLOATING BALLOON CANNOT BE INFLATED. AS PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON 10FEB2023, IT WAS STATED THAT THE DEVICE WAS USED ON THE PATIENT AND THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01170 | 2023-02-24 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT THE NEXT DAY AFTER USE, THE CATHETER BALLOON SPONTANEOUSLY FELL OUT OF THE PATIENT. IT WAS STATED THAT A TEAR WAS OBSERVED NEAR THE TIP OF THE CATHETER FROM WHICH THE WATER LEAKED AND THE CATHETER FELL OFF. ACCORDING TO THE INQUIRY SHEET ATTACHED, THE BALLOON RUPTURE WAS OBSERVED. IT WAS STATED THAT NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER AND THE CATHETER WAS PULLED RARELY. PER IBC FOLLOW UP INFORMATION RECEIVED ON 16-FEB-2023,THE INITIAL EVENT DESCRIPTION WAS INCORRECT. CHECKING THE RETURNED SAMPLE, THERE WAS A PINHOLE OBSERVED ON THE CATHETER SHAFT. SINCE THE BALLOON RUPTURE WAS NOT OBSERVED TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01175 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED AND INFLATED WITH 10ML WITHOUT COMPLICATION. A COUPLE HOURS LATER, IT WAS NOTED THAT THE FOLEY SPONTANEOUSLY EXPELLED. THE FOLEY WAS TAKEN TO THE BATHROOM AND INFLATED WITH 10ML NORMAL SALINE (NS). LIQUID WAS DRIPPING FROM THE END OF THE CATHETER, UNSURE IF IT WAS URINE IN TUBING. WITHIN 10 MINUTES, THE FOLEY BALLOON WAS COMPLETELY DEFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01180 | 2023-02-24 | 2023-02-04 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION BUT DID NOT SAY IT WAS FROM THE PUREWICK FEMALE EXTERNAL CATHETER. PATIENT HAD BEEN USING THE PUREWICK PRODUCT FOR MORE THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-02
1018233-2023-01177 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT THE INSERTED INDWELLING FOLEY CATHETER FELL OUT 2 MINUTES AFTER INSERTION, BALLOON WAS INFLATED AFTER THE FOLEY WAS INSERTED. BALLOON WAS REINFLATED AFTER IT FELL OUT OF PATIENT AND FLUID WAS LEAKING FROM BALLOON. IT WAS ALSO STATED THAT THE BALLOON DID NOT INFLATE AT ALL BECAUSE OF THE LEAK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01176 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165PL08 | 0165PL08 |  | IT WAS REPORTED THAT 8FRENCH FOLEY CATHETERS USED IN A PEDIATRIC CODE, FAILED WHEN ATTEMPTING TO INFLATE THE BALLOON DURING THE INSERTION ATTEMPTS. WHEN ANOTHER CATHETER WAS REMOVED AND INSPECTED, IT WAS NOTED THAT THERE WERE MULTIPLE HOLES IN THE CATHETER ITSELF AND THE BALLOON WAS UNABLE TO INFLATE. ANOTHER CATHETER WAS INSPECTED THAT WAS SENT DOWN FROM THE PEDIATRIC FLOOR AND THE SAME ISSUE WAS NOTED WITH THEIR SUPPLY AS WELL. THE 10FRENCH FOLEY CATHETERS WERE NOT NOTED TO HAVE MALFUNCTIONED, BUT COULD NOT BE USED AS THEY WERE TOO SMALL FOR THE PATIENT. THIS ISSUE CAUSED THE INTUBATED PEDIATRIC PATIENT TO NOT HAVE A FOLEY CATHETER IN PLACE PRIOR TO BEING FLOWN OUT TO ANOTHER HOSPITAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01161 | 2023-02-24 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT 10 ML OF WATER WAS INJECTED INTO THE CATHETER BALLOON WHEN THE FOLEY CATHETER WAS PLACED FOR A HOME CARE PATIENT, BUT ONLY 8 ML OF WATER REMAINED WHEN THE CATHETER WAS REMOVED 27 DAYS LATER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
16435201 | 2023-02-24 | 2023-01-17 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | 1758SI16 | 1758SI16 |  | PATIENT'S FOLEY CATHETER WAS ORDERED TO BE REMOVED. WHEN ATTEMPTING TO REMOVE THE CATHETER, RESISTANCE WAS MET AND WAS UNABLE TO BE REMOVED. ONCE REMOVED, THE BALLOON OF THE CATHETER APPEARED TO BE "ROLLED" EVEN THOUGH THE BALLOON WAS EMPTIED PRIOR TO REMOVAL. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01184 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 8FRENCH FOLEY CATHETERS USED IN A PEDIATRIC CODE, FAILED WHEN ATTEMPTING TO INFLATE THE BALLOON DURING THE INSERTION ATTEMPTS. WHEN ANOTHER CATHETER WAS REMOVED AND INSPECTED, IT WAS NOTED THAT THERE WERE MULTIPLE HOLES IN THE CATHETER ITSELF AND THE BALLOON WAS UNABLE TO INFLATE. ANOTHER CATHETER WAS INSPECTED THAT WAS SENT DOWN FROM THE PEDIATRIC FLOOR AND THE SAME ISSUE WAS NOTED WITH THEIR SUPPLY AS WELL. THE 10FRENCH FOLEY CATHETERS WERE NOT NOTED TO HAVE MALFUNCTIONED, BUT COULD NOT BE USED AS THEY WERE TOO SMALL FOR THE PATIENT. THIS ISSUE CAUSED THE INTUBATED PEDIATRIC PATIENT TO NOT HAVE A FOLEY CATHETER IN PLACE PRIOR TO BEING FLOWN OUT TO ANOTHER HOSPITAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01187 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0174L24 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS OR THEIR MEDICAL APPLICATION(S)" BECAUSE OF "UNCLEAR INFORMATION AND INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONALS" IN RELATION TO THE BARD® 3-WAY 5CC BIO CATH® HYDROGEL COATED LATEX FOLEY CATHETER, FRENCH SIZE 24. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01185 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THERE WAS A HOLE JUST DISTAL TO THE THREE-WAY PORT OF THE FOLEY CATHETER, SO THE IRRIGANTS WERE SQUIRTING OUT. STATED THAT THE SAMPLE WAS NOT AVAILABLE SINCE IT WAS WITH THE PATIENT, AND IT WAS DISCARDED. PER FOLLOW UP VIA EMAIL RECEIVED ON (B)(6) 2023, IT WAS STATED THAT THEY HAD THREE INCIDENTS AT THE HOSPITAL WHERE 22FRENCH SILICON 3 WAYS CATHETERS WERE PLACED. ONCE THE CBI WAS STARTED IT WAS NOTICED THAT WATER WAS GUSHING OUT OF THE SIDES DUE TO A HOLE. IN TERMS OF IMPACT, THE HOLE IN THE CATHETER WOULD BREAK STERILITY. IT WAS ALSO STATED THAT IT WAS PAINFUL FOR PATIENTS TO HAVE CATHETERS REPLACED ESPECIALLY ONE AS LARGE AS A 22FRENCH CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01186 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO ¿THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) DOES NOT PROVIDE SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS OR THEIR MEDICAL APPLICATION(S)" BECAUSE "ILLEGIBLE INFORMATION , MISSING INFORMATION , LACK OF MAINTENANCE DOCUMENTATION OR GUIDELINES, AND INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONALS" IN RELATION TO THE BIOCATH® FOLEY CATHETER, 3-WAY CATHETER, 30 ML BALLOON, NO FRENCH SIZE SPECIFIED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01188 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO ¿THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) DOES NOT PROVIDE SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS OR THEIR MEDICAL APPLICATION(S)" BECAUSE "ILLEGIBLE INFORMATION (PR#7127838), MISSING INFORMATION, LACK OF MAINTENANCE DOCUMENTATION OR GUIDELINES, AND INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONALS" IN RELATION TO THE BIOCATH® FOLEY CATHETER, 3-WAY CATHETER, 30 ML BALLOON, NO FRENCH SIZE SPECIFIED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01190 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0174L24 |  | IT WAS REPORTED THAT THE PHYSICIAN IN AN ONLINE SURVEY STATED THAT NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE INACCURATE INFORMATION, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO BARD 3-WAY 5CC BIOCATH HYDROGEL COATED LATEX FOLEY CATHETER, FRENCH SIZE 24. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01183 | 2023-02-24 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT 8FRENCH FOLEY CATHETERS USED IN A PEDIATRIC CODE, FAILED WHEN ATTEMPTING TO INFLATE THE BALLOON DURING THE INSERTION ATTEMPTS. WHEN ANOTHER CATHETER WAS REMOVED AND INSPECTED, IT WAS NOTED THAT THERE WERE MULTIPLE HOLES IN THE CATHETER ITSELF AND THE BALLOON WAS UNABLE TO INFLATE. ANOTHER CATHETER WAS INSPECTED THAT WAS SENT DOWN FROM THE PEDIATRIC FLOOR AND THE SAME ISSUE WAS NOTED WITH THEIR SUPPLY AS WELL. THE 10FRENCH FOLEY CATHETERS WERE NOT NOTED TO HAVE MALFUNCTIONED, BUT COULD NOT BE USED AS THEY WERE TOO SMALL FOR THE PATIENT. THIS ISSUE CAUSED THE INTUBATED PEDIATRIC PATIENT TO NOT HAVE A FOLEY CATHETER IN PLACE PRIOR TO BEING FLOWN OUT TO ANOTHER HOSPITAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01217 | 2023-02-25 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0174L24 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS OR THEIR MEDICAL APPLICATION(S)" BECAUSE OF "UNCLEAR INFORMATION AND INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONALS" IN RELATION TO THE BARD® 3-WAY 5CC BIO CATH® HYDROGEL COATED LATEX FOLEY CATHETER, FRENCH SIZE 24. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01215 | 2023-02-25 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THESE MAGIC 3 GO MALE COUDE INTERMITTENT CATHETERS CAUSED PAIN AND NO URINE CAME OUT OF THE CATHETER. NO ADDITIONAL INFORMATION PROVIDED. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01207 | 2023-02-25 | 2021-04-29 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TEMPORARY PACING ELECTRODE CATHETER WAS KINKED. THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01213 | 2023-02-25 | 2019-01-30 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TEMPORARY PACING ELECTRODE CATHETER BALLOON WAS RUPTURED. AS PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023,USER EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS" AS MENTIONED IN SOURCE DOCUMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01210 | 2023-02-25 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT WATER DID NOT ENTER IN THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01206 | 2023-02-25 | 2022-02-22 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE BALLOON LEAKED AIR. ALSO STATED THAT THE PATIENT EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01205 | 2023-02-25 | 2022-01-14 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE BALLOON LEAKED AIR. ALSO STATED THAT THE PATIENT HAS BEEN EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023,USER EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01208 | 2023-02-25 | 2021-04-29 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TEMPORARY PACING ELECTRODE CATHETER WAS KINKED. THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01209 | 2023-02-25 | 2020-05-07 | I, n, j, u, r, y | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE TEMPORARY PACING ELECTRODE CATHETER WAS KINKED. THE PATIENT WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01216 | 2023-02-25 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L22 |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, "THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS OR THEIR MEDICAL APPLICATIONS." BECAUSE OF "INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONALS AND INADEQUATE OR INSUFFICIENT TRAINING" IN RELATION TO THE AQUAMATIC® BIOCATH® FOLEY CATHETER, STANDARD LENGTH, 5 ML BALLOON, FRENCH SIZE 22. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00555 | 2023-02-27 | 2022-09-05 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6CF INT W SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8716001 |  | IT WAS REPORTED THAT SIX MONTHS AND EIGHTEEN DAYS POST PORT PLACEMENT, DURING REGULAR PORT CHECK AS PART OF A FOLLOW-UP, THE CATHETER TIP WAS ALLEGEDLY FOUND TO BE SPONTANEOUSLY REPOSITIONED IN THE JUGULAR VEIN. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 05/2022). | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00558 | 2023-02-27 | 2022-11-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THE EXTENSION TUBING HAD A PROBLEM, UNABLE TO BE ENGAGED. SINCE THERE WAS NO SPARE PART AVAILABLE, THE CATHETER WAS SCRAPPED. NO OTHER INFORMATION WAS PROVIDED. (B)(6) 2023 - THE RETURNED SAMPLES EXHIBITED DAMAGE ON THE LOCKING ARMS OF THE PROXIMAL CONNECTOR. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE (IF AVAILABLE), PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A POOR CONNECTION IN A GROSHONG NEXT CONNECTOR IS CONFIRMED AND WAS DETERMINED TO BE MANUFACTURING RELATED. FOUR 4 FR TWO-PIECE GROSHONG NEXT CONNECTORS AND ONE STIFFENING STYLET WITH FLUSHING HUB ATTACHED WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED SOME USE RESIDUE ON THE RETURNED SAMPLES. THE CONNECTORS WERE RETURNED NOT ASSEMBLED, AND THE BLUE INTERNAL COMPRESSION SLEEVE OF EACH DISTAL CONNECTOR PIECE APPEARS TO HAVE BEEN ADVANCED. A MICROSCOPIC OBSERVATION REVEALED THE LOCKING ARMS OF THE PROXIMAL CONNECTOR PIECES WERE VERY SLIGHTLY DAMAGED. AN ATTEMPT WAS MADE TO ASSEMBLE AND THEN DISASSEMBLE THE CONNECTORS. AN AUDIBLE CLICK COULD BE HEARD WHEN EACH CONNECTOR WAS ASSEMBLED; HOWEVER, EACH CONNECTOR WAS FOUND TO BE ABLE TO BE PULLED APART BY HAND WITH SOME FORCE. THE DISTANCE BETWEEN THE LOCKING ARMS OF THE PROXIMAL CONNECTOR PIECES WAS MEASURED AND EACH WAS FOUND TO BE TOO WIDE AND OUT OF SPECIFICATION. THE MANUFACTURING FACILITY WAS NOTIFIED OF THIS EVENT, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. A BATCH HISTORY REVIEW (BHR) OF REFX2635 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01256 | 2023-02-27 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 33426 | 33426 |  | IT WAS REPORTED THAT THE PATIENT ORDERED TO HAVE A 26FR URETHRAL CATHETER EXCHANGED. THE CATHETER WAS EXCHANGED USING STERILE TECHNIQUE. URINE RETURN WAS OBSERVED. UPON INFLATING BALLOON STERILE WATER SPRAYED OUT OF THE PORT ABOVE WHERE THE SYRINGE WAS ATTACHED, AND THE REGISTERING NURSE OBSERVED THAT THAT PIECE WAS DRY ROTTED, SO THE CATHETER WAS PROMPTLY REMOVED FROM THE PATIENT. AFTER REMOVAL, THE PACKAGING WAS INSPECTED FOR AN EXPIRATION DATE BY MULTIPLE REGISTERED NURSES. THERE WAS NO EXPIRATION DATE ON THE CATHETER OR PACKAGING. EXTERNAL CATHETER PLACED ON THE PATIENT UNTIL ANOTHER CATHETER COULD BE OBTAINED. ALTHOUGH NO PATIENT HARM OCCURRED, THE PATIENT WOULD HAVE TO HAVE A SECOND INSERTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00562 | 2023-02-27 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED "DOCTOR ATTEMPTED TO INSERT A PICC LINE. THE FIXED GUIDEWIRE INSIDE THE ACTUAL CANNULA/LINE WAS ¿CRUMPLED¿ AND HAD NOT PUSHED THROUGH THE CANNULA SO THAT THE LINE COULD NOT BE INSERTED, AND PROCESS HAD TO BE ABORTED- THIS WOULD CAUSE A DELAY IN TREATMENT." (B)(6) 2023 - THE STYLET OF THE RETURNED SAMPLE EXHIBITED MULTIPLE KINKS. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE ANALYSIS, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A KINKED STYLET IS CONFIRMED AND WAS DETERMINED TO BE LIKELY RELATED TO REINSERTION OF THE STYLET BACK INTO THE SEPTUM. ONE CUT SECTION OF A 3CG STYLET WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED THE STYLET APPEARED CUT FROM THE PROXIMAL SEGMENT WHILE THE REST OF THE CATHETER WAS AGGRESSIVELY KINKED ALONG THE ENTIRETY OF THE STYLET. THE RETURNED STYLET WAS MEASURED TO BE APPROXIMATELY 76.5 CM LONG. A MICROSCOPIC OBSERVATION REVEALED BLOOD USE RESIDUES ALONG THE CATHETER, AND A CONFIRMED CUT ON THE PROXIMAL SIDE OF THE STYLET. THE POLYIMIDE TUBING AND MAGNETS APPEARED UNREMARKABLE. THE COMPLAINT OF A KINKED STYLET IS CONFIRMED AND LIKELY CAUSED BY THE REMOVAL AND REINSERTION OF THE STYLET FROM THE CATHETER VIA THE SEPTUM ON THE CATHETER HUB. THE STYLET IS NOT DESIGNED TO BE REMOVED AND REINSERTED INTO THE INJECTION SITE ATTACHED TO THE T-CONNECTOR. A BATCH HISTORY REVIEW (BHR) OF REGN0211 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01251 | 2023-02-27 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE MAGIC 3 GO MALE COUDE INTERMITTENT CATHETERS WERE TWISTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
3006260740-2023-00560 | 2023-02-27 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 18G X 10 CM WITH GUARDIVA AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F118108PT |  | IT WAS REPORTED BY THE CUSTOMER "GUIDEWIRE DID NOT DEPLOY AND WAS FOUND COILED IN THE BACK OF THE DEVICE HANDLE." 02/16/2023 - THE RETURNED SAMPLE EXHIBITED A DAMAGED CATHETER TIP. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE ANALYSIS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF GUIDEWIRE FAILURE TO ADVANCE WAS INCONCLUSIVE DUE TO THE CONDITION OF THE RETURNED SAMPLE. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 18GA X 10CM POWERGLIDE PRO MIDLINE CATHETER ASSEMBLY. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. THE CATHETER HAD BEEN ADVANCED AND THE SAFETY MECHANISM WAS ENGAGED. THE CATHETER REMAINED ATTACHED TO THE ADVANCER. THE CATHETER EXHIBITED CURVED SHAPE MEMORY ALONG THE PROXIMAL HALF OF THE CATHETER SHAFT. THE SAMPLE WAS RECEIVED PARTIALLY DISASSEMBLED. THE NEEDLE/SAFETY/CATHETER ASSEMBLY WAS COMPLETELY DETACHED FROM THE HOUSING. THE GUIDEWIRE ADVANCER WAS LOOSE WITHIN THE HOUSING AND WAS NOT ENGAGED WITH THE GUIDEWIRE. MICROSCOPIC INSPECTION OF THE CATHETER REVEALED DEFORMATION OF THE TIP. THE EDGES OF THE CATHETER TIP WERE BUCKLED AND FLARED OUTWARD. MICROSCOPIC INSPECTION THE GUIDEWIRE REVEALED IT TO BE INTACT. SLIGHT CURVED SHAPE MEMORY WAS OBSERVED WITHIN THE COIL/CORE REGION. MICROSCOPIC INSPECTION OF THE GUIDEWIRE ADVANCER REVEALED WEAR MARKS WITHIN THE COUPLER REGION. THE DISASSEMBLED CONDITION OF THE SAMPLE UPON RECEIPT PREVENTED DETERMINATION OF THE STATE OF THE GUIDEWIRE DURING ATTEMPTED DEVICE USE. CONSEQUENTLY THIS COMPLAINT IS INCONCLUSIVE AT THIS TIME. THE WEAR MARKS ON THE COUPLER REGION OF THE GUIDEWIRE ADVANCER INDICATED THAT THE WIRE WAS INITIALLY PROPERLY ASSEMBLED WITH THE ADVANCER. THE GUIDEWIRE AND CATHETER DEFORMATION WERE CONSISTENT WITH RESISTANCE ENCOUNTERED DURING ATTEMPTED DEVICE INSERTION, WHICH MAY HAVE CONTRIBUTED TO THE REPORTED EVENT AS GUIDEWIRE ADVANCEMENT AGAINST RESISTANCE CAN CAUSE IT TO BECOME DISENGAGED FROM THE ADVANCER. A BATCH HISTORY REVIEW (BHR) OF REGU2409 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00556 | 2023-02-27 | 2022-11-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THE EXTENSION TUBING HAD A PROBLEM, UNABLE TO BE ENGAGED. SINCE THERE WAS NO SPARE PART AVAILABLE, THE CATHETER WAS SCRAPPED. 02/17/2023 - THE RETURNED SAMPLES EXHIBITED DAMAGE ON THE LOCKING ARMS OF THE PROXIMAL CONNECTOR. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE (IF AVAILABLE), PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A POOR CONNECTION IN A GROSHONG NXT CONNECTOR IS CONFIRMED AND WAS DETERMINED TO BE MANUFACTURING RELATED. FOUR 4 FR TWO-PIECE GROSHONG NXT CONNECTORS AND ONE STIFFENING STYLET WITH FLUSHING HUB ATTACHED WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED SOME USE RESIDUE ON THE RETURNED SAMPLES. THE CONNECTORS WERE RETURNED NOT ASSEMBLED, AND THE BLUE INTERNAL COMPRESSION SLEEVE OF EACH DISTAL CONNECTOR PIECE APPEARS TO HAVE BEEN ADVANCED. A MICROSCOPIC OBSERVATION REVEALED THE LOCKING ARMS OF THE PROXIMAL CONNECTOR PIECES WERE VERY SLIGHTLY DAMAGED. AN ATTEMPT WAS MADE TO ASSEMBLE AND THEN DISASSEMBLE THE CONNECTORS. AN AUDIBLE CLICK COULD BE HEARD WHEN EACH CONNECTOR WAS ASSEMBLED; HOWEVER, EACH CONNECTOR WAS FOUND TO BE ABLE TO BE PULLED APART BY HAND WITH SOME FORCE. THE DISTANCE BETWEEN THE LOCKING ARMS OF THE PROXIMAL CONNECTOR PIECES WAS MEASURED AND EACH WAS FOUND TO BE TOO WIDE AND OUT OF SPECIFICATION. THE MANUFACTURING FACILITY WAS NOTIFIED OF THIS EVENT, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. A BATCH HISTORY REVIEW (BHR) OF REFX2635 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00559 | 2023-02-27 | 2022-11-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THE EXTENSION TUBING HAD A PROBLEM, UNABLE TO BE ENGAGED. SINCE THERE WAS NO SPARE PART AVAILABLE, THE CATHETER WAS SCRAPPED. NO OTHER INFORMATION WAS PROVIDED. 02/17/2023 - THE RETURNED SAMPLES EXHIBITED DAMAGE ON THE LOCKING ARMS OF THE PROXIMAL CONNECTOR. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE FOURTH DEVICE. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE (IF AVAILABLE), PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A POOR CONNECTION IN A GROSHONG NXT CONNECTOR IS CONFIRMED AND WAS DETERMINED TO BE MANUFACTURING RELATED. FOUR 4 FR TWO-PIECE GROSHONG NXT CONNECTORS AND ONE STIFFENING STYLET WITH FLUSHING HUB ATTACHED WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED SOME USE RESIDUE ON THE RETURNED SAMPLES. THE CONNECTORS WERE RETURNED NOT ASSEMBLED, AND THE BLUE INTERNAL COMPRESSION SLEEVE OF EACH DISTAL CONNECTOR PIECE APPEARS TO HAVE BEEN ADVANCED. A MICROSCOPIC OBSERVATION REVEALED THE LOCKING ARMS OF THE PROXIMAL CONNECTOR PIECES WERE VERY SLIGHTLY DAMAGED. AN ATTEMPT WAS MADE TO ASSEMBLE AND THEN DISASSEMBLE THE CONNECTORS. AN AUDIBLE CLICK COULD BE HEARD WHEN EACH CONNECTOR WAS ASSEMBLED; HOWEVER, EACH CONNECTOR WAS FOUND TO BE ABLE TO BE PULLED APART BY HAND WITH SOME FORCE. THE DISTANCE BETWEEN THE LOCKING ARMS OF THE PROXIMAL CONNECTOR PIECES WAS MEASURED AND EACH WAS FOUND TO BE TOO WIDE AND OUT OF SPECIFICATION. THE MANUFACTURING FACILITY WAS NOTIFIED OF THIS EVENT, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. A BATCH HISTORY REVIEW (BHR) OF REFX2635 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11 | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00557 | 2023-02-27 | 2022-11-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THE EXTENSION TUBING HAD A PROBLEM, UNABLE TO BE ENGAGED. SINCE THERE WAS NO SPARE PART AVAILABLE, THE CATHETER WAS SCRAPPED. NO OTHER INFORMATION WAS PROVIDED. 02/17/2023 - THE RETURNED SAMPLES EXHIBITED DAMAGE ON THE LOCKING ARMS OF THE PROXIMAL CONNECTOR. IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE (IF AVAILABLE), PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A POOR CONNECTION IN A GROSHONG NXT CONNECTOR IS CONFIRMED AND WAS DETERMINED TO BE MANUFACTURING RELATED. FOUR 4 FR TWO-PIECE GROSHONG NXT CONNECTORS AND ONE STIFFENING STYLET WITH FLUSHING HUB ATTACHED WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED SOME USE RESIDUE ON THE RETURNED SAMPLES. THE CONNECTORS WERE RETURNED NOT ASSEMBLED, AND THE BLUE INTERNAL COMPRESSION SLEEVE OF EACH DISTAL CONNECTOR PIECE APPEARS TO HAVE BEEN ADVANCED. A MICROSCOPIC OBSERVATION REVEALED THE LOCKING ARMS OF THE PROXIMAL CONNECTOR PIECES WERE VERY SLIGHTLY DAMAGED. AN ATTEMPT WAS MADE TO ASSEMBLE AND THEN DISASSEMBLE THE CONNECTORS. AN AUDIBLE CLICK COULD BE HEARD WHEN EACH CONNECTOR WAS ASSEMBLED; HOWEVER, EACH CONNECTOR WAS FOUND TO BE ABLE TO BE PULLED APART BY HAND WITH SOME FORCE. THE DISTANCE BETWEEN THE LOCKING ARMS OF THE PROXIMAL CONNECTOR PIECES WAS MEASURED AND EACH WAS FOUND TO BE TOO WIDE AND OUT OF SPECIFICATION. THE MANUFACTURING FACILITY WAS NOTIFIED OF THIS EVENT, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. A BATCH HISTORY REVIEW (BHR) OF REFX2635 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01246 | 2023-02-27 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT STERILIZED WATER DID NOT COME OUT DURING CUFF CHECK AND SO THE USE WAS DISCONTINUED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01244 | 2023-02-27 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE STERILIZED WATER DID NOT ENTER DUE TO DAMAGE IN THE INFLATION VALVE DURING PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01219 | 2023-02-27 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63612 |  | IT WAS REPORTED THAT SOME OF THE WATER SACHETS EITHER HAD NO WATER IN THEM OR VERY LITTLE AND IT WAS NOT ENOUGH TO COAT THE INTERMITTENT CATHETER EVENLY. STATED THAT THE PATIENT WOULD RATHER USE THE WELLSPECT LOFRICS AS THEIR MAIN CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS / AIR IN LINES. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTERMITTENT CATHETER HYDROPHILIC COATED SILICONE HOW TO PREPARE AND USE YOUR HYDROSILTM CATHETER FOR UROLOGICAL USE ONLY. INTENDED FOR USE BY PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION AND MEASUREMENT. THE DEVICES ARE PASSED TO THE URINARY BLADDER VIA THE URETHRA. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. 1. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. 2. PRIOR TO OPENING THE SEALED CATHETER POUCH, APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. 3. WET THE CATHETER BY HOLDING THE PACKAGE WITH THE PRINTED SIDE UP AND TIP THE PACKAGE END-TO-END THREE TO SIX TIMES TO WET THE CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. 4. PEEL OPEN THE PACK AT THE FUNNEL/HANDLE END JUST ENOUGH TO EXPOSE 5 CM (2¿) OF THE CATHETER, OR FOR PRODUCTS WITH AN INSERTION SLEEVE, EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 5. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. 6. WASH YOUR HANDS AGAIN. 7. USING THE CATHETER . WITH INSERTION SLEEVE (GRIPPER): A. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. B. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. C. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. WITHOUT INSERTION SLEEVE (GRIPPER): A. HOLD THE HANDLE/FUNNEL END AND REMOVE THE CATHETER FROM THE PACKAGING. B. GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA AND ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. 8. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. 9. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER." | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01218 | 2023-02-27 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0174L20 |  | IT WAS REPORTED THAT THE PHYSICIAN IN AN ONLINE SURVEY STATED THAT NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE INACCURATE INFORMATION, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO BARD 3-WAY 5CC BIOCATH HYDROGEL COATED LATEX FOLEY CATHETER, FRENCH SIZE 20. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01220 | 2023-02-27 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0174L20 |  | IT WAS REPORTED THAT THE PHYSICIAN IN AN ONLINE SURVEY STATED THAT NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE INACCURATE INFORMATION, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO BARD 3-WAY 5CC BIOCATH HYDROGEL COATED LATEX FOLEY CATHETER, FRENCH SIZE 24. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00109 | 2023-02-27 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 014 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | US141502H15 | US141502H15 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE BELOW THE KNEE VIA AN IPSILATERAL APPROACH, THE PTA BALLOON ALLEGEDLY RUPTURED AT 6 ATM. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 10/2024), G3. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01225 | 2023-02-27 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT STERILIZED WATER LEAKED DID NOT ENTER INTO THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01222 | 2023-02-27 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTERMITTENT ROCHESTER CATHETERS | UNKNOWN MAGIC 3 CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT THE URINE DRAIN BAGS WITH THE BUILT-IN STRAW WERE MORE DIFFICULT TO USE THAN NORMAL. IF THE STRAW BENDS WHILE ADVANCING IT, THEN GETTING IT TO ADVANCE FAR ENOUGH TO EMPTY THE URINE CAN GET TO BE VERY DIFFICULT. CUSTOMER WANTED TO KNOW IF THERE WAS A PARTICULAR WAY TO USE THE DRAIN BAG BEYOND THEIR CURRENT METHODS. NO HARM OR INJURY HAS OCCURRED FROM DIFFICULTY OF USE EXPERIENCE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OUTSIDE DIAMETER IS TOO LARGE ROUGH OR IRREGULAR SHAPE PROTRUSIONS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-02
2020394-2023-00108 | 2023-02-27 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE TIP OF THE CATHETER ALLEGEDLY SEPARATED. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETER PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETER PRODUCTS ARE IDENTIFED. MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED MATERIAL SEPARATION ISSUE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRATION DATE: 08/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01228 | 2023-02-27 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE BALLOON PART OF THE FOLEY CATHETER DURING THE PRETEST AND THE BALLOON WAS RUPTURED OR TEAR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01229 | 2023-02-27 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 33426 | 33426 |  | IT WAS REPORTED THAT THE PATIENT ORDERED TO HAVE A 26FR URETHRAL CATHETER EXCHANGED. THE CATHETER WAS EXCHANGED USING STERILE TECHNIQUE. URINE RETURN WAS OBSERVED. UPON INFLATING BALLOON STERILE WATER SPRAYED OUT OF THE PORT ABOVE WHERE THE SYRINGE WAS ATTACHED, AND THE REGISTERING NURSE OBSERVED THAT PIECE WAS DRY ROTTED, SO THE CATHETER WAS PROMPTLY REMOVED FROM THE PATIENT . AFTER REMOVAL, THE PACKAGING WAS INSPECTED FOR AN EXPIRATION DATE BY MULTIPLE REGISTERED NURSES. THERE WAS NO EXPIRATION DATE ON THE CATHETER OR PACKAGING. EXTERNAL CATHETER PLACED ON THE PATIENT UNTIL ANOTHER CATHETER COULD BE OBTAINED. ALTHOUGH NO PATIENT HARM OCCURRED, THE PATIENT WOULD HAVE TO HAVE A SECOND INSERTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01231 | 2023-02-27 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THERE WAS SOMETHING LIKE A SCRATCH ON THE BALLOON PART OF THE CATHETER, AND THE USE WAS DISCONTINUED. WHEN THE PRODUCT WAS CHECKED, IT WAS FOUND THAT THE BALLOON PART WAS DAMAGED. | THE REPORTED EVENT WAS CONFIRMED HOWEVER THE CAUSE WAS UNKNOWN. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. VISUAL INSPECTION NOTED IRREGULAR BURST WITHOUT MISSING PIECES. INTRODUCED WATER INTO THE INFLATION LUMEN AND DID NOT EXPERIENCE ANY OBSTRUCTIONS TO IMPEDE FLOW. MICROSCOPIC EXAMINATION FOUND NO CONDITIONS THAT COULD BE ASSOCIATED WITH THE REPORTED EVENT. THE EXACT CAUSE ON HOW AND WHEN PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR) AND "HIGH MECHANICAL STABILITY TIME (MST) OF INCOMING LATEX" OR "COAGULANT DIP 2- DIPPING SPEED TOO SLOW". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01238 | 2023-02-27 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 39305 | 39305 |  | IT WAS REPORTED THAT THE PATIENT FELT A LAYER OF SKIN WAS COMING OFF TOO ,WHEN THE MALE EXTERNAL CATHETER TOOK IT OFF FROM PENIS. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-02
1018233-2023-01239 | 2023-02-27 | 2023-02-06 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER LF HC | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WAS IN THE HOSPITAL WITH BAD URINARY TRACT INFECTION DOCTOR WAS STATING COULD BE FROM THE PRODUCT BUT PATIENT WAS USING THE PUREWICK FEMALE EXTERNAL CATHETERS FOR MORE THAN 12 HOURS. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. | THE REPORTED ISSUE WAS CONFIRMED AS USE RELATED ISSUE AS THE PATIENT USED THE WICK MORE THAN 12 HOURS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿USER UNAWARE OF TIME LIMITATION OR FORGETS THE PATIENT IS USING THE DEVICE.". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE SETUP: CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE DRYDOCTM VACUUM STATION, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE DRYDOCTM VACUUM STATION USER GUIDE FOR FURTHER INFORMATION. USING STANDARD SUCTION TUBING, CONNECT THE PUREWICKTM FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER. PERI-CARE AND PLACEMENT: PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. REMOVAL: TO REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER, FULLY SEPARATE THE LEGS, GLUTEUS, AND LABIA. TO AVOID POTENTIAL SKIN INJURY UPON REMOVAL, GENTLY PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER DIRECTLY OUTWARD. ENSURE SUCTION IS MAINTAINED WHILE REMOVING THE PUREWICKTM FEMALE EXTERNAL CATHETER. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. MAINTENANCE: REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. INDICATION FOR USE: THE PUREWICKTM FEMALE EXTERNAL CATHETER IS INTENDED FOR NON-INVASIVE URINE OUTPUT MANAGEMENT IN FEMALE PATIENTS. CONTRAINDICATIONS: PATIENTS WITH URINARY RETENTION. WARNINGS: DO NOT USE THE PUREWICKTM FEMALE EXTERNAL CATHETER WITH BEDPAN OR ANY MATERIAL THAT DOES NOT. ALLOW FOR SUFFICIENT AIRFLOW. TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER. HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE. EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON. DO NOT USE BARRIER CREAM ON THE PERINEUM WHEN USING THE PUREWICKTM FEMALE EXTERNAL CATHETER.BARRIER CREAM MAY IMPEDE SUCTION. PROCEED WITH CAUTION IN PATIENTS WHO HAVE UNDERGONE RECENT SURGERY OF THE EXTERNAL UROGENITAL TRACT. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW. PUREWICKTM FEMALE EXTERNAL CATHETER. MAINTAIN SUCTION UNTIL THE PUREWICKTM FEMALE EXTERNAL CATHETER IS FULLY REMOVED FROM THE PATIENT TO AVOID URINE BACKFLOW. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. PRIOR TO CONNECTING THE PUREWICKTM FEMALE EXTERNAL CATHETER TO HOSPITAL WALL SUCTION TUBING, VERIFY SUCTION FUNCTION BY COVERING THE OPEN END OF THE SUCTION TUBING WITH ONE HAND AND OBSERVING THE PRESSURE DIAL. IF THE PRESSURE DOES NOT INCREASE WHEN THE LINE IS COVERED, VERIFY THAT THE TUBING IS SECURED, CONNECTED, AND NOT KINKED. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-02
1018233-2023-01227 | 2023-02-27 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THERE WAS A HOLE IN THE BIFURCATED PART OF THE FUNNEL OF THE FOLEY CATHETER DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01290 | 2023-02-28 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PULLED OUT SPONTANEOUSLY DURING ITS INDWELLING. STATED THAT THERE WAS A CRACK IN THE BALLOON PART. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01294 | 2023-02-28 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165814 | 165814 |  | IT WAS REPORTED THAT FOUR FOLEY CATHETERS WERE BROKEN AND WERE RETURNED BY THE PATIENT. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 14DEC2022, ALL 4 PROBES WERE USED ON THE PATIENT AND THE BALLOONS WENT OFF EACH TIME. THE PATIENT WAS AN ALZHEIMER'S PATIENT AND THEREFORE WAS DEPENDENT ON A BLADDER PROBE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 03JAN2023, STATED THAT THE CUSTOMER MEANT THAT PRESSURE DISAPPEARED AS THE WATER WENT OUT BY THE "BALLOON WENT OFF" STATEMENT. PER SAMPLE RECEIVED ON 07FEB2023, STATED THAT THE CUSTOMER RETURNED TWO ADDITIONAL USED SAMPLES FOR THIS COMPLAINT. | THE REPORTED EVENT WAS CONFIRMED HOWEVER THE CAUSE WAS UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINHOLE IN BALLOON OR CUFF. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER. VISUAL INSPECTION NOTED THAT THE CATHETER APPEARED TO BE TIED INTO A KNOT ON THE SHAFT AND THERE WAS A BALLOON BURST MEASURING 4.085" WITH NO MISSING PIECES. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE WHICH STATES, "BALLOON MUST NOT BE TORN". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER. 2. ACCESSORIES CLOSED DRAINAGE BAG. (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER. WATER SOLUBLE LUBRICANT; ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION, TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS, GLOVES, AND STATLOCK FOLEY. [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE: (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE: (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD. WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS: 1) MALFUNCTION: - CATHETER KINKING, DAMAGE, RUPTURE, - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE, - OCCLUSION OF CATHETER INNER LUMENS, - ENCRUSTATION, - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE, - DEVICE DAMAGE DUE TO INAPPROPRIATE USE, - FAILURE TO MEASURE TEMPERATURE, AND - IMPROPER TEMPERATURE INDICATION. 2) ADVERSE EVENTS: - URINARY-TRACT INFECTION, - HEMORRHAGE, HEMATURIA, - ALLERGY REACTION TO THE DEVICE, - CALCULUS FORMATION, - EDEMA, - PAIN, - DISCOMFORT, - INJURY OF BLADDER OR URETHRAL, - URETHRITIS, URINARY INCONTINENCE, AND - RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00572 | 2023-02-28 | 2022-11-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT "STATLOCK BROKE." (B)(6) 2023 - THE RETURNED STATLOCK WAS FOUND TO BE BROKEN INTO MULTIPLE PIECES. IT IS UNCONFIRMED IF THE BREAKS HAPPENED UPON OPENING THE PACKAGE OR DURING USE ON A PATIENT. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE ANALYSIS, PATIENT SEVERITY, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BROKEN STATLOCK IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE IV ULTRA PEDIATRIC STATLOCK DEVICE AND ONE SEALED IV ULTRA PEDIATRIC STATLOCK KIT WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED THE RETAINER OF THE STATLOCK DEVICE RETURNED OUTSIDE OF ANY PACKAGING WAS BROKEN INTO MULTIPLE PIECES. A MICROSCOPIC OBSERVATION REVEALED THE BREAK SURFACES OF THE RETAINER WERE JAGGED AND UNEVEN WITH SOME MATERIAL WHITENING. IT IS UNKNOWN HOW THIS DAMAGE OCCURRED; HOWEVER, POSSIBLE CAUSES INCLUDE FORCE APPLIED DURING SHIPPING, STORAGE, HANDLING, OR USE. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00573 | 2023-02-28 | 2022-12-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 18 GAUGE POWERGLIDE PRO MIDLINE CATHETER 10CM BASIC KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6F118100 |  | IT WAS REPORTED BY THE CUSTOMER POWERGLIDE PRO MIDLINE MALFORMATION OF THE METAL GUIDE. ADDITIONAL INFO. RECEIVED (B)(6) 2023: "THERE WAS NO HARM TO THE PATIENT BECAUSE WHEN WE REALIZED THAT THE GUIDE WAS DAMAGED, THE CATHETER WAS NOT IMPLANTED." | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE ANALYSIS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GUIDEWIRE WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 18GA X 10CM POWERGLIDE PRO MIDLINE CATHETER ASSEMBLY. USAGE RESIDUES WERE ABUNDANT THROUGHOUT THE SAMPLE. THE DEVICE WAS FULLY ASSEMBLED WITH THE CATHETER OVERLAYING THE NEEDLE SHAFT. A COMPLETE BREAK WAS OBSERVED IN THE CORE WIRE AND THE COIL WIRE WAS SEVERELY ELONGATED. THE WELD TIP WAS INTACT AT THE DISTAL END OF THE ELONGATED WIRE. MICROSCOPIC INSPECTION OF THE BREAK IN THE CORE WIRE REVEALED A GRANULAR FRACTURE SURFACE. A REGION OF INCREASED LUSTER WAS OBSERVED LEADING INTO THE FRACTURE SITE. MICROSCOPIC INSPECTION OF THE NEEDLE REVEALED DEFORMATION ALONG THE PROXIMAL EDGE OF THE BEVEL. THE FRACTURE FEATURES AND NEEDLE BEVEL DEFORMATION WERE CONSISTENT WITH DAMAGE CAUSED BY CONTACT BETWEEN THE GUIDEWIRE AND THE NEEDLE SHAFT. SUCH DAMAGE CAN OCCUR IF THE WIRE IS INSERTED AT A STEEP ANGLE OR AGAINST RESISTANCE AND IF THE WIRE IS RETRACTED AGAINST THE NEEDLE. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. A BATCH HISTORY REVIEW (BHR) OF REGP4491 SHOWED 1 OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11 | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00563 | 2023-02-28 | 2023-01-31 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT ONE YEAR THREE MONTHS AND FIVE DAYS POST PORT PLACEMENT VIA AN INTERNAL JUGULAR VEIN FOR PANCREATIC CANCER INSIDE THE BREAST, THE PATIENT ALLEGEDLY EXPERIENCED SWELLING AROUND THE PORT BODY, WHEN ANTIEMETICS WAS ADMINISTERED PRIOR TO CHEMOTHERAPY. IT WAS FURTHER REPORTED THAT, THE PORT WAS REMOVED DUE TO SWELLING. REPORTEDLY, A CRACK WAS ALLEGEDLY FOUND AT APPROXIMATELY 8 CM DISTAL TO THE PORT BODY ON THE REMOVED CATHETER, AND LEAKAGE HAS OCCURRED FROM THE CRACK. THE PROCEDURE WAS COMPLETED REPLACING ANOTHER DEVICE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 08/2023). | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01289 | 2023-02-28 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT AFTER INDWELLING IN THE PATIENT, NO URINE FLOWED AND THE FOLEY CATHETER BALLOON WAS FOUND TO BE CRACKED, SO IT WAS REMOVED AND REPLACED WITH A NEW CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00565 | 2023-02-28 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716000J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA THE SUBCLAVIAN VEIN, BLOOD RETURN COULD NOT BE ALLEGEDLY CONFIRMED. IT WAS FURTHER REPORTED THAT THE HUB OF THE INTRODUCER NEEDLE WAS ALLEGEDLY CRACKED. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED IN SECTION HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS IS IDENTIFIED IN COMMON DEVICE NAME AND PMA/510 K. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 07/2023). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00566 | 2023-02-28 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT DURING PREPARATION OF A PORT PLACEMENT PROCEDURE, THE SHEATH WAS ALLEGEDLY BROKEN. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 04/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00575 | 2023-02-28 | 2023-02-06 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT WAS ADMITTED TO DEPARTMENT OF RESPIRATORY MEDICINE, PNEUMONIA, DUE TO POOR PERIPHERAL VEINS, GROSHONG PICC WAS INSERTED IN THE BASILIC VEIN OF THE RIGHT UPPER EXTREMITY ON FEBRUARY 6, INSERTED 35CM, EXPOSED 3CM, EVALUATED BEFORE CATHETERIZATION CATHETER, PREFLUSH AND TEST VALVE FUNCTION ARE NORMAL, NO ABNORMALITY IS FOUND IN THE SUPPORT GUIDE WIRE, THERE IS NO RESISTANCE WHEN THE CATHETER IS DELIVERED IN THE CATHETER, IT IS INSERTED SMOOTHLY, NO VIOLENT CATHETER DELIVERY, AFTER THE PUNCTURE IS SUCCESSFUL, THE SUPPORT GUIDE WIRE IS WITHDRAWN, AND THERE IS NO TRACTION ADHESION THE SUPPORT GUIDE WIRE WAS WITHDRAWN SMOOTHLY, BUT THE LENGTH OF THE SUPPORT GUIDE WIRE WAS NOT OBSERVED, AND THE WITHDRAWAL INSPECTION WAS PERFORMED. THE BLOOD RETURNED SMOOTHLY, AND THE NORMAL SALINE FLUSHED THE CATHETER SMOOTHLY, AND NO ABNORMALITY WAS FOUND. THE NURSE IN THE WARD FOUND THAT ABOUT PART OF THE GUIDE WIRE WAS EXPOSED IN THE PATIENT'S ABDOMEN, AND INFORMED THE DOCTOR. THE DOCTOR PERFORMED A GUIDE WIRE FOREIGN BODY REMOVAL OPERATION, AND TOOK OUT A SUPPORT GUIDE WIRE OF ABOUT 37CM. THE SURFACE OF THE GUIDE WIRE WAS NOT TWISTED AND DISCOUNTED. THE PATIENT'S VITAL SIGNS WERE STABLE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE (IF AVAILABLE), PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BROKEN STYLET IS CONFIRMED. ONE SEGMENT OF A STIFFENING STYLET WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED THE PROXIMAL END OF THE STYLET INCLUDING THE FLUSHING HUB WAS MISSING. A MICROSCOPIC OBSERVATION REVEALED A BREAK AT THE PROXIMAL END OF THE RETURNED STYLET SEGMENT. THIS BREAK WAS OBSERVED TO BE AT A SLIGHT ANGLE, BUT FLAT AND SMOOTH WITH DISTINCT EDGES. STRIATIONS WERE OBSERVED ON THE SURFACE OF THE BREAK, WHICH WAS ALSO OBSERVED TO BE LUSTROUS. THE OBSERVED CHARACTERISTICS OF THE BREAK IN THE RETURNED STYLET INDICATE THAT THE STYLET WAS CUT BY A SHARP BLADED INSTRUMENT SUCH AS SCISSORS. THE REPORTED EVENT COULD NOT BE RELATED TO A DEFICIENCY IN PRODUCT MANUFACTURE OR DEFICIENCY WHILE UNDER CORRECT PRODUCT USE. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01292 | 2023-02-28 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT EVEN THOUGH SUCTION PRESSURE WAS NOT APPLIED THE STERILIZED WATER DID NOT DRAIN AND URINE WAS FOUND MIXED WITH THE STERILE WATER THAT WAS DRAINED. IT WAS STATED THAT URINE LEAKAGE FROM THE INFLATION VALVE WAS CONFIRMED . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00571 | 2023-02-28 | 2023-02-08 | I, n, j, u, r, y | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY (4F) SL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4174118 |  | IT WAS REPORTED BY THE CUSTOMER, "THERE WAS A 'SERIOUS ADVERSE EVENT WITH A PATIENT WITH A 4FR MONOLUNEN PICC CATHETER'. THEY INDICATE THAT WHEN THE CATHETER IS PREPARED, IT IS IRRIGATED WITH A SYRINGE PRE-FILLED WITH NORMAL SALINE SOLUTION. AT THE TIME OF IRRIGATION, THE SOLUTION IS OBSERVED TO LEAK FROM THE CATHETER INSERTION SITE; WHEN THE INSERTION SITE IS REASSESSED, THE CATHETER MOVED OUT AND WAS FOUND FRACTURED / "BROKEN" AT A HEIGHT OF 3 CM. THE REST OF THE CATHETER REMAINS IN THE VEIN WITHOUT BEING ABLE TO BE REMOVED. THE PATIENT WAS SENT TO THE EMERGENCY DEPARTMENT WHERE A CHEST X-RAY IS TAKEN, REVEALING THE CATHETER IN THE PROPER POSITION, FRACTURED NEAR THE INSERTION SITE AND WITH MILD SIGNS OF DEEP VEIN THROMBOSIS AT ARM LEVEL; THE PATIENT WAS TAKEN TO ANGIOGRAPHY TO REMOVE THE CATHETER, WHICH WAS APPROACHED VIA THE FEMORAL ROUTE. THE NURSE REPORTS THAT AT THE TIME OF REMOVAL, IT FRACTURES AGAIN, LEAVING IT IN TWO PARTS: 1 THAT IS REMOVED IMMEDIATELY AND THE SECOND PART TRAVELS TO THE PULMONARY VEIN. FROM WHICH IT IS SUCCESSFULLY EXTRACTED. THE PATIENT SUFFERED NO FURTHER COMPLICATIONS. ADDITIONAL INFORMATION RECEIVED ON 13.FEB.2023: PATIENT ARRIVED IN THE OUTPATIENT CHEMOTHERAPY SERVICE ON (B)(6) 2023 FOR "HEALING PROCEDURE". WHEN THE CLINICIAN CHANGED THE NEEDLE-FREE CONNECTOR AND FLUSHED THE CATHETER, THE PATIENT COMPLAINED OF PAIN AND BURNING FEELING. THE CLINICIAN NOTED SALINE SOLUTION LEAKING FROM THE INSERTION SITE, AND PROCEEDS TO LIFT THE ADHESIVE AND CHECK IRRIGATION AGAIN WITH A PREFILLED SYRINGE. SHE IDENTIFIES A SPONTANEOUS RUPTURE OF THE CATHETER WITH 3CM CATHETER EXTERNAL AND 32CM INTERNAL. IN ADDITION TO THE X-RAY TAKEN IN THE EMERGENCY DEPARTMENT, A DOPPLER WAS ALSO DONE. THE DOPPLER ULTRASOUND MAKES AN IMPRESSION DIAGNOSE FINDINGS COMPATIBLE WITH DEEP VEIN THROMBOSIS COMPROMISING THE ASSOCIATED RIGHT AXILLARY VEIN. THE PATH OF THE REMNANT OF THE PERIPHERALLY INSERTED CENTRAL VENOUS CATHETER EXTENDS FROM THE BRACHIAL VEIN TO THE EVALUABLE SEGMENT OF THE RIGHT SUBCLAVIAN VEIN, WITH ADEQUATE PERMEABILITY OF THESE TWO VENOUS STRUCTURES. WITHOUT CLINICAL SIGNS OF THROMBOSIS, THE VASCULAR SURGEON WAS CONSULTED, THEN ON (B)(6) 2023 HE WAS TAKEN TO ANGIOGRAPHY TO REMOVE THIS FOREIGN BODY ON (B)(6) 2023. PER MEDICAL RECORDS RECEIVED, A SMALL FRAGMENT WAS LEFT IN THE UPPER BRANCH OF THE RIGHT PULMONARY VEIN AND ANOTHER FRAGMENT IN THE RIGHT BASILIC VEIN, "INSIDE THE THROMBOSED VEIN". THE PATIENT WAS DISCHARGED FROM THE FACILITY ON FEBRUARY 9 WITH AN ANTICOAGULANT TREATMENT PRESCRIPTION FOR 3 MONTHS WITH APIXABAN 10 MG AND FOLLOW UP WITH VASCULAR SURGERY IN TWO WEEKS. AT THE MOMENT HE IS WELL AND AT HOME. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A COMPLETE BREAK/SPLIT IN THE CATHETER SHAFT WAS CONFIRMED. TWO POWERPICC CATHETER SEGMENTS WERE RETURNED FOR EVALUATION. ONE OF THE CATHETER SEGMENTS MEASURED TO BE 10 CM IN LENGTH AND THE SECOND SEGMENT MEASURED TO BE 1.8 CM IN LENGTH. THE PROXIMAL END OF THE CATHETER CONTAINING THE LUER HUB WAS NOT RETURNED. BOTH CATHETER SEGMENTS CONTAINED JAGGED BREAK SURFACES AT BOTH ENDS AND CONTAINED DRIED USAGE RESIDUES. A PHOTO SAMPLE AND VIDEO DEPICTING A 4 FR SL POWERPICC CATHETER WERE ALSO PROVIDED. THE PHOTOGRAPH SHOWED THE PROXIMAL SECTION OF THE CATHETER SECURED TO THE PATIENT¿S ARM USING A STATLOCK STABILIZATION DEVICE AND THE VIDEO WAS OF THE EXTRACTION OF A SMALL PART OF THE RETAINED CATHETER SHAFT. THE EXTRACTED CATHETER FRAGMENT CONTAINED A WAVY SHAPE MEMORY. THE RETURNED SAMPLE CORRESPONDED WITH THE PROVIDED MEDIA SAMPLES. MICROSCOPIC INSPECTION OF THE BREAK SURFACES REVEALED THEM TO BE ROUGH AND UNEVEN. THE 1.8 CM CATHETER SEGMENT WAS OBSERVED TO CONTAIN A PARTIAL SPLIT IN THE MID SECTION OF THE SEGMENT. THE SPLIT EDGES CONTAINED MATERIAL STRESS WHITENING AND A GRANULAR BREAK SURFACE. THE CATHETER WAS CUT NEAR THE DISTAL END OF THE CATHETER SEGMENT IN ORDER TO MEASURE ITS DIMENSIONAL PROPERTIES. THE CATHETER WALL THICKNESS WAS FOUND TO BE WITHIN MANUFACTURING SPECIFICATION. BASED ON THE CHARACTERISTICS OF THE RETURNED SAMPLE AND INFORMATION PROVIDED, LIKELY CONTRIBUTING FACTORS INCLUDE REPEATED KINKING OF THE CATHETER LEADING TO MATERIAL FATIGUE AND DAMAGE DURING HANDLING OR USAGE OF THE DEVICE. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. SINCE COMPLETE BREAKS AND PARTIAL SPLITS WERE FOUND TO BE PRESENT IN THE RETURNED SAMPLES, THE COMPLAINT IS CONFIRMED. THE LOCATION OF THE DAMAGE SUGGESTED THAT CATHETER INSERTION AND SECUREMENT TECHNIQUES MAY HAVE BEEN CONTRIBUTING FACTORS. EVALUATION FINDINGS ARE IN SECTION H.11. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01285 | 2023-02-28 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT EVEN THOUGH SUCTION PRESSURE WAS NOT APPLIED THE STERILIZED WATER DID NOT DRAIN AND URINE WAS FOUND MIXED WITH THE STERILE WATER THAT WAS DRAINED. IT WAS STATED THAT URINE LEAKAGE FROM THE INFLATION VALVE WAS CONFIRMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01291 | 2023-02-28 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT EVEN THOUGH SUCTION PRESSURE WAS NOT APPLIED THE STERILIZED WATER DID NOT DRAIN AND URINE WAS FOUND MIXED WITH THE STERILE WATER THAT WAS DRAINED. IT WAS STATED THAT URINE LEAKAGE FROM THE INFLATION VALVE WAS CONFIRMED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00564 | 2023-02-28 | 2023-02-01 | I, n, j, u, r, y | Manufacturer report | VACCESS CT POWER-INJECTABLE IMPLANTABLE PORT, CHRONOFLEX, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 7480000 | 7480000 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT TO THE LEFT SIDE OF THE NECK, THE PATIENT ALLEGEDLY EXPERIENCED STIFFNESS AND SORENESS TO THE LEFT SIDE OF THE NECK WITH NO REDNESS, SWELLING OR PAIN AROUND THE AREA. IT WAS FURTHER REPORTED THAT DURING ULTRASOUND EXAMINATION, THE PATIENT WAS DIAGNOSED WITH BLOOD CLOT ON THE LEFT SIDE OF NECK AND PORT WAS FOUND TO BE THE SOURCE OF THE BLOOD CLOT. REPORTEDLY, THE PATIENT WAS TREATED WITH ANTICOAGULANT MEDICATION FOR BLOOD THINNING. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 08/2024). | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00567 | 2023-02-28 | 2022-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN PORT SINGLE LUMEN | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0603880C |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA RIGHT INTERNAL JUGULAR VEIN, THE SHEATH ALLEGEDLY DEFORMED WHILE DELIVERING OVER THE GUIDE WIRE. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED IN SECTION HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 6.6F PRODUCTS ARE IDENTIFIED IN COMMON DEVICE NAME AND PMA/510K. MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE BARDPORT MRI L/P IMPLANTABLE PORT, ONE CATH-LOCK LOADED TO A CATHETER, TWO INTRODUCER NEEDLES, ONE RIGHT-ANGLE NON-CORING NEEDLE, ONE CATH-LOCK, ONE SYRINGE, ONE 7.0FR PEEL-APART SHEATH AND VESSEL DILATOR, AND ONE TUNNELER WERE RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, THREE ELECTRONIC PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MATERIAL INTEGRITY ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. HOWEVER THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED DEFORMATION ISSUE AS THE DISTAL TIP OF THE VESSEL DILATOR WAS NOTED TO BE DEFORMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRY DATE: 06/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01265 | 2023-02-28 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE VALVE OF THE FOLEY CATHETER WAS FAULTY. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 16FEB2023, STATED THAT THE DISTRICT NURSE AND USER COULDN¿T DEFLATE THE BALLOON WITH A 10ML LEUR LOCK SYRINGE. THEY INSTILLED A FURTHER 8MLS OF STERILE WATER INTO THE BALLOON AND THEN TRIED TO WITHDRAW FLUID WITHOUT ANY LUCK. THE USER INSERTED A NEEDLE AND SYRINGE ABOVE THE BALLOON VALVE TO TRY AND DEFLATE BUT COULD ONLY GET 2 MLS OUT. THE USER GAVE UP AND LEFT AFTER ARRANGING AN URGENT CYSTOSCOPY WITH UROLOGY. APPROXIMATELY 30 MINUTES LATER THE PATIENT NOTICED THAT THE CATHETER HAD FALLEN OUT. THE USER CHECKED THE BALLOON ON THE CATHETER AFTER IT FELL OUT AND FOUND THAT THEY COULD INFLATE IT WITH AIR BUT COULD NOT DEFLATE IT. HOWEVER, AS SOON AS THE USER TOOK THE SYRINGE OFF THE BALLOON DEFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
2020394-2023-00114 | 2023-02-28 | 2022-12-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 PTA DILATATION CATHETER | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQF7584 | CQF7584 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE CATHETER ALLEGEDLY LEAKED DURING INFLATION OF THE BALLOON. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE CONQUEST 40 PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, NO ANOMALIES WERE NOTED TO THE LUERS, BIFURCATE OR GLUE FILLETS. DURING FUNCTIONAL EVALUATION, AN IN-HOUSE PRESTO INFLATION DEVICE WAS USED TO INFLATE THE BALLOON AND WATER WAS NOTED LEAKING FROM THE CATHETER SHAFT. UNDER MICROSCOPIC EXAMINATION, A TEAR WAS NOTED TO THE CATHETER SHAFT. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED LEAK AND IDENTIFIED CATHETER TEAR AS A TEAR WAS SEEN ON THE CATHETER SHAFT UNDER MICROSCOPIC EXAMINATION FROM WHICH WATER WAS NOTED LEAKING. A DEFINITIVE ROOT CAUSE FOR THE REPORTED LEAK AND IDENTIFIED CATHETER TEAR COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. EXPIRY DATE: 09/2023. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01266 | 2023-02-28 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® FOLEY TRAY WITH BARDEX® ALL SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 897216 | 897216 |  | IT WAS REPORTED THAT THE FOLEY CATHETER KITS CAME WITH A 10ML SALINE SYRINGE TO INFLATE THE BALLOON WITH INSTRUCTIONS TO INFLATE ALL 10MLS. THEIR STUDENTS FOLLOWED THESE INSTRUCTIONS, HOWEVER, THE BALLOONS ALL RUPTURED INSIDE OF THEIR MANNEQUINS. WHEN THEY WENT TO ASSIST AND NOTED THAT THE FOLEY BALLOONS INCLUDED WITH THE KITS WERE ONLY 5ML CAPACITY. THEREFORE, THERE WERE MISLEADING OR INACCURATE INSTRUCTIONS AND INCOMPATIBLE SUPPLIES PROVIDED WITH THIS KIT AND THEY HAD TO DISASSEMBLE THE MANNEQUINS AND REMOVE THE SALINE NOW INSIDE OF THEM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00569 | 2023-02-28 | 2020-08-19 | I, n, j, u, r, y | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | BARD PERIPHERAL VASCULAR, INC. | 5397190 | 5397190 |  | IT WAS REPORTED THROUGH RESULTS OF A CLINICAL TRIAL THAT TWO DAYS POST INDEX PROCEDURE, THE SUBJECT EXPERIENCED AN ADVERSE EVENT OF A CLOTTED CATHETER. IT WAS FURTHER REPORTED THAT THE ADVERSE EVENT WAS DETERMINED TO BE ASSOCIATED WITH DEVICE FAILURE AND POSSIBLY RELATED TO STUDY DEVICE. REPORTEDLY, THE CATHETER WAS REMOVED DUE TO THROMBOSIS. THE PROCEDURE WAS COMPLETED BY EXCHANGING ANOTHER CATHETER. HOWEVER, THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 11/2021). | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00568 | 2023-02-28 | 2023-01-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT ONE MONTH AND FIVE DAYS POST PORT PLACEMENT VIA THE LEFT INTERNAL JUGULAR VEIN, THE CATHETER LUMEN WAS FOUND TO BE NARROWED DUE TO MOVEMENT OF THE PATIENT'S NECK. IT WAS REPORTED THAT THE PATIENT COMPLAINED OF NECK PAIN DUE TO MEDICINE LEAKAGE WHILE USING THE PORT SYSTEM. REPORTEDLY, THE PORT SYSTEM WAS REMOVED AND LEAKAGE FROM THE CATHETER WAS FOUND AFTER REMOVAL THE PORT SYSTEM. THE PROCEDURE WILL BE COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED IN SECTION HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN COMMON DEVICE NAME AND PMA/510 K. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 01/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-02
16451860 | 2023-02-28 | 2022-10-26 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | DAVOL FEMALE CATH KIT | COLLECTOR, URINE, PEDIATRIC, FOR INDWELLING CATHETER | C. R. BARD, INC. | 0035380 | 0035380 |  | CATHETER IN THE FEMALE STRAIGHT CATHETER KIT IS VERY LOOSE INSIDE OF CAP. CATHETER USED TO BE HELD IN PLACE INSIDE CAP WHICH WOULD ALLOW IT TO BE PULLED FORWARD BUT WOULD PREVENT IT FROM SLIDING BACK IN TUBE OR FALLING OUT. DANGER IS THAT THE CATHETER COULD BE ADVANCED TO FAR FORWARD INTO THE URETHRA AND BECOME DIFFICULT TO REMOVE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01272 | 2023-02-28 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT THE FOLEY CATHETER HAD THE PRE-PRINTED TUBE BAND PRINTED IN THE WRONG SPOT , THEY ATTACHED AN IMAGE OF THE PACKAGING AS WELL. IT WAS ALSO STATED THAT THEY HAD ANOTHER TRAY FROM THE SAME CASE OF TRAYS THAT WAS MISSING ONE OF THE SHEETING CLIPS., IT CAME OUT OF THE PACKAGING WITH ONLY 2 INSTEAD OF 3. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
16452016 | 2023-02-28 | 2022-12-23 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARD® URINE METER FOLEY TRAY WITH LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | 942216 | 942216 |  | INSERTION OF INDWELLING FOLEY CATHETER WAS ORDER SECONDARY TO (2/2) PATIENT URINARY RETENTION AND THE NEED FOR 3 PREVIOUS STRAIGHT CATHETERIZATIONS TO DRAIN BLADDER. INSERTED INDWELLING FOLEY PER PROTOCOL. PICKED UP BAG TO HANG ON SIDE OF BED IT WAS NOTED THE BAG WAS WET. DRIED THE BAG AND UPON CLOSE INSPECTION IT WAS FOUND TO BE LEAKING URINE. NO VISIBLE HOLES NOTED. NOTIFIED COVERING PHYSICIAN. WAS ADVISED TO OPEN SEALED CONNECTION BETWEEN THE CATHETER AND DRAINAGE BAG TO EXCHANGE THE BAG ONLY. CLEANED THE TUBING AT CONNECTION SITE WITH CHG WIPES, APPLIED STERILE GLOVES AND EXCHANGED THE DRAINAGE BAG. |  | N | Y |  |  | 2023-02
1018233-2023-01279 | 2023-02-28 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE INFLATION VALVE WAS DAMAGED WHEN THE PRODUCT WAS OPENED AND STERILIZED WATER WAS INJECTED INTO THE PATIENT AFTER THE FOLEY CATHETER WAS PLACED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00570 | 2023-02-28 | 2023-02-01 | I, n, j, u, r, y | Manufacturer report | M.R.I. HARD BASE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0604520 | 0604520 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE DISTAL END OF THE CHEMO PORT CATHETER ALLEGEDLY FOUND TO BE FRACTURED IN THE THIRD CYCLE OF CHEMOTHERAPY. IT WAS FURTHER REPORTED THAT THE FRACTURED CATHETER MIGRATED TO THE HEART REGION. REPORTEDLY THE MIGRATED SEGMENT WAS REMOVED THROUGH ANGIOGRAM. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 07/2026) THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01277 | 2023-02-28 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414 |  | IT WAS REPORTED THAT THE PATIENT COMPLAINED OF DIFFICULTY IN URINATING, AND FOUND THAT THE FOLEY CATHETER HAD FALLEN OFF UPON CHECKING ON THE NEXT DAY AFTER PLACEMENT. THE EVENT OCCURRED ONE DAY AFTER PLACEMENT. STATED THAT THE BALLOON RUPTURE OR TEAR WAS OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BALLOON. IT WAS STATED THAT THE CATHETER WAS NOT PULLED AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01288 | 2023-02-28 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165L16 | 0165L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER TUBING CLOSEST TO THE PATIENT OFTEN BECAME DISCOLORED AND KIND OF SHRIVELED BEFORE THE MONTH WAS UP. IT WAS STATED THAT THE TUBING HAD A DISCOLORED FLUID YELLOWISH AND THE BULB WAS IN A VERY POOR CONDITION AND CAUSED TRAUMA AND BLEEDING WHEN WITHDRAWN FROM URETHRA. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
1018233-2023-01274 | 2023-02-28 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT THE FOLEY BALLOON HAD A RIDGE WHEN DEFLATED. THE RIDGE MADE THE DISCONTINUATION VERY UNCOMFORTABLE FOR THEIR MALE PATIENTS. SOMETIMES REMOVAL WAS DIFFICULT AND ADDITIONAL FORCE HAD TO BE USED TO PULL IT PAST THE MEATUS. IT WAS ALSO STATED THAT THE NURSE PLACED THE FOLEY, AND DID NOT PRE-INFLATE THE BALLOON. THIS PARTICULAR CASE WAS A 16 FRENCH FOLEY. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-02
3006260740-2023-00593 | 2023-03-01 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM SLIMPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0605550 | 0605550 |  | IT WAS REPORTED THAT SOMETIME PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN IN SHOULDER. HOWEVER, THERE IS NO INFORMATION AVAILABLE REGARDING ADDITIONAL INTERVENTION OR MEDICATION TO TREAT SHOULDER PAIN. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00588 | 2023-03-01 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED "SET CRACKING AT THE DISTAL END LUER LOCK." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01353 | 2023-03-01 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HEMATURIA FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 2551H22 |  | IT WAS REPORTED THAT THE STERILIZATION BAG WAS TORN. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. BASED ON THE EVALUATION, A SMALL TEAR WAS OBSERVED AT THE OUTER PACKAGING OF THE SAMPLE. THE TEAR WAS AT OPEN AREA AND HAS AFFECTED BOTH PLASTIC AND PAPER SIDE OF THE PACKAGE. NO DAMAGE OR TEAR WAS OBSERVED ON THE POLYSLEEVE. THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. HOWEVER, THE POTENTIAL ROOT CAUSE COULD BE DUE TO OPERATOR ERROR (ROUGH HANDLING), USER RELATED (ROUGH HANDLING). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00590 | 2023-03-01 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPIC 4F S/L 3CG BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 22184118 |  | IT WAS REPORTED THAT "CLAMP ON POWER PICC BROKEN. PATIENT WENT HOME WITH CADD IN SITU BUT NO CLAMP. WHEN HE RETURNS FOR CADD OFF WE WILL NOT BE ABLE TO CLAMP PICC. DEVICE HAS BEEN IN SITU FOR A WHILE. THERE WAS NO EXPOSURE TO BODILY FLUIDS AND NO ERRONEOUS RESULTS." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFW2340 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00591 | 2023-03-01 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED BY THE CUSTOMER "ON (B)(6) 2022 LINE FLUSHED WELL AND HAD BLOOD RETURN WHEN HOLDING LIGHT TRACTION ON EXTERNAL PORTION OF LINE BUT OCCLUDED WHEN TRACTION REMOVED. SECURACATH REMOVED BUT PROBLEM PERSISTED WITH STATLOCK IN PLACE. AFTER X-RAY CONFIRMED KINK, LINE RETRACTED 3-4 CM AND THEN BRISK BLOOD RETURN." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED ON 2/28/2023: CUSTOMER REPORTS THEY WERE ABLE TO USE THE PICC UNTIL THE END OF THERAPY AFTER IT WAS RETRACTED 3CM. THE EVENT DID NOT OCCUR DURING AN URGENT OR LIFE-THREATENING MEDICAL CIRCUMSTANCE AND DID NOT RESULT IN A CLINICALLY SIGNIFICANT DELAY IN THERAPY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00592 | 2023-03-01 | 2021-12-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000819 |  | IT WAS REPORTED BY THE CUSTOMER "(B)(6) 2022 PT REPORTED THAT HER PICC WOULD OCCLUDE FREQUENTLY AND THAT SHE WOULD HAVE TO ADDUCT HER RIGHT ARM WHILE KEEPING THE PALMS UP TOWARDS THE CEILING. DRESSING CHANGED RN STARTED NS BOLUS AND NO ISSUES OBSERVED. (B)(6) REPORTS PICC OCCLUSION ISSUES. X-RAY REQUESTED. ATTEMPTED PULLING LINE BACK TO ELIMINATE LOOP X 2. PICC WAS REMOVED ON 1/6 AND NEW PICC ORDER PLACED." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: CATHETER HAD BEEN SECURED WITH SECURACATH. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DIFFICULTY INFUSING THROUGH THE CATHETER WAS CONFIRMED BUT THE EXACT CAUSE IS UNKNOWN. THREE CONED DOWN RADIOGRAPHIC IMAGES OF THE PATIENT¿S RIGHT ARM AND THE PROXIMAL PORTION OF THE PICC WERE RETURNED FOR EVALUATION OF THIS COMPLAINT. THE IMPLICATED PRODUCT WAS A 4FR S/L POWERPICC CATHETER. THERE WAS A LOOP AND POTENTIALLY A KINK PRESENT IN THE CATHETER SHAFT. THE LOOP APPEARED TO BE IN THE SUBCUTANEOUS SOFT TISSUES AND THE POSSIBLE KINK WAS LOCATED AT THE ENTRANCE SITE INTO THE VEIN. A COMPRESSION SECUREMENT DEVICE WAS SEEN IN THE IMAGES, PROXIMAL OF THE CATHETER LOOP AND KINK. IT WAS REPORTED THAT FLOW THROUGH THE CATHETER WAS CONTINGENT UPON THE PATIENT¿S ARM POSITION, INDICATING A POSITIONAL OCCLUSION. THE OCCLUSION WAS LIKELY CAUSED BY THE LOOP/KINK IN THE CATHETER. POSSIBLE CONTRIBUTING FACTORS COULD INCLUDE THE PATIENT¿S ANATOMY (MUSCLES AND FASCIA), CATHETER PLACEMENT, OR CATHETER SECUREMENT. AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT. THE REPORT OF DIFFICULTY INFUSING THE CATHETER, SECONDARY TO CATHETER KINKING, HAS BEEN DOCUMENTED AND WILL BE MONITORED AS PART OF COMPLAINT TRENDING. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11 | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00121 | 2023-03-01 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQF75124 | CQF75124 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY HAD PINHOLE RUPTURED AT 28 ATM. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: B5, D4 (EXPIRATION DATE: 09/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00598 | 2023-03-01 | 2022-11-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED ON 11/30 PATIENT PULLED OUT PICC. NO NEW PICC WAS INSERTED. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGQ3272 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00595 | 2023-03-01 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET (4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000819 |  | IT WAS REPORTED BY THE CUSTOMER THAT "SOME SCISSORS FROM THE KIT THAT LEFT DISCOLORATION ON THE STAT SEAL DISC WHEN IT WAS TRIMMED. " ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER: "THE DISCOLORATION WAS BROWN. A NEW DISC AND SCISSORS WERE OBTAINED BUT PICC HAD ALREADY BEEN PLACED AND THE SCISSORS WERE USED TO TRIM CATHETER. PATIENT IS BEING MONITORED FOR POTENTIAL COMPLICATIONS. PATIENT IS BEING TREATED FOR MSSA BACTEREMIA AND PULMONARY VALVE ENDOCARDITIS." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DISCOLORATION IS CONFIRMED; HOWEVER, THE ROOT CAUSE COULD NOT BE DETERMINED. THE PRODUCT RETURNED FOR EVALUATION WAS A PAIR OF SCISSORS. MICROSCOPIC INSPECTION OF THE INSIDE AND OUTSIDE OF THE SCISSOR EDGES REVEALED SPOTS OF A REDDISH-BROWN RESIDUE, WHICH APPEARED CONSISTENT WITH MATERIAL OXIDATION. ALTHOUGH A REDDISH-BROWN RESIDUE WAS OBSERVED ON THE RETURNED SAMPLES, IT IS UNKNOWN HOW THE RESIDUE ACCUMULATED ON THE RETURNED SAMPLE. CONSEQUENTLY, THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00596 | 2023-03-01 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PPICC SOLO 4F SL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001027 |  | IT WAS REPORTED BY THE CUSTOMER THE GOWN TORE DURING PROCEDURE. THEY HAD TO OPEN ANOTHER KIT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A TORN GOWN IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE SURGICAL GOWN WAS RETURNED FOR EVALUATION WITHOUT PRODUCT PACKAGING. AN INSPECTION OF THE GOWN REVEALED DAMAGE NEAR THE ARM SLEEVE. THE DAMAGED SECTION WAS OBSERVED TO BE A TEAR NEAR THE SLEEVED REGION. THE EDGES WERE OBSERVED TO BE ROUGH AND THE CHARACTERISTICS WERE CONSISTENT WITH TEARING OF THE MATERIAL. ADDITIONAL DAMAGE ON THE GOWN WAS NOTED AND APPEARS TO BE A SLIT. THE CHARACTERISTICS OF THE SPLIT EDGES ARE INDICATIVE TO POTENTIAL EXPOSURE TO A SHARP INSTRUMENT. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE DAMAGE DURING HANDLING AND SHARP INSTRUMENT DAMAGE. SINCE DAMAGE TO THE GOWN WAS FOUND TO BE PRESENT, THE COMPLAINT IS CONFIRMED BUT THE EXACT ROOT CAUSE(S) REMAINS UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00599 | 2023-03-01 | 2023-01-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH 15 CM PRE-CURVE KIT | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5554150 |  | IT WAS REPORTED BY THE CUSTOMER, THE WIRE ACTUALLY BROKE WITHIN THE NEEDLE. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGP0146 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00600 | 2023-03-01 | 2022-12-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D |  | IT WAS REPORTED BY THE CUSTOMER "12/20 PICC POSITIONAL AND KINK SHOWN ON XRAY. ATTEMPTED TO PULSATILE FLUSH CATHETER TO FIX KINK BUT NOT SUCCESSFUL. LINE ONLY HAS BLOOD RETURN WHEN RIGHT ARM IS LIFTED UP AND HAND UP BY RIGHT EAR. PICC REMOVED AND NEW ONE PLACED." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED 2/28/2023: IT WAS REPORTED BY THE CUSTOMER "ORIGINAL CATHETER WAS SECURED WITH SECURACATH." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00601 | 2023-03-01 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED BY THE CUSTOMER "ON DAY OF INSERTION PICC WITHOUT BLOOD RETURN AND DIFFICULTY FLUSHING. LINE PULLED BACK (TIP IN ATRIUM) BUT LINE CONTINUED TO HAVE ISSUES. LINE DECLOTTED ABOUT EVERY 2 DAYS UNTIL (B)(6) 2023 WHEN LINE REWIRED BY IR (PT COULD NOT TOLERATE BEDSIDE PROCEDURE). ON 1/31 PATIENT WAS REFERRED TO ONC PICC TEAM FOR EVALUATION OF EXCHANGED IR PICC. LINE FOUND KINKED UNDER STATSEAL RING PULLED BACK AND REDRESSED." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED 2/28/2023: ADMISSION DIAGNOSIS IS MALIGNANT NEOPLASM OF ENDOMETRIUM. CATHETER WAS SECURED WITH SECURACATH. CUSTOMER REPORTS THERE WAS NO LONG-TERM HARM FOR THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00597 | 2023-03-01 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PPICC SOLO 4F SL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001027 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE INTRODUCER WAS NOT ABLE TO PENETRATE SKIN WITHOUT BENDING OR BREAKING. THEY OPENED ANOTHER INSERTION KIT. IT WAS STATED THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGW0938 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00583 | 2023-03-01 | 2023-01-20 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT ONE YEAR TEN MONTHS AND SIXTEEN DAYS POST PORT DEVICE REMOVAL FOR A PATIENT HAVING RECTAL CANCER, THE CATHETER WAS ALLEGEDLY FRACTURED NEAR THE INSERTION SITE INTO THE SUBCLAVIAN VEIN. IT WAS FURTHER IDENTIFIED THAT THE PORT SHOWED CATHETER PERFORATION NEAR THE SUBCLAVIAN VEIN INSERTION SITE, LEAKAGE OF FLUID ALLEGEDLY HAPPENED FROM LEFT CHEST PORT ACCESS SITE WHILE TRIED TO REMOVE THE PRODUCT. FURTHERMORE, THE PATIENT ALLEGEDLY EXPERIENCED DISCOMFORT DURING INFUSIONS AND PAIN WITH PALPATION AROUND THE AREA OF THE PORT. IT WAS FURTHER REPORTED THAT THE PORT WAS ATTEMPTED TO BE REMOVED BY OUTPATIENT SURGERY, BUT CATHETER PERFORATION BECAME A COMPLETE CATHETER FRACTURE WITH A RETAINED FRAGMENT LOCATED INSIDE THE PATIENT¿S BODY. THE PATIENT WAS THEN TRANSFERRED TO VASCULAR INTERVENTIONAL RADIOLOGY FOR RETAINED FOREIGN BODY RETRIEVAL. THE USER FACILITY REPORTED THAT THE FRACTURED CATHETER FRAGMENT HAD NOT EMBOLIZED INTO THE HEART OR PULMONARY ARTERIES; HOWEVER, THE DECISION WAS MADE TO SNARE THE FRACTURED CATHETER AND REMOVE IT FROM THE FEMORAL VEIN ACCESS SITE. THE PATIENT WAS REPORTED TO BE WELL POST-PROCEDURE. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER IN THREE SEGMENTS WAS RETURNED FOR EVALUATION. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, TEN ELECTRONIC PHOTOS WERE PROVIDED FOR REVIEW. ALL PHOTOS SHOW ONE (1) POWERPORT ISP MRI IMPLANTABLE PORT WITH A SMALL SEGMENT OF CATHETER STILL ATTACHED AND TWO (2) OTHER GROSHONG CATHETER SEGMENTS PLACED IN A BIOHAZARD BAG. FUNCTIONAL, GROSS VISUAL, TACTILE, AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION, AND FLUID LEAK ISSUES. A COMPLETE COMPOUND BREAK WAS NOTED, ABOUT 14 CM FROM THE DISTAL END, WHICH MANIFESTS VARIATION IN THE TEXTURE OF THE BREAK SURFACES SUCH THAT CERTAIN REGIONS SHOW SMOOTHNESS/ROUNDNESS AND GRANULARITY IN OTHERS. THE LUMEN IS COMPRESSED AT BOTH ENDS OF THIS BREAK AND PROJECTIONS ARE NOTED, MAKING THE BREAK SIGNIFICANTLY UNEVEN. SOME LONGITUDINAL CRACKING WAS ALSO NOTED. UPON GENTLY STRETCHING EACH CATHETER, ABNORMAL ELASTICITY WAS FELT THROUGHOUT. KINKS WERE NOTED THROUGHOUT BOTH CATHETER SEGMENTS. FROM THE ANALYSIS OF THE SAMPLE AS A WHOLE, WEAR AND KINKING WERE ALSO IDENTIFIED. THE REPORTED DIFFICULTY OF REMOVAL CANNOT BE DIRECTLY CONFIRMED FROM THE SAMPLES RETURNED. THE IDENTIFIED BREAK CONTAINS CHARACTERISTICS TYPICAL OF FLEXURAL FATIGUE, WHICH IS DUE TO THE REPETITIVE, CYCLIC KINKING OF THE CATHETER. SUCH KINKING MAY HAVE PHYSIOLOGICAL, PLACEMENT, USAGE OR MECHANICAL CONTRIBUTING FACTORS. IT APPEARS THAT THE BREAK WAS ACCELERATED BY THE PROCESS OF REMOVAL FROM THE PATIENT. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 04/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01357 | 2023-03-01 | 2023-02-08 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WAS WAKING UP WET. THE PAD UNDER THE PATIENT WAS WET WHEN THE CUSTOMER CHANGED THE PUREWICK FEMALE EXTERNAL CATHETER. THE WICK DID NOT STAY INTACT AT TOP. THE CUSTOMER PUT SOMETHING OVER THE TOP TO HELP BUT THE PATIENT WAS STILL WET. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. PER CUSTOMER VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT PATIENT HAS URINARY TRACT INFECTION AND WAS CURRENTLY ON AN ANTIBIOTIC. PER CUSTOMER VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE PATIENT SOUGHT MEDICAL INTERVENTION AND WAS PRESCRIBED AMOXICILLIN ANTIBIOTIC. URINARY TRACT INFECTION WAS NOT CONFIRMED TO HAVE BEEN CAUSED FROM USING PUREWICK. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS: TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00594 | 2023-03-01 | 2023-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DLMBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED BY THE CUSTOMER "ON (B)(6) 2023 CALLED TO ASSESS PATIENT LUA DL PICC, DUE TO OCCLUSION AND POSITIONAL SINCE LAST NIGHT, SEEN BY VAT RN, NOTED WITH BRISK BLOOD RETURN ON BOTH LUMEN, DRESSING IS INTACT NO VISIBLE KINK, CPN AND IV FLUIDS FLOWING VERY GOOD WHEN ARM IS ABDUCT AND ADDUCT , BUT WHEN ASK PATIENT TO RAISED HIS ARM ABOVE HIS HEAD , IV PUMP STARTED TO BEEP AND OCCLUSION DETECTED, POSSIBLE CATHETER ISSUE AS SEEN IN CXR ,ABOUT 7-8 CM ABOVE THE SITE NOTED INTRA LUMEN (POSSIBLE KINK/LINE FRACTURE)." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED 2/28/2023: IT WAS REPORTED BY THE CUSTOMER "CATHETER WAS SECURED WITH SECURACATH." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01342 | 2023-03-01 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 165812 | 165812 |  | IT WAS REPORTED THAT THERE WERE 3 INCIDENTS WHERE THEY WERE USING SILICONE FOLEY AFTER REMOVING THE WATER, THE BALLOON SHOWED EVIDENCE OF CUFFING. THE CUSTOMER HAD SAVED ONE SAMPLE AND REQUESTED A BOX TO PICK IT UP FOR REVIEW. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED ONE PHOTO SAMPLE RECEIVED WHICH INDICATED MUSHROOMING ON THE CATHETER BALLOON. THEREFORE, PRODUCT DOES NOT MEET SPECIFICATIONS WHICH STATES "MISSING/INCORRECT COMPONENTS ARE NOT ALLOWED." ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE BALLOON MATERIAL DOES NOT SHRINK QUICKLY ENOUGH. THE PRODUCT WAS USED FOR PATIENT DIAGNOSTIC OR TREATMENT. THE PRODUCT WAS INFLUENCED BY THE REPORTED EVENT. DHR REVIEW IS NOT REQUIRED AS NO LOT NUMBER WAS PROVIDED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. ALLOW THE BALLOON TO DEFLATE SLOWLY ON ITS OWN. DO NOT ASPIRATE OR MANUALLY ACCELERATE THE DEFLATION OF THE BALLOON. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00118 | 2023-03-01 | 2023-01-01 | I, n, j, u, r, y | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG120164 | ATG120164 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE THROUGH FEMORAL ARTERY, THE PTA BALLOON ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT WHILE RETRACTING BALLOON THROUGH THE INTRODUCER SHEATH BALLOON GOT STUCK AT THE DISTAL TIP OF SHEATH. REPORTEDLY THE NEW ACCESS SITE WAS USED TO REMOVE THE BALLOON. THE PATIENT IS REPORTED TO BE STABLE. | ADDITIONAL INFORMATION WAS RECEIVED, AND THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE REPORTABLE AS SERIOUS INJURY. MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ATLAS GOLD PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, THE DISTAL TIP AND BALLOON WERE NOTED COMPLETELY DETACHED FROM THE CATHETER. THE BALLOON WAS ALSO INVERTED WITH FRAYED FIBERS. THE INNER GUIDEWIRE LUMEN WAS COMPLETED DETACHED FROM THE OUTER CATHETER AND WAS NOTED TO HAVE BUNCHING THROUGHOUT. BOTH THE INNER GUIDEWIRE LUMEN AND OUTER CATHETER WERE NOTED TO BE KINKED THROUGHOUT. THE MARKER BANDS WERE STILL PRESENT. NO FUNCTIONAL TESTING WAS PERFORMED DUE TO THE CONDITION OF THE SAMPLE. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED BALLOON RUPTURE AND SHEATH REMOVAL DIFFICULTY AS NO FUNCTIONAL TESTING COULD BE PERFORMED DUE TO THE CONDITION OF THE RETURNED SAMPLE. HOWEVER, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED DETACHMENT AS THE SAMPLE WAS RETURNED DETACHED WITH THE INNER GUIDEWIRE LUMEN, OUTER CATHETER AND BALLOON ALL SEPARATE. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED BALLOON RUPTURE, SHEATH REMOVAL DIFFICULTY AND IDENTIFIED DETACHMENT COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01341 | 2023-03-01 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 175814 | 175814 |  | IT WAS REPORTED THAT THE BALLOON ON THE FOLEY CATHETER DID NOT INFLATE SYMMETRICALLY DURING TRAINING. DURING TRAINING FOR THE BARD FOLEY TRAYS IN BARCELONA, A HEALTHCARE PERSONNEL POINTED OUT THAT THE BALLOON HAD NOT INFLATED SYMMETRICALLY. ALSO STATED THAT IT APPEARS THAT THE BALLOON WOULD NOT SIT IN THE PATIENT¿S BLADDER TO ENSURE CORRECT DRAINAGE. THE CATHETER HAD BEEN INSERTED INTO AN ANATOMICAL MODEL. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE MET RELEVANT SPECIFICATIONS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITH ORIGINAL PACKAGING), FOLEY CATHETERIZATION TRAY. VISUAL INSPECTION OF THE SAMPLE NOTED THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE BALLOON CONCENTRICITY WAS OBSERVED TO BE 60:40. WITH THE SYRINGE ATTACH THE BALLOON DEFLATED PASSIVELY WITH NO ISSUES. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01305 | 2023-03-01 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE PATIENT WOULD LIKE TO RETURN MATERIAL# 51814 SINCE IT DID NOT WORK FOR THEM. THE MAGIC 3 GO FEMALE INTERMITTENT CATHETERS WERE DIFFICULT TO INSERT. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿OUTSIDE DIAMETER IS TOO LARGE OR ROUGH OR IRREGULAR SHAPE PROTRUSIONS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INDICATIONS FOR USE THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, OR IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS.¿ THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01303 | 2023-03-01 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® FOLEY TRAY WITH BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | 899616 | 899616 |  | IT WAS REPORTED THAT THERE WAS A LEAK FROM THE FOLEY TRAY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "IMPROPER GROMMET FIT WITHIN BARREL RESULTING IN BLOW-BY/ SPILLAGE OF WATER". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DIRECTIONS FOR USE: VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. OPEN CSR WRAP TO FORM STERILE FIELD AND PLACE UNDERPAD BENEATH PATIENT, PLASTIC SIDE DOWN. PUT ON CUFFED GLOVES AND POSITION DRAPE ON PATIENT. OPEN CATHETER LUBRICATING JELLY. REMOVE TOP TRAY AND OPEN PLASTIC POUCH SURROUNDING CATHETER. LUBRICATE CATHETER. PREPARE PATIENT WITH POVIDONE-IODINE SWAB PACKET. PROCEED WITH CATHETERIZATION IN USUAL MANNER. TO INFLATE CATHETER, SIMPLY INSERT TIP OF WATER-FILLED SYRINGE GENTLY INTO VALVE (DO NOT OVERPENETRATE) AND DEPRESS PLUNGER. INSTILL ENTIRE AMOUNT OF STERILE WATER. POSITION HANGER ON BEDSIDE RAIL NEAR THE FOOT OF THE BED USING SHEETING CLIP TO SECURE DRAINAGE TUBE TO DRAW SHEET. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. PERIODIC OBSERVATIONS OF THIS SYSTEM SHOULD BE MADE TO ENSURE THAT URINE IS FLOWING FREELY. IF A STANDING COLUMN OF URINE IS OBSERVED, CHECK FOR CORRECT POSITIONING OF BAG AND THEN FOR A PHYSICAL OBSTRUCTION. IF CORRECT POSITIONING OR REMOVAL OF PHYSICAL OBSTRUCTION DOES NOT ALLOW FREE FLOW, THE BAG MAY HAVE TO BE CHANGED. TAMPER-EVIDENT SEAL DISCONNECT INSTRUCTIONS: IF NECESSARY TO DISCONNECT CATHETER AND DRAINAGE TUBE CONNECTOR, FLEX CATHETER AS ILLUSTRATED, GRASP TAB AND SLOWLY PULL ALONG PERFORATIONS UNTIL SECTION IS REMOVED. DIRECTIONS FOR USING BARD® EZ-LOK¿ SAMPLING PORT: BARD® EZ-LOK¿ SAMPLING PORT ACCEPTS A LUER-LOCK OR SLIP TIP SYRINGE. KINK DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT (AT APPROXIMATELY 80-100 DEGREE ANGLE). PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. NOTE: IMPROPER PENETRATION TECHNIQUE COULD CAUSE FORMATION OF A DROP OF URINE ON THE SURFACE OF THE SAMPLING PORT. PERIODIC OBSERVATION OF THE SAMPLING PORT IS RECOMMENDED. ASPIRATE DESIRED VOLUME OF URINE. UNKINK TUBING AND SEND SPECIMEN TO LABORATORY." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01306 | 2023-03-01 | 2022-12-29 | D, e, a, t, h | Manufacturer report | NBIH¿ TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 007406P | 007406P |  | IT WAS REPORTED THAT FOLLOWING TRANSCATHETER IMPLANTATION OF AN AORTIC VALVE, THE PATIENT EXPERIENCED AN ATRIOVENTRICULAR BLOCK, AND A TEMPORARY PACING ELECTRODE (TPE) CATHETER WAS PLACED. THE TPE CATHETER USED WAS A LARGER CALIBER THAN USUAL (CONTEXT OF LABORATORY RUPTURE). A 7F WAS USED INSTEAD OF THE USUAL 5F, BUT THE PATIENT WAS OKAY AT THE TIME OF DISCHARGE FROM THE BLOCK. THERE WAS TAMPONADE AT DAY 1 FOLLOWING THE TRANSCATHETER IMPLANTATION OF THE AORTIC VALVE WITH PRESENCE OF A FIBRILLATION ARTIALE DROITE (AF D; RIGHT-SIDED ATRIAL FIBRILLATION) BREACH. THE PATIENT WAS TRANSFERRED TO THE OPERATING ROOM 3 HOURS LATER FOR THE REMOVAL OF THE PROBE AND TREATMENT. THERE WAS AN UNFAVORABLE SITUATION ON (B)(6) 2022 AT 7 PM, AND THE PATIENT EXPIRED. IT WAS REPORTED THAT A POTENTIAL CAUSE OF THE PATIENT¿S DEATH WAS RELATED TO THE STIMULATION PROBE. 7F STIMULATION PROBES ARE NO LONGER USED AT THIS FACILITY. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL FROM THE INTERNATIONAL BUSINESS CENTERS (IBC) ON 03-FEB-2023, IT WAS REPORTED THAT THE DEVICE SEEMED VERY RIGID TO THE SURGEON. ADDITIONAL CLINICAL FOLLOW UP INFORMATION WAS RECEIVED VIA EMAIL ON 23-FEB-2023 FROM THE IBC. FOLLOWING THE TRANSCATHETER IMPLANTATION OF AN AORTIC VALVE, THE PATIENT EXPERIENCED THE ONSET OF HIGH-GRADE ATRIOVENTRICULAR BLOCK, AND A TPE CATHETER WAS PLACED. A 7F CATHETER HAD TO BE USED BECAUSE THE FACILITY DID NOT HAVE ANY 5F TPE CATHETERS AVAILABLE. A 7F WAS OBTAINED FOR USE FROM ANOTHER ESTABLISHMENT. FOLLOWING PLACEMENT OF THE TPE CATHETER, THE PATIENT WAS DISCHARGED TO THE INTENSIVE CARE UNIT (ICU). SEVERAL HOURS AFTER THE DEVICE WAS PLACED, THE PATIENT EXPERIENCED THE ONSET OF SHOCK WITH TAMPONADE, WHICH REQUIRED EMERGENCY BYPASS SURGERY FOR REPAIR OF A 2CM RIGHT VENTRICULAR WOUND. ON (B)(6) 2022, THE PATIENT EXPIRED, AND THE DEATH WAS REPORTED AS A RESULT OF MULTIVISCERAL FAILURE AT DAY 1. THE FOLLOWING IS LISTED AS THE CAUSE OF DEATH ON THE PATIENT¿S DEATH CERTIFICATE: OCCURRENCE OF TAMPONADE, EMERGENCY SURGICAL DRAINAGE, REFRACTORY SHOCK COMPLICATED BY POLYVISCERAL FAILURE, ASYSTOLE. PER ADDITIONAL CLINICAL FOLLOW UP INFORMATION RECEIVED VIA EMAIL ON 16-MAR-2023 FROM THE IBC, THERE WERE NO DIFFICULTIES ENCOUNTERED WHILE INSERTING THE TPE DEVICE INTO THE RIGHT VENTRICULAR POINT. THE PATIENT LATER EXPERIENCED TAMPONADE DUE TO A 2CM RIGHT VENTRICULAR WOUND. IT WAS REPORTED THE WOUND POSSIBLY OCCURRED DUE TO BOTH THE RIGIDITY AND SIZE OF THE TPE DEVICE. | THE REPORTED EVENT IS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED AND FURTHER INVESTIGATION WAS NOT CONCLUSIVE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT POTENTIAL ROOT CAUSES COULD BE, "MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE", "USER ERROR", "MOISTURE FROM WATER TEST (MP5709) AND FLUSH", "SPLICE BANDS / CIRCUIT WIRES TOUCHING WITHIN MOLD, CIRCUIT WIRES TOUCHING WITHIN SHAFT", "IMPROPER SPLICE / IMPROPER CONTACT BETWEEN CIRCUIT WIRE AND ELECTRODES", "SHORT AT SLOT", "OPERATOR ERROR". THE DEVICE WAS USED FOR TREATMENT PURPOSES. AS THE DEVICE WAS NOT RETURNED IT WAS UNKNOWN IF IT HAD MET ALL RELEVANT SPECIFICATIONS OR RESULTED IN THE REPORTED EVENT. NO MANUFACTURING ISSUES OR NON-CONFORMANCES WERE NOTED DURING REVIEW OF THE DHR THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "THE TEMPORARY PACING ELECTRODE CATHETER INFORMATION FOR USE, (B)(6) REVISION 8, WAS REVIEWED AND FOUND TO BE ADEQUATE AS IT STATES: "READ THIS DOCUMENT IN ITS ENTIRETY PRIOR TO USE." "INAPPROPRIATE ELECTRICAL CONNECTIONS, E.G. INTO A WALL SOCKET, MAY POSE SERIOUS RISK OF ADVERSE HEALTH CONSEQUENCES OR DEATH." "THIS DEVICE SHOULD BE USED ONLY BY OR UNDER THE SUPERVISION OF PHYSICIANS TRAINED IN THE TECHNIQUES OF TRANSVENOUS INTRACARDIAC STUDIES AND TEMPORARY PACING." "THE RISKS OF USING TEMPORARY PACING CATHETERS INCLUDE THOSE RISKS RELATED TO HEART CATHETERIZATION, SUCH AS THROMBOEMBOLISM, PERFORATION, TAMPONADE, AND INFECTION. THE INDUCTION OF AN UNINTENDED ARRHYTHMIA IS A KNOWN COMPLICATION." "VISUALLY INSPECT THE CATHETER, UNDER STERILE CONDITIONS, FOR KINKS IN THE CATHETER SHAFT, INTEGRITY OF THE CONNECTOR, CONDITION OF ELECTRODES, AND ANY OTHER DAMAGE." "INSERTION INSTRUCTIONS USING A NEEDLE CANNULA 1. OPEN THE PACKAGE AND PLACE THE CONTENTS ON A STERILE FIELD. 2. PREP THE SKIN AT THE SITE OF INSERTION AND INJECT A LOCAL ANESTHETIC. 3. REMOVE THE PROTECTIVE GUARD FROM THE NEEDLE CANNULA. 4. ENTER THE VEIN WITH THE NEEDLE CANNULA. SIMULTANEOUS ASPIRATION INTO A SYRINGE WILL HELP CONFIRM VESSEL ENTRY. 5. REMOVE THE SYRINGE AND THE NEEDLE. 6. IF USING AN OPEN-LUMEN CATHETER, FLUSH THE CATHETER WITH A HEPARINIZED SOLUTION. REMOVE ANY STYLETTE PRIOR TO INSERTION. 7. USING THE AID OF FLUOROSCOPY OR AN ECG MONITOR, ADVANCE THE CATHETER THROUGH THE CANNULA TO THE DESIRED POSITION. IF USING A BALLOON CATHETER, INFLATE THE BALLOON WHEN THE CATHETER IS IN THE RIGHT ATRIUM. PLEASE NOTE THAT THE BALLOON CAN BE INFLATED OR DEFLATED ONLY WHEN THE STOPCOCK IS PARALLEL TO THE CATHETER SHAFT. DO NOT PULL THE CATHETER BACK THROUGH THE CANNULA AS IT MAY CAUSE DAMAGE TO THE CATHETER. 8. IF USING A BALLOON CATHETER, DEFLATE THE BALLOON AFTER THE CATHETER HAS REACHED THE DESIRED LOCATION. 9. TEST THE PACING CHARACTERISTICS FOR OPTIMAL PACING. 10. PULL THE CANNULA BACK AND SECURE IT TO THE PROXIMAL END OF THE CATHETER. 11. SECURE THE ELECTRODE CATHETER IN PLACE AT THE INSERTION SITE." "ELECTRICAL CONNECTIONS FOR MEASURING INTRACARDIAC ECGS 1. INSERT ADAPTER PIN INTO STANDARD 2 MM (0.080¿) PIN RECEPTACLE OF EQUIPMENT. IF EQUIPMENT CONTAINS A LOCKING MECHANISM SUCH AS A COLLET OR THUMBSCREW, TIGHTEN DOWN ONTO ADAPTER. LEAVE AFFIXED TO EQUIPMENT. WARNING: INAPPROPRIATE ELECTRICAL CONNECTIONS, E.G. INTO A WALL SOCKET, MAY POSE SERIOUS RISK OF ADVERSE HEALTH CONSEQUENCES OR DEATH. 2. THREAD LEADS OF CATHETER THROUGH THE ADAPTER EYELET. CONNECT THE NEGATIVE JACK (MARKED ¿DISTAL¿) TO THE V-LEAD OF THE ECG, AND THE POSITIVE JACK (UNMARKED) TO THE POSITIVE TERMINAL OF THE EXTERNAL PULSE GENERATOR. ELECTRICAL CONNECTIONS FOR PACING 1. INSERT ADAPTER PIN INTO STANDARD 2 MM (0.080¿) PIN RECEPTACLE OF EQUIPMENT. IF EQUIPMENT CONTAINS A LOCKING MECHANISM SUCH AS A COLLET OR THUMBSCREW, TIGHTEN DOWN ONTO ADAPTER. LEAVE AFFIXED TO EQUIPMENT. WARNING: INAPPROPRIATE ELECTRICAL CONNECTIONS, E.G. INTO A WALL SOCKET, MAY POSE SERIOUS RISK OF ADVERSE HEALTH CONSEQUENCES OR DEATH. 2. THREAD LEADS OF CATHETER THROUGH THE ADAPTER EYELET. CONNECT THE NEGATIVE JACK (MARKED DISTAL) TO THE NEGATIVE TERMINAL OF THE EXTERNAL PULSE GENERATOR, AND THE POSITIVE JACK (UNMARKED) TO THE POSITIVE TERMINAL OF THE PULSE GENERATOR." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00580 | 2023-03-01 | 2023-01-31 | I, n, j, u, r, y | Manufacturer report | UNKNOWN DIALYSIS | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | UNK DIALYSIS |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED INFECTION. IT WAS FURTHER REPORTED THAT THE CATHETER WAS REMOVED DUE TO POSSIBLE BIOFILM. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: ADDITIONAL INFORMATION WAS RECEIVED, THE AFFECTED DEVICES WERE NOT BD/BARD DEVICES AND HENCE THIS RECORD WILL NOT BE CONSIDER AS A COMPLAINT. HOWEVER, SINCE AN INITIAL MDR WAS SUBMITTED, THE FILE WILL REMAIN ASSESSED AS A SERIOUS INJURY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01309 | 2023-03-01 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI12 | 0165SI12 |  | IT WAS REPORTED THAT THE STERILE WATER WAS UNABLE TO BE INJECTED INTO THE FOLEY CATHETER DUE TO WATER LEAKAGE FROM SOMEWHERE DURING THE PRETEST. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. THE AFFECTED SYRINGE WAS MANUFACTURED BY HANSCENT. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER ¿ HANSCENT. THE DEFORMATION AT THE TIP OF THE SYRINGE IS THE ROOT CAUSE THAT CONTRIBUTED TO THE LEAK; HOWEVER, THE ROOT CAUSE OF THE DEFORMATION IS UNKNOWN. NO FAILURES WERE FOUND WITHIN THE SUPPLIER INVESTIGATIONS. BOTH FLEX AND HANSCENT SHOWED NO CLEAR EVIDENCE THAT THE FAILURES OCCURRED DURING THEIR PROCESSES. HANSCENT DID NOT FIND ANY ROOT CAUSE OF THE LEAKAGE. ALL PROCESSES WITHIN HANSCENT WERE UNDER CONTROL AND NO DEFORMATION AND LEAKAGE WERE FOUND IN PRODUCTION RETAIN SAMPLES, INVENTORY, AND SIMULATION PRODUCTION PRODUCTS, NO LEAKAGE RECORD IN DHR. AS BD WILL NO LONGER BE USING HANSCENT AND FLEX FOR THE SYRINGES, THIS INVESTIGATION WILL BE CONCLUDED AS CAUSE UNKNOWN. PRIME CARE WILL BECOME BD¿S QUALIFIED SUPPLIER. DHR REVIEW IS NOT REQUIRED. A LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01310 | 2023-03-01 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE SAMPLE OF THE MAGIC 3 GO MALE HYDROPHILIC INTERMITTENT CATHETER WAS TOO HARD TO INSERT, AND THE PATIENT HAD A LITTLE BLEEDING WHEN REMOVING THE CATHETER AS THE TIP WAS BLUNT. THE CATHETERS THAT WERE TAPERED WORKED MUCH BETTER FOR THE PATIENT. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PLC FAILURE ,OPERATOR ERROR". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00581 | 2023-03-01 | 2022-09-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN DIALYSIS | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | UNK DIALYSIS |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY HAD INADEQUATE FUNCTIONING OF THE VENOUS ACCESS. THERE WAS NO REPORTED PATIENT INJURY. | ADDITIONAL INFORMATION WAS RECEIVED, THE DEVICES WERE NOT BD/BARD DEVICES AND HENCE THIS RECORD WILL NOT BE CONSIDERED AS A COMPLAINT. HOWEVER, SINCE AN INITIAL MDR WAS SUBMITTED, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01304 | 2023-03-01 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165814 | 165814 |  | IT WAS REPORTED THAT FOUR FOLEY CATHETERS WERE BROKEN AND WERE RETURNED BY THE PATIENT . AS PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 14DEC2022, ALL 4 PROBES WERE USED ON THE PATIENT AND THE BALLOONS WENT OFF EACH TIME. THE PATIENT WAS AN ALZHEIMER'S PATIENT AND THEREFORE WAS DEPENDENT ON A BLADDER PROBE. AS PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 03JAN2023, STATED THAT THE CUSTOMER MEANT THAT PRESSURE DISAPPEARED AS THE WATER WENT OUT BY THE "BALLOON WENT OFF" STATEMENT. AS PER SAMPLE RECEIVED ON 07FEB2023, STATED THAT THE CUSTOMER RETURNED TWO ADDITIONAL USED SAMPLES FOR THIS COMPLAINT . | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01311 | 2023-03-01 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT STERILE WATER WAS UNABLE TO ENTER THE CATHETER. THERE WAS NO PROBLEM OBSERVED IN THE CATHETER DURING INFLATION. THE FUNNEL DID NOT SWELL AND NO ABNORMALITY OBSERVED ON THE SHAFT. THE WATER WAS INJECTED GENTLY. USUALLY, THE BALLOON CAN BE INFLATED WITH IN FEW SECONDS. ADDITIONAL INFORMATION WAS RECEIVED VIA IBC ON 16FEB2023, STATED THAT "THE EVENT WAS REPORTED TO HAVE OCCURRED AFTER INSERTION TO THE PATIENT. THE PRODUCT WAS USED ON THE PATIENT." | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. SAMPLE WAS EVALUATED AND FOUND TEAR ON SHAFT NEAR BIFURCATION AREA THAT ALLOW WATER TO LEAK OUT. THE TEAR LOOKS LIKE IT HAD BEEN EXPOSED TO SHARP OBJECT. THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE (EXAMPLE: USER RELATED/OPERATOR ERROR/ROUGH HANDLING OR EXPOSE TO SHARP OBJECT/MECHANICAL FORCE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. (THE URETHRA MAY BE INJURED) (2) DO NOT PULL THE CATHETER HARD. (THE BLADDER/URETHRA MAYBE INJURED) 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER (WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED). CONTRAINDICATIONS 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE ON PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01313 | 2023-03-01 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT EVEN THOUGH SUCTION PRESSURE WAS NOT APPLIED THE STERILIZED WATER DID NOT DRAIN AND URINE WAS FOUND MIXED WITH THE STERILE WATER THAT WAS DRAINED. IT WAS STATED THAT URINE LEAKAGE FROM THE INFLATION VALVE WAS CONFIRMED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE TORN RUBBERIZE ON RUBBERIZE LAYER THAT MAY LEAD TO LUMEN-TO-LUMEN LEAK. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01319 | 2023-03-01 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER CAME OUT. ALSO STATED THAT PATIENT NOTICED THAT AFTER 6 HOURS OF PLACEMENT, IT WAS SPONTANEOUSLY REMOVED. AT THAT TIME, NO STERILIZED WATER WAS CONTAINED IN THE BALLOON. ALSO STATED THEY DO NOT DO PRE-TEST, BUT WHEN PUT AIR IN AND CHECKED IT AFTER THE INCIDENT IT INFLATED ONCE, BUT THE AIR QUICKLY ESCAPED, SO THEY GUESSED THE BALLOON WAS DAMAGED. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE PUNCH DEPTH TOO LARGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01338 | 2023-03-01 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THE PATIENT STATED IT WAS DRY AND DRAGS UNCOMFORTABLE WHILE USING THE MAGIC 3 GO HYDROPHILIC MALE INTERMITTENT CATHETER. IT WAS STATED NEXT MORNING THE PATIENT DID SEE SMALL AMOUNT OF BLOOD AND FIND IT UNHANDY TO USE ALSO STATED DID NOT REACH OUT TO A DOCTOR. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01323 | 2023-03-01 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT DURING USE IN THE WARD, LEAKAGE HAS OCCURRED. CHECKING THE CATHETER, THERE WAS A TEAR FOUND ON THE BALLOON. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01321 | 2023-03-01 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L20 |  | IT WAS REPORTED THAT THERE WAS A DAMAGE FOUND ON THE BALLOON DURING A PRETEST. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 24 EVERY 2FR 1. BALLOON CATHETER [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. THE SURFACE OF CATHETER IS COATED FIRST WITH A TRACE AMOUNT OF METALLIC SILVER, AND THEN WITH POLYURETHANE ONTO THAT, TO INHIBIT PROLIFERATION OF BACTERIA THAT MAY ADHERE TO THE CATHETER. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLE-LESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. 1) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) 2) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) 3) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01320 | 2023-03-01 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SIJ16 |  | IT WAS REPORTED THAT THE DRAINAGE FUNNEL PORTION OF THE FOLEY CATHETER WAS DIRTY. ALSO STATED THAT IT WAS DIRTY OR COLORED CATHETER. PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON 21FEB2023, STATED THAT THERE WAS A BLACK STAIN. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. EVALUATION ON THE CATHETER AND OBSERVED BLACK FOREIGN MATERIAL SCATTERED ON THE FUNNEL PART OF THE CATHETER. THE BLACK FOREIGN MATERIAL WAS EMBEDDED INSIDE FUNNEL DIPPING. THIS COMPLAINT IS CONFIRMED MANUFACTURING RELATED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE CROSS CONTAMINATION OF FUNNEL SLURRY DIP SOLUTION AND WATER DURING FUNNEL DIP PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01327 | 2023-03-01 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT URINE WAS LEAKING ON THE SIDE OF THE FOLEY CATHETER AND THE INCIDENT OCCURRED DURING USE. ALSO STATED THAT MEDI CARRIER GOT A REPORT THAT OF ALL OF THOSE TYPES OF CATHETERS WERE LEAKING. ALSO TRIED TO CHANGE SIZE BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAS BEEN AN ISSUE FOR A FEW MONTHS. MEDI CARRIER STATED THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE WAS FOR 12CH AND 14CH COULD BE MORE SIZES EFFECTED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO LOW OR HIGH BUILD UP GAUGE WHICH CAUSES THE DIAMETER OF THE CATHETER TO BECOME LARGE OR SMALL. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00582 | 2023-03-01 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT ONE YEAR ELEVEN MONTHS AND NINE DAYS POST PORT PLACEMENT PROCEDURE VIA RIGHT INTERNAL JUGULAR VEIN, THE CIRCUMFERENTIAL CRACKS WAS ALLEGEDLY FOUND NEAR TIP OF THE REMOVED CATHETER. IT WAS FURTHER REPORTED THAT WHEN ANTIEMETICS WERE ADMINISTERED, PAIN AND SWELLING WERE OBSERVED AROUND THE PORT BODY. THE DEVICE WAS REMOVED FROM THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY. | THE ASTRAL DEVICE WAS RETURNED TO RESMED FOR AN INVESTIGATION. REVIEW OF THE DEVICE DATA LOGS CONFIRMED THE REPORTED COMPLAINT, HOWEVER, PERFORMANCE TESTING COULD NOT REPRODUCE THE REPORTED COMPLAINT. BASED ON ALL AVAILABLE EVIDENCE AND COMPLAINT INVESTIGATIONS OF A SIMILAR NATURE, THE INVESTIGATION DETERMINED THE REPORTED COMPLAINT WAS DUE TO AN INTERMITTENT CONNECTION WITH THE BATTERY. RESMED¿S RISK ANALYSIS FOR THIS FAILURE MODE CONCLUDES THAT THE RISK IS ACCEPTABLE. RESMED REFERENCE#: (B)(4). | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01331 | 2023-03-01 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165L16 | 0165L16 |  | IT WAS REPORTED THAT THE STAFF WAS HAVING TROUBLE WITH THE FOLEY CATHETERS. THE CATHETERS WERE NOT DRAINING PROPERLY, AND THEY HAVE TO MOVE THE BAG TO GET THE URINE TO DRAIN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: SALT ACCUMULATION)/BLOCK DRAINAGE LUMEN/NO DRAINAGE EYE). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE 1) SECURE THE FOLEY CATHETER. USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED. 2) MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. 3) MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. 4) KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. 5) EMPTY THE COLLECTION BAG REGULARLY USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH. FOLEY CATHETER REMOVAL 1) TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TOMAKE THE SYRINGE ¿STICK¿ IN THE VALVE 2) ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY 3) USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION 4) IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL 5) SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01333 | 2023-03-01 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT URINE WAS LEAKING ON THE SIDE OF THE FOLEY CATHETER, AND THE INCIDENT OCCURRED DURING USE. ALSO STATED THAT MEDI CARRIER GOT A REPORT THAT THOSE TYPES OF CATHETERS WERE LEAKING. ALSO TRIED TO CHANGE SIZE BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAS BEEN AN ISSUE FOR A FEW MONTHS. MEDI CARRIER STATED THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE WAS FOR 12CH AND 14CH COULD BE MORE SIZES EFFECTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE BAD FIT WITH SHAFT/NO DRAINAGE EYE/ POORLY SEATED BALLOON /BLADDER NECK INTERFACE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX.. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01328 | 2023-03-01 | 2023-02-08 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK URINE COLLECTION SYSTEM FAILED TROUBLESHOOT WITH THE NEW CANISTER. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. PER ADDITIONAL INFORMATION RECEIVED VIA SAMPLE FORM ON 08FEB2023, IT WAS REPORTED THAT THE PATIENT DEVELOPED A URINARY TRACT INFECTION AND BELIEVE DUE TO THE FAILURE OF SYSTEM TO REMOVE URINE AND ENDED UP STUFFING PAPER TOWELS IN THEIR CROTCH TO HOLD IT. STATED THAT THE PHYSICIAN VISIT AND GAVE MEDICATION. ALSO, STATED THAT THE PUREWICK FEMALE EXTERNAL CATHETER WOULD NOT STAY POSITIONED, IF IT COULD BE IN THE FIRST PLACE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS: DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE. RECOMMENDATIONS: ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00584 | 2023-03-01 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7957505J |  | IT WAS REPORTED THAT THE CATHETER WAS BROKEN. THERE WAS NO REPORTED PATIENT INJURY. FEBRUARY 14, 2023 ¿ ADDITIONAL INFORMATION WAS RECEIVED. BROKEN CATHETER WAS DISCOVERED WHEN THE NURSE CAME TO CHECK IT OUT. THE SEVERED CATHETER HAS BEEN RECOVERED AND IS NOT IN A THREATENING SITUATION. THE PICC LINE WAS REPLACED. THERE WERE NO PATIENT COMPLICATIONS DUE TO THE FRACTURED CATHETER. CATHETER FRACTURE WAS OUTSIDE OF THE PATIENT¿S BODY. NO DEBRIS WERE RETAINED WITHIN THE PATIENT. NO SURGERY WAS REQUIRED TO REMOVE THE CATHETER WAS THE FRACTURE WAS OUTSIDE OF THE PATIENT¿S BODY, SIMPLY PULLED OUT THE CATHETER. PATIENT WAS RECEIVING NUTRIENTS THROUGH THE PICC. PICC WAS SECURED WITH STATLOCK AND TRANSPARENT DRESSING, THE AREA AROUND THE INSERTION SITE OF THE CATHETER WAS COVERED WITH DRESSING AT 35 CM INSERTION, BUT THE STATLOCK PART WAS NOT COVERED WITH DRESSING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BREAK IN THE CATHETER IS CONFIRMED AND WAS DETERMINED TO BE RELATED TO TENSILE FAILURE. ONE 5 FR D/L GROSHONG CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED LIGHT USE RESIDUES THROUGHOUT THE CATHETER, AND THE CATHETER WAS FRACTURED AT THE 54 CM DEPTH MARKER. A TACTILE OBSERVATION OF THE DEVICE REVEALED TENSILE WEAKNESS IN THE PROXIMAL SIDE OF THE BREAK WITH WHITENING OBSERVED CIRCUMFERENTIALLY ALONG THE CATHETER. A MICROSCOPIC OBSERVATION REVEALED A FINELY GRANULAR FRACTURE SURFACE AND WHITENING ON THE PROXIMAL PIECE OF THE BREAK ONCE TENSILE FORCES WERE APPLIED. THE COMPLAINT OF A BREAK IN THE CATHETER IS CONFIRMED AND WAS DETERMINED TO BE DUE TO TENSILE FAILURE. CATHETER SECUREMENT TECHNIQUES MAY HAVE CONTRIBUTED TO THIS FAILURE, AS THE BREAK OCCURRED AT A POTENTIAL SECUREMENT AREA FOR THE CATHETER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01326 | 2023-03-01 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE SIDE OF THE FUNNEL WHEN INJECTING. ALSO STATED THAT FUNNEL HAD A CRACK ON SIDE. PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON 22FEB2023, STATED THAT THE WATER LEAKED OCCURRED FROM THE FUNNEL PART. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. INSPECTED THE EXTERIOR OF THE SAMPLE AND FOUND TEAR ON THE INFLATION FUNNEL. THE TEAR MODE LOOKS LIKE ORIGINATED SHARP OBJECT FROM OUTSIDE AND STRETCHED OUT THROUGH THE LATEX LAYER UNTIL THE INNER SIDE, CREATING A CHANNEL. INFLATED BALLOON WITH WATER AND FOUND WATER LEAKING FROM THE TEAR. THE TEAR CAUSE THE WATER TO LEAK. EXACT CAUSE OF SAMPLE CONDITION COULD NOT BE DETERMINED AND COULD NOT ASSIGN A MANUFACTURING RELATED PROCESS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01317 | 2023-03-01 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT A KNOT WAS FOUND ON THE CATHETER UPON REMOVING THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A ROOT CAUSE FOR THIS EVENT COULD BE KINKED LUMEN. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01330 | 2023-03-01 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123620A | 123620A |  | IT WAS REPORTED THAT THE BALLOON WATER WAS UNABLE TO ESCAPE. THE PATIENT STATUS/ INTERVENTION WAS CUT OFF. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE THIN RUBBERIZE WALL THAT CAUSING COLLAPSE/PINCH INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT AND MIGHT BE CONTRIBUTED BY USER RELATED. EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON BURST. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE PATIENT USE ONLY. DO NOT REUSE AND RESTERILIZE. FOR UROLOGICAL USE ONLY. USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PROTOCOL, THE MAY BE CUT OFF. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01404 | 2023-03-02 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT PLACED AN INDWELLING FOLEY CATHETER FOR THE DRAINAGE OF URINE ON (B)(6), 2023. WHEN THE CATHETER WAS PLANNED TO BE REMOVED AT AROUND 8:00 ON (B)(6), 2023, THE LIQUID IN THE BALLOON COULD NOT BE DRAWN OUT AND THE PATIENT FELT UNCOMFORTABLE. THEN THE MIDDLE PART OF THE CATHETER WAS CUT OFF AND PULLED SLOWLY, SO THAT THE LIQUID IN THE BALLOON COULD FLOW OUT SLOWLY. THE CATHETER WAS PULLED OUT, AND THE PATIENT HAD NO SERIOUS ADVERSE EFFECTS EXCEPT THE DISCOMFORT DUE TO DELAYED REMOVAL. IF THE CATHETER COULD NOT BE REMOVED SUCCESSFULLY, IT MIGHT HAVE BEEN REMOVED WITH ASSISTANCE OF CYSTOSCOPY OR CYSTOTOMY, AND THE DAMAGE WOULD HAVE BE SERIOUS. AS PER DESCRIPTION OF EVENT CAUSE ANALYSIS, PRODUCT QUALITY WAS NOT UP TO STANDARD. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01406 | 2023-03-02 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS LEAKING URINE ON THE SIDE AND THE INCIDENT OCCURRED DURING USE. MED CARRIER GOT A REPORT THAT ALL OF THOSE TYPES OF CATHETERS WERE LEAKING. THEY HAVE TRIED TO CHANGE SIZE BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAD BEEN AN ISSUE FOR A FEW MONTHS AND THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE WAS FOR 12CH,14CH AND 16 (B)(6) COULD BE MORE SIZES EFFECTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE BUBBLE VOID ON RUBBERIZE LAYER. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01407 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE URINE LEAKAGE ON THE SIDE OF THE FOLEY CATHETER AND INCIDENT OCCURRED DURING USE. MEDI CARRIER GOT A REPORT THAT ALL OF THOSE TYPES OF CATHETERS WERE LEAKING. THEY HAVE TRIED TO CHANGE SIZE, BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAS BEEN AN ISSUE FOR A FEW MONTHS, AND THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE IS FOR 12CH, 14CH, AND 16 COULD BE MORE SIZES EFFECTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO WRONG PRODUCT SIZE SELECTED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00625 | 2023-03-02 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE GUIDE WIRE WAS NOT SMOOTH AND THERE WERE ABNORMAL BULGES ON IT. NO OTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGN0444 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01405 | 2023-03-02 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE URINE LEAKAGE ON THE SIDE OF THE FOLEY CATHETER AND THE INCIDENT OCCURRED DURING USE. MEDI CARRIER GOT A REPORT THAT ALL OF THOSE TYPES OF CATHETERS WERE LEAKING. THEY HAVE TRIED TO CHANGE SIZE BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAD BEEN AN ISSUE FOR FEW A MONTHS AND THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE WAS FOR 12CH , 14CH AND 16 COULD BE MORE SIZES EFFECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00613 | 2023-03-02 | 2023-01-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | IT WAS REPORTED THAT "DURING CATHETER PLACEMENT, BURRS WERE FOUND WITH THE STYLET INSIDE THE POWERPICC CATHETER AND THE SURFACE WAS NOT SMOOTH. ADDITIONALLY, FOREIGN MATTERS WERE ADHERED TO THE SURFACE OF THE GUIDE WIRE." IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF RESIDUES PRESENT ON THE STYLET IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. TWO OPEN SINGLE LUMEN 4 FR POWERPICC CATHETERS KITS AND CATHETERS WITH STIFFENING STYLET WERE RETURNED FOR EVALUATION. THE PRODUCT LABELS INDICATED LOT: REGQ4691. EACH SAMPLE WAS INSPECTED INDIVIDUALLY AND BOTH SAMPLES CONTAINED VISIBLE MATERIAL RESIDUES IN THE STYLET SECTIONS EXTENDING PROXIMAL TO THE YELLOW SEPTUM ON THE T-LOCK STYLET ASSEMBLIES. THE STYLETS WERE REMOVED FROM THE CATHETER IN ORDER TO INSPECT THE ENTIRE LENGTHS. THE RESIDUES WAS FOUND TO DISTRIBUTED THROUGHOUT THE COMPLETE LENGTH OF THE STYLET. MICROSCOPIC INSPECTION OF THE RESIDUES REVEALED IT TO BE A SOLID, OPAQUE MATERIAL. THE MATERIAL APPEARED TO BE THE HYDROPHILIC COATING APPLIED TO THE STYLETS DURING THE MANUFACTURING APPLICATION PROCESS. A REVIEW OF THE MANUFACTURING RECORDS DID NOT REVEAL ANY EVIDENCE TO SUGGEST A MANUFACTURING RELATED ROOT CAUSE. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE USE OF NON-BD COMPONENTS AND RETRACTION OF THE STYLET AGAINST RESISTANCE. SINCE RESIDUES WERE NOTED TO BE PRESENT ON THE STYLET, THE COMPLAINT IS CONFIRMED; HOWEVER, THE EXACT FACTORS RESULTING THE EVENT REMAIN UNKNOWN AT THIS TIME. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00612 | 2023-03-02 | 2023-01-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | IT WAS REPORTED THAT "DURING CATHETER PLACEMENT, BURRS WERE FOUND WITH THE STYLET INSIDE THE POWERPICC CATHETER AND THE SURFACE WAS NOT SMOOTH. ADDITIONALLY, FOREIGN MATTERS WERE ADHERED TO THE SURFACE OF THE GUIDE WIRE." IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGQ4691 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGQ4691) HAVE BEEN REPORTED FROM THE SAME FACILITY IN CHINA. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00628 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PER-Q-CATH PLUS 2F SINGLE-LUMEN PICC WITH STIFFENING STYLET BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4132105 |  | IT WAS REPORTED THAT THE CENTRAL VENOUS ACCESS CATHETER CAME WITH HOLES SO IT WAS NOT USED IN A PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAKING CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS FOUR PHOTOGRAPHS WHICH DEPICTED A 2FR S/L PER-Q-CATH CATHETER. THE STYLET WAS INLAID THROUGH THE CATHETER AND THE T-LOCK ASSEMBLY WAS ATTACHED TO THE LUER ADAPTER. THE PRE-FLUSH HANG TAG WAS VISIBLE ON THE STYLET. BLOOD RESIDUE WAS VISIBLE ON THE CATHETER. THE CATHETER SHAFT EXHIBITED SERPENTINE DEFORMATION NEAR THE MOLDED JOINT. THE CATHETER SHAFT APPEARED BUCKLED AND COMPRESSED. LEAKAGE WAS NOT OBVIOUS IN THE SUBMITTED PHOTOGRAPH; HOWEVER, THE SERPENTINE DEFORMATION WAS CONSISTENT WITH DAMAGE CAUSED BY STYLET MANIPULATION PRIOR TO WETTING, WHICH CAN RESULT IN LEAK-CAUSING CATHETER DAMAGE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01403 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE URINE LEAKAGE ON THE SIDE OF THE FOLEY CATHETER AND INCIDENT OCCURRED DURING USE. MEDI CARRIER GOT A REPORT THAT ALL OF THOSE TYPES OF CATHETERS WERE LEAKING. THEY HAVE TRIED TO CHANGE SIZE BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAS BEEN AN ISSUE FOR A FEW MONTHS AND THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE IS FOR 12CH,14CH, AND 16 COULD BE MORE SIZES EFFECTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE BAD FIT WITH SHAFT/NO DRAINAGE EYE/ POORLY SEATED BALLOON /BLADDER NECK INTERFACE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00623 | 2023-03-02 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD MIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT "HATES OUR INTRODUCER IN THE MIDLINE TRAYS. WE SHOULD MAKE THEM LIKE THE TELEFLEX INTRODUCER. HAS HAD ONE BUCKLE ON HIM AND HAS TO NICK ALL THE TIME. NEVER USED TO NEED TO ON OURS." ADDITIONAL INFORMATION PROVIDED (B)(6) 2023: THE PEEL AWAY PORTION OF THE BD IS RAISED. BECAUSE IT IS RAISED WHEN IT MEETS THE SKIN THERE IS A GREATER LEVEL OF FORCE REQUIRED TO PENETRATE THE PATIENTS SKIN. SOMETIMES THE FORCE REQUIRED IS SO GREAT THAT THE RAISED PORTION WILL CUP UP MAKING IT EVEN MORE DIFFICULT TO IMPOSSIBLE TO PENETRATE THE SKIN. IF BD ENGINEERS COULD DESIGN THE PEEL AWAY PORTION OF THE INTRODUCER SO IT IS FLUSH TO THE LONGER PORTION THAT THE PEEL AWAY PORTION COVERS IT WOULD REQUIRE LESS FORCE TO INSERT AND THEREFORE LESS PRESSURE AND DISCOMFORT FOR THE PATIENT. ALSO, IT WOULD REQUIRE LESS NICKS TO THE PATIENT'S SKIN DECREASING BLEEDING AND RISK FOR INFECTION." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00630 | 2023-03-02 | 2023-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000375 |  | IT WAS REPORTED THAT THE ORANGE NEEDLE IN PICC PACK BROKE AND PINGED UP NEARLY CAUSING A NEEDLE STICK INJURY WHILST BEING ACTIVATED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED HYPODERMIC NEEDLE IS CONFIRMED AND WAS DETERMINED TO BE SUPPLIER RELATED. THE PRODUCT RETURNED FOR EVALUATION WAS A 25GA HYPODERMIC NEEDLE SAFETY MECHANISM. MICROSCOPIC INSPECTION OF THE SAFETY MECHANISM REVEALED THE ADHESIVE ON THE NEEDLE SHAFT FAILED TO ADHERE TO THE PLASTIC ON THE LUER CONNECTOR, LEADING TO NEEDLE DETACHMENT. LIGHT USE RESIDUE WAS OBSERVED ON THE NEEDLE. THE ADHESIVE THAT FAILED TO ADHERE TO THE PLASTIC ON THE LUER CONNECTOR LIKELY OCCURRED DURING DEVICE MANUFACTURE. THIS COMPLAINT WILL BE FORWARDED TO THE MANUFACTURING SITE FOR FURTHER EVALUATION. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00611 | 2023-03-02 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT POWER MIDLINE CATHETER 4F SL | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001029 |  | IT WAS REPORTED BY THE CUSTOMER THAT "INTRODUCER WILL NOT PENETRATE THE PATIENT'S SKIN WITHOUT BENDING OR BREAKING. CLINICIAN HAD TO OPEN ANOTHER KIT." NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGX0443 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00617 | 2023-03-02 | 2022-12-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED ON JUNE 30, THE PICC CATHETER WAS PLACED IN OUR HOSPITAL SUCCESSFULLY, THE BLOOD EXTRACTION WAS NORMAL, THE BUILT-IN 39CM, THE PICC HEAD WAS POSITIONED AT T7, THE NURSES' EDUCATION AND MAINTENANCE AND DAILY PRECAUTIONS. THE CATHETER WAS MAINTAINED NORMALLY. ON (B)(6) 2022, AFTER THE TREATMENT OF THE PATIENT, THE PICC CATHETER WAS PULLED OUT IN THE INFUSION ROOM OF THE FACILITY. THE NURSE ROUTINELY DISINFECTED THE PATIENT'S WOUND AND EXTRACTED THE CATHETER. WHEN THE CATHETER WAS PULLED OUT 15CM, IT WAS FOUND THAT THE CATHETER WAS BROKEN AND ONLY 1/3 OF THE CATHETER WAS CONNECTED WITH THE TRANSVERSE SECTION, AND THE SECTION OF THE FRACTURE WAS NEAT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, COMPLAINT AND BATCH HISTORY, APPLICABLE PREVIOUS INVESTIGATION(S), LABELING, APPLICABLE MANUFACTURING RECORDS, SAMPLE ANALYSIS AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A CATHETER SPLIT IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE PHOTO SAMPLE OF A 4 FR POWERPICC SOLO CATHETER WAS RETURNED FOR EVALUATION. THE CATHETER IS IMPLANTED WITHIN THE PATIENT. THE CATHETER APPEARED TO HAVE BEEN RETRACTED UP TO THE 17 CM DEPTH MARKER. A CIRCUMFERENTIAL SPLIT IS NOTED IN THE CATHETER AT SECTION NEAR THE INSERTION SITE. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE MATERIAL FATIGUE CAUSED BY REPEATED KINKING AND DAMAGE DURING HANDLING. SINCE A SPLIT IN THE CATHETER WAS OBSERVED IN THE IMAGE, THE COMPLAINT IS CONFIRMED BUT THE EXACT CONTRIBUTING FACTORS REMAIN UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00619 | 2023-03-02 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER 5F RADIOLOGY BASIC TRAY (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3295355 |  | IT WAS REPORTED THE PATIENT REQUESTING ASSISTANCE IDENTIFYING A 5CM CLEAR "SHEATH" ON THE OUTSIDE OF HER REMOVED POWERPICC SOLO CATHETER. STATES SHE IS IMMUNOCOMPROMISED AND HAS BEEN GETTING HER PICC REMOVED AND REPLACED EVERY 84 DAYS. STATES SHE IS ALSO AN EPIDEMIOLOGIST AND HAS BEEN RECEIVING IVIG INFUSIONS FOR MANY YEARS. SHE IS CONCERNED ON WHAT COULD BE ON THE OUTSIDE OF THE CATHETER AND IF IT DUE TO USER ERROR DURING PLACEMENT AS SHE HAS NEVER SEEN THIS BEFORE. HAS NO COMPLAINTS OR SYMPTOMS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A FOREIGN MATERIAL ON THE CATHETER WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. A SERIES OF FOUR PHOTOGRAPHS OF THE IMPLICATED 5FR D/L POWERPICC SOLO CATHETER WERE RETURNED FOR EVALUATION. A CLEAR FOREIGN MATERIAL WAS SEEN AROUND THE POWERPICC CATHETER SHAFT. THE FOREIGN MATERIAL APPEARED TO BE APPROXIMATELY 5CM LONG AND WAS LOCATED FROM ABOUT THE 13CM DEPTH MARKING THROUGH THE 18CM DEPTH MARKING. THE THICKNESS OF THE CLEAR MATERIAL APPEARED UNEVEN. THE ENDS OF THE CLEAR MATERIAL CONTAINED SOME REDDISH DISCOLORATION, AND A FIBROUS TEXTURE WAS SEEN ON ONE END OF THE MATERIAL. WITHOUT RECEIPT OF THE PHYSICAL SAMPLE, THE FOREIGN MATERIAL COULD NOT BE CONCLUSIVELY IDENTIFIED. POSSIBLE CONTRIBUTING FACTORS FOR THE FOREIGN MATERIAL COULD INCLUDE CATHETER HANDLING (E.G., CONTAMINATION ON THE CATHETER) OR BIOLOGICAL MATERIAL ON THE CATHETER (E.G., FIBRIN SHEATH). A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00626 | 2023-03-02 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE GUIDE WIRE WAS NOT SMOOTH AND THERE WERE ABNORMAL BULGES ON IT. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGN0444 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00632 | 2023-03-02 | 2023-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 18G X 10 CM WITH GUARDIVA AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F118108PT |  | IT WAS REPORTED BY THE CUSTOMER THAT "THE INSERTION OF THE CATHETER "FELT FUNKY - WORKED FINE BUT SAFETY DIDN'T DEPLOY - IT'S STILL UP IN THE HOUSING." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED: PLACED A POWER GLIDE, AND WHEN PULLING BACK ON THE HOUSING TO DISENGAGE FROM THE DEVICE, THE SAFETY DID NOT COVER NEEDLE, IT STAYED IN HOUSING. " | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE ANALYSIS, PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF SAFETY MECHANISM FAILURE TO ADVANCE WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 18GA X 10CM POWERGLIDE PRO MIDLINE CATHETER ASSEMBLY. THE CATHETER HAD BEEN ADVANCED AND WAS NOT RETURNED FOR EVALUATION. THE CATHETER ADVANCER WINGS WERE ALSO NOT RETURNED. THE SAFETY REMAINED AT THE BASE OF THE NEEDLE SHAFT AND WAS NOT ENGAGED OVER THE NEEDLE TIP. MICROSCOPIC INSPECTION OF THE NEEDLE SHAFT AND THE SAFETY MECHANISM WAS UNREMARKABLE. AN ATTEMPT TO ADVANCE THE SAFETY MECHANISM WAS SUCCESSFUL AND UNREMARKABLE. NO RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE SAFETY AND IT SECURED OVER THE NEEDLE TIP AS INTENDED. THE SAFETY MECHANISM FAILED TO ADVANCE WITH THE CATHETER AS INTENDED; HOWEVER, NO RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE SAFETY DURING SAMPLE EVALUATION. THESE OBSERVATIONS SUGGESTED THAT THE SAFETY WAS NOT PROPERLY ENGAGED WITH THE CATHETER ADVANCER PRIOR TO CATHETER PLACEMENT; HOWEVER, THE ADVANCER ITSELF WAS NOT AVAILABLE FOR INSPECTION AND NO OTHER EVIDENCE WAS OBSERVED TO INFORM HOW THE SAFETY MECHANISM AND CATHETER ADVANCER BECAME DETACHED. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00634 | 2023-03-02 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGROSHONG CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000517 |  | THE CUSTOMER REPORTED THAT THE VALVED POWERPICC BROKE, IN 2 PLACES, AT 30 CM (THEREFORE INSIDE THE VESSEL) WITH CHEMO IN PROGRESS. THE PATIENT DID NOT HAVE THE EXTRAVASATION. CATHETER WAS INSERTED 3 WEEKS AGO IN BASILICA VEIN ON THE RIGHT SIDE WITH ULTRASOUND GUIDED POSITIONING WITH ECG MONITORING AND RADIOGRAPHIC CONTROL POST INSERTION. PICC WAS INSERTED 40CM. IT WAS NOT A DIFFICULT PLACEMENT, WITHOUT COMPLICATIONS DURING IMPLANTATION OR IMMEDIATE POST-IMPLANTATION. DAMAGED PICC WAS REMOVED AND REPLACED WITH A 4FR POWERPICC SOLO ON (B)(6) 2023. ADDITIONAL INFORMATION RECEIVED 2/16/2023: CUSTOMER REPORTS THERE WAS A "DOUBLE FISSURE (NON-FRACTURE) UNDER THE SKIN" APPROXIMATELY 10CM FROM THE INSERTION POINT. NO PIECE WAS RETAINED IN THE PATIENT. THERE WAS NO INJURY OR COMPLICATION, AND THE EVENT DID NOT RESULT IN A CLINICALLY SIGNIFICANT DELAY IN TREATMENT. THE EVENT DID NOT OCCUR DURING AN URGENT OR LIFE-THREATENING MEDICAL SITUATION. THE CATHETER HAD BEEN SECURED WITH SECURACATH. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00635 | 2023-03-02 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED BY THE CUSTOMER, "SEVEN DAYS AFTER INSERTION LINE KINKED. ON (B)(6) 2023 DISCUSSED WITH UNIT RN PICC LINE NOW DRAWING BLOOD AND VERY HARD TO FLUSH. PICC WAS DECLOTTED ON (B)(6) 2023. PATIENT AND MOM STATE LINE IS VERY POSITIONAL AND OCCLUDING OFTEN. RECOMMEND CHEST XRAY WITH LEFT ARM INCLUDED TO CHECK FOR KINKS IN LINE. PICC WAS REWIRED." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED ON 2/28/2023: IT WAS REPORTED BY THE CUSTOMER "CATHETER WAS SECURED WITH SECURACATH." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00633 | 2023-03-02 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER (WITH 70CM NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174355 |  | THE CUSTOMER REPORTED A PRESSURE PROBLEM WITH POWER PICC. A CT STAFF MEMBER REPORTED THAT THERE ARE PROBLEMS WITH CONTRAST INJECTIONS WITH ULTRAVIST® AND PICC. THE CONTRAST MEDIUM IS VERY STICKY. THE INFUSION LINE FROM THE INJECTOR IS SCREWED TO THE NEEDLE-FREE CONNECTOR. FIRST RINSE WITH NACL 0.9% AND THEN INJECT A VOLUME OF 50-300ML OF THE CONTRAST MEDIUM UNDER PRESSURE. THE FLOW IS 4ML/SEC. AFTER A FEW SECONDS THE PRESSURE IN THE SYSTEM RISES TO 8BAR AND THE DEVICE REGULATES. HOWEVER, SOMETHING ALWAYS CAME OUT OF THE PICC AND I DID NOT HAVE ANY KINK IN IT. I TESTED WITH AND WITHOUT A NEEDLE-FREE CONNECTOR, ALMOST UNCHANGED VALUES. ADDITIONAL INFORMATION: THE CATHETER DOESN¿T LEAK AND WITH NACL THERE WAS NO PROBLEM. BUT WITH THE CONTRAST, THE PUMP SHOWED AFTER A MOMENT A HIGH PRESSION AND STOP TO WORK. THE MACHINE WAS PUT ON 4ML/SECOND, BUT EVERY TIME THEY TRIED IT DOESN¿T WORK MEANS THE PRESSION GOT HIGHER AND HIGHER AND THEN STOPPED THE PROCEDURE. THE PICC IS IN THE PATIENT. THE SOLUTION FOR THE PROCEDURE WAS THEN, THAT THE PATIENT GOT AND PVC LIKE VENFLON PRO SAFETY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF AN OCCLUSION WARNING DURING INFUSION AT 4ML/SEC WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. TWO PHOTOGRAPHS OF THE SCREEN ON A POWER INJECTOR WERE RETURNED FOR EVALUATION. A WARNING FOR AN OCCLUSION WAS SEEN ON ONE OF THE SCREENS WHILE THE OTHER SCREEN SHOWED THAT THE PRESSURE WAS REACHING APPROXIMATELY 8 BARS WHILE INFUSED AT A RATE OF 4ML/SEC. IT WAS REPORTED THAT SOME CONTRAST WAS SEEN INFUSING THROUGH THE CATHETER DURING POWER INJECTION, WHICH WOULD INDICATE EITHER A PARTIAL OCCLUSION OR A BUILDUP OF PRESSURE DUE TO THE INFUSATE. THE SUBMITTED PHOTOGRAPH DID NOT CONTAIN ENOUGH INFORMATION TO IDENTIFY A SPECIFIC ROOT CAUSE. POSSIBLE CONTRIBUTING FACTORS COULD INCLUDE A KINK/BEND IN THE CATHETER, A RESTRICTION IN ONE OF THE DEVICES (E.G. A CONNECTOR, EXTENSION TUBING, CATHETER, ETC.), A PARTIAL OCCLUSION (EITHER BIOLOGICAL OR PRECIPITATE) IN THE CATHETER OR AUXILIARY DEVICE, OR A VISCOUS INFUSATE. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE IFU STATES, ¿VIGOROUSLY FLUSH THE POWERPICC® CATHETER USING A 10 ML OR LARGER SYRINGE AND STERILE NORMAL SALINE PRIOR TO AND IMMEDIATELY FOLLOWING THE COMPLETION OF POWER INJECTION STUDIES. IN ADDITION, LOCK EACH LUMEN OF THE CATHETER WITH HEPARINIZED SALINE. USUALLY ONE ML PER LUMEN IS ADEQUATE. THIS WILL ENSURE THE PATENCY OF THE POWERPICC® CATHETER AND PREVENT DAMAGE TO THE CATHETER. RESISTANCE TO FLUSHING MAY INDICATE PARTIAL OR COMPLETE CATHETER OCCLUSION. DO NOT PROCEED WITH POWER INJECTION STUDY UNTIL OCCLUSION HAS BEEN CLEARED.¿ AND ¿CONTRAST MEDIA SHOULD BE WARMED TO BODY TEMPERATURE PRIOR TO POWER INJECTION. WARNING: FAILURE TO WARM CONTRAST MEDIA TO BODY TEMPERATURE PRIOR TO POWER INJECTION MAY RESULT IN CATHETER FAILURE.¿ H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00610 | 2023-03-02 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PPICC SOLO 5F DL FT 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | IT WAS REPORTED BY THE CUSTOMER THAT "INTRODUCER WILL NOT PENETRATE SKIN WITHOUT BENDING OR BREAKING." CLINICIAN OPENED A NEW INSERTION KIT. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGW2650 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00636 | 2023-03-02 | 2023-01-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC 4F SL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT "WITH PICC INSERTION: SECOND ATTEMPT LEFT BRACHIAL VEIN CATHETER INSERTED WITHOUT DIFFICULTY BUT FLUSHED WITH RESISTANCE AND WAS TIGHT TO REPOSITION SO REMOVED." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED ON 2/28/2023: CATHETER SECURED WITH SECURACATH. A NEW PICC WAS PLACED. PATIENT DIAGNOSIS WAS DETERMINING BETWEEN APLASTIC ANEMIA VS. LEUKEMIA VS. VIRAL PROCESS ADDITIONAL INFORMATION RECEIVED 5/25/2023: CATHETER INSERTED 9CM ABOVE THE ANTECUBITAL FOSSA AND 4.5CM LEFT EXTERNAL TO THE PATIENT. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00126 | 2023-03-02 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRUE DILATATION CATHETER | BALLOON VALVULOPLASTY CATHETER | BARD PERIPHERAL VASCULAR, INC. | 0184511 | 0184511 |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE, THE BALLOON ALLEGEDLY SUCKED IN AIR ON PREPARATION. THERE WAS NO PATIENT CONTACT. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 08/2025), G3. H11: B5, H6 (PATIENT, DEVICE). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00122 | 2023-03-02 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DORADO | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | DR4064 | DR4064 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE VIA AN IPSILATERAL ANTEGRADE APPROACH, THE PTA BALLOON ALLEGEDLY RUPTURED AT 18 ATM ON THE FIRST INFLATION ATTEMPT. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE DORADO PTA DILATATION CATHETER WAS RETURNED FOR EVALUATION. ON VISUAL EVALUATION OF THE DEVICE, NO OTHER SPECIFIC ANOMALIES WERE NOTED. ON FUNCTIONAL EVALUATION, THE DEVICE WAS INFLATED WITH AN IN-HOUSE PRESTO INFLATION DEVICE AT 16 ATM, WITH LEAKS NOTED AT THE CATHETER BOND JOINT. IT WAS FURTHER OBSERVED UNDER THE MICROSCOPE THAT A CIRCUMFERENTIAL BREAK WAS NOTED AT THE PROXIMAL GLUE JOINT, WHICH CAUSED THE LEAK. NO EVIDENCE OF BALLOON RUPTURE WAS NOTED DUE TO THE CONDITION OF THE DEVICE, AND FUNCTIONAL TESTING COULDN¿T BE FURTHERED DUE TO THE BOND JOINT FAILURE. THEREFORE, THE INVESTIGATION FOR THE REPORTED BALLOON RUPTURE REMAINS INCONCLUSIVE, AS WATER LEAKED FROM THE BOND JOINT DURING THE EVALUATION AND NO EVIDENCE OF A BALLOON RUPTURE COULD BE NOTED DUE TO THE CONDITION OF THE DEVICE. THE INVESTIGATION WAS CONFIRMED FOR THE IDENTIFIED BOND JOINT BREAK BECAUSE WATER LEAKED FROM THE CATHETER BOND JOINT DURING FUNCTIONAL TESTING. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BALLOON RUPTURE AND IDENTIFIED BOND JOINT BREAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRATION DATE: 06/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00123 | 2023-03-02 | 2022-07-18 | I, n, j, u, r, y | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80144 | ATG80144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE BALLOON ALLEGEDLY SHEARED OFF COMPLETELY DURING INFLATION IN THE PATIENT'S SUBCLAVIAN VEIN AT 16 ATM AND THE ENTIRETY OF THE BALLOON GOT SEPARATED FROM THE SHAFT, AND THE BALLOON GOT STUCK IN TO PATIENTS VEIN. IT WAS FURTHER REPORTED THAT, A STENT WAS PUT TO PREVENT THE SHEARED OFF BALLOON FROM MIGRATING FURTHER INTO THE CIRCULATORY SYSTEM AND THIS HAS RESULTED IN AN OCCLUSION OF THE SUBCLAVIAN VEIN. REPORTEDLY, THE PATIENT WAS DIAGNOSED WITH CONSTANT BLOOD CLOTS WHICH MIGHT REQUIRE ADDITIONAL SURGICAL INTERVENTION. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE ATLAS GOLD DILATATION CATHETER WAS RETURNED FOR EVALUATION. THE BALLOON WAS NOTED TO HAVE DETACHED FROM THE CATHETER AND WAS NOT RETURNED FOR EVALUATION. A GLUE BULLET WAS NOT NOTED SEATED IN THE POSITION. NO OTHER SPECIFIC ANOMALIES WERE NOTED DURING THE VISUAL EVALUATION. NO FURTHER TESTING WAS PERFORMED DUE TO THE NATURE OF THE COMPLAINT. AS A RESULT, THE BALLOON DETACHMENT INVESTIGATION WAS CONFIRMED, AS THE BALLOON WAS COMPLETELY DETACHED DURING THE VISUAL EVALUATION AND WAS NOT RETURNED FOR EVALUATION. HOWEVER, THE INVESTIGATION REMAINS INCONCLUSIVE FOR THE REPORTED BALLOON ENTRAPMENT AS THE CONDITIONS OF USE COULDN¿T BE REPRODUCED UNDER LABORATORY CONDITIONS. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BALLOON DETACHMENT AND BALLOON ENTRAPMENT COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00124 | 2023-03-02 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQF7574 | CQF7574 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: B5, D4 (EXPIRATION DATE: 07/2025), G3. H11: B3. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00604 | 2023-03-02 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. ULTRA SLIMPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0605640 | 0605640 |  | IT WAS REPORTED THAT FOUR MONTHS AND TWENTY SIX DAYS POST PORT PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY DETACHED. IT WAS FURTHER REPORTED THAT, THE DEVICE WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE BARD MRI ULTRA SLIMPORT IMPLANTABLE PORT, ONE CATHETER AND ONE CATH-LOCK WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE, DIMENSIONAL AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. PRODUCT PACKAGING AND LABEL WERE NOT RETURNED. A COMPLETE DIAGONAL BREAK WAS NOTED ON THE PROXIMAL END OF THE CATHETER RETURNED. THE EDGES OF THE COMPLETE DIAGONAL BREAK ON THE PROXIMAL END OF THE RETURNED CATHETER WERE NOTED TO BE UNEVEN AND THE SURFACE WAS NOTED TO BE GLOSSY WITH STRIATIONS THROUGHOUT. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED CATHETER DISCONNECTION ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT ARE UNKNOWN. BASED ON THE MEASUREMENTS RECORDED DURING SAMPLE EVALUATION AND APPLICABLE DRAWINGS, NO MEASUREMENTS RECORDED COULD BE CONFIRMED TO BE OUT OF TOLERANCE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 05/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01364 | 2023-03-02 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 129416M | 129416M |  | IT WAS REPORTED THAT THE TEMPERATURE SENSING CATHETER SENSING FUNCTION WAS NOT WORKING VERY LONG AND ALSO THE TUBING WAS COMING IN KINKED RIGHT AT WHERE IT CONNECTS TO THE URINE METER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "SENSOR WIRE TOO WEAK". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: " VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. PROPER TECHNIQUES FOR URINARY CATHETER INSERTION: INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. NOTE: COMPATIBLE WITH APPROPRIATE 400-SERIES TEMPERATURE MONITORS. INTERCHANGEABILITY +0.2 DEG C AT 37 DEG C. AS WITH ALL TEMPERATURE PROBES , IN THE PRESENCE OF RF ENERGY SOURCES, LOCAL HEATING, TEMPERATURE ERRORS AND PROBE DAMAGE MAY OCCUR. IN MEDICAL USE, UNPLUG THE TEMPERATURE SENSING CATHETER AT THE TEMPERATURE MONITOR BEFORE ACTIVATING ELECTROSURGICAL OR OTHER TYPES OF DIRECT COUPLED RF ENERGY SOURCES." | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01367 | 2023-03-02 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | SILICON FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165814 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON/CUFF DIFFUSES AND SHRINKS AFTER SHORT PERIOD. ALSO STATED THAT THE BALLOON DIFFUSES WITH 2-6ML DURING A PERIOD OF APPROXIMATELY 2 WEEKS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01369 | 2023-03-02 | 2022-12-29 | I, n, j, u, r, y | Manufacturer report | NBIH¿ TEMPORARY PACING ELECTRODE CATHETER | TEMPORARY PACING ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 007406P | 007406P |  | IT WAS REPORTED THAT FOLLOWING TRANSCATHETER IMPLANTATION OF AN AORTIC VALVE, THE PATIENT EXPERIENCED AN ATRIOVENTRICULAR BLOCK, AND A TEMPORARY PACING ELECTRODE (TPE) CATHETER WAS PLACED. THE TPE CATHETER USED WAS A LARGER CALIBER THAN USUAL (CONTEXT OF LABORATORY RUPTURE). A 7F WAS USED INSTEAD OF THE USUAL 5F, BUT THE PATIENT WAS OKAY AT THE TIME OF DISCHARGE FROM THE BLOCK. THERE WAS TAMPONADE AT DAY 1 FOLLOWING THE TRANSCATHETER IMPLANTATION OF THE AORTIC VALVE WITH PRESENCE OF A FIBRILLATION ARTIALE DROITE (AF D; RIGHT-SIDED ATRIAL FIBRILLATION) BREACH. THE PATIENT WAS TRANSFERRED TO THE OPERATING ROOM 3 HOURS LATER FOR THE REMOVAL OF THE PROBE AND TREATMENT. THERE WAS AN UNFAVORABLE SITUATION ON (B)(6) 2022 AT 7 PM, AND THE PATIENT EXPIRED. IT WAS REPORTED THAT A POTENTIAL CAUSE OF THE PATIENT¿S DEATH WAS RELATED TO THE STIMULATION PROBE. 7F STIMULATION PROBES ARE NO LONGER USED AT THIS FACILITY. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL FROM THE INTERNATIONAL BUSINESS CENTERS (IBC) ON 03-FEB-2023, IT WAS REPORTED THAT THE DEVICE SEEMED VERY RIGID TO THE SURGEON. ADDITIONAL CLINICAL FOLLOW UP INFORMATION WAS RECEIVED VIA EMAIL ON 23-FEB-2023 FROM THE IBC. FOLLOWING THE TRANSCATHETER IMPLANTATION OF AN AORTIC VALVE, THE PATIENT EXPERIENCED THE ONSET OF HIGH-GRADE ATRIOVENTRICULAR BLOCK, AND A TPE CATHETER WAS PLACED. A 7F CATHETER HAD TO BE USED BECAUSE THE FACILITY DID NOT HAVE ANY 5F TPE CATHETERS AVAILABLE. A 7F WAS OBTAINED FOR USE FROM ANOTHER ESTABLISHMENT. FOLLOWING PLACEMENT OF THE TPE CATHETER, THE PATIENT WAS DISCHARGED TO THE INTENSIVE CARE UNIT (ICU). SEVERAL HOURS AFTER THE DEVICE WAS PLACED, THE PATIENT EXPERIENCED THE ONSET OF SHOCK WITH TAMPONADE, WHICH REQUIRED EMERGENCY BYPASS SURGERY FOR REPAIR OF A 2CM RIGHT VENTRICULAR WOUND. ON (B)(6) 2022, THE PATIENT EXPIRED, AND THE DEATH WAS REPORTED AS A RESULT OF MULTIVISCERAL FAILURE AT DAY 1. THE FOLLOWING IS LISTED AS THE CAUSE OF DEATH ON THE PATIENT¿S DEATH CERTIFICATE: OCCURRENCE OF TAMPONADE, EMERGENCY SURGICAL DRAINAGE, REFRACTORY SHOCK COMPLICATED BY POLYVISCERAL FAILURE, ASYSTOLE. PER ADDITIONAL CLINICAL FOLLOW UP INFORMATION RECEIVED VIA EMAIL ON 16-MAR-2023 FROM THE IBC, THERE WERE NO DIFFICULTIES ENCOUNTERED WHILE INSERTING THE TPE DEVICE INTO THE RIGHT VENTRICULAR POINT. THE PATIENT LATER EXPERIENCED TAMPONADE DUE TO A 2CM RIGHT VENTRICULAR WOUND. IT WAS REPORTED THE WOUND POSSIBLY OCCURRED DUE TO BOTH THE RIGIDITY AND SIZE OF THE TPE DEVICE. | THE REPORTED EVENT IS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED AND FURTHER INVESTIGATION WAS NOT CONCLUSIVE. NO CHANGES ON MANUFACTURING PROCESS CONTRIBUTING TO THE EVENT REPORTED WERE NOTED. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT POTENTIAL ROOT CAUSES COULD BE, "LATEX MODULUS TOO HIGH", "OPERATOR ERROR", "P-20 ERROR, (DUE TO MISHANDLING), OPERATOR / SET-UP / MIXER, ERROR". THE DEVICE WAS USED FOR TREATMENT PURPOSES. AS THE DEVICE WAS NOT RETURNED IT IT UNKNOWN IF IT HAD MET ALL RELEVANT SPECIFICATIONS OR RESULTED IN THE REPORTED EVENT. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® TEMPORARY PACING CATHETERS ARE DESIGNED TO TRANSMIT AN ELECTRICAL SIGNAL FROM AN EXTERNAL PULSE GENERATOR TO THE HEART OR FROM THE HEART TO A MONITORING DEVICE. WHEN AN INTERNAL LUMEN IS PRESENT (OTHER THAN THE ONE USED FOR BALLOON INFLATION), IT MAY BE USED FOR FLUID INFUSION, PRESSURE MONITORING, OR BLOOD SAMPLING. EXCESSIVE BENDING, TORQUING, OR KINKING OF THE ELECTRODE CATHETER MAY CAUSE DAMAGE TO THE CATHETER INCLUDING DAMAGE TO INTERNAL WIRES. WHEN USING A BALLOON CATHETER, USE CARE WHEN REMOVING THE PROTECTIVE SLEEVE FROM THE DISTAL PORTION OF THE CATHETER. FORCED REMOVAL OF THIS PROTECTIVE SLEEVE MAY RESULT IN DAMAGE TO THE BALLOON AND THE CATHETERS STRUCTURAL INTEGRITY. INSPECTION INSTRUCTIONS: INSPECT THE STERILE PACKAGE CAREFULLY FOR DAMAGE DURING TRANSIT OR STORAGE. DO NOT USE THE CATHETER IF THE PACKAGE IS DAMAGED. VISUALLY INSPECT THE CATHETER, UNDER STERILE CONDITIONS, FOR KINKS IN THE CATHETER SHAFT, INTEGRITY OF THE CONNECTOR, CONDITION OF ELECTRODES, AND ANY OTHER DAMAGE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01372 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS PLUS CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE CUSTOMER WAS HAVING ISSUES WITH THE INTERMITTENT CATHETER NOT RELEASING THE URINE. CUSTOMER FELT LIKE IT WAS BLOCKED, AND THE URINE WOULD NOT COME OUT. CUSTOMER ALSO MENTIONED THAT THEY EXPERIENCED DIFFICULTY INSERTING THE CATHETER ALL THE WAY. STATED THAT THEY WERE RECEIVED 2 IODINE IN ONE OF THE KITS AND RECEIVED A KIT WHERE THERE WAS NO IODINE IN IT. CUSTOMER RECEIVED INFECTION TWICE IN ONE YEAR DUE TO REUSING THE HOSE OR RUNNING OUT OF SUPPLY. CUSTOMER COULD NOT CONFIRM IF IT WAS A BD PRODUCT AND SAID IT WAS A LONG WHILE AGO AND COULD NOT REMEMBER. ALSO STATED CUSTOMER WIFE AND DAD USED THIS PRODUCT AS WELL AND EXPERIENCED SOME ISSUES. NO SPECIFICATION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01373 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT AFTER DRAINING THE WATER FROM THE BALLOON, IT WAS DIFFICULT TO PULL IT OUT, BUT AFTER PULLING IT OUT WITH FORCE, THERE WAS SOMETHING LIKE FOLDS AT THE BALLOON PART. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01374 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE CUSTOMER WAS HAVING ISSUES WITH THE INTERMITTENT CATHETER NOT RELEASING THE URINE. CUSTOMER FELT LIKE IT WAS BLOCKED, AND THE URINE WOULD NOT COME OUT. THE CUSTOMER ALSO MENTIONED THAT THEY EXPERIENCED DIFFICULTY INSERTING THE CATHETER ALL THE WAY. STATED THAT THEY WERE RECEIVED 2 IODINE IN ONE OF THE KITS AND RECEIVED A KIT WHERE THERE WAS NO IODINE IN IT. CUSTOMER RECEIVED INFECTION TWICE IN ONE YEAR DUE TO REUSING THE HOSE OR RUNNING OUT OF SUPPLY. CUSTOMER COULD NOT CONFIRM IF IT WAS A BD PRODUCT AND SAID IT WAS A LONG WHILE AGO AND COULD NOT REMEMBER. ALSO STATED CUSTOMER WIFE AND DAD USED THIS PRODUCT AS WELL AND EXPERIENCED SOME ISSUES. NO SPECIFICATION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT OPERATION". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. BASED ON PATIENT CODE (B)(6) THE PRODUCT DOES NOT APPEAR TO HAVE BEEN USED. HOWEVER, THE PRODUCT IS INTENDED TO BE USED FOR TREATMENT PURPOSE, BUT IT IS UNKNOWN IF THERE IS A RELATIONSHIP BETWEEN THE REPORTED EVENT AND THE DEVICE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW IS NOT REQUIRED BECAUSE A REVIEW OF THE LABEL COULD NOT HAVE PREVENTED THIS ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01380 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS LEAKING URINE ON THE SIDE AND THE INCIDENT OCCURRED DURING USE. MED CARRIER GOT A REPORT THAT ALL OF THOSE TYPES OF CATHETERS WERE LEAKING. THEY HAVE TRIED TO CHANGE SIZE BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAD BEEN AN ISSUE FOR A FEW MONTHS AND THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE WAS FOR 12CH,14CH AND 16 COULD BE MORE SIZES EFFECTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE POOR BALLOON SHAPE. A REVIEW OF THE DEVICE LABELING HAVE BEEN CONDUCTED FOR INVESTIGATION PURPOSE. A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01385 | 2023-03-02 | 2023-02-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT AFTER THE PATIENT WAS PLACED IN THE INTENSIVE CARE UNIT, A CRACK WAS FOUND IN THE CATHETER WHEN MOVED TO THE HIGH CARE UNIT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE.A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE MANUFACTURING RELATED DUE TO OPERATOR ERROR (NO TIP GLUING). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01387 | 2023-03-02 | 2023-02-10 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTIONS. THE PATIENT DID NOT ALLEGE THE PUREWICK FEMALE EXTERNAL CATHETER AS THE CAUSE. DOCTOR ADVISED PATENT TO START BACK AFTER THE URINARY TRACT INFECTION CLEARED. THE PATIENT HAD BEEN USING THIS PRODUCT FOR MORE THAN 90 DAYS. IT WAS UNKNOWN IF THE PUREWICK FEMALE EXTERNAL CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: EXPERIENCING SKIN IRRITATION BREAKDOWN AT THE SITE. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01390 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS THAT THE CUSTOMER HAD ANOTHER ISSUE WITH A SILICONE FOLEY CATHETER BEING REMOVED. REGISTER NURSE STATED THE BALLOON DID NOT APPEAR TO BE COMPLETELY DEFLATED AND A RIDGE WERE NOTED IN THE BALLOON UPON REMOVAL. THERE WAS NO BLOOD NOTED BUT PATIENT EXPERIENCED PAIN ON URINATION. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. ALLOW THE BALLOON TO DEFLATE SLOWLY ON ITS OWN. DO NOT ASPIRATE OR MANUALLY ACCELERATE THE DEFLATION OF THE BALLOON. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01388 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS LEAKING URINE ON THE SIDE AND THE INCIDENT OCCURRED DURING USE. MED CARRIER GOT A REPORT THAT ALL OF THOSE TYPES OF CATHETERS WERE LEAKING. THEY HAVE TRIED TO CHANGE SIZE, BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAD BEEN AN ISSUE FOR A FEW MONTHS, AND THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE WAS FOR 12CH (UNK),14CH AND 16CH (MYDP6271) COULD BE MORE SIZES EFFECTED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. VISUAL AND FUNCTIONAL INSPECTION OBSERVED THE EXTERIOR AND FUNCTIONALITY OF THE SAMPLE AND DID NOT FIND ANY EVIDENCE OF A FAILURE THAT WOULD SUPPORT THE REPORTED EVENT. THE CATHETER SHAFT MEASUREMENT MET SPECIFICATION. BASED ON THE SAC MANUFACTURING PROCESS, A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE LOW LATEX VISCOSITY. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01389 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE URINE LEAKAGE ON THE SIDE OF THE FOLEY CATHETER AND THE INCIDENT OCCURRED DURING USE. MEDI CARRIER GOT A REPORT THAT ALL OF THOSE TYPES OF CATHETERS WERE LEAKING. THEY HAVE TRIED TO CHANGE SIZE BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAD BEEN AN ISSUE FOR FEW A MONTHS AND THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE WAS FOR 12CH, 14CH AND 16 COULD BE MORE SIZES EFFECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00608 | 2023-03-02 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN HEMODIALYSIS/APHERESIS CATHETER, DUAL-LUMEN,13.5F | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 0600692 | 0600692 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE PEEL AWAY SHEATH ALLEGEDLY FAILED TO LOCK AT THE HUB WHERE THE BLUE AREA CONNECTS. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 12/2023). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01396 | 2023-03-02 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC 3 GO MALE COUDE CATHETER | COUDE INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD SOME BLEEDING WITH THE INTERMITTENT CATHETER. PATIENT STATED THE CATHETERS WERE TURNED LIKE A HOCKEY STICK AND THEY MADE THEM BLEED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR'. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, AND IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01402 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPECIALTY TUBES & CATHETERS | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50000000 | UNK |  | IT WAS REPORTED THAT THE ARCTIC SUN DEVICE WAS COOLING PATIENT TO 36C AND HAD FOLEY PROBE IN PLACE. FOLEY PROBE WAS NOT READING AT ALL, SO THEY MOVED TO A RECTAL PROBE. THEY HAVE AN ESOPHAGEAL PROBE 36.4C AND A RECTAL PROBE 37.4C. THEY WERE SWAPPING BACK AND FORTH BETWEEN TWO DIFFERENT DEVICES AND THE TWO DIFFERENT PROBES. NURSE WANTED TO KNOW WHICH TEMPERATURE WAS ACCURATE. MIS EXPLAINED THAT THE DIFFERENCE IN MODALITIES WOULD CREATE SMALL DIFFERENCES IN TEMPERATURE NOT TO MENTION DIFFERENT SITES WERE THESE PROBES WERE LOCATED. PATIENT WAS GENERATING HEAT EARLIER AND MEDICATIONS WERE ADMINISTERED. NURSE DENIED ANY CURRENT HEAT GENERATION. THE PATIENT TEMPERATURE WAS 37C CURRENTLY WITH RECTAL, WATER TEMPERATURE WAS 8.2C, FLOW RATE WAS 2.3 L/M AND 4 PADS CONNECTED WITH SOME EXPOSED ABDOMEN NOTED. MIS ENCOURAGED ADDITION OF UNIVERSAL PAD TO ABDOMEN TO ASSIST WITH SURFACE AREA COVERAGE AND TEMPERATURE CONTROL. NURSE NOTED THEY COULD UTILIZE BAIR HUGGERS OR BLANKETS FOR COUNTER WARMING IF PATIENT BEGINS GENERATING HEAT AGAIN. MIS SUGGESTED CHOOSING ONE DEVICE AND ONE PROBE TO RUN THERAPY. SWAPPING BACK AND FORTH CAN CREATE ISSUES WITH MANAGING PATIENT TEMPERATURE. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND THAT THERE WAS NO ALLEGATION AGAINST PRODUCT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00609 | 2023-03-02 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PPICC SOLO 4F SL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001027 |  | IT WAS REPORTED BY THE CUSTOMER THAT "CUT IN TABLE DRAPE. HAD TO OPEN ANOTHER KIT BECAUSE OF STERILITY ISSUE." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED DRAPE IS INCONCLUSIVE DUE TO POOR SAMPLE CONDITION. ONE TABLE DRAPE WAS RETURNED FOR EVALUATION WITHOUT PRODUCT PACKAGING. THE DRAPE WAS EXPANDED IN ORDER TO INSPECT THE MATERIAL. A SLIT IN THE DRAPE WAS OBSERVED TO BE PRESENT. THE EDGES OF THE SPLIT APPEARED TO BE SLIGHTLY COURSE. BASED ON THE CHARACTERISTICS OF THE RETURNED SAMPLE, POSSIBLE CONTRIBUTING FACTORS INCLUDE DAMAGE DURING HANDLING, SHARP INSTRUMENT DAMAGE, AND EXPOSURE TO A SHARP EDGE. THE CONDITION OF THE COMPONENT PRIOR TO KIT OPENING AND HANDLING IS UNKNOWN; THEREFORE, THE COMPLAINT COULD NOT BE CONFIRMED AND REMAINS INCONCLUSIVE AT THIS TIME. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00638 | 2023-03-02 | 2022-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED "LINE PLACED (B)(6) 2022. 1/5 DECLOT X 2 NOT SUCCESSFUL ON (B)(6) DURING REMOVAL PICC RN COULD FEEL A KINK JUST RELEASE AS PICC WAS WITHDRAWN. XRAY IMAGE DOES NOT SHOW ARM OF PATIENT BUT OF NOTE TIP POSITIONING IS SHORT." IT WAS STATED DEVICE WAS SECURED WITH CATHETER WAS SECURED WITH SECURACATH. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01381 | 2023-03-02 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE URINE LEAKAGE ON THE SIDE OF THE FOLEY CATHETER AND INCIDENT OCCURRED DURING USE. MEDI CARRIER GOT A REPORT THAT ALL OF THOSE TYPES OF CATHETERS WERE LEAKING. THEY HAVE TRIED TO CHANGE SIZE BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAS BEEN AN ISSUE FOR A FEW MONTHS AND THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE IS FOR 12CH, AND 16 COULD BE MORE SIZES EFFECTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO OPERATOR ERROR, USER ERROR, INCORRECT STRIPPING CAUSING TEAR. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01371 | 2023-03-02 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT SHAFT PART INFLATED. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE CATHETER WAS EVALUATED AND OBSERVED A LUMP ON THE SHAFT OF THE CATHETER. THE LUMP LOCATED AT 11.5CM FROM THE FUNNEL. MEASURED THE THICKNESS OF THE CATHETER (BOTH NORMAL SHAFT AND LUMP PART) AND THE THICKNESS OF LUMP PART WAS NOT WITHIN THE SPECIFICATION. INTRODUCED WATER INTO THE CATHETER AND THE CATHETER WAS ABLE TO INFLATE AND DEFLATE. DISSECTED THE LUMP PART OF THE CATHETER AND OBSERVED A HOLE IN THE FINISH DIPPING LAYER OF THE CATHETER. THE REPORTED EVENT IS CONFIRMED AS MANUFACTURING RELATED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO LATEX DIP OUT TOO SLOW CAUSED LATEX RUN BACK RESULTING IN LUMP. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00606 | 2023-03-02 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT FOUR YEARS THREE MONTHS AND FIFTEEN DAYS POST PORT PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY DAMAGED NEAR APPROXIMATELY 17 CM FROM THE TIP OF THE CATHETER. REPORTEDLY, THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE, AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED CATHETER FRACTURE, WEAR AND DEFORMATION ISSUES, AS A COMPOUND BREAK WAS NOTED APPROXIMATELY 6.7CM FROM THE DISTAL END OF THE CATH-LOCK. THE EDGES OF THE COMPOUND BREAK ON THE ATTACHED CATHETER WERE NOTED TO BE UNEVEN AND SURFACE WAS NOTED TO BE ROUND AND SMOOTH. HOWEVER THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED MATERIAL INTEGRITY PROBLEM AS THE MORE SPECIFIC DAMAGE SUCH AS FRACTURE WAS IDENTIFIED DURING INVESTIGATION. THE IDENTIFIED BREAK IS TYPICAL OF FLEXURAL FATIGUE, WHICH IS DUE TO THE REPETITIVE, CYCLIC KINKING OF THE CATHETER. SUCH KINKING MAY HAVE PHYSIOLOGICAL, PLACEMENT, USAGE OR MECHANICAL CONTRIBUTING FACTORS. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 10/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00657 | 2023-03-03 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR (RADEL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770352 |  | IT WAS REPORTED THE LOCATION FOR PICC LINES ARE NOT ACCURATE. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00660 | 2023-03-03 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG* TCS SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770131 |  | IT WAS REPORTED THE ECG READINGS ARE NOT CORRECT. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00665 | 2023-03-03 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8173118 |  | IT WAS REPORTED THAT THE CATHETER WITH FLUID LEAKAGE THROUGH WALLS. CATHETER WAS REPLACED. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00667 | 2023-03-03 | 2022-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY (6F) (TRIPLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4386118 |  | IT WAS REPORTED "PICC CATHETER INSTALLED ON (B)(6) 2022, AFTER PLATELET TRANSFUSION, IT SHOWED RESISTANCE IN THE RED PATHWAY, WHERE THERE WAS AN ATTEMPT TO CLEAR THE OBSTRUCTION IN THE AFTERNOON WITH A 10 ML SYRINGE. AFTER CLEARING THE OBSTRUCTION, THE PATIENT REPORTED A COLD SENSATION, AT NIGHT AT 10 PM THE CATHETER PRESENTED LEAKAGE BEING ISOLATED VIA THE RED ROUTE AND THE OTHER ROUTES BEING USED, THE CATHETER BEING REMOVED AND PERIPHERAL VENOUS ACCESS PUNCTURED. WHEN THE CATHETER WAS REMOVED, TWO HOLES WERE OBSERVED IN THE MEDIAL THIRD WHEN WASHING WAS CARRIED OUT IN THE RED ROUTE." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFS0141 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00666 | 2023-03-03 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | IT WAS REPORTED THAT DURING CATHETER PLACEMENT, BURRS WERE FOUND WITH THE STYLET IN THE POWER CATHETER AND IT WAS NOT SMOOTH. FOREIGN MATTERS WERE ADHERED TO THE SURFACE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGS0766 SHOWED TWO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGS0766) HAVE BEEN REPORTED FROM THE SAME FACILITY IN CHINA. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00668 | 2023-03-03 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER 5F BASIC TRAY DUAL-LUMEN W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7957505J |  | IT WAS REPORTED THAT THE STATLOCK PEELED OFF COMPLETELY. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00131 | 2023-03-03 | 2023-02-02 | I, n, j, u, r, y | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80144 | ATG80144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE AT THE RIGHT BRACHIOCEPHALIC VEIN FOR STENOSIS IN FISTULA THROUGH CEPHALIC VEIN, THE PTA BALLOON ALLEGEDLY RUPTURED CIRCUMFERENTIALLY AT 10 ATM IN SUPERIOR VENA CAVA. IT WAS FURTHER REPORTED THAT THE BALLOON ALLEGEDLY DETACHED FROM THE SHAFT AND WOULD NOT DEFLATE. REPORTEDLY, THE DEVICE WAS DIFFICULT TO REMOVE THRU THE SHEATH. FURTHER, THE DETACHED BALLOON AND TIP STAYED INSIDE THE VESSEL WHICH REQUIRED VENOTOMY TO REMOVE IT. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BALLOON RUPTURE, BALLOON DETACHMENT, BALLOON ENTRAPMENT AND DIFFICULT TO REMOVE FROM PATIENT COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRATION DATE: 08/2025). THE DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00637 | 2023-03-03 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED BY THE CUSTOMER THAT "UNABLE TO DECLOT LINE ON REMOVAL REALIZED IT WAS KINKED. LINE PLACED (B)(6) 2022. (B)(6) DECLOT X 2 NOT SUCCESSFUL ON (B)(6) DURING REMOVAL PICC RN COULD FEEL A KINK JUST RELEASE AS PICC WAS WITHDRAWN. XRAY IMAGE DOES NOT SHOW ARM OF PATIENT BUT OF NOTE TIP POSITIONING IS SHORT." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED 2/28/2023: IT WAS REPORTED BY THE CUSTOMER "CATHETER WAS SECURED WITH SECURACATH." ADDITIONAL INFORMATION RECEIVED 5/25/2023: 3CM OF CATHETER LEFT EXTERNAL WHEN INSERTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00671 | 2023-03-03 | 2022-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PICC HAS GREAT BLOOD RETURN AND FLUSH WHEN PATIENT'S ARM IS ABDUCTED AWAY FROM BODY. BOTH LUMENS OF PICC ARE NOT PATENT WHEN ARM IS BY THE PATIENT'S SIDE. THIS VAT RN CONSULTED WITH PICC RN, ON HOW TO TROUBLESHOOT POSITIONAL PICC. PICC RN RECOMMENDED PULLING PICC 0.5CM. THIS VAT RN PULLED PICC 0.5CM DURING DRESSING CHANGED. PICC HAS GOOD BLOOD RETURN AND FLUSH IN BOTH LUMENS. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REFQ1088 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00676 | 2023-03-03 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405J |  | IT WAS REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, BACK FLOW OCCURRED INTO THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00675 | 2023-03-03 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED VIA THE RIGHT BASILIC VEIN AND THE CATHETER WAS SECURED IN THE SUTURE WING ON (B)(6) 2023. ON (B)(6) 2023, INFUSION WAS DIFFICULT. THEREFORE, SECUREMENT OF THE SUTURE WING WAS LOOSENED. WHEN BLOOD RETURN WAS CONFIRMED, BLOOD COULD NOT BE ASPIRATED SMOOTHLY. THERE WAS NO REPORTED PATIENT INJURY. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00651 | 2023-03-03 | 2023-01-30 | I, n, j, u, r, y | Manufacturer report | POWER PORT MRI ISP, 8 FR. GROSHONG, INT W SP, ATTACHABLE SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808560 |  | IT WAS REPORTED THAT ELEVEN MONTHS AND SEVEN DAYS POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY FOUND FRACTURED. IT WAS FURTHER REPORTED THAT, THE CATHETER FRAGMENTS ALLEGEDLY MIGRATED TO THE PULMONARY ARTERY. REPORTEDLY, THE SUBCUTANEOUS PORTION OF THE DAMAGED CATHETER FRAGMENTS WERE REMOVED BY PERFORMING SURGICAL EXTRACTION AS OUTPATIENT PROCEDURE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT MRI ISP IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED CATHETER FRACTURE AND MATERIAL SEPARATION ISSUE AS A COMPLETE DIAGONAL BREAK WAS NOTED AT THE DISTAL END OF THE ATTACHED CATHETER AND THE EDGES OF THE COMPLETE DIAGONAL BREAK WERE NOTED TO BE JAGGED. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MIGRATION ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 05/2022). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00624 | 2023-03-03 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 18G X 10 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F318108PT |  | IT WAS REPORTED THAT NEEDLE DETACHED FROM THE POWERGLIDE DEVICE. NO PATIENT IMPACT. ADDITIONAL INFORMATION PROVIDED: PATIENT IS SOMEWHAT OBESE AND WAS NEEDING A MIDLINE FOR TRANSFER TO ANOTHER FACILITY. THERE WAS NO EMERGENT MEDICAL SITUATION OR INTERVENTION NEEDED BY THIS EVENT. THE NEEDLE REMAINED IN THE CATHETER AND I WAS ABLE TO RECOGNIZE THE FAILURE AND REMOVED THE DEVICE INTACT AT THAT TIME. GAUZE AND PRESSURE WAS HELD IN PLACE FOR A TIME FOR THE VEIN TO STOP BLEEDING TO RESUME ASSESSMENT FOR NEW PLACEMENT. THE ONLY HARM OR INJURY THAT OCCURRED WAS THE NEED FOR AN ADDITIONAL 'POKE' TO ACHIEVE ACCESS THAT WAS NEEDED. I DON'T RECALL ANY ISSUES WITH THE PATIENT, I BELIEVE THAT SHE MAY NOT HAVE BEEN AWARE OF THE ISSUE. THERE WAS NO INTERRUPTION OF TREATMENTS OR MEDICATIONS. THE MIDLINE PLACEMENT WAS FOR REMOVAL OF A CVC AND THEN PATIENT WAS TO BE TRANSFERRED TO ANOTHER FACILITY. NO ADDITIONAL INTERVENTION OR OTHERWISE WAS REQUIRED DUE TO THE POWERGLIDE FAILURE. PLACEMENT WAS ACHIEVED WITH THE SECOND ATTEMPT, NO FURTHER ACTION WAS REQUIRED OR NEEDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGW2635 SHOWED TWO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00669 | 2023-03-03 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1806050J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, SALINE SOLUTION ALLEGEDLY COULD NOT BE INFUSED THROUGH THE INTRODUCER NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE INTRODUCER NEEDLE, ONE STRAIGHT NON-CORING NEEDLE, ONE RIGHT ANGLE NON-CORING NEEDLE, ONE J-TIP GUIDEWIRE, ONE STRAIGHT TIP GUIDEWIRE, TWO GUIDEWIRE HOOPS, ONE 6.5FR INTRODUCER PEEL-APART SHEATH AND VESSEL DILATOR, ONE 5.0FR DILATOR, ONE TUNNELER, ONE CATH-LOCK, ONE VEIN PICK AND ONE SEALED SAFETY INFUSION SET WERE RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC VISUAL AND FUNCTIONAL EVALUATIONS WERE PERFORMED ON THE RETURNED DEVICE. NO CRACKS WERE OBSERVED THROUGHOUT THE INTRODUCER NEEDLE HUB AND NO LEAKS WERE OBSERVED THROUGHOUT THE INFUSION TEST. THE INVESTIGATION IS INCONCLUSIVE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 12/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00659 | 2023-03-03 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG PLUS TCS | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770351 |  | IT WAS REPORTED THE SENSOR IS NOT ACCURATELY TRACKING PICC. VERIFIED PICC HAS FLUSHED. VERIFIED TABLET HAS BEEN POWER CYCLED TRIED MULTIPLE PICCS. VERIFIED CALIBRATION. NO OTHER UNIT'S AVAILABLE. VERIFIED THEY'RE USING THE TABLET. EMAILED TM TO SEE IF SHE WOULD BE ABLE TO SWING BY WITH HER DEMO UNIT TO TEST IT WITH THEIR TABLET. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00670 | 2023-03-03 | 2022-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED, VAT CALLED TO TROUBLESHOOT PICC JUST EXCHANGED THIS MORNING. UPON ASSESSMENT, PICC COMPLETELY OCCLUDED AND UNABLE TO FLUSH OR DRAW BLOOD. XRAY IMAGE OBTAINED SHOWING SIMILAR KINKING OF CATHETER WITH PREVIOUS LINE. ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023: CATHETER WAS SECURED WITH SECURACATH. PATIENT BEING TREATED FOR SMALL BOWEL OBSTRUCTION WITH A MEDICAL HISTORY OF JEJUNAL ATRESIA AND MALROTATION S/P EX-LAP, RESECTION OF JEJUNAL ATRESIA WITH JEJUNO-JEJUNOSTOMY, AND LADD'S PROCEDURE IN INFANCY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00655 | 2023-03-03 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA CATHETER WITH SHERLOCK 3CG TPS STYLET 5F TRIPLE-LUMEN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000823 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE LINE NOT DRAWING FAILED TROUBLESHOOTS, POSITIONAL. (B)(6) 2023 MULTIPLE DECLOTS REPORTED VAT ASKED TO ASSESS PICC. CHANGED DRESSING NO KINKING NOTED UNDER DRESSING. BEDSIDE RN DID NOTE THAT THE LINE FLUSHES EASIER AND GETS BLOOD RETURN WHEN THE ARM IS ABDUCTED FROM THE BODY. (B)(6) 2023 PT WAS DIURESED AND ARM IS NO LONGER EDEMATOUS. LOOSE SKIN NOTED. REPORTS THAT THE PUMP ALARMS OCCLUSION WITH INFUSION AND THAT THE BLOOD RETURN IS NOT RELIABLE. PICC LINE WAS ASSESSED BY THIS RN. ALL LUMENS WITH BRISK BLOOD RETURN AND FLUSHES EASILY WHEN LINE WAS ASSESSED WITH PATIENT'S LEFT ARM IN DIFFERENT POSITIONS. IT APPEARS THAT WITH THE EDEMA RESOLVING AND THE STEEP INSERTION ANGLE IS THE CAUSE OF THE SHARP TURN SEEN ON THE X-RAY. NP REQUESTED THAT VAT PULL THE LINE BACK TO SEE IF THE TURN WOULD BE RESOLVED. LINE WAS PULLED BACK TO 7CM EXTERNAL FROM (FROM 5 CM). AFTER DRESSING WAS APPLIED AND ARM PLACED BACK IN A NEUTRAL POSITION, ALL LUMENS CONTINUE TO FLUSH EASILY WITH BRISK BLOOD RETURN. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REFX5392 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00642 | 2023-03-03 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833730 | 5833730 |  | IT WAS REPORTED THAT TWO DAYS POST DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY OCCLUDED. IT WAS FURTHER REPORTED THAT THE CATHETER WAS REPLACED BY A FEMORAL CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00650 | 2023-03-03 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT "ISSUES WITH INTRODUCER IN THE PICC TRAYS. HAS HAD ONE BUCKLE ON HIM AND HAS TO NICK ALL THE TIME. NEVER USED TO NEED TO." ADDITIONAL INFORMATION PROVIDED (B)(6) 2023: IT WAS STATED "THE PEEL AWAY PORTION IS RAISED. BECAUSE IT IS RAISED WHEN IT MEETS THE SKIN THERE IS A GREATER LEVEL OF FORCE REQUIRED TO PENETRATE THE PATIENTS SKIN. SOMETIMES THE FORCE REQUIRED IS SO GREAT THAT THE RAISED PORTION WILL CUP UP MAKING IT EVEN MORE DIFFICULT TO IMPOSSIBLE TO PENETRATE THE SKIN. IF BD ENGINEERS COULD DESIGN THE PEEL AWAY PORTION OF THE INTRODUCER SO IT IS FLUSH TO THE LONGER PORTION THAT THE PEEL AWAY PORTION COVERS IT WOULD REQUIRE LESS FORCE TO INSERT AND THEREFORE LESS PRESSURE AND DISCOMFORT FOR THE PATIENT. ALSO, IT WOULD REQUIRE LESS NICKS TO THE PATIENT'S SKIN DECREASING BLEEDING AND RISK FOR INFECTION." A SPECIFIC NUMBER OF AFFECTED DEVICES WAS NOT PROVIDED BY THE COMPLAINANT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01429 | 2023-03-03 | 2023-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT WAS HAVING DIFFICULTY INSERTING THE MAGIC 3 GO MALE INTERMITTENT CATHETER ALSO STATED SOME TIMES THEY WOULD GO IN AND SOME TIMES IT WOULD NOT . | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "ROUGH OR IRREGULAR SHAPE PROTRUSIONS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA INSTRUCTIONS FOR USE: REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00648 | 2023-03-03 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 20G X 10 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F320108PT |  | IT WAS REPORTED BY THE CUSTOMER THAT "WHILE INSERTING POWERGLIDE AND ENGAGING THE SAFETY MECHANISM, THE SMALL NEEDLE COVER DID NOT CAP THE NEEDLE. INSTEAD THE NEEDLE CAME OUT UNCOVERED AND THE SMALL CLEAR HOUSING STAYED ATTACHED TO THE PINK WINGS OF THE PLUG IN THE END OF THE POWERGLIDE HUB". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00649 | 2023-03-03 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 18G X 10 CM WITH GUARDIVA AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F118108PT |  | IT WAS REPORTED BY THE CUSTOMER THAT "PER THE UNIT MANAGER THE DEVICE REACHED THE PATIENT-MIDLINE IV DEVICE THAT THE SAFETY CAP/SHEATH DID NOT DEPLOY." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGX0599 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00647 | 2023-03-03 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT POWERPICC 4F SL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000631 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THERE WAS A 'MUSHROOM' ON PICC INTRODUCER. UNABLE TO PLACE INTRODUCER." ADDITIONAL INFORMATION RECEIVED FEBRUARY 15, 2023: "THERE WAS NO HARM CAUSED TO THE PATIENT. PICC LINE WAS BEING PLACED FOR LONG TERM ANTIBIOTIC THERAPY. AS I ATTEMPTED TO ADVANCE INTRODUCER AFTER SKIN WAS NICKED, THERE WAS NO ADVANCEMENT. I TOOK IT OFF THE WIRE AND INSPECTED IT AND NOTICED THERE WAS A SMALL LIP/BURR ON THE GRAY PORTION. I HAD TO OBTAIN ANOTHER INTRODUCER TO COMPLETE THE PICC." | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGQ1683 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGQ1683) HAVE BEEN REPORTED FROM THE SAME FACILITY. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00646 | 2023-03-03 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY THE CUSTOMER THAT UPON ACTIVATING THE PUSH BUTTON SAFETY, THE NEEDLE FAILED TO RETRACT INTO THE SAFE POSITION. INSERTION WAS ABORTED AND THE CLINICIAN DISPOSED OF THE DEVICE IN SHARPS DISPOSAL. NO CLINICIAN OR PATIENT HARM REPORTED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGZ1729 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00641 | 2023-03-03 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK APHERESIS PORT ACCESS SYS |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT ALLEGEDLY BLED THROUGH THE CATHETER. HOWEVER, THERE IS NO INFORMATION REGARDING TREATMENT PROVIDED FOR BLEEDING. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3, H6 (METHOD) H11: B5 H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00644 | 2023-03-03 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED, LINE WAS REMOVED DUE TO IT BEING POSITIONAL AND MULTIPLE KINKS IN CATHETER. RECOMMEND REPLACEMENT OF PICC SINCE PICC IS EXTREMELY POSITIONAL WITH ARM MOVEMENTS AND TIP IS IN MID SVC WHICH IS NOT IDEAL LOCATION FOR CPN. POSITIONAL ON LINE REMOVAL NOTED MULTIPLE KINKS IN CATHETER. NO XRAY IMAGE OF ARM OBTAINED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGX2173 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01423 | 2023-03-03 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT ON (B)(6) 2023, WHEN THE NURSE INFUSING THE FOLEY CATHETER, IT WAS FOUND THAT THE RESISTANCE THE END OF THE WATER INJECTION TUBE WAS TOO HIGH, AND THE LIQUID COULD NOT BE INJECTED. THE SAME SITUATION WAS NOT FOUND AFTER THE CATHETER WAS REPLACED, AND NO ADVERSE EFFECTS WERE CAUSED TO THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01425 | 2023-03-03 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WHEN THE FOLEY CATHETER WAS REMOVED, THE BALLOON COULD NOT BE DEFLATED AND WAS DIFFICULT TO REMOVE. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT S POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE COLLAPSED LUMEN AND BLOCKED LUMEN. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01426 | 2023-03-03 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE-SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WAS A FEELING OF RESISTANCE WHEN REMOVING THE FOLEY CATHETER, AND IT WAS DIFFICULT TO REMOVE. ALSO STATED THAT PATIENT HAD BEEN USING THIS FOR MANY YEARS, BUT THIS WAS THE FIRST TIME THAT IT WAS DIFFICULT TO REMOVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01417 | 2023-03-03 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | URETHRAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE CUSTOMER WAS HAVING ISSUES WITH THE INTERMITTENT CATHETER NOT RELEASING THE URINE. CUSTOMER FELT LIKE IT WAS BLOCKED, AND THE URINE WOULD NOT COME OUT. THE CUSTOMER ALSO MENTIONED THAT THEY EXPERIENCED DIFFICULTY INSERTING THE CATHETER ALL THE WAY. STATED THAT THEY WERE RECEIVED 2 IODINE IN ONE OF THE KITS AND RECEIVED A KIT WHERE THERE WAS NO IODINE IN IT. CUSTOMER RECEIVED INFECTION TWICE IN ONE YEAR DUE TO REUSING THE HOSE OR RUNNING OUT OF SUPPLY. CUSTOMER COULD NOT CONFIRM IF IT WAS A BD PRODUCT AND SAID IT WAS A LONG WHILE AGO AND COULD NOT REMEMBER. ALSO STATED CUSTOMER WIFE AND DAD USED THIS PRODUCT AS WELL AND EXPERIENCED SOME ISSUES. NO SPECIFICATION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01415 | 2023-03-03 | 2023-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | TEMP-SENSING CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT A NORMOTHERMIA PATIENT TEMPERATURE CURRENTLY AT 39C ON THE ARCTIC SUN DEVICE. THEY SWAPPED OUT DEVICES BECAUSE PATIENT TEMPERATURE WAS FLUCTUATING ON FIRST DEVICE AND RECEIVED ERRATIC PATIENT TEMPERATURE ALERT. ESOPHAGEAL PROBE IN USE AND PLACEMENT VERIFIED, BUT IT WAS STILL NOT CORRELATING WITH OTHER PROBE AND NOTICING FLUCTUATIONS WITH THIS DEVICE, TOO. SOMETIMES IT WOULD REGISTER SIMILAR AS OTHER PROBE, BUT MOSTLY STAYS AROUND 39C. OTHER PROBE SHOWS PATIENT TEMPERATURE 35.6C. NOTED THAT THE ISSUE COULD EITHER BE THE TEMPERATURE CABLES ATTACHED TO THE PROBE AND INTO THE DEVICE OR THE ACTUAL PROBE ITSELF. MIS SUGGESTED PLUGGING THE PROBE INTO A DIFFERENT DEVICE SUCH AS A BEDSIDE MONITOR TO SEE IF IT CORRELATES. IF FLUCTUATIONS AND DISCREPANCIES STILL NOTED, THEN THEY NEED TO SWAP OUT THE PROBE. OTHERWISE, THEY CAN TRY REPLACING THE TEMPERATURE IN CABLE. MIS SUGGESTED HAVE A COUPLE TEMPERATURE CABLES ON HAND IN CASE THIS HAPPENED IN THE FUTURE. MIS ENCOURAGED NURSE TO CALL BACK WITH ANY QUESTIONS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01413 | 2023-03-03 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE URINE LEAKAGE ON THE SIDE OF THE FOLEY CATHETER AND THE INCIDENT OCCURRED DURING USE. MEDI CARRIER GOT A REPORT THAT ALL OF THOSE TYPES OF CATHETERS WERE LEAKING. THEY HAVE TRIED TO CHANGE SIZE, BUT THE ISSUE WITH LEAKAGE REMAINS. THIS HAD BEEN AN ISSUE FOR A FEW A MONTH, AND THERE WAS AN ISSUE WITH THE BARDIA SILICONE ELASTOMER CATHETERS. THE ISSUE WAS FOR 12CH, 14CH AND 16 COULD BE MORE SIZES EFFECTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE POOR BALLOON SHAPE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01410 | 2023-03-03 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT AN INDWELLING FOLEY CATHETER WAS PLACED AFTER THE OPERATION ON (B)(6) 2023 AND RETAINED THE BLADDER IRRIGATION. AT 23:50 ON (B)(6), THE PATIENT'S CATHETER BALLOON RUPTURED AND THE CATHETER FELL OUT. THE DOCTOR REPLACED A NEW CATHETER WITH GOOD FIXATION AT THE URETHRAL ORIFICE FOR THE PATIENT AND CONTINUED BLADDER IRRIGATION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT IF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00643 | 2023-03-03 | 2023-01-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED BY THE CUSTOMER "R BRACHIAL ATTEMPTED 10 CM ABOVE ANTECUBITAL FOSSA: LINE SUCCESSFULLY PLACED AND CONFIRMED WITH ECG; HOWEVER UPON PLACING DRESSING RED LUMEN VERY DIFFICULT TO FLUSH. DRESSING REMOVED AND ASSESSED; WITH NO VISIBLE KINKS: CXR TAKEN TO ASSESS TIP LOCATION, AND KINK NOTED IN ARM. ATTEMPTED TO RETRACT LINE TO RELEASE KINK BUT UNSUCCESSFUL. NEW PICC LINE PLACED." NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGV0508 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. THE COMPLAINTS FOR THIS BATCH NUMBER HAVE BEEN REPORTED FROM THE SAME FACILITY. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00680 | 2023-03-04 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1859660 | 1859660 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, WHEN PEELING THE SHEATH AWAY, THE PLASTIC RING ALLEGEDLY BROKE IN TWO PIECES AND FELL OFF OF WHITE PLASTIC PART OF SHEATH MAKING IT DIFFICULT FOR RADIOLOGIST. IT WAS FURTHER REPORTED THAT THE RADIOLOGIST MANAGED TO FINISH THE PROCEDURE BY INSERTING THE IMPLANTED VASCULAR ACCESS DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE FULLY PEELED 10.0F PEEL-APART SHEATH WERE RETURNED FOR EVALUATION. GROSS VISUAL EVALUATION WERE PERFORMED. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED DIFFICULT SEPARATION, FRACTURE, AND MATERIAL SEPARATION AND THE IDENTIFIED DEFORMATION ISSUE AS TWISTING AND DEFORMATION WAS NOTED THROUGHOUT BOTH SHEATH SHAFTS AND THE CLEAR RING WAS BROKEN AWAY FROM THE TAB. MANUFACTURING SITE EVALUATION OF THE SAMPLE FOUND THAT THE TOP CAP WAS DETACHED IN ONE OF THE HALVES OF T-HANDLE, THE VALVE BELONGING TO THIS HALF OF THE LID WAS ABSENT IN THE SHEATH. THEREFORE, THE CAUSE OF THESE CONDITIONS IS RELATED TO THE MANUFACTURING/PROCESSING OF THE DEVICE. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00681 | 2023-03-04 | 2022-12-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT TWO YEARS THREE MONTHS AND ONE DAYS POST PLACEMENT, DURING NORMAL SALINE FLUSHING, THE PATIENT ALLEGEDLY EXPERIENCED SWELLING AT THE SITE NEAR CLAVICLE AND ALONG THE TUNNEL, UP TO INSERTION SITE. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY HAD LEAKAGE OF NORMAL SALINE BECAUSE OF CATHETER BREAKAGE. HOWEVER, THERE WAS NO MEDICAL INTERVENTION PROVIDED FOR THE SWELLING AND LET IT BE ABSORBED. REPORTEDLY, THE CATHETER WAS REMOVED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 06/2024). DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00679 | 2023-03-04 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT THE APPROXIMATELY THREE YEARS EIGHT MONTHS AND EIGHT DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, THE PORT ALLEGEDLY HAD LEAK. IT WAS FURTHER REPORTED THAT THE CRACK WAS ALLEGEDLY FOUND AT APPROXIMATELY 10.5 CM FROM THE TIP OF THE REMOVED CATHETER. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THE PROCEDURE WAS COMPLETED BY REPLACING ANOTHER PORT SYSTEM. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00138 | 2023-03-04 | 2022-12-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SEEKER | SUPPORT CATHETER | BARD PERIPHERAL VASCULAR, INC. | SK9035M | SK9035M |  | IT WAS REPORTED THAT PRIOR TO A RECANALIZATION PROCEDURE, THE LAST FOUR INCHES OF THE CATHETER BETWEEN THE MARKERS ALLEGEDLY SEPARATED AFTER REMOVING IT FROM THE PLASTIC HOOP TO FLUSH. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE SEEKER CROSSING SUPPORT CATHETER WAS RECEIVED AND EVALUATED. ON VISUAL EVALUATION, THE SAMPLE WAS RETURNED INTO TWO SEGMENTS. THE TWO SEGMENTS OF THE CATHETER WERE THE HUB AND REST OF THE SHAFT. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED TO THE CATHETER. THEREFORE, BASED ON THE FINDINGS, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED DETACHMENT ISSUE AS THE CATHETER DETACHED. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DEVICE DAMAGED PRIOR TO USE AS NO EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED DEVICE DAMAGED PRIOR TO USE AND THE IDENTIFIED DETACHMENT ISSUES COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. DEVICE EXPIRY DATE: 09/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00139 | 2023-03-04 | 2023-01-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80142 | ATG80142 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY DID NOT EXPAND AND THE SHAPE OF THE BALLOON WAS WEIRD. IT WAS FURTHER REPORTED THAT THERE WAS DIFFICULTY IN RETRACTING THE DEVICE THRU THE INTRODUCER SHEATH. REPORTEDLY, THE BALLOON ALLEGEDLY DETACHED AND HAIR STRINGS WAS NOTED AROUND THE BALLOON. THERE WAS NO REPORTED PATENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ATLAS GOLD PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, THE BALLOON WAS OBSERVED TO BE INVERTED WITH UNRAVELING FIBERS, AND DETACHED FROM THE CATHETER BUT STILL ATTACHED TO THE INNER GUIDEWIRE LUMEN THAT WAS COMPLETELY EXPOSED. THE GLUE BULLET WAS NOT SEATED IN THE CORRECT POSITION AND MARKER BANDS WERE STILL PRESENT. NO ANOMALIES NOTED TO THE LUERS/Y-BODY. NO FUNCTIONAL TESTING WAS PERFORMED DUE TO THE CONDITION OF THE SAMPLE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED UNRAVELED MATERIAL AS UNRAVELING FIBERS WERE NOTED ON THE DEVICE. THE INVESTIGATION IS ALSO CONFIRMED FOR THE IDENTIFIED BREAK AS THE BALLOON WAS SEEN DETACHED FROM THE CATHETER AT ONE END BUT STILL ATTACHED TO THE INNER GUIDEWIRE LUMEN. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED UNEVEN INFLATION AND REMOVAL DIFFICULTY AS NO FUNCTIONAL TESTING COULD BE PERFORMED DUE TO THE CONDITION OF THE RETURNED DEVICE. ADDITIONALLY, THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED DETACHMENT AS NO COMPLETE DETACHMENT WAS NOTED ON THE DEVICE. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED UNEVEN INFLATION, REMOVAL DIFFICULTY, UNRAVELED MATERIAL, DETACHMENT AND IDENTIFIED BREAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 12/2025), G3, H6 (DEVICE, METHOD) . H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00134 | 2023-03-04 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE TIP OF THE CATHETER ALLEGEDLY SEPARATED AFTER ONE HUNDRED AND TWENTY SECONDS. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED MATERIAL SEPARATION ISSUE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 04/2024), G3, H6 (METHOD). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00132 | 2023-03-04 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8064 | VA8064 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE BALLOON WAS INFLATED AND THE PTA BALLOON CATHETER ALLEGEDLY HAD SMALL HOLE AT THE PROXIMAL PORTION OF THE SHAFT WHERE THE DILUTED CONTRAST ALLEGEDLY LEAKED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE VACCESS DILATATION CATHETER WAS RETURNED FOR EVALUATION. NO SPECIFIC ANOMALIES WERE NOTED DURING THE VISUAL EVALUATION. DURING THE FUNCTIONAL TESTING, NEGATIVE PRESSURE WAS APPLIED TO DEFLATE THE BALLOON COMPLETELY. THEN THE CATHETER GUIDE WIRE LUMEN AND IN-HOUSE INTRODUCER SHEATH WERE FLUSHED. BY USING THE IN-HOUSE GUIDE WIRE THE CATHETER WAS PASSED THROUGH IT WITHOUT ANY ISSUE, THE BALLOON WAS INSERTED THROUGH THE SHEATH AND INFLATED WITH THE IN-HOUSE PRESTO. PIN HOLE RUPTURE WAS NOTED ON THE CATHETER SHAFT PROXIMALLY NEAR STRAIN RELIEF. UNDER THE MICROSCOPIC OBSERVATION A TEAR WAS NOTED TO THE CATHETER SHAFT. BALLOON WAS RETRACTED WITH NO ISSUES. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. THEREFORE THE INVESTIGATION WAS CONFIRMED FOR THE REPORTED CATHETER SHAFT PUNCTURE AND IDENTIFIED TEAR, AS A PIN-HOLE TEAR WAS NOTED ON THE CATHETER SHAFT UPON INFLATING THE BALLOON, THE ANOMALY WAS ALSO CONFIRMED UNDER MICROSCOPIC OBSERVATION. A DEFINITIVE ROOT CAUSE FOR THE REPORTED CATHETER SHAFT PUNCTURE AND IDENTIFIED TEAR COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 06/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00678 | 2023-03-04 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602680 | 0602680 |  | IT WAS REPORTED THAT APPROXIMATELY TEN MONTHS POST PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY HAD PAIN DURING INFUSION AND EXPERIENCED SWELLING AROUND THE PORT RESERVOIR SITE. IT WAS FURTHER REPORTED THAT THE ULTRASOUND SHOWED FLUID LEAK AROUND THE PORT. FURTHERMORE, A CONTRAST INJECTION WAS PERFORMED UNDER FLUORO AND THERE WAS NO FIBRIN SHEATH, BUT CONTRAST WAS NOTED AROUND THE PORT RESERVOIR. REPORTEDLY, THE PORT WAS REMOVED, AND IT WAS EXPECTED A HOLE IN THE SEPTUM INSTEAD, THE SEPTUM APPEARED INTACT, BUT THE PORT MATERIAL BALLOONED AT THE BACK OF THE PORT AND GRADUALLY LEAKED AROUND THE SILICONE. THE PROCEDURE WAS COMPLETED BY REPLACING WITH ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE PORT INFUSION SET WITH A LOADED SYRINGE LOADED TO A BARDPORT IMPLANTABLE PORT ATTACHED TO A CATHETER WAS RECEIVED FOR EVALUATION AND ONE IMAGE WAS PROVIDED FOR REVIEW. VISUAL, GROSS VISUAL, MICROSCOPIC, FUNCTIONAL AND TACTILE EVALUATIONS WERE PERFORMED. THE BOTTOM PORTION OF THE PORT BODY WAS NOTED TO BALLOON UPON INFUSION AND LEAK WAS NOTED WHILE THE IN-HOUSE INFUSION SET WITH THE LOADED SYRINGE WAS INSERTED TO THE PORT SEPTUM. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FLUID LEAK, MATERIAL PROTRUSION AND STRETCHED ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01449 | 2023-03-06 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT WATER HAS LEAKED OUT OF THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01436 | 2023-03-06 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE CUSTOMER WAS HAVING ISSUES WITH THE INTERMITTENT CATHETER NOT RELEASING THE URINE. CUSTOMER FELT LIKE IT WAS BLOCKED, AND THE URINE WOULD NOT COME OUT. CUSTOMER ALSO MENTIONED THAT THEY EXPERIENCED DIFFICULTY INSERTING THE CATHETER ALL THE WAY. STATED THAT THEY WERE RECEIVED 2 IODINE IN ONE OF THE KITS AND RECEIVED A KIT WHERE THERE WAS NO IODINE IN IT. CUSTOMER RECEIVED INFECTION TWICE IN ONE YEAR DUE TO REUSING THE HOSE OR RUNNING OUT OF SUPPLY. CUSTOMER COULD NOT CONFIRM IF IT WAS A BD PRODUCT AND SAID IT WAS A LONG WHILE AGO AND COULD NOT REMEMBER. ALSO STATED CUSTOMER WIFE AND DAD USED THIS PRODUCT AS WELL AND EXPERIENCED SOME ISSUES. NO SPECIFICATION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01463 | 2023-03-06 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILIZED WATER LEAKED FROM THE BALLOON PART WHEN WATER WAS INJECTED DURING THE CUFF CHECK. THE EVENT OCCURRED DURING PRETEST. THERE WAS NO BALLOON TEAR OR BURST OBSERVED. NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. PER FOLLOW UP INFORMATION RECEIVED ON (B)(6)2023, STATED THAT THERE WAS A CRACK ON THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01451 | 2023-03-06 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT UPON INJECTING WATER AFTER THE CATHETER PLACEMENT TO THE PATIENT, WATER LEAKED FROM THE CATHETER SHAFT. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS INFLATED WITH WATER AND OBSERVED WATER LEAKING THRU DRAINAGE FUNNEL. DISSECTED THE CATHETER AND OBSERVED TEAR ORIGINATED FROM INSIDE AT INFLATION FUNNEL AREA. TEAR WAS EXAMINED UNDER MICROSCOPE AND NOTED TO BE ROUNDED AND NOT SMOOTH AND IDENTIFIED THIS IS RELATED TO MANUFACTURING. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO POOR STRIPPING OR ROUGH STRIPPING. A REVIEW OF THE MANUFACTURING PROCESS INDICATES THAT THE PROCESS IS CAPABLE OF PRODUCING OF THIS TYPE OF DEFECT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER SHAPE, CONFIGURATION AND PRINCIPLES: BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS. SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY INTENDED USE & EFFECT- EFFICACY: THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. DIRECTIONS FOR USE: 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR. (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION CATHETER KINKING, DAMAGE, RUPTURE DIFFICULTY OR FAILURE TO REMOVE THE DEVICE OCCLUSION OF CATHETER INNER LUMENS ENCRUSTATION ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS URINARY-TRACT INFECTION HEMORRHAGE, HEMATURIA ALLERGY REACTION TO THE DEVICE CALCULUS FORMATION EDEMA PAIN DISCOMFORT INJURY OF BLADDER OR URETHRAL URETHRITIS, URINARY INCONTINENCE RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00685 | 2023-03-06 | 2023-01-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT FOUR DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN FOR THE TOTAL PARENTERAL NUTRITION ADMINISTRATION, LEAKAGE ALLEGEDLY OCCURRED FROM THE SITE WHICH THE PORT NEEDLE WAS PUNCTURED. IT WAS FURTHER REPORTED THAT, X-RAY EXAMINATION WAS PERFORMED AND REVEALED THAT TOTAL PARENTERAL NUTRITION AND BLOOD TRANSFUSION WERE LEAKED NEAR THE PORT BODY. THE PATIENT EXPERIENCED PAIN AND THE PORT SYSTEM WAS REMOVED TO TREAT PAIN. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED FLUID LEAK ISSUE AS THE PORT BODY WAS PATENT TO BOTH INFUSION AND ASPIRATION WITHOUT ISSUE AND NO LEAKS WERE OBSERVED THROUGHOUT BOTH TESTS. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 09/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01446 | 2023-03-06 | 2023-02-13 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE 14FR MAGIC 3 HYDROPHILIC MALE INTERMITTENT CATHETER WAS LARGER THAN THE 14FR MAGIC 3 GO INTERMITTENT CATHETER AND HAD CAUSED URINARY TRACT INFECTION. THE REPRESENTATIVE ADVISED THAT THE MAGIC 3 14FR SHOULD NOT BE LARGER BUT OFFERED TO ATTEMPT TO FIT THE PATIENT INTO ANOTHER CATHETER, THE PATIENT DECLINED AND PREFERRED THE MAGIC 3 GO CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DONT REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01431 | 2023-03-06 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE LEAKS OCCURRED IN THE FOLEY CATHETER AFTER SIX DAYS OF USE. THE USER NORMALLY USED THIS FOR A MONTH. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 20FEB2023, STATED THAT THE LOT NUMBER OF THE DEFECTIVE CATHETER WAS MYEY1619 (BATCH# MYEY1619). THE PATIENT HAD ALREADY DISPOSED THE OTHER PACKAGING, SO THE LOT NUMBERS WERE UNKNOWN (BATCH# UNK). CLARIFIED THAT THE PACKAGE NORMALLY LASTS A MONTH, IF IT DID NOT LEAK. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE BAD FIT WITH SHAFT/NO DRAINAGE EYE/ POORLY SEATED BALLOON /BLADDER NECK INTERFACE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01440 | 2023-03-06 | 2023-02-13 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE PATIENT THAT AFTER USING THE INTERMITTENT CATHETER 3 TIMES THEY ENDED UP IN THE HOSPITAL WITH BLADDER INFECTION. THE PATIENT STATED THAT THEY WERE IN THE HOSPITAL FOR THE DAY THEY WAS SENT HOME WITH ANTIBIOTICS. THE PATIENT HAD A LOT OF BLOODY URINE. THE PATIENT STATED THEY GOT BLADDER INFECTION WITH MIGHT BE MORE THAN TWO YEARS AGO WITH A DIFFERENT CATHETER, BUT IT WAS NOT WITH THE MAGIC 3 GO. LET THE PATIENT KNOW THEY COULD TRY ANOTHER HYDROPHILIC CATHETER. IT WAS UNKNOWN IF THE INTERMITTENT CATHETER CONTRIBUTED TO THE BLADDER INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE; THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | Y | N | LAY USER/PATIENT |  | 2023-03
16484891 | 2023-03-06 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP INTERMITTENT CATHETER TRAY | CATHETER, UROLOGICAL | C. R. BARD, INC. | INTS16 | INTS16 |  | AFTER REMOVING STRAIGHT CATHETER FROM PATIENT, THE STAFF BROUGHT CATHETER TO THE TRASH AND CATHETER SPONTANEOUSLY DISCONNECTED FROM TEAL HUB ON CATHETER BAG, EJECTING URINE ON STAFF'S CHEST AND NECK AREA. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00677 | 2023-03-06 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT LOW-PROFILE IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0655870 | 0655870 |  | IT WAS REPORTED THAT ONE MONTH AND EIGHT DAYS POST PORT PLACEMENT, THE TEAR ALLEGEDLY OCCURRED AT ABOUT FIFTY CENTIMETER AND WAS IDENTIFIED UNDER FLUOROSCOPY USING A CONTRAST AGENT. IT WAS FURTHER REPORTED THAT, THE PATIENT ALLEGEDLY EXPERIENCED ARM PAIN AND WARM AT FIRST DUE TO TEAR. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE X-PORT MRI ISP ATTACHED TO A CATHETER IN TWO SEGMENTS WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE, AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, ONE ELECTRONIC IMAGE WAS PROVIDED FOR REVIEW. THE IMAGE SHOW AN ARM WITH POSSIBLE CONTRAST AROUND THE HUMERUS AREA AND A TEAR IN THE TUBING COULD OVER TIME IF THE TUBING CROSSES A POINT OF FLEXION. BASED ON THE IMAGE REVIEW THE REPORTED EVENT CANNOT BE CONFIRMED. HOWEVER, THE WHOLE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE, AS A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED AT THE DISTAL END OF THE CATHETER ATTACHED AND A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED ON THE PROXIMAL END OF THE DISTAL CATHETER SEGMENT. UPON INFUSION, A LEAK WAS OBSERVED FROM THE PARTIAL CIRCUMFERENTIAL BREAK ON THE CATHETER. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00148 | 2023-03-06 | 2023-01-31 | I, n, j, u, r, y | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80144 | ATG80144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE SUBCLAVIAN VEIN, AFTER THE BALLOON WAS INFLATED, THE BALLOON SLID TOWARDS THE HEART AND THE PHYSICIAN APPLIED SLIGHT FORCE TO MAINTAIN THE BALLOON SHAFT. IT WAS FURTHER REPORTED THAT WHEN PULLING THE BALLOON CATHETER, THE BALLOON AND THE SHAFT WERE ALLEGEDLY SEPARATED. FURTHERMORE, THE BALLOON ALLEGEDLY RUPTURED IN THE LESION. REPORTEDLY, THE SHAFT WAS TAKEN OUT FROM THE PATIENT AND THE PHYSICIAN USED A BARE METAL STENT TO APPOSE THE RUPTURED BALLOON MATERIAL IN THE LESION TO COMPLETE THE PROCEDURE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ATLAS GOLD PTA DILATATION CATHETER WAS RETURNED FOR EVALUATION. THE BALLOON PORTION WAS DETACHED ENTIRELY FROM THE CATHETER, EXPOSING THE INNER GUIDE WIRE LUMEN AND WAS NOT RETURNED FOR EVALUATION. THE GLUE BULLET WAS NOT SEATED CORRECTLY. NO OTHER SPECIFIC ANOMALIES WERE NOTED DURING THE VISUAL EVALUATION. NO FUNCTIONAL TESTING WAS PERFORMED DUE TO THE NATURE OF THE COMPLAINT AND THE CONDITION OF THE SAMPLE. THEREFORE, THE INVESTIGATION WAS CONFIRMED FOR REPORTED BALLOON DETACHMENT AS THE RETURNED SAMPLE CONFIRMED THE BALLOON NOTED TO HAVE DETACHED FROM THE CATHETER, EXPOSING THE INNER GUIDE WIRE LUMEN, AND WAS NOT RETURNED FOR EVALUATION. HOWEVER, THE INVESTIGATION FOR THE REPORTED BALLOON RUPTURE REMAINS INCONCLUSIVE AS THE FUNCTIONAL TESTING COULDN¿T BE PERFORMED AND NO OBJECTIVE OF THE BALLOON RUPTURE COULD BE OBSERVED DUE TO THE CONDITION OF THE DEVICE. A DEFINITIVE ROOT CAUSE FOR THE REPORTED BALLOON DETACHMENT AND BALLOON RUPTURE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 11/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00141 | 2023-03-06 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U42003H20RX | U42003H20RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE FOR RIGHT LOWER EXTREMITY STENOSIS IN POSTERIOR TIBIAL, WHILE TRYING TO REMOVE THE PTA BALLOON, THE WIRE AND THE BALLOON BECAME ALLEGEDLY STUCK INSIDE THE SHEATH. IT WAS FURTHER REPORTED THAT ANOTHER WIRE WAS INSERTED TO KEEP THE ACCESS AND PULLED OUT THE BALLOON, WIRE AND THE SHEATH. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 01/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01457 | 2023-03-06 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE 14FR MAGIC 3 HYDROPHILIC MALE INTERMITTENT CATHETER WAS LARGER THAN THE 14FR MAGIC 3 GO INTERMITTENT CATHETER AND HAD CAUSED URINARY TRACT INFECTION. REPRESENTATIVE ADVISED THAT THE MAGIC 3 14FR SHOULD NOT BE LARGER BUT OFFERED TO ATTEMPT TO FIT PATIENT INTO ANOTHER CATHETER, PATIENT DECLINED AND PREFERRED THE MAGIC 3 GO CATHETER. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER. CAUTION: FEDERAL (USA) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN." | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01438 | 2023-03-06 | 2023-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THESE MAGIC 3 GO HYDROPHILIC MALE INTERMITTENT CATHETER DID NOT HAVE ENOUGH LUBRICITY. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." | N | Y | LAY USER/PATIENT |  | 2023-03
2020394-2023-00144 | 2023-03-06 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE TIP OF THE CATHETER ALLEGEDLY DETACHED. IT WAS FURTHER REPORTED THAT CATHETER WAS STILL ATTACHED TO THE CORE WIRE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN COMMON DEVICE NAME AND PMA/510K. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 10/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00686 | 2023-03-06 | 2022-02-07 | I, n, j, u, r, y | Manufacturer report | POWERPICC 4F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174155 |  | CUSTOMER REPORTED THAT THE DEVICE WAS PULLED OUT ACCORDING TO PROTOCOL WITHOUT RESISTANCE, DURING THE COURSE IT WAS PULLED OUT 38 CM AS INSERTED, BUT IT WAS NOTICED THAT THE TIP WAS BROKEN. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REEN1256 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01456 | 2023-03-06 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD ISSUES WITH PLACEMENT AND WAS FRUSTRATED BECAUSE THE INSTRUCTION WAS NOT CLEAR ABOUT HOW TO PLACE THE PUREWICK FEMALE EXTERNAL CATHETER. THEY EMAILED LINK TO VIDEOS AND PICTURE OF INSERTION AND HOW THE WICKS WERE PLACED AS WELL AS INSTRUCTIONS. THE PATIENT HAD BEEN USING THIS PRODUCT FOR LESS THAN 90 DAYS. PER ADDITIONAL INFORMATION RECEIVED FROM LIBERATOR ON (B)(6)2023 , STATED THAT THE PATIENT HAD LEAKING WHILE USING PUREWICK FEMALE EXTERNAL CATHETER. THE REPRESENTATIVE WENT OVER PROVIDE DIFFERENT TIPS WITH THE PATIENT. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01443 | 2023-03-06 | 2023-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WHEN THE FOLEY CATHETER WAS REMOVED, THE BALLOON COULD NOT BE DEFLATED AND WAS DIFFICULT TO REMOVE. ALSO STATED THAT THE BALLOON PART WAS HARD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01442 | 2023-03-06 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT PATIENT WAS IN MAINTENANCE PHASE, BUT WATER TEMPERATURE CONTINUES TO INCREASE AND STAY WARM ON THE ARCTIC SUN DEVICE. NURSE NOTED THAT PATIENT FEELS WARMER THAN WHAT ESOPHAGEAL PROBE WAS READING, AND THEY HAVE BEEN RECEIVING ERRATIC PATIENT TEMPERATURE ALERTS. NURSE TRIED FLEXING AND UNPLUGGING CABLE ALREADY WITH NO CHANGE. AXILLARY TEMPERATURE READS 100F, BUT ESOPHAGEAL PROBE READS 96F. MIS EXPLAINED THAT THE DEVICE WAS RESPONDING TO PATIENT TEMPERATURE WHICH WAS WHY WATER TEMPERATURE WAS INCREASING. NURSE CONFIRMED THAT THEY HAVE VERIFIED PLACEMENT. MIS ASKED NURSE TO CONNECT ESOPHAGEAL PROBE TO BEDSIDE MONITOR TO VERIFY PROBE REGISTERING, BUT NURSE STATED THAT THEY DID HAVE THAT OPTION AVAILABLE. NURSE NOTED THERE WAS EXPECTED DISCREPANCIES BETWEEN AN AXILLARY TEMPERATURE AND ESOPHAGEAL TEMPERATURE BECAUSE ONE WAS A SURFACE TEMPERATURE VERSUS A TRUE CORE. MIS SUGGESTED THEY SWAP OUT THE TEMPERATURE IN CABLE TO SEE IF THAT WAS THE CULPRIT. THEY WOULD CONTACT BIOMED FOR ANOTHER CABLE AND CALL BACK IF NEEDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01430 | 2023-03-06 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165PL10 | 0165PL10 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OFF DUE TO THE BALLOON BURST AT MIDNIGHT OF THE DAY IT WAS PLACED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00687 | 2023-03-07 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT HAD ISSUES WITH THROAT AND FACE, AND ALLEGEDLY EXPERIENCED BREAKING OUT OF SORES WITH PUS AND SWELLING IN THROAT. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING TREATMENT OR MEDICATION TO TREAT SORES WITH PUS AND SWELLING IN THROAT. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
MW5115544 | 2023-03-07 | 2023-02-14 | I, n, j, u, r, y | Voluntary report | LONG POLYURETHANE BARD MIDLINE CATHETER 4 FRENCH PICC LINE | CATHETER, INTRAVASCULAR, THERAPEUTIC, LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. |  |  |  | ASSESSED MIDLINE THAT IS LEAKING AT THE HUB. FLUSHED MIDLINE IN RIGHT UPPER EXTREMITY WITH SALINE AND IT IS LEAKING AT THE DISTAL END OF THE CATHETER. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01495 | 2023-03-07 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THERE WAS A CRACK ON THE CATHETER BALLOON. THE CATHETER SPONTANEOUSLY FELL OFF AND THERE WAS A DAMAGE OBSERVED ON THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00695 | 2023-03-07 | 2023-02-06 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808000 | 1808000 |  | IT WAS REPORTED THAT FOUR YEARS TWO MONTHS AND TWENTY-FIVE DAYS POST PORT PLACEMENT PROCEDURE, THE CATHETER TUBING ALLEGEDLY BROKEN IN THE CLEAR ADAPTOR PIECE WHERE THE PORT MEETS THE CATHETER AND ATTACHMENT. IT WAS FURTHER REPORTED THAT A PIECE OF THE CATHETER WAS UNABLE TO BE REMOVED FROM THE PATIENT. THE PATIENT WAS UNABLE TO HAVE THE NEW PORT PLACED IN THE CHEST AND REQUIRED TO HAVE A DIFFERENT TYPE OF PORT PLACED IN THE ARM. PATIENT HAD EXTRAVASATION TO THE SITE DURING AN INJECTION. REPORTEDLY, SURGERY WAS PERFORMED TO REMOVE THE PORT. THE PATIENT IS STABLE. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRY DATE: 04/2020. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00694 | 2023-03-07 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, KIT, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1809601 | 1809601 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY HAD PAINS AND SPASMS IN HER NECK THAT RADIATE UP INTO HER JAW AND TONGUE. THERE WAS NO REPORTED PATIENT INJURY. | H10: ADDITIONAL INFORMATION WAS RECEIVED, THERE WAS NO PROBLEM WITH THE DEVICE AND HENCE THIS RECORD WILL NOT BE CONSIDERED AS A COMPLAINT. HOWEVER, SINCE AN INITIAL MDR WAS SUBMITTED, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRATION DATE: 09/2020). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00703 | 2023-03-07 | 2023-01-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY HEALTHCARE PROFESSIONAL "PATIENT ATTENDED FOR TREATMENT. PICC FLUSHED. LEAKAGE NOTED TO EXTERNAL PART OF LINE JUST PASSED THE WINGED AREA. PICC LINE WAS REMOVED AND WAS ABLE TO GET PERIPHERAL CANNULA TO PROCEED WITH TREATMENT. THE PATIENT IS RECEIVING PACLITAXOL. PICC WAS LOCATED IN THE BASILLIC VEIN AND EXIT SITE MEASURED 19.5CM (5CM)." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS THE MANUFACTURING LOT NUMBER PROVIDED BY THE COMPLAINANT IS INVALID. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00702 | 2023-03-07 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY HEALTHCARE PROFESSIONAL "PICC FRACTURE NOTED WHEN FLUSHING WITH SALINE, FRACTURE NOTED. ATTENDED LH AND REMOVED. HOLE NOTED AT 4CM MARK. DOES NOT APPEAR TO HAVE BEEN BENT. PICC WAS INSERTED IN 23.5CM (9CM) AND USED FOR SACT." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFY1412 SHOWED NINE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00699 | 2023-03-07 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN INSERTING THE GUIDE WIRE AND REMOVING THE STYLET FOLLOWING SUCCESSFUL PUNCTURE, THE CUSTOMER FOUND BARBS WITH THE TAIL END OF THE GUIDE WIRE, MAKING IT IMPOSSIBLE TO WITHDRAW THE INTRODUCER NEEDLE. THE OPERATOR OPENED A NEW SELDINGER FOR RE-PUNCTURE. IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGQ2352 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGQ2352) HAVE BEEN REPORTED FROM THE SAME FACILITY IN [COUNTRY]. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00701 | 2023-03-07 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN INSERTING THE GUIDE WIRE AND REMOVING THE STYLET FOLLOWING SUCCESSFUL PUNCTURE, THE CUSTOMER FOUND BARBS WITH THE TAIL END OF THE GUIDE WIRE, MAKING IT IMPOSSIBLE TO WITHDRAW THE INTRODUCER NEEDLE. THE OPERATOR OPENED A NEW SELDINGER FOR RE-PUNCTURE. IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE FIFTH DEVICE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGQ2352 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGQ2352) HAVE BEEN REPORTED FROM THE SAME FACILITY IN [COUNTRY]. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01475 | 2023-03-07 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER HAD ISSUE WITH INDWELLING CATHETER. IT WAS STATED THAT THE CATHETER WAS PLACED ON 9 MONTHS FEMALE IN OPERATING ROOM ON (B)(6) AND REMOVED ON (B)(6) BY A BEDSIDE NURSE. DOCUMENTATION STATED THE CATHETER WAS 6FR FILLED WITH 3CC AND TOOK 2 ATTEMPTS (PR#(B)(4)). UPON REMOVING NURSE DEFLATED THE BULB PASSIVELY AND GOT 3CC BACK, WHEN ATTEMPTED TO PULL THE CATHETER NURSE MET RESISTANCE. WHEN THE CATHETER WAS REMOVED IT WAS NOTED THAT THE BULB HAD FOLDED BACK AND HAD FORMED A RIDGE. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON (B)(6) 2022, NO MEDICAL INTERVENTION WAS REQUIRED. PER FOLLOW-UP VIA EMAIL ON (B)(6) 2023, IT WAS REPORTED THAT THE CUSTOMER HAD ANOTHER ISSUE WITH AN INDWELLING CATHETER ON (B)(6). THE NURSE PLACED AN IN-AND-OUT CATHETER (THE ONE WITH A BAG ATTACHED) TO OBTAIN A SPECIMEN AND RELIEVE BLADDER PRESSURE. APPROXIMATELY 700 CC OF URINE WERE RETURNED. UPON TRYING TO REMOVE THE CATHETER, RESISTANCE WAS FELT, SO THE BEDSIDE NURSE STOPPED AND CALLED THEM. THEY TRIED PUSHING THE CATHETER BACK IN (WITHOUT RESISTANCE) AND APPLIED MORE LUBRICATION BUT WERE NOT ABLE TO REMOVE THE CATHETER, SO THE PHYSICIAN TEAM WAS NOTIFIED, UROLOGY WAS CONSULTED, AND THEY CAME TO THE BEDSIDE AND REMOVED THE CATHETER. IT WAS ALSO STATED THAT THE INTERMITTENT CATHETER PRE-MADE KIT HAD BEEN OPENED. THE IODINE SWAB PACK IN THE KIT CONTAINED ONLY IODINE AND DID NOT CONTAIN ANY SWABS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON (B)(6) 2023, IT WAS REPORTED THAT THE CUSTOMER HAD ANOTHER ISSUE WITH THE FOLEY CATHETER. STATED THAT THE BEDSIDE NURSE ATTEMPTED TO PLACE A CATHETER, WHEN GOING TO INFLATE THE BALLOON MET A LOT OF RESISTANCE AND WAS UNABLE TO INFLATE DESPITE BEING IN BLADDER, AS URINE WAS FILLING BAG. STAFF REMOVED THIS CATHETER AND ATTEMPTED AGAIN, UNABLE TO INFLATE THAT WAS AS WELL. THEN CONSULTED UROLOGY AND THEY CAME TO THE BEDSIDE AND PLACED A FOLEY. STAFF HAVE ATTACHED A PICTURE OF THE FIRST CATHETER AFTER REMOVAL, THEY TRIED TO RE-INFLATE IT AND ONLY ONE SIDE INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00700 | 2023-03-07 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN INSERTING THE GUIDE WIRE AND REMOVING THE STYLET FOLLOWING SUCCESSFUL PUNCTURE, THE CUSTOMER FOUND BARBS WITH THE TAIL END OF THE GUIDE WIRE, MAKING IT IMPOSSIBLE TO WITHDRAW THE INTRODUCER NEEDLE. THE OPERATOR OPENED A NEW SELDINGER FOR RE-PUNCTURE. IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE FOURTH DEVICE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGQ2352 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGQ2352) HAVE BEEN REPORTED FROM THE SAME FACILITY IN [COUNTRY]. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00688 | 2023-03-07 | 2023-01-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER, SINGLE-LUMEN, 4.2F | CHRONIC CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 0600520 | 0600520 |  | IT WAS REPORTED THAT SOME TIME POST CATHETER PLACEMENT PROCEDURE, THE INFUSION SYSTEM THAT WAS CONNECTED WITH THE CHRONIC CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE CATHETER HAS TO BE CUT ALLEGEDLY AND REPAIRED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE.EXPIRY DATE: 09/2025. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00698 | 2023-03-07 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN INSERTING THE GUIDE WIRE AND REMOVING THE STYLET FOLLOWING SUCCESSFUL PUNCTURE, THE CUSTOMER FOUND BARBS WITH THE TAIL END OF THE GUIDE WIRE, MAKING IT IMPOSSIBLE TO WITHDRAW THE INTRODUCER NEEDLE. THE OPERATOR OPENED A NEW SELDINGER FOR RE-PUNCTURE. IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGQ2352 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGQ2352) HAVE BEEN REPORTED FROM THE SAME FACILITY IN [COUNTRY]. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00696 | 2023-03-07 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS CT POWER-INJECTABLE IMPLANTABLE PORT, CHRONOFLEX, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 7480000 | 7480000 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, GUIDEWIRE ALLEGEDLY GOT STUCK INSIDE THE PATIENT. GUIDEWIRE WAS PULLED FORCEFULLY TO REMOVE FROM PATIENT, A SIGNIFICANT AMOUNT OF TISSUE WAS OBSERVED ON THE DISTAL TIP. THE PROCEDURE WAS COMPLETED USING THE SAME DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 08/2024). DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00157 | 2023-03-07 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U415024RX | U415024RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A CALCIFIED TARGET LESION IN THE SUPERFICIAL FEMORAL ARTERY, THE PTA BALLOON ALLEGEDLY RUPTURED AT 6 ATM. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 04/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00156 | 2023-03-07 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A CALCIFIED TARGET LESION IN THE POPLITEAL ARTERY, THE DEVICE WAS ALLEGEDLY UNABLE TO CROSS THE LESION. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY DETACHED AFTER FIVE MINUTES ACTIVATION. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED. MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED FAILURE TO ADVANCE AND MATERIAL SEPARATION ISSUES COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. EXPIRATION DATE: 08/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00154 | 2023-03-07 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE OF A CALCIFIED TARGET LESION, THE TIP OF THE CATHETER ALLEGEDLY SEPARATED AFTER TEN SECONDS ACTIVATION. IT WAS FURTHER REPORTED THAT THE DEVICE STILL ATTACHED TO THE CORE WIRE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE CROSSER CTO RECANALIZATION CATHETER WAS RECEIVED FOR EVALUATION. THE DISTAL TIP WAS NOTED TO BE SEPARATED FROM THE OUTER CATHETER BUT STILL ATTACHED TO THE INNER CORE WIRE. UNDER MICROSCOPIC EXAMINATION A TEAR WAS NOTED TO THE OUTER CATHETER. THE OUTER CATHETER WAS NOTED TO BE FRAYED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED MATERIAL SEPARATION AND THE IDENTIFIED MATERIAL TEAR AND MATERIAL FRAY ISSUES. A DEFINITIVE ROOT CAUSE FOR THE REPORTED MATERIAL SEPARATION AND THE IDENTIFIED MATERIAL TEAR AND MATERIAL FRAY ISSUES COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRY DATE: 10/2024), G3, H6 (DEVICE. METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01469 | 2023-03-07 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119108 | 119108 |  | IT WAS REPORTED THAT STERILIZED WATER LEAKED FROM AROUND THE BASE OF THE PART WHERE STERILE WATER WAS INJECTED. IT WAS INSERTED INTO THE URETHRAL OPENING ONCE BUT WAS REMOVED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 01MAR2023 STATED THAT THE WATER LEAKED AFTER INSERTION AND THE BALLOON COULD NOT BE INFLATED. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01479 | 2023-03-07 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998606 | 998606 |  | IT WAS REPORTED THAT WHEN PREPARING THE BALLOON FOR DILATATION DURING OPERATION, THE ASSISTANT FOUND THAT WATER COULD NOT BE INJECTED INTO THE PRESSURE PUMP AND JUDGED THAT IT MAY BE DUE TO BALLOON DAMAGE. THE SPARE KIT WAS THEREFORE UNPACKED TO PROCEED TO THE OPERATION. THEREAFTER, THE OPERATION WAS DONE SUCCESSFULLY, CAUSING NO ADVERSE IMPACT ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00690 | 2023-03-07 | 2022-12-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6903230 |  | IT WAS REPORTED THAT DURING THE PREPARATION OF A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY CRACKED AND THE NURSES SAW IT BEFORE CONNECTING THE PATIENT TO THE MACHINE. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 11/2023). DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01484 | 2023-03-07 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316M | 119316M |  | IT WAS REPORTED THAT PORT ON THE FOLEY CATHETER BROKE OFF. STATED THAT THE PORT WAS NOT BEING ACCESSED IT JUST POPPED OFF AND NO PATIENT HARM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01491 | 2023-03-07 | 2023-02-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PUREWICK DISPOSABLE | UNKNOWN PUREWICK CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT HAD SUCH A BAD EXPERIENCE IN THE HOSPITAL URINE WAS ALL OVER OF THE PLACE AND SPENT 35 DAYS IN THE HOSPITAL AND GOT OUT WITH A URINARY TRACT INFECTION DUE TO THE PUREWICK FEMALE EXTERNAL CATHETER PRODUCT. NO ADDITIONAL INFORMATION PROVIDE . IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01490 | 2023-03-07 | 2023-02-14 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT STATED THE MAGIC3 HYDROPHILIC MALE INTERMITTENT CATHETER WAS CAUSED BLADDER INFECTION AND GETTING A URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01489 | 2023-03-07 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0102L16 | 0102L16 |  | IT WAS REPORTED THAT THE PATIENT HAD A HARD TIME WITH USING A SPECIFIC FOLEY CATHETER (BATCH#NGGV2931). THEY WERE PHYSICALLY DIFFERENT AND EXTREMELY DIFFICULT, PAINFUL TO USE, PATIENT PREFERS (BATCH# NGGW3097) AND THE PATIENT COULD POSSIBLY GET SENT SPECIFIC ONES AND THE PATIENT WOULD LIKE TO KEEP USING THIS CATHETER. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "TORTUOUS ANATOMY (E.G. ENLARGED PROSTATE) / INADEQUATE DESIGN". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE THE BALLOON TO BURST. CAUTION: DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. SINGLE PATIENT USE ONLY. DO NOT REUSE. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP TIP SYRINGE. DO NOT USE NEEDLE. RECOMMENDED INFLATION CAPACITIES 3CC BALLOON: USE 5ML STERILE WATER 5CC BALLOON: USE 10ML STERILE WATER 15CC BALLOON: USE 20ML STERILE WATER 20CC BALLOON: USE 25ML STERILE WATER 30CC BALLOON: USE 35ML STERILE WATER 40CC BALLOON: USE 45ML STERILE WATER 75CC BALLOON: USE 80ML STERILE WATER DO NOT EXCEED RECOMMENDED CAPACITIES. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE, AND FEDERAL LAWS AND REGULATIONS. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01494 | 2023-03-07 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT AFTER CHECKING THE CUFF OF THE FIXED WATER, A FOLEY CATHETER WAS INSERTED INTO THE PATIENT, BUT IT WAS REMOVED ONCE WITHOUT URINATION. WHEN ANOTHER URINARY CATHETER WAS INSERTED, IT WAS REINSERTED BECAUSE URINATION WAS CONFIRMED. HOWEVER, A NEW ARTICLE WAS USED BECAUSE URINATION WAS NOT OBSERVED. THE NEWLY USED PRODUCT WAS USED AS IT WAS BECAUSE URINATION WAS PERFORMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00697 | 2023-03-07 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN INSERTING THE GUIDE WIRE AND REMOVING THE STYLET FOLLOWING SUCCESSFUL PUNCTURE, THE CUSTOMER FOUND BARBS WITH THE TAIL END OF THE GUIDE WIRE, MAKING IT IMPOSSIBLE TO WITHDRAW THE INTRODUCER NEEDLE. THE OPERATOR OPENED A NEW SELDINGER FOR RE-PUNCTURE. IT WAS REPORTED THIS OCCURRED WITH FIVE DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGQ2352 SHOWED FOUR OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. THE COMPLAINTS FOR THIS LOT NUMBER (REGQ2352) HAVE BEEN REPORTED FROM THE SAME FACILITY IN [COUNTRY]. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00691 | 2023-03-07 | 2023-01-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | A710962 | A710962 |  | IT WAS REPORTED THAT DURING THE PREPARATION OF A PORT PLACEMENT PROCEDURE, THE PORT WAS ALLEGEDLY NOT ABLE TO BE ACCESSED USING TWO DIFFERENT NEEDLES. IT WAS FURTHER REPORTED THAT ON BOTH ATTEMPTS, FLUID WAS ALLEGEDLY NOT ABLE TO BE WITHDRAWN OR FLUSHED. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01527 | 2023-03-08 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THERE WAS A CRACK WHEN THE BALLOON WAS INJECTED AS A TRIAL BEFORE USE ON A PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01545 | 2023-03-08 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THERE WAS A HOLE IN THE FOLEY CATHETER BALLOON AND STERILIZED WATER LEAKED DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01518 | 2023-03-08 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WOKE UP WET USING THE PUREWICK FEMALE EXTERNAL CATHETER. CUSTOMER SAID THAT SOME URINE WENT INTO THE CANISTER AND MOST WENT ON THEM. REPRESENTATIVE SUGGESTED PLACEMENT AND CHECKING FOR INNER TUBE DISPLACEMENT. PER ADDITIONAL INFORMATION ON (B)(6) 2023, PATIENT STILL HAD WETNESS GOING TO THE BED. CUSTOMER STATED THAT PATIENT WAS ADVISED TO PINCH THE CATHETER. CUSTOMER STATED THAT THEY DID NOT GET A BROCHURE ON HOW THE PATIENT SHOULD LAY OR SLEEP. REPRESENTATIVE ADVISED OF POSITIONING OF THE WICK AND THAT THERE HAD TO BE ENOUGH LABIA TO COVER THE WICK. ALSO ADVISED THAT PATIENT SHOULD BE IMMOBILE AND NOT TURN OR TOSS OR THE PATIENT COULD EXPERIENCE LEAKAGE. PATIENT HAD BEEN USING THE PUREWICK PRODUCT FOR LESS THAN 90 DAYS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01528 | 2023-03-08 | 2023-02-15 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER LF HC | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WAS CURRENTLY IN THE HOSPITAL WITH A URINARY TRACT INFECTION. THE CAREGIVER SAID RECENTLY RECEIVED PUREWICK URINE COLLECTION SYSTEM, BUT CLAIMED THE BOX MANUFACTURE BECAUSE THE BOX WAS OPENED WHEN RECEIVED THE ORDER. THE CAREGIVER TOOK PICTURES OF THE BOX BEFORE USING THE SUPPLIES BUT INSTEAD OF CALLING TO REPORT THE ISSUE BEFORE THAT THEY USED THE PUREWICK ON THE PATIENT. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01521 | 2023-03-08 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THERE WAS A HOLE FOUND IN THE FOLEY CATHETER BALLOON WHEN CHECKED BEFORE USE. THE BALLOON WAS FOUND TO BE CRACKED BEFORE USE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00729 | 2023-03-08 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FLEXXICON DIALYSIS CATHETER KIT 11F X 20CM (SHORT-TERM) (DUAL-LUMEN) | CATHETER, SUBCLAVIAN | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5678200 |  | IT WAS REPORTED THAT THERE ARE BLOOD CLOTS FOUND WITHIN CATHETER AFTER 2 DAYS OF THE INSERTION. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: CUSTOMER REPORTED THE CLOT WAS INSIDE THE CATHETER. THE CUSTOMER FOUND THE PUMP MACHINE KEPT ALARMING, SO THEY CHECKED THE CATHETER AND FOUND IT CLOTTED. THEY REMOVED THE CLOTTED CATHETER AND PLACED A NEW ONE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REFT3234 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01515 | 2023-03-08 | 2022-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT AFTER PLACING THE CATHETER ON THE PATIENT, THE PATIENT'S TEMPERATURE WAS NOT DISPLAYED ON THE MONITOR DURING TEMPERATURE MEASUREMENT. WHEN A NEW CATHETER OF ANOTHER TRAY WAS USED WITHOUT CHANGING THE SENSOR CABLE, THE PATIENT'S TEMPERATURE WAS DISPLAYED CORRECTLY. THIS SUGGESTS THAT THE SENSOR CABLE WAS UNLIKELY TO BE THE CAUSE OF NO DISPLAY PROBLEM, AND THE CATHETER MIGHT HAVE ANY DEFECT. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, POTENTIAL ROOT CAUSES FOR THIS FAILURE MODE COULD BE, ¿LONG SHAFT / SHORT THERMISTOR¿, "DAMAGED THERMISTOR", AND "CATHETERS ARE EXCESSIVE PULLED IN CLEANING." THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE." "DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT." "THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES." | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01563 | 2023-03-08 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE WATER IN THE FOLEY CATHETER BALLOON WAS UNABLE TO BE REMOVED AFTER 7 DAYS OF PLACEMENT. A SMALL AMOUNT OF WATER WAS ASPIRATED BY SYRINGE. THE USER TRIED TO DEFLATE THE BALLOON BY SPONTANEOUS OR GENTLE DEFLATION BY SYRINGE, MANUAL ASPIRATION BY SYRINGE AND PUMPING. IT WAS STATED THAT THE USER TRIED TO DEFLATE USING THE SYRINGE FOR A MINUTE AND THE BALLOON WAS PARTIALLY DEFLATED, BUT THE CATHETER COULD BE REMOVED. USUALLY, THE BALLOON CAN BE DEFLATED WITHIN AROUND 5 TO 10 SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00732 | 2023-03-08 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK IV ULTRA PED CATHETER RETAINER | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | IV0500 |  | IT WAS REPORTED THAT THE RN WAS REMOVING A MIDLINE DRESSING AND PREPARING TO REMOVE STATLOCK. STATLOCK NOTED TO BE BROKEN/SHATTERED ON THE BLUE PLASTIC COMPONENT THAT HOLDS CATHETER IN PLACE. UNABLE TO IDENTIFY CAUSE OF BREAKAGE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00733 | 2023-03-08 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER 22GX1.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1221250 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE NURSE HIT THE SAFETY AND THE NEEDLE WOULD NOT RETRACT AFTER USE. THERE WAS NO HARM REPORTED. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGV2119 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01523 | 2023-03-08 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PRODUCT WAS IMPOSSIBLE TO BE WITHDRAWN, THE BALLOON WAS UNABLE TO ASPIRATE WATER, THE LUMEN WAS OCCLUDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00736 | 2023-03-08 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2275108 |  | IT WAS REPORTED BY CLINICAL TEAM THAT WHEN THEY WERE FLUSHING THE LINE, IT WAS LEAKING (LIKELY WHERE THE GUIDE WIRE GOES THROUGH THE CAP) AND DRAWING AIR WHEN THEY PULLED BACK ON THE SYRINGE. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGY0762 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00735 | 2023-03-08 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROVPPSOLO 4FDL MBP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1294108D2 |  | IT WAS REPORTED BY THE CUSTOMER THAT "SHE FEELS LIKE WITH DUAL LUMEN 4 FRENCH PICC LINE THE STOPPER HOLE IS TOO LARGE AND COULD BE SUCKING IN AIR DURING TROUBLE SHOOTING WHILE PLACING A PICC. WHEN FLUSHING THEY SEE FLUID LEAK OUT OF THE STOPPER." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00714 | 2023-03-08 | 2023-01-09 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT THREE MONTHS AND FOURTEEN DAYS POST PORT PLACEMENT, THE CATHETER TIP ALLEGEDLY DISLOCATED FROM VENA CAVA TO VENA JUGULARIS. REPORTEDLY, THE PORT WAS REMOVED. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 11/2023. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00731 | 2023-03-08 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT THE CATHETER PLACEMENT WAS PERFORMED BY THE HEAD NURSE OF THE HEMATOLOGY DEPARTMENT. DURING THE PLACEMENT, AFTER THE INTRODUCER NEEDLE WAS INSERTED SUCCESSFULLY AND THE GUIDE WIRE WAS ENTERED, THE OPERATOR PLANNED TO CROSS THE DISTAL END OF THE GUIDE WIRE INTO THE DILATOR OF THE MICRO-INTRODUCER SHEATH, BUT FAILED. THE PROXIMAL END OF THE DILATOR WAS KINKED IN THE PROCESS OF CONNECTION. A NEW KIT OF INTRODUCER SHEATH WAS UNPACKED. BUT THE DISTAL END OF THE GUIDE WIRE IN THE FIRST KIT WAS STILL UNABLE TO BE CROSSED THROUGH THE DILATOR OF THE NEW MICRO-INTRODUCER SHEATH. THE CATHETER WAS EVENTUALLY PLACED SUCCESSFULLY WITHOUT THE USE OF ANY DILATOR. THE PROCESS TOOK ABOUT 1 HOUR IN TOTAL. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REFV2294 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01524 | 2023-03-08 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT IT WAS DIFFICULT TO DEFLATE DURING PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00719 | 2023-03-08 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK IV ULTRA PED CATHETER RETAINER | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | IV0500 |  | IT WAS REPORTED BY THE CUSTOMER THAT RN ATTEMPTED TO CHANGE T-CONNECTOR AND NOTED A CRACK STAT-LOCK. UNABLE TO IDENTIFY CAUSE OF BREAKAGE. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01514 | 2023-03-08 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE BALLOON WAS INFLATED WITHOUT ANY PROBLEMS DURING THE CUFF CHECK BEFORE INSERTION, BUT IT WAS WITHDRAWN SPONTANEOUSLY AFTER INSERTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01519 | 2023-03-08 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE CATHETER FELL OUT WHEN THE PATIENT STOOD UP, AND WHEN THE BALLOON WAS CHECKED, IT WAS NOT INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00717 | 2023-03-08 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602660 | 0602660 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY HAD PUNCTURE. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY LEAKED WHILE FLUSHING. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 10/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
16501514 | 2023-03-08 | 2023-01-20 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | LUBRI-SIL I.C. COMPLETE CARE TEMP-SENSING FOLEY CATHETER TRAY AND URINE METER | CATHETER, UROLOGICAL (ANTIMICROBIAL) AND ACCESSORIES | C. R. BARD, INC. | 319516AM | 319516AM |  | DURING INSERTION OF FOLEY CATHETER, RN NOTICED PIECE OF HAIR INSIDE THE STERILE GLOVES PACKAGING. THIS RN DISCARDED ENTIRE FOLEY CATHETER KIT AND USED A COMPLETELY NEW ONE. INFORMATION # (B)(4) PACKAGING WITH LOT# AND GLOVES/HAIR SEQUESTERED. MANUFACTURER RESPONSE FOR STERILE GLOVES IN TEMP SENSING FOLEY KIT, STERILE GLOVES IN TEMP SENSING FOLEY KIT (PER SITE REPORTER) [REDACTED DATE] - EMAILED BD REP TO REQUEST AN RMA/MAILER. |  | N | Y |  |  | 2023-03
16501432 | 2023-03-08 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC SV CATHETER WITH SHERLOCK 3CG TIP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 1274118 | 1274118 |  | THE "SHERLOCK" DEVICE USED TO PLACE BEDSIDE PICC LINES ISN'T READING THE TIP OF THE PICC CORRECTLY. THIS MEANS THE PATIENT HAS TO HAVE A CHEST X-RAY TO DETERMINE THE PICC TIP. THE DEPARTMENT IS VASCULAR ACCESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
16501230 | 2023-03-08 | 2023-02-26 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARD URINE METER FOLEY CATHETER TRAY WITH BARDEX LUBRICATH FOLEY CATHETER | CATHETER, COUDE | C. R. BARD, INC. | 304916 | 304916 |  | SAMPLING PORT FELL OFF DRAINAGE BAG. |  | N | Y | OTHER |  | 2023-03
3006260740-2023-00713 | 2023-03-08 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | AIRGUARD VALVED INTRODUCER, 16.5F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 5513165 | 5513165 |  | IT WAS REPORTED THAT PRIOR TO A CATHETER PLACEMENT PROCEDURE, SALINE ALLEGEDLY LEAKED OUT OF THE SYSTEM WHEN SYRINGE WAS ATTACHED TO THE BLUE END OF THE DILATOR LOCK. IT WAS FURTHER REPORTED THAT THE SALINE ALLEGEDLY SQUIRTED OUT OF A SMALL HOLE IN THE CIRCULAR MOLDING OF THE BLUE SECTION. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 04/2023). DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00710 | 2023-03-08 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THIS PATIENT HAD THE CATHETER PLACED IN FACILITY ON (B)(6). PRIOR TO INSERTION, THE HEAD NURSE OF THE CARING CLINIC PRE-FLUSHED THE CATHETER AND FOUND LIQUID LEAKING FROM THE DISTAL END. HOWEVER, THE HEAD NURSE CHOSE TO TRIM THE CATHETER BY 1 CM AND INSERTED IT IN THIS PATIENT. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFW2825 SHOWED TWO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00709 | 2023-03-08 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC DL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED "DURING CLINICAL ASSESSMENT, FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. POWERPICC DL PLACED IN UPPER ARM." NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00704 | 2023-03-08 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY HEALTHCARE PROFESSIONAL, "PICC LINE INSERTED 230119. DURING TREATMENT 230209, HOLE IS DISCOVERED ON THE TUBE 28 MM PROXIMAL TO THE WINGS AFTER GIVEN TREATMENT. DISCOVERED THAT IT WAS WET IN THE BED AND UNDER THE TEGADERM. WHEN FLUSHING, DROPS ARE VISIBLE FROM THE HOSE. PREPARATIONS GIVEN WERE CYCLOPHOSPHAMIDE, DOXORUBICIN, VINCRISTINE AND RITUXIMAB. THERE WAS NO IMPACT ON THE PATIENT OTHER THAN DELAYED TREATMENT AND SPILLAGE OF CYTOSTATICS." | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGU1734 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00723 | 2023-03-08 | 2023-02-10 | I, n, j, u, r, y | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | A710962 | A710962 |  | IT WAS REPORTED THAT SOME TIME POST PORT PLACEMENT, THE PATIENT IS CURRENTLY GETTING IV FLUIDS, ANTIBIOTICS AND BEING TREATED FOR SEPSIS. PATIENT CURRENT STATUS IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DEVICE EXPIRY DATE: 04/2021. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01506 | 2023-03-08 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI18 |  | IT WAS REPORTED THAT THE CATHETER BALLOON WAS DAMAGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01501 | 2023-03-08 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SILICONE FOLEY CATHETER | UNKNOWN SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT THE CUSTOMER HAD ISSUE WITH INDWELLING CATHETER. IT WAS STATED THAT THE CATHETER WAS PLACED ON 9MONTHS FEMALE IN OPERATING ROOM ON (B)(6) AND REMOVED ON (B)(6) BY A BEDSIDE NURSE. DOCUMENTATION STATED THE CATHETER WAS 6FR FILLED WITH 3CC AND TOOK 2 ATTEMPTS. UPON REMOVING NURSE DEFLATED THE BULB PASSIVELY AND GOT 3CC BACK, WHEN ATTEMPTED TO PULL THE CATHETER NURSE MET RESISTANCE. WHEN THE CATHETER WAS REMOVED IT WAS NOTED THAT THE BULB HAD FOLDED BACK AND HAD FORMED A RIDGE. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON 24DEC2022, NO MEDICAL INTERVENTION WAS REQUIRED. PER FOLLOW-UP VIA EMAIL ON 17JAN2023, IT WAS REPORTED THAT THE CUSTOMER HAD ANOTHER ISSUE WITH AN INDWELLING CATHETER ON 16JAN. THE NURSE PLACED AN IN-AND-OUT CATHETER (THE ONE WITH A BAG ATTACHED) TO OBTAIN A SPECIMEN AND RELIEVE BLADDER PRESSURE. APPROXIMATELY 700 CC OF URINE WERE RETURNED. UPON TRYING TO REMOVE THE CATHETER, RESISTANCE WAS FELT, SO THE BEDSIDE NURSE STOPPED AND CALLED THEM. THEY TRIED PUSHING THE CATHETER BACK IN (WITHOUT RESISTANCE) AND APPLIED MORE LUBRICATION BUT WERE NOT ABLE TO REMOVE THE CATHETER, SO THE PHYSICIAN TEAM WAS NOTIFIED, UROLOGY WAS CONSULTED, AND THEY CAME TO THE BEDSIDE AND REMOVED THE CATHETER. IT WAS ALSO STATED THAT THE INTERMITTENT CATHETER PRE-MADE KIT HAD BEEN OPENED. THE IODINE SWAB PACK IN THE KIT CONTAINED ONLY IODINE AND DID NOT CONTAIN ANY SWABS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON 14FEB2023, IT WAS REPORTED THAT THE CUSTOMER HAD ANOTHER ISSUE WITH THE FOLEY CATHETER. STATED THAT THE BEDSIDE NURSE ATTEMPTED TO PLACE A CATHETER, WHEN GOING TO INFLATE THE BALLOON MET A LOT OF RESISTANCE AND WAS UNABLE TO INFLATE DESPITE BEING IN BLADDER, AS URINE WAS FILLING BAG. STAFF REMOVED THIS CATHETER AND ATTEMPTED AGAIN, UNABLE TO INFLATE THAT WAS AS WELL. THEN CONSULTED UROLOGY AND THEY CAME TO THE BEDSIDE AND PLACED A FOLEY. STAFF HAVE ATTACHED A PICTURE OF THE FIRST CATHETER AFTER REMOVAL, THEY TRIED TO RE-INFLATE IT AND ONLY ONE SIDE INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00707 | 2023-03-08 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERGLIDE | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THE WIRE WAS MISSING FROM THE BLOOD CHAMBER, WHEN PULLING BACK THE HOUSING FROM THE CATHETER. THE WIRE APPEARED TO BE AT THE BOTTOM OF THE INSIDE OF THE CATHETER HOUSING. ADDITIONAL INFORMATION RECEIVED 3/1/2023 - IT WAS REPORTED BY THE CUSTOMER "WE IMMEDIATELY DISCARDED IN THE SHARPS TO PREVENT ACCIDENTAL STICK INJURY." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00721 | 2023-03-08 | 2023-02-16 | I, n, j, u, r, y | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER (20G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202252 |  | IT WAS REPORTED THAT THE CLINICIAN WAS ATTEMPTING PLACEMENT OF A 20G 2.25¿ ACCUCATH IN LEFT FOREARM. HE VISUALIZED NEEDLE TIP IN CENTER OF VEIN AND ADVANCED FULL SYSTEM, THEN BEGAN DEPLOYING GUIDEWIRE, THEN HIT RESISTANCE WITH GUIDEWIRE ABOUT HALFWAY THROUGH DEPLOYMENT. AFTER HITTING RESISTANCE, HE BEGAN TO RETRACT GUIDEWIRE AND IT BECAME STUCK. HE WAS UNABLE TO RETRACT THE GUIDEWIRE ANY SIGNIFICANT AMOUNT (UPON RESISTANCE UNABLE TO BEGIN RETRACTION). THE CLINICIAN'S FOCUS WAS ON PATIENT SAFETY AND PRESERVING INTEGRITY OF GUIDEWIRE. HE TRIED DIFFERENT WAYS TO LOOSEN BUT UNABLE AND THE ACCUCATH WAS STILL RESISTANT TO RETRACTION OF GUIDEWIRE. THEN, TO ELIMINATE COMPONENTS, HE PRESSED THE SAFETY BUTTON TO RETRACT THE NEEDLE (NEEDLE RETRACTED SUCCESSFULLY). CLINICIAN PULLED REMAINING COMPONENTS (CATHETER AND GUIDEWIRE) AS A SYSTEM OUT OF PATIENT TISSUE, AS GENTLY AS POSSIBLE, AND ALMOST REMOVED BOTH COMPONENTS ENTIRELY HOWEVER JUST BEFORE CATHETER/GUIDEWIRE WAS FULLY OUT, HE FELT A SNAP. UPON INSPECTION AFTER REMOVAL, THE COIL TIP LOOKED BLUNT AS THOUGH THE COIL TIP HAD SNAPPED OFF. HE THEN WENT BACK WITH ULTRASOUND TO EVALUATE THE SITE AND IDENTIFIED WHAT LOOKED LIKE THE SHINY COIL TIP PORTION REMAINING IN DERMAL TISSUE, SUPERFICIAL TO VEIN. IT APPEARED AS WHITE DOT. CLINICIAN WAS UNABLE TO SEE IT OR FEEL IT BY PALPATION BUT COULD VISUALIZE IT UNDER ULTRASOUND. CLINICIAN REPORTED TO ICU ATTENDING AND ER EDUCATOR. INTENSIVIST EVALUATED AND REGARDED IT AS POTENTIALLY EQUIVALENT TO DERMAL STAPLE THAT WOULD SHED WITH TIME. ADDITIONAL INFORMATION RECEIVED: NO HARM OCCURRED DESPITE THE MALFUNCTION IN THE PRODUCT. NO SECUREMENT WAS USED BECAUSE THE DEVICE WAS NOT PLACED, AND THERE WAS NO EXPLANT DATE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01543 | 2023-03-08 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILIZED WATER DID NOT COME OUT WHEN THE USER INFLATED THE FOLEY CATHETER BALLOON DURING THE LEAK TEST AND TRIED TO RETURN THE STERILIZED WATER TO THE SYRINGE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. . | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00720 | 2023-03-08 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PICC CATHETER LEAKED LIQUID, LEADING TO FAILURE OF THE PUNCTURE. THE CATHETER WAS THEN REMOVED AND A NEW KIT OF CATHETER WAS OPENED. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REFZ3490 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01511 | 2023-03-08 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT FIXATION WATER WAS UNABLE TO REMOVE. THE EVENT WAS FOUND DURING PRETEST AND THE DEVICE WAS NOT USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01516 | 2023-03-08 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THESE MAGIC 3 GO MALE HYDROPHILIC INTERMITTENT CATHETERS WERE VERY HARD TO INSERT, AND THE PATIENT HAD SOME BLEEDING WHEN THEY TOOK THEM OUT. THE PATIENT STATED THAT IT DID NOT HAVE EXTRA LUBRICANT ON THEM. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00727 | 2023-03-08 | 2023-02-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT POST PORT PLACEMENT PROCEDURE, THE PORT WAS NOT INFECTED AND THE PORT WAS WORKING AS INTENDED. THERE WAS NO REPORTED PATIENT INJURY. | H10: ADDITIONAL INFORMATION RECEIVED THAT THE PORT WAS NOT INFECTED AND IT WAS WORKING AS INTENDED. THUS THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A SERIOUS INJURY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00725 | 2023-03-08 | 2023-02-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT POST PORT PLACEMENT PROCEDURE, THE PORT WAS NOT INFECTED AND THE PORT WAS WORKING AS INTENDED. THERE WAS NO REPORTED PATIENT INJURY. | H10: ADDITIONAL INFORMATION RECEIVED THAT THE PORT WAS NOT INFECTED AND IT WAS WORKING AS INTENDED. THUS THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A SERIOUS INJURY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01542 | 2023-03-08 | 2023-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT STERILE WATER LEAKED, AND THE FOLEY CATHETER FELL OUT. IT WAS UNCLEAR WHETHER THE LEAKAGE WAS FROM THE BALLOON PART OR NEAR THE VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00706 | 2023-03-08 | 2023-01-23 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO 4F SINGLE-LUMEN CATHETER FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6194108 |  | IT WAS REPORTED BY CUSTOMER "GUIDE WIRE GOT TWISTED UNTIL IT FORMED A KNOT, WHILE IT WAS BEING ADVANCED INTO PATIENT VESSEL. ONCE INTRODUCER WAS INSERTED, CLINICIAN WAS UNABLE TO REMOVE THE WIRE. AFTER SEVERAL FAILED ATTEMPTS, A CHEST RAY WAS ORDERS THAT SHOWED THE KNOT IN THE WIRE. STAFF HAD TO SEEK GUIDANCE FROM IR CONSULTANT WHO GAVE ADVICE VIA VIDEO CALL. THE WHOLE PROCESS TO REMOVE THE GUIDEWIRE TOOK OVER AN HOUR." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, COMPLAINT AND BATCH HISTORY, APPLICABLE PREVIOUS INVESTIGATION(S), APPLICABLE MANUFACTURING RECORDS, PHOTO ANALYSIS AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT THAT THE GUIDEWIRE BECAME KNOTTED UPON ITSELF WAS CONFIRMED. FIVE PHOTOGRAPHS OF THE FLOPPY DISTAL TIP OF THE IMPLICATED GUIDEWIRE WERE RETURNED FOR EVALUATION. RESIDUAL MATERIAL, WHICH APPEARED CONSISTENT WITH CLINICAL USE, WAS SEEN ALONG THE WIRE. THE FLOPPY DISTAL TIP OF THE GUIDEWIRE WAS KNOTTED UPON ITSELF. NO DAMAGE WAS SEEN TO THE DISTAL WELD TIP OF THE WIRE AND NO BREAKS WERE SEEN IN THE WIRE MATERIAL. A SMALL SECTION OF COILS, IN THE COIL WIRE, WERE SLIGHTLY OUT OF ALIGNMENT WHICH LIKELY OCCURRED DUE TO THE TIGHT RADIUS OF THE KNOT. IN ADDITION, A SLIGHTLY BENT SHAPE MEMORY WAS SEEN IN THE COILED SECTION OF THE GUIDEWIRE PROXIMAL OF THE KNOT. DUE TO THE EVIDENCE OF USE ON THE GUIDEWIRE, IT APPEARED THAT THE WIRE INADVERTENTLY FOLDED OVER ITSELF AND WAS TWISTED INTO A KNOT DURING THE PROCEDURE. HAD THE KNOT EXISTED PRIOR TO USE, THE WIRE WOULD NOT HAVE FIT THROUGH THE NEEDLE. UNKNOWN FACTORS, SUCH AS TORTUOUS VASCULATURE OR A DIFFICULT PLACEMENT MAY HAVE CONTRIBUTED TO THIS EVENT. A REVIEW OF THE INSPECTION RECORDS AND QUALITY/MANUFACTURING CONTROLS FOR THE IMPLICATED LOT INDICATED NO ISSUES WITH THE MANUFACTURING PROCESS. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00724 | 2023-03-08 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) SL (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED THAT DURING PLACEMENT OF THE POWERPICC SOLO IN THIS PATIENT, THE OPERATOR PERFORMED PUNCTURE SUCCESSFULLY AND INSERTED THE GUIDE WIRE. DURING REMOVAL OF THE INTRODUCER NEEDLE, IT GOT STUCK BY THE DISTAL END OF THE GUIDE WIRE AND COULD NOT BE REMOVED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGT2990 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00722 | 2023-03-08 | 2023-01-13 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK TLS TECHNOLOGY 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9295108D |  | IT WAS REPORTED VIA MEDWATCH THAT THEY WERE UNABLE TO PASS THE PICC THROUGH THE BRACHIAL PAST 11CM. IT WAS DIFFICULT TO REMOVE THE LINE AS THOUGH THE VESSEL AS VASO SPASMED. AFTER WAITING APPROXIMATELY 10 SECONDS, THE PATIENT WAS REPOSITIONED, AND CLINICIAN ATTEMPTED TO PASS 11 CM UNSUCCESSFULLY. BASILIC AND CEPHALIC VEINS ARE TOO SMALL FOR THE PICC ON THE RIGHT ARM SO THE CLINICIAN MOVED TO THE LEFT SIDE. STERILE FIELD AS MAINTAINED, AND CLINICIAN FOUND APPROPRIATELY SIZED BASILIC VEIN AND CONTINUED THE PROCEDURE. CATHETER PASSED EASILY. CATHETER DRESSED PER HOSPITAL PROTOCOL. CUSTOMER REPORTS A SMALL WIRE FRAGMENT WAS PRESENT MEDIAL TO THE PROXIMAL SHAFT OF THE RIGHT HUMERUS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A BROKEN STYLET WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE TLS STYLET. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND CONFIRMED TO BE BROKEN. THE FOLLOWING OBSERVATIONS WERE MADE WHICH WERE CONSISTENT WITH THIS FAILURE TYPE: MICROSCOPIC EXAMINATION REVEALED DEFORMATION OF THE STYLET TUBING AT MAGNET EDGE(S) MICROSCOPIC EXAMINATION REVEALED LOCALIZED DELAMINATION OF THE STYLET TUBING AT THE DAMAGE SITE(S), INDICATIVE OF TUBING KINKING/FOLDING KINKING OF THE SURROUNDING STYLET TUBING, LOCATED AT MAGNET INTERFACE(S) THIS ADDITIONAL STYLET KINKING RESULTED IN A GENERAL CURL IN A SIMILAR DIRECTION MEASUREMENTS OF THE STYLET TUBING, STYLET CORE WIRE, AND MAGNET WERE TAKEN AND CONFIRMED TO BE WITHIN SPECIFICATIONS. ALTHOUGH THE BREAK IN THE WIRE COULD BE CONFIRMED, THE ROOT CAUSE COULD NOT BE DETERMINED THROUGH EXAMINATION OF THE RETURNED SAMPLE. THE OBSERVED FEATURES WERE INDICATIVE OF CIRCUMSTANCES WHICH INCLUDED STYLET KINKING, LIKELY DURING THE INSERTION PROCESS. IT APPEARED THAT ADDITIONAL UNIDENTIFIED FACTORS MAY HAVE ALSO CONTRIBUTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00718 | 2023-03-08 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG PLUS TCS | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770351 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE P-WAVE TECHNOLOGY WAS NOT CONSISTENTLY CAPTURING THEIR PEAKED P-WAVES ACCURATELY, CAUSING THEM TO LEAVE TOO MUCH EXTERNAL LENGTH OR ADVANCING TOO FAR IN WITH CATHETER ON THE SMALL TABLET. SAID TECHNOLOGY WAS INCONSISTENT CONFIRMING PLACEMENT, AS SOMETIMES IT WOULD WORK PROPERLY. BUT DUE TO THEIR LACK OF CONFIDENCE IN THE TECHNOLOGY, THEY STILL ORDERED CXR TO CONFIRM TIP PLACEMENT. | THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01591 | 2023-03-09 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE INTERFACE BETWEEN THE FUNNEL AND THE SHAFT OF THE CATHETER WAS DAMAGED (CATHETER BREAKAGE) AND URINE LEAKED. PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON 06MAR2023, STATED THAT THERE WAS A TEAR IN THE CATHETER SHAFT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01592 | 2023-03-09 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT UPON ATTEMPTING TO INJECT STERILE WATER INTO THE CATHETER DURING A PRETEST, THE WATER LEAKED FROM THE FUNNEL PART DUE TO A CRACK ON THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00754 | 2023-03-09 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERGLIDE | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED BY THE CUSTOMER THEY ARE EXPERIENCING AN INCREASE IN KINKING IN ALL GAUGES AND LENGTHS OF POWERGLIDE. CUSTOMER STATES THAT THE KINKING IS RESOLVED WHEN A DIFFERENT (PIVC) STATLOCK IS UTILIZED. CUSTOMER FEELS THAT THERE IS A CHANGE IN THE ANGLE OF THE POWERGLIDE STATLOCK. THE CUSTOMER REPORTED THIS IS A RECURRING EVENT HOWEVER THE EXACT QUANTITY OR DEVICE SPECIFICATIONS WERE NOT PROVIDED TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01595 | 2023-03-09 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT AFTER INSERTION, THE BALLOON BURST AND THE FOLEY CATHETER WAS REMOVED. ALSO STATED THAT BALLOON RUPTURE OR TEAR OCCURRED AFTER 1 HOUR OF PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01598 | 2023-03-09 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT UPON INJECTING WATER INTO THE FOLEY CATHETER BALLOON, THE WATER LEAKED. IT WAS UNKNOWN WHERE IT IS COMING FROM. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01601 | 2023-03-09 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE INTERMITTENT HURTS WHEN THE PATIENT PULLED IT OUT, THEY WERE NOT. THE PATIENT STATED THAT THESE INTERMITTENT CATHETERS WERE NOT LUBED ENOUGH, AND THEY MIGHT NEED ANOTHER. NO MEDICAL INTERVENTION WAS REPORTED. PER CLARIFICATION MAIL ON (B)(6)2023, PATIENT SAID IT HURTS WHEN THEY PULLED IT OUT, THEY WERE NOT LUBED ENOUGH. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE HYDROPHILIC COATED SILICONE INSTRUCTIONS FOR USE INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00755 | 2023-03-09 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT DURING INSERTION OF THE PICC CATHETER, THE GUIDE WIRE WAS TOO THICK TO PASS THROUGH THE INTRODUCER SHEATH. AFTER TRIMMED, IT STILL COULD NOT PASS THE INTRODUCER SHEATH. THEREFORE, A NEW CANNULA KIT WAS USED AND THE GUIDE WIRE STILL COULD NOT PASS. THIS PATIENT EXPRESSED GREAT DISSATISFACTION. THE THICKER PORTION OF THE GUIDE WIRE AT THE DISTAL END WAS THEN CUT OFF AND THE CATHETER WAS PLACED SUCCESSFULLY. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFY3459 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. THE COMPLAINTS FOR THIS BATCH NUMBER HAVE BEEN REPORTED FROM THE SAME FACILITY IN CHINA. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01602 | 2023-03-09 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165820 | 165820 |  | IT WAS REPORTED THAT THE CATHETER BALLOON HAD BEEN RUPTURED BEFORE USE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00756 | 2023-03-09 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT DURING INSERTION OF THE PICC CATHETER, THE GUIDE WIRE WAS TOO THICK TO PASS THROUGH THE INTRODUCER SHEATH. AFTER TRIMMED, IT STILL COULD NOT PASS THE INTRODUCER SHEATH. THEREFORE, A NEW CANNULA KIT WAS USED AND THE GUIDE WIRE STILL COULD NOT PASS. THIS PATIENT EXPRESSED GREAT DISSATISFACTION. THE THICKER PORTION OF THE GUIDE WIRE AT THE DISTAL END WAS THEN CUT OFF AND THE CATHETER WAS PLACED SUCCESSFULLY. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFY3459 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. THE COMPLAINTS FOR THIS BATCH NUMBER HAVE BEEN REPORTED FROM THE SAME FACILITY IN CHINA. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00753 | 2023-03-09 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 4808560 | 4808560 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE DEVICE ALLEGEDLY BROKE AND MIGRATED INTO PATIENT BODY. REPORTEDLY, THE DEVICE REMAINED IN THE PATIENT BODY. THE CURRENT PATIENT STATUS WAS UNKNOWN. | H10: ADDITIONAL INFORMATION WAS RECEIVED, AND THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE REPORTABLE AS SERIOUS INJURY. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00752 | 2023-03-09 | 2023-02-01 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808060 | 1808060 |  | IT WAS REPORTED THAT APPROXIMATELY TWO WEEKS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED CHEST PAIN. IT WAS FURTHER REPORTED THAT ON CHEST IMAGING IT WAS FOUND THAT THE CATHETER PORTION OF CHEST PORT ALLEGEDLY BROKE OFF AND DETACHED. REPORTEDLY, THE TIP WAS ALLEGEDLY MIGRATED FROM SUPERIOR VENA CAVA TO RIGHT DISTAL ATRIUM/VENTRICULAR JUNCTION. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01584 | 2023-03-09 | 2022-11-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE PATIENT INDWELLED 18-GAUGE THREE-LUMEN 30 ML BALLOON FOLEY CATHETER ON (B)(6) 2022. AFTER GIVING BLADDER IRRIGATION AND CLAMPING THE CATHETER FOR 30 MINUTES, THE NURSE FOUND THAT THE PATIENT'S CATHETER FELL OUT AND DISCHARGED ABOUT 50 ML OF PINK LIQUID ON ITS OWN. UPON INSPECTION, THE BALLOON OF THE CATHETER WAS RUPTURED AND THE PATIENT HAD NO COMPLAINTS OF DISCOMFORT . THEN THE PATIENT WAS RE-INDWELLED WITH A BARD DISPOSABLE 18-GAUGE THREE-LUMEN 30 ML BALLOON FOLEY CATHETER. THE CATHETER BALLOON WAS ROUTINELY INSPECTED BEFORE INSERTING, AND IT WAS FOUND THAT THE BALLOON COULD NOT BE COMPLETELY DEFLATED AFTER FILLING WITH SALINE . THEREFORE, ANOTHER NEW CATHETER WAS USED AND THE PROCESS WENT SMOOTHLY. AS PER DESCRIPTION OF EVENT CAUSE ANALYSIS, STATED THAT THE CAUSE WAS UNKNOWN, WAITING FOR URINARY CATHETER INSPECTION AND REASONS FOR PRODUCT QUALITY. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 24FEB2023, CLARIFIED THAT THE DEFECT FOR THE SECOND CATHETER WAS FOUND BEFORE USING IN THE PATIENT BECAUSE THE NURSE WAS USED TO DOING PRE-FILING THE BALLOON TO TEST THE FUNCTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). A REVIEW OF THE DEVICE LABELING HAVE BEEN CONDUCTED FOR INVESTIGATION PURPOSE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEM OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATER ALLERGY." SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX". | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01600 | 2023-03-09 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT FOLEY CATHETER BALLOON BURST INSIDE BLADDER, APPROXIMATELY 1 MONTH FROM INSERTION. THE USER REINSERTED A NEW CATHETER. AS PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 23FEB2023, THERE ISN¿T ANY MISSING PIECES OF THE BALLOON. | THE REPORTED EVENT IS INCONCLUSIVE DUE TO CONDITION OF THE SAMPLE RECEIVED. VISUAL EVALUATION NOTED ONE PHOTO SAMPLE RECEIVED SHOWING A CATHETER WITH BALLOON WAS BURST. NO FUNCTIONAL TESTING COULD BE PERFORMED AS ONLY PHOTOS ARE PROVIDED AS SAMPLE. IT IS UNKNOWN IF THE PRODUCT MEETS SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01585 | 2023-03-09 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT A CRACK WAS FOUND IN THE FOLEY CATHETER SHAFT, AND URINE LEAKED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01577 | 2023-03-09 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165814 | 165814 |  | IT WAS REPORTED THAT THE 2 OUT OF 3 FOLEY CATHETERS WOULD NOT DEFLATE AND ISSUE WAS FOUND 2 WEEKS AGO. PER FOLLOW-UP VIA EMAIL ON 26JAN2023, BALLOON WOULD NOT DEFLATE WHEN CHECKED BEFORE INSERT AND 2 CATHETERS FROM THIS LOT PERFORMED THE SAME AND NO HARM WAS DONE. PER NOTIFICATION FROM UCC ON 16FEB2023, IT WAS NOTED THAT THE SILICONE FOLEY CATHETER WITH BATCH# NGGV4482 WAS RETURNED FOR EVALUATION. PER ADDITIONAL INFORMATION RECEIVED VIA SAMPLE FORM ON (B)(6) 21FEB2023, IT WAS REPORTED THAT WHEN TESTING THE FOLEY BALLOONS TO BE SURE THEY DEFLATE, NEITHER OF THESE WOULD DEFLATE. STATED THAT THEY WERE NOT USED AND OTHER CATHETERS IN THE SHIPMENT WERE NOT DEFECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
16512039 | 2023-03-09 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® LUBRICATH® TEMPERATURE SENSING CATHETER | CATHETER, COUDE | C. R. BARD, INC. | A900216 | A900216 |  | AFTER MOVING PATIENT TO A BED IN THE POST-ANESTHESIA CARE UNIT, SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER SLID OUT OF PATIENT WHEN NURSE WAS MOVING FOLEY BAG. IT WAS OBSERVED THAT THE BALLOON WAS NOT INFLATED AT THE END OF THE CASE. THE FOLEY BALLOON WAS INFLATED AT BEGINNING OF CASE BECAUSE THERE WAS URINE RETURN. ONE OF THE CIRCULATING REGISTERED NURSES (RNS) CHECKED THE FOLEY BALLOON TO SEE IF IT HAD A LEAK. IT WAS FOUND THAT THE BALLOON WAS INTACT BUT THE INJECTION PORT TO INFLATE THE BALLOON WAS LEAKING SLOWLY. |  | N | Y |  |  | 2023-03
3006260740-2023-00738 | 2023-03-09 | 2023-02-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 18G X 10 CM WITH GUARDIVA AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F118108PT |  | IT WAS REPORTED BY THE CUSTOMER THAT "SAFETY ON POWERGLIDE PRO FAILED TO ENGAGE." CUSTOMER REPORTS THERE WAS NO PATIENT IMPACT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGU2409 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00744 | 2023-03-09 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 22 GAUGE POWERGLIDE PRO MIDLINE CATHETER 8CM INTERMEDIATE KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6F122081 |  | IT WAS REPORTED, THE CATHETER WAS BLOCKED WITH THE GUIDEWIRE, WHILE IN THE PATIENT'S VEIN. THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. THE WIRE DID NOT BREAK INTO MULTIPLE PIECES, NO FRAGMENTS WERE LEFT IN THE PATIENT, WIRE WAS COMPLETELY REMOVED INTACT. SMALL SKIN INCISION WAS MADE TO RETRIEVE CATHETER AND GUIDEWIRE STUCK IN SUBCUTANEOUS TISSUE. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFY1175 SHOWED THREE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00745 | 2023-03-09 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 22 GAUGE POWERGLIDE PRO MIDLINE CATHETER 8CM INTERMEDIATE KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6F122081 |  | IT WAS REPORTED, THE CATHETER WAS BLOCKED WITH THE GUIDEWIRE, WHILE IN THE PATIENT'S VEIN. THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. THE WIRE DID NOT BREAK INTO MULTIPLE PIECES, NO FRAGMENTS WERE LEFT IN THE PATIENT, WIRE WAS COMPLETELY REMOVED INTACT. SMALL SKIN INCISION WAS MADE TO RETRIEVE CATHETER AND GUIDEWIRE STUCK IN SUBCUTANEOUS TISSUE. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFY1175 SHOWED THREE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00746 | 2023-03-09 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK IV ULTRA PED CATHETER RETAINER | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | IV0500 |  | IT WAS REPORTED BY THE CUSTOMER "THERE ARE MULTIPLE INSTANCES OF THE BLUE PLASTIC COMPONENT BREAKING WHEN ALREADY IN PLACE ON A CATHETER AS WELL AS WHEN PLACING ON A CATHETER." NO OTHER INFORMATION WAS PROVIDED. THE CUSTOMER REPORTED MULTIPLE EVENTS BUT THE EXACT NUMBER OF EVENTS WAS NOT PROVIDED TO BD. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00750 | 2023-03-09 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F WITH MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1295108D |  | IT WAS REPORTED BY THE CUSTOMER "SHE IS HAVING ISSUES FLUSHING INTO THE PICC AND CANNOT GET A BLOOD RETURN. THE PATIENT HAD THE CATHETER PLACED (B)(6) 2023. CUSTOMER STATES SHE HAS REPOSITIONED THE PATIENT MULTIPLE TIMES. SHE TOOK THE NEEDLELESS CONNECTOR OFF OF THE PICC AND TRIED TO FLUSH DIRECTLY INTO THE PICC BUT WAS STILL UNSUCCESSFUL. SHE CALLED THE DOCTOR AND GOT AN ORDER FOR AN X-RAY. THE PICC WAS X-RAYED AND CONFIRMED IT IS PLACED PROPERLY. SHE STATES THE PATIENT RETURNED TO HER AND SHE IS STILL UNABLE TO FLUSH INTO THE PICC OR GET A BLOOD RETURN. CUSTOMER DID DESCRIBE THE PATIENT'S INTERNAL ANATOMY AS 'WEIRD'." MS&S RESPONSE: "EXPLAINED THAT BECAUSE THE PICC IS SO NEW IT IS UNLIKELY THAT FIBRIN HAS OCCLUDED THE LINE AT THIS TIME. CALLER CONFIRMS SHE IS FLUSHING WITH A 10ML SYRINGE AND THERE ARE NO CLAMPS ON THE DEVICE. DISCUSSED THAT IT IS POSSIBLE THAT THE CATHETER IS KINKED AT OR NEAR THE AXILLA AREA, POSSIBLY RELATED TO THE PATIENT'S ANATOMY. SUGGESTED THE PICC MAY NEED TO BE EVALUATED AGAIN VIA DYE STUDY TO DETERMINE IF THE PATIENT'S ANATOMY WILL SUPPORT A PICC LINE. IT IS POSSIBLE THAT THE PICC LINE ITSELF IS DAMAGED OR NOT WORKING AS INTENDED. EXPLAINED THAT WE WILL FORWARD THIS REPORT TO OUR PRODUCT COMPLAINTS TEAM." | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGQ0646 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00743 | 2023-03-09 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 22 GAUGE POWERGLIDE PRO MIDLINE CATHETER 8CM INTERMEDIATE KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6F122081 |  | IT WAS REPORTED, THE CATHETER WAS BLOCKED WITH THE GUIDEWIRE, WHILE IN THE PATIENT'S VEIN. THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. THE WIRE DID NOT BREAK INTO MULTIPLE PIECES, NO FRAGMENTS WERE LEFT IN THE PATIENT, WIRE WAS COMPLETELY REMOVED INTACT. SMALL SKIN INCISION WAS MADE TO RETRIEVE CATHETER AND GUIDEWIRE STUCK IN SUBCUTANEOUS TISSUE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFY1175 SHOWED THREE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00737 | 2023-03-09 | 2023-02-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FLEXXICON DIALYSIS CATHETER KIT 11F X 20CM (SHORT-TERM) (DUAL-LUMEN) | CATHETER, SUBCLAVIAN | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5678200 |  | THE CUSTOMER REPORTS THAT "FREQUENTLY" THERE ARE BLOOD CLOTS FOUND WITHIN CATHETER AFTER 2 DAYS OF THE INSERTION. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REFT3234 SHOWED ONE OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01580 | 2023-03-09 | 2022-11-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE PATIENT INDWELLED 18-GAUGE THREE-LUMEN 30 ML BALLOON FOLEY CATHETER ON (B)(6) 2022. AFTER GIVING BLADDER IRRIGATION AND CLAMPING THE CATHETER FOR 30 MINUTES, THE NURSE FOUND THAT THE PATIENT'S CATHETER FELL OUT AND DISCHARGED ABOUT 50 ML OF PINK LIQUID ON ITS OWN. UPON INSPECTION, THE BALLOON OF THE CATHETER WAS RUPTURED AND THE PATIENT HAD NO COMPLAINTS OF DISCOMFORT. THEN THE PATIENT WAS RE-INDWELLED WITH A BARD DISPOSABLE 18-GAUGE THREE-LUMEN 30 ML BALLOON FOLEY CATHETER. THE CATHETER BALLOON WAS ROUTINELY INSPECTED BEFORE INSERTING, AND IT WAS FOUND THAT THE BALLOON COULD NOT BE COMPLETELY DEFLATED AFTER FILLING WITH SALINE . THEREFORE, ANOTHER NEW CATHETER WAS USED AND AND THE PROCESS WENT SMOOTHLY. AS PER DESCRIPTION OF EVENT CAUSE ANALYSIS, STATED THAT THE CAUSE WAS UNKNOWN, WAITING FOR URINARY CATHETER INSPECTION AND REASONS FOR PRODUCT QUALITY.AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 24FEB2023, CLARIFIED THAT THE DEFECT FOR THE SECOND CATHETER WAS FOUND BEFORE USING IN THE PATIENT BECAUSE THE NURSE WAS USED TO DOING PRE-FILING THE BALLOON TO TEST THE FUNCTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATER ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00740 | 2023-03-09 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC HF 5F TL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3385118QJ |  | IT WAS REPORTED THAT THE INTRODUCER SHEATH WAS UNABLE TO BE PEELED AWAY AND UNABLE TO BE REMOVED FROM THE CATHETER. THEREFORE, THE TIP OF THE INTRODUCER SHEATH WAS FORCIBLY PEELED AWAY AND THE PLACEMENT PROCEDURE WAS COMPLETED. THERE WAS NO REPORTED PATIENT INJURY. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00742 | 2023-03-09 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REPLACEMENT CONNECTOR FOR 4F GROSHONG NXT CLEARVUE PICC WITH EXTENSION | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7812400 |  | IT WAS REPORTED THAT THE CATHETER WAS TRIMMED AND CONNECTED AND THEN WAS FLUSHED. LIQUID LEAKED FROM THE TRANSPARENT PORTION OF THE CATHETER AT THE CONNECTOR. THE CATHETER PART NEAR THE CONNECTOR WAS TRIMMED AND A NEW CONNECTOR WAS USED FOR CONNECTION. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REDT2758 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00739 | 2023-03-09 | 2023-02-17 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO HF CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1395108QD |  | IT WAS REPORTED BY THE CUSTOMER THAT THEY "HAD AN INCIDENT WITH A PICC LINE ON FRIDAY. THE PICC LINE SOMEHOW BECAME KNOTTED INSIDE THE PATIENT¿S VEIN AND HAD TO BE REMOVED IN CATH LAB." ADDITIONAL INFORMATION RECEIVED: A DETAILED DESCRIPTION OF THE EVENT: PICC LINE BEING PLACED FOR VASOPRESSOR THERAPY. PICC DIDN'T APPEAR TO BE ¿DROPPING¿ RIGHT, APPEARED TO BE FLIPPING AND CURLING. PULLED WIRE BACK TO RESET SHERLOCK. TRIED TO ADVANCE WIRE BUT COULDN'T GET THE WIRE IN COMPLETELY, FELT RESISTANCE. UNABLE TO DRAW BLOOD OR FLUSH AS WELL. ATTEMPTED TO PULL THE PICC LINE BACK BUT FELT RESISTANCE WHEN TRYING TO LINE OUT. GOT AN XRAY TO VIEW WHAT WAS HAPPENING WITH THE LINE AND SAW THAT THE LINE WAS KNOTTED INSIDE VESSEL. PHYSICIAN TOOK PATIENT TO CATH LAB TO REMOVE LINE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) OF REGT2196 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00751 | 2023-03-09 | 2023-02-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT ALLEGEDLY INFECTED. THE CURRENT PATIENT STATUS IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01573 | 2023-03-09 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT THE PATIENT STATED THE INTERMITTENT CATHETERS CAUSED PROBLEMS AND HAD DIFFICULTY WITHDRAWING. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "TOO LONG, PATIENT INSERTS CATHETER TOO FAR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE HYDROPHILIC COATED SILICONE. INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
2020394-2023-00163 | 2023-03-09 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQ5064 | CQ5064J |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A HIGHLY CALCIFIED TARGET LESION, THE PTA BALLOON ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 07/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01571 | 2023-03-09 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER COULD NOT BE INJECTED INTO THE FOLEY CATHETER DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01569 | 2023-03-09 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE INFLATION VALVE OF THE FOLEY CATHETER WHEN INJECTING STERILE WATER. THE CATHETER WAS NOT USED IN PATIENTS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01572 | 2023-03-09 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT AN ATTEMPT WAS MADE TO INFLATE THE FOLEY CATHETER BALLOON DURING THE TEST, BUT THE STERILE WATER DID NOT ENTER AND IT DID NOT INFLATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01567 | 2023-03-09 | 2023-01-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER SHORT TIP 3 WAY CONTINUOUS IRRIGATION | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS UNABLE TO INFLATED. THE EVENT OCCURRED AFTER CATHETER INSERTION INTO THE PATIENT AND THE INFLATION FUNNEL DID NOT SWELL. USER STATED THAT THE CATHETER WAS DIFFICULT TO INFLATE AS THEY WERE UNABLE TO PUSH THE PLUNGER AT ALL. THERE WAS NO ABNORMALITY OBSERVED ON THE SHAFT OF FOLEY CATHETER AND NO PROBLEM WAS OBSERVED ON THE CATHETER DURING INFLATION. THE WATER WAS INJECTED GENTLY. USUALLY, THE BALLOON COULD BE INFLATED WITHIN AROUND 5 TO 10 SECONDS. AS PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON 27FEB2023, IT WAS STATED THAT THE SYRINGE WAS NOT RETURNED, SO IT COULD NOT CONFIRM SYRINGE HAS SUCH PROBLEM OR NOT. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "SINGLE USE ONLY: [WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: (PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES]: BARDIA®BIOCATH®FOLEYCATHETER IS A BALLOON CATHETER.THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS.SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX. SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO 26 EVERY 2FR. BALLOONCATHETER: INFLATION VALVE; IRRIGATION FUNNEL; DRAINAGE FUNNEL; INFLATION FUNNEL; BALLOON. CROSS SECTION OF CATHETER: DRAINAGE LUMEN; IRRIGATION LUMEN; INFLATION LUMEN; CAP. ACCESSORIES: SYRINGE PREFILLED WITH STERILE WATER; WATER-SOLUBLE LUBRICANT; (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.); [INTENDED USE & EFFECT-EFFICACY]; THIS DEVICE ISDESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE]: METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. TO DEFLATE BALLOON AND REMOVE CATHETER, INSERTA LUERTIP(NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATERTO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. PRECAUTIONS FOR USE: WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE.[THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATEURINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNELAND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS]: PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS. EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OFTHEBALLOON AND/OR THE CATHETERIS MISSING,CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON,TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES: NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.); BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAYREMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. NON-RUPTURE METHOD: ATTACH LUER TIP SYRINGE TO THE INFLATIONVALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER.IF SITUATIONWOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) IF SITUATION WOULDN'T BE IMPROVED WITH3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) BALLOON-RUPTURE METHOD: INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) MINERAL OIL. IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER.(FIG.6) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUMUNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) MALFUNCTION AND ADVERSE EVENTS: MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE; DIFFICULTY OR FAILURE TO REMOVE THE DEVICE; OCCLUSION OF CATHETER INNER LUMENS; ENCRUSTATION; ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE; DEVICE DAMAGE DUE TO INAPPROPRIATE USE ADVERSE EVENTS: URINARY-TRACT INFECTION; HEMORRHAGE, HEMATURIA; ALLERGY REACTION TO THE DEVICE; CALCULUS FORMATION; EDEMA; PAIN; DISCOMFORT; INJURY OF BLADDER OR URETHRAL; URETHRITIS, URINARY INCONTINENCE; RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]: STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01564 | 2023-03-09 | 2023-02-14 | I, n, j, u, r, y | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165812 | 165812 |  | IT WAS REPORTED THAT THE FOLEY CATHETER IF CUFFING WHEN BALLOON WAS DEFLATED CAUSING TRAUMA AND PAIN TO THEIR PEDIATRIC PATIENTS. THIS WAS THE THIRD OCCURRENCE REPORTED IN APPROXIMATELY 3 WEEKS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00165 | 2023-03-09 | 2022-12-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U3513076 | U3513076 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE DISTAL PORTION OF THE DEVICE ALLEGEDLY HAD A LEAK. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ULTRAVERSE 035 PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, NO KINKS WERE NOTED TO THE CATHETER SHAFT AND NO ANOMALIES TO THE LUERS/Y-BODY. DURING FUNCTIONAL EVALUATION, AN IN-HOUSE PRESTO INFLATION DEVICE WAS USED TO INFLATE THE BALLOON AND WATER WAS NOTED STREAMING FROM THE PROXIMAL END. UNDER MICROSCOPIC EXAMINATION, A PINHOLE RUPTURE WAS NOTED TO THE BALLOON. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED LEAK AND IDENTIFIED PINHOLE BALLOON RUPTURE AS A PINHOLE RUPTURE WAS NOTED TO THE BALLOON UNDER MICROSCOPIC EXAMINATION FROM WHICH WATER WAS SEEN LEAKING. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED LEAK AND IDENTIFIED PINHOLE BALLOON RUPTURE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. EXPIRY DATE: 03/2025. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00164 | 2023-03-09 | 2022-12-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DORADO | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | DR13556 | DR13556 |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON WAS ALLEGEDLY DIFFICULT TO BE REMOVED FROM PACKAGE. THE PROCEDURE WAS COMPLETED USING ANOTHER BALLOON. THERE WAS NO PATIENT CONTACT. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE DORADO PTA DILATATION CATHETER LOADED ONTO A PACKAGING HOOP WAS RETURNED FOR EVALUATION. A BREAK WAS NOTED AT THE GLUE JOINT. NO OTHER SPECIFIC DAMAGES WERE NOTED DURING THE VISUAL EVALUATION. THE STYLET WAS REMOVED FROM THE DISTAL TIP OF THE BALLOON WITH SLIGHT RESISTANCE. ALTHOUGH THE STYLET WAS REMOVED WITH SLIGHT RESISTANCE FROM THE CATHETER DURING THE FUNCTIONAL TESTING, A BREAK AT THE BOND JOINT AT THE CATHETER WAS NOTED DURING THE VISUAL EVALUATION; THEREFORE, THE INVESTIGATION WAS CONFIRMED FOR THE REPORTED DIFFICULTY IN REMOVING THE STYLET FROM THE BALLOON CATHETER AND THE IDENTIFIED CATHETER BOND JOINT BREAK. A DEFINITIVE ROOT CAUSE FOR THE REPORTED DIFFICULT TO REMOVE STYLET FROM THE BALLOON CATHETER AND IDENTIFIED CATHETER BOND JOINT BREAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRY DATE: 12/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01570 | 2023-03-09 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE FOLEY CATHETER DURING PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
MW5115677 | 2023-03-10 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C.R. BARD, INC. | A942216 | A942216 |  | A 16FR FOLEY CATHETER PLACED INTRAOPERATIVE WITHOUT DIFFICULTY AROUND 1230PM. THE PATIENT WAS FLIPPED PRONE AND THE FOLEY CATHETER WAS INTACT, SECURED AND DRAINING CLEAR YELLOW URINE. CRNA EMPTIED APPROX 200-300CC. AT ABOUT 1500, IT WAS NOTED THAT THE FOLEY CATHETER HAD FALLEN OUT OF PATIENT'S PENIS AND WAS ON THE FLOOR UNDER THE JACKSON TABLE. WHEN I EXAMINED THE BALLOON OF THE FOLEY, IT WAS DEFLATED AND APPEARED TO HAVE A HOLE IN IT. I USED A SYRINGE TO TRY TO ADD AIR TO THE BALLOON AND IT WAS IN FACT BROKEN. |  | N | Y |  |  | 2023-03
3006260740-2023-00766 | 2023-03-10 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR (RADEL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770352 |  | IT WAS REPORTED THE SHERLOCK HAVING ISSUES ACCURATELY TRACKING PICC. FLUSHED AND CALIBRATED SENSOR. VERIFIED OTHER SENSOR IS WORKING FINE ON SAME SR8. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01650 | 2023-03-10 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166L18 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE FOLEY CATHETER. NO OTHER INFORMATION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00765 | 2023-03-10 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE SENSOR IS INTERMITTENTLY GIVING INCORRECT READINGS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF SENSOR IS INTERMITTENTLY GIVING INCORRECT READINGS WAS CONFIRMED. THE SENSOR WAS CONNECTED TO A TEST SR8, AND THE ECG WAS INTERMITTENTLY FUNCTIONING. THE ROOT CAUSE OF THE REPORTED FAILURE WAS IDENTIFIED AS A LOOSE CONNECTION BETWEEN THE RIBBON CABLE AND THE ECG PCB. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01643 | 2023-03-10 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167L24 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01646 | 2023-03-10 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L24 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01645 | 2023-03-10 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167L24 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. NO OTHER INFORMATION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01632 | 2023-03-10 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE CATHETER PLACEMENT WAS CONFIRMED AFTER INSERTION TO THE PATIENT BUT 4 DAYS AFTER PLACEMENT, THE CATHETER FELL OUT OF THE PATIENT. THE POSSIBILITY OF HOLES HAS BEEN POINTED OUT, BUT IT WAS SAID THAT THEY HAVE NOT CONFIRMED THE PRESENCE OR ABSENCE OF HOLES BY RE-INJECTING WATER AT THE SITE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01641 | 2023-03-10 | 2023-02-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT EXPERIENCED A URINARY TRACT INFECTION AND BLADDER INFECTION IN HOSPITAL USING PUREWICK FEMALE EXTERNAL CATHETER. PATIENT SAID IT WAS BECAUSE THE NURSES WERE NOT CHANGING THE WICK OFTEN. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION AND BLADDER INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
16520995 | 2023-03-10 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDEX LUBRI-SIL ALL-SILICONE FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | 175816 | 175816 |  | UPON AN ATTEMPT TO REMOVE A MALE PATIENT'S URINARY CATHETER, THERE WAS NOTED TO BE A DEFECT IN THE DISTAL MOST PART OF THE BALLOON OF THE CATHETER. WHEN THE 10CCS OF WATER WERE REMOVED FROM THE BALLOON, THE SILICONE APPEARS NOT TO HAVE RETRACTED ALL THE WAY BACK INTO ORIGINAL POSITION, AND CREATED A SILICONE FOLD. THE PRESENCE OF THE FOLD CAUSED THE CATHETER TO BECOME LODGED JUST INSIDE THE PT'S URINARY MEATUS, PREVENTING NURSING STAFF FROM BEING ABLE TO REMOVE THE LAST BIT OF THE FOLEY CATHETER ALL THE WAY. PRIMARY NURSE LEFT CATHETER IN PLACE AND NOTIFIED THE INTENSIVIST. ON CALL UROLOGY NOTIFIED. PATIENT PULLED IT OUT HIMSELF. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01619 | 2023-03-10 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT MENTIONED SOME OF THE INTERMITTENT CATHETERS DID NOT HAD OUTLET HOLES. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01616 | 2023-03-10 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L22 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 03MAR2023, STATED THAT THE CUSTOMER WAS AN EXISTING CUSTOMER. CONFIRMED WITH CUSTOMER THAT THIS WAS THE FIRST TIME THEY WERE EXPERIENCING THIS ISSUE. EVENT WAS OBSERVED BY AN NURSE COMPLETING STOCKADE IN THE THEATRE STOCK ROOM. PRODUCT WAS STORED ON THE SHELF IN THE PACKAGING AS IT IS SEEN IN PHOTO, NOT IN A BOX OR CONTAINER, LOOSE ON THE SHELF. THE QUANTITY HAD BEEN CONFIRMED WITH THE CUSTOMER. AFTER OBSERVING EACH ITEM IT APPEARS THAT THE WATERMARKS AND RIPS WERE ONLY ON THE PACKAGING AND DID NOT GO THROUGH TO THE CATHETER. AFTER REVIEWING THE ITEMS, THE REPRESENTATIVE STATED THEIR OPINION THAT IT WAS CLEAR THAT THE ITEMS HAD BEEN STORED ON A SHELVING UNIT (PRESUMABLY WIRE RACKS AS IT WAS USUAL FOR THEATRE STORAGE). THERE WAS MOISTURE MARKS ON THE PACKAGING (BROWN MARKS) AND SMALL TEARS. THEATRE STORE ROOMS WERE CLIMATE CONTROLLED AND HAD BEEN KNOWN TO BE PROBLEMATIC FOR HUMIDITY/MOISTURE. THIS IS WHY STOCK WAS ALWAYS TAKEN OUT OF CARDBOARD PACKAGING AND GENERALLY STORED PER ITEM ON SHELVING. RECOMMENDATION COULD BE TO SUGGEST THEM STORING CATHETERS IN A PLASTIC BOX TO AVOID FUTURE REOCCURRENCE. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE HAS BEEN EVALUATED AND OBSERVED THERE WAS PRESENT SIGN OF WATERMARK AT OUTER POUCH OF CATHETER. THE WATERMARK APPEARS TO BE CORROSIVE AND CAUSES THE PACKAGING TO BE PERFORATED. NO ABNORMALITY ON INNER PACKAGE OF CATHETER. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE CONTRIBUTED TO USER-MISHANDLING THE PRODUCT. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01618 | 2023-03-10 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 806510 | 806510 |  | IT WAS REPORTED THAT THE CUSTOMER THAT ORDERED PART 806510 THROUGH MCKESSON. THE DESCRIPTION IN THEIR CATALOGUE STATED CATHETER, FOLEY SILICONE 5CC 10FR (12/CS) WHEN THE CUSTOMER OPENED THE BOX, THE PACKAGES INSIDE OF IT STATED, THEY WERE 3CC. THE LOT ON THE CASE OF CATHETERS WAS NGGW4222 AND THE LOT NUMBER ON THE ACTUAL CATHETER WAS NGGX3331. THE CUSTOMER WAS CONCERNED THAT THERE MIGHT HAVE BEEN AN ERROR ON BD¿S PART WHEN PACKAGING THIS ITEM, SINCE THE ITEM STATED IT WAS A 5CC, BUT THE REFERENCE NUMBER THAT MATCH WHAT THEY ORDERED AND WHAT WAS IN THE PACKAGE STATED IT WAS A 3CC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01615 | 2023-03-10 | 2023-02-17 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE CUSTOMER HAD IRRITATION THAT WAS LIKE AN INFLAMED HAIR FOLLICLE PIMPLE. THEY SAW DOCTOR AND WAS TREATING WITH ANTIBIOTICS AND CREAM. ALSO CUSTOMER HAD ABRASION AND WOULD NOT BE USING FOR A WHILE. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01613 | 2023-03-10 | 2023-02-17 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD DEVELOPED AN IRRITATION AND A RASH FROM USING THE PUREWICK FEMALE EXTERNAL CATHETER. PATIENT SAID THAT THE CATHETER WOULD NOT STAY IN PLACE AND SO THE RUIN GOT ALL OVER THE PATIENT¿S BED SHEETS, CLOTHES AND ENTIRE BODY. PATIENT HAD BEEN USING DIAPER TO TRY TO KEEP IN PLACE AND SAID THAT THE DOCTOR TOLD THEM NOT TO USE AND WAS PRESCRIBED WITH SOME CREAM. STATED THAT THEY WOULD LIKE A REFUND AND REPRESENTATIVE ADVISED PATIENT LIBERATOR MEDICAL SERVICE WOULD NOT ACCEPT THE RETURN OF THE PRODUCT DUE TO THE PRODUCT BEING OPENED AND USED. HOWEVER, WE COULD ACCEPT BACK THE ONE BOX OF WICKS THAT WAS NOT OPENED OR USED. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01644 | 2023-03-10 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0174L20 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. PER FOLLOW UP INFORMATION RECEIVED FROM IBC ON 03MAR2023, IT APPEARS TO ME AFTER OBSERVING EACH ITEM THAT THE ¿WATERMARKS¿ AND RIPS ARE ONLY ON THE PACKAGING AND DO NOT GO THROUGH TO THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE HAS BEEN EVALUATED AND OBSERVED THERE WAS PRESENT SIGN OF WATERMARK AT OUTER POUCH OF CATHETER. THE WATERMARK APPEARS TO BE CORROSIVE AND CAUSES THE PACKAGING TO BE PERFORATED. NO ABNORMALITY ON INNER PACKAGE OF CATHETER. BASED ON FEEDBACK FROM USER, THE WATERMARK COULD BE FROM DISINFECTANT SPRAYING TO CLEAN SURFACES DURING THE COVID YEARS THAT CAN CONTRIBUTED AS PER REPORTED EVENT. HOWEVER, THE EXACT CAUSED OF HOW AND WHEN PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE CONTRIBUTED TO USER-MISHANDLING THE PRODUCT. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01617 | 2023-03-10 | 2023-02-17 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD BEEN REUSED THE PUREWICK FEMALE EXTERNAL CATHETER ALSO HAD WOUND SO COULD NOT USE PUREWICK BECAUSE OF URINARY TRACT INFECTION FROM REUSED WICKS. THE REPRESENTATIVE ADVISED THE PATIENT NOT TO REUSE BECAUSE THAT COULD LEAD TO INFECTION. AND THE PATIENT WAS AWARE BUT HAD NOT SEEN DOCTOR. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. PER FOLLOW UP VIA PHONE ON 08MAR2023, IT WAS REPORTED THAT THE CUSTOMER SOUGHT MEDICAL ATTENTION AND PROVIDER DID NOT CONFIRM WICKS CAUSED URINARY TRACT INFECTION. CUSTOMER WAS STILL USING DEVICE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01611 | 2023-03-10 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166L24 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01612 | 2023-03-10 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT THE VACUTAINER NOT WORKING IN THE URINE SPECIMEN KIT IN BOTH (B)(6) HOSPITALS. ALSO RECENTLY IN (B)(6) HOSPITAL THEY HAD A BALLOON FAILURE IN AN INSERTION KIT . CUSTOMER MENTIONED THE PACKAGING FOR THE PROVON WIPES, DO WE HAD NEW POSTERS WITH NEW COLOR SCHEME, AS STAFF WERE ASKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01625 | 2023-03-10 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166L18 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE FOLEY CATHETER. NO OTHER INFORMATION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE, A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01629 | 2023-03-10 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166L24 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01633 | 2023-03-10 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, THE STERILE WATER WAS UNABLE TO REMOVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01634 | 2023-03-10 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT USED A 22 GAUGE TRIPLE LUMEN URINARY CATHETER FOR URINARY DRAINAGE AFTER SURGERY ON (B)(6) 2023. ON (B)(6) 2023, AT 01:05 PM, THE CATHETER FELL OUT DUE TO A RUPTURED BALLOON, WHICH WAS REPORTED TO THE DOCTORS, AND A NEW CATHETER WAS REINTRODUCED. THE URINARY CATHETER WAS UNOBSTRUCTED. THIS OPERATION INCREASED THE PHYSICAL PAIN, PSYCHOLOGICAL BURDEN AND FINANCIAL COST INCURRED BY THE PATIENT. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 24FEB2023, IT WAS STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. NO MEDICAL INTERVENTION WAS REQUIRED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW CAUSING THIN SAC. A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL (B)(6), THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL (B)(6). SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY". THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01635 | 2023-03-10 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER FIXATION WAS CONFIRMED AFTER PLACEMENT IN THE PATIENT BLADDER BUT THE CATHETER FELL OUT OF THE PATIENT AFTER TWO DAYS LATER. THE BALLOON WAS AGAIN INJECTED WITH 10 ML OF WATER, BUT THE WATER LEAKED OUT IMMEDIATELY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01614 | 2023-03-10 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT DURING A SURGERY, THE DEVICE WAS PLACED BY 2 NURSES AND CONFIRMED 10 ML OF WATER INJECTION INTO THE FOLEY CATHETER. ALSO STATED 7 DAYS AFTER PLACEMENT, SPONTANEOUS REMOVAL OF THE CATHETER OCCURRED IN THE WARD, AND IT WAS CONFIRMED THAT 2ML OF STERILE PURIFIED WATER REMAINED IN THE BALLOON AT THAT TIME. IT WAS ALSO NOTED THAT THERE WAS NO BALLOON RUPTURE OR TEAR OCCURRED AND WATER LEAKED FROM UNKNOWN LOCATION AND BECAUSE THE EVENT WAS CATHETER FELL OUT AND DID NOT TRY TO REPRODUCE THE EVENT. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER SHAPE, CONFIGURATION AND PRINCIPLES: BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY INTENDED USE & EFFECT- EFFICACY: THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01624 | 2023-03-10 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L16 |  | IT WAS REPORTED THAT WATER WAS SPONTANEOUSLY REMOVED FROM THE FOLEY CATHETER BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00772 | 2023-03-12 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY TWO YEARS POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN. IT WAS FURTHER REPORTED THAT THE PORT SYSTEM HAS NOT BEEN USED FOR APPROXIMATELY HALF A YEAR. REPORTEDLY, THE DISTAL CATHETER SEGMENT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00771 | 2023-03-12 | 2022-04-17 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1608062 |  | IT WAS REPORTED THAT TWO WEEKS AND FOUR DAYS POST A PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN AND SWELLING. IT WAS FURTHER REPORTED THAT VENOUS ULTRASOUND WAS PERFORMED, WHICH CONFIRMED OCCLUSIVE THROMBUS WITHIN RIGHT INTERNAL JUGULAR VEIN. FURTHERMORE, ANGIOGRAM PORT CHECK WAS PERFORMED WHICH SHOWED RIGHT INTERNAL JUGULAR DEEP VEIN THROMBOSIS. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED, AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION; MEDICAL RECORD WAS PROVIDED FOR REVIEW. THE MEDICAL RECORD ALLEGES THAT THE PATIENT UNDERWENT RIGHT INTERNAL JUGULAR VEIN PORT PLACEMENT PROCEDURE. APPROXIMATELY, EIGHTEEN DAYS LATER, THE PATIENT EXPERIENCED RIGHT NECK SWELLING AND PAIN. A CT WAS SHOWED THROMBOSIS OF THE LOWER PORTION OF THE RIGHT INTERNAL JUGULAR VEIN. ANGIOGRAM STUDY SHOWED NO COMPLICATION OF THE PORT AND NO LEAK OF CONTRAST. THERE WAS NO FIBRIN SHEATH AND PORT FUNCTION WELL. THEREFORE, IT CAN BE CONFIRMED THAT THE PATIENT EXPERIENCED NECK SWELLING, PAIN AND THROMBOSIS AT RIGHT IJV. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO FLUSH AND SUCTION ISSUES AS NO SUCH PROBLEMS RELATED TO THE PORT WAS REPORTED DURING ANGIOGRAM STUDY, HOWEVER, THE RELATIONSHIP TO THE PORT IS UNKNOWN AS NO PHOTO EVIDENCE WAS PROVIDED. FURTHERMORE, THE CLINICAL CONDITION ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3, H6 (METHOD). H11: D4 (UNIQUE IDENTIFIER (UDI) #). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00773 | 2023-03-12 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 7707540 | 7707540 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE PORT WAS ALLEGEDLY FRACTURED. IT WAS FURTHER REPORTED THAT THE PORT WAS REMOVED FROM PATIENT. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 10/2024). DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00777 | 2023-03-12 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER WITH SHERLOCK TLS STYLET 5F MAXIMAL BARRIER TRAY DL 45CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9827508D |  | IT WAS REPORTED BY THE CUSTOMER "A PATIENT WITH A 5F DL GROSHONG NXT PICC HAS HAD THIS CATHETER FOR A LONG TIME. RECENTLY, THE RED HUB OF THE PICC "FELL OFF" AND THE PATIENT REPORTED THAT SHE "CLEANED IT WITH ALCOHOL AND STUCK IT BACK ON". CALLER IN PHARMACY IS LOOKING FOR A REPAIR KIT FOR THE RED LEG. RESPONSE: VERIFIED THAT 9827508D IS A 5F DL GROSHONG NXT PICC. EXPLAINED THE NXT IS NOT POWER INJECTABLE. INFORMED THE NXT IS MADE OF SILICONE AND HAS A MAX OF 25 PSI. INFORMED THAT THERE IS A RISK OF INFECTION DUE TO THE HUB FALLING OFF AND PATIENT CLEANING AND PUTTING IT BACK ON. INFORMED HIM THAT THE NXT CAN BE REPAIRED SINCE IT IS NOT POWER INJECTABLE. PROVIDED VALID AND AVAILABLE CODES FOR REPAIR." | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00778 | 2023-03-12 | 2023-01-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET MICROINTRODUCER 4F SL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174118 |  | IT WAS REPORTED THAT "ON (B)(6) PATIENT CAME INTO ED WITH MALFUNCTIONING SINGLE LUMEN PICC WHICH HAD BEEN INSERTED ON (B)(6). ED STAFF WAS UNABLE TO TPA. PATIENT STATED THE 'PICC NEVER WORKED RIGHT'. THEY ATTEMPTED TO TPA AND REQUESTED CHEST X-RAY INCLUDING ARM. KINK NOTED ON X-RAY. KINK WOULD NOT PULL STRAIGHT BUT KEPT RE-KINKING UNDER THE SKIN AT THE 9CM MARK. AT THAT POINT, PICC WAS NO LONGER CENTRAL AND IT WAS PULLED. PIV STARTED AND PATIENT TO GO TO OUTPATIENT INFUSION FOR REPLACEMENT." | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01663 | 2023-03-12 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER CREEVY MODEL | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 01267524 | 01267524 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON COULD NOT BE INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01660 | 2023-03-12 | 2023-01-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT HAD AN INDWELLING CATHETER AFTER THE OPERATION ON (B)(6)2023. THE OPERATION WAS NORMAL, BUT THE BALLOON RUPTURED LATER. CONSIDERING THAT THE QUALITY OF THE CATHETER WAS UNQUALIFIED, A NEW CATHETER WAS REPLACED AND RE-INDWELLED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 03MAR2023,THERE WAS NO MISSING PIECES OF BALLOON | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01661 | 2023-03-12 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THE FOLEY CATHETER WAS INDWELLED AFTER THE OPERATION ON (B)(6), AND THE BLADDER IRRIGATION WAS FLUSHED CONTINUOUSLY. ON (B)(6), THE CATHETER FELL OUT, AND THE BALLOON RUPTURE WAS FOUND AFTER CHECKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00781 | 2023-03-12 | 2021-10-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH DIALYSIS CATHETER KIT 12F X 15CM | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5553150 |  | IT WAS REPORTED THAT THE GUIDE WIRE GOT STUCK, UNABLE TO PASS. NO OTHER INFORMATION WAS PROVIDED. 02/21/2023- THE RETURNED GUIDEWIRE WAS CONFIRMED TO BE BROKEN AND THE OUTER COIL WIRE BECAME UNRAVELED. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE ANALYSIS, PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A STUCK GUIDEWIRE WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE OPEN 12 FR SLIM-CATH NIAGARA KIT CONTAINING A GUIDEWIRE STUCK WITHIN AN INTRODUCER NEEDLE. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE CAUSED BY RETRACTION OF THE GUIDEWIRE AGAINST THE BEVEL OF THE INTRODUCER NEEDLE. THE RETURNED PRODUCT SAMPLE WAS EVALUATED, AND THE GUIDEWIRE WAS CONFIRMED TO BE BROKEN WHICH ALLOWED THE OUTER COIL WIRE TO BECOME UNRAVELED. MICROSCOPIC EXAMINATION OF THE FRACTURE SITES REVEALED THE FOLLOWING: NARROWING OF THE WIRE CROSS-SECTION NEAR THE FRACTURE SITE, WHICH IS A CHARACTERISTIC FEATURE OF A STRONG PULL ON THE WIRE. DAMAGE TO THE INSIDE EDGE OF THE INTRODUCER NEEDLE WHICH CAN OCCUR IF THE GUIDEWIRE IS FORCEFULLY RETRACTED AGAINST THE NEEDLE, DAMAGING THE SHAPE OF THE SHARPENED BEVEL BIOLOGICAL MATERIAL WAS ALSO SEEN ON THE WIRE WHICH MAY HAVE CONTRIBUTED TO THE OBSERVED GUIDEWIRE FRACTURE AS RETRACTION OF THE WIRE TOGETHER WITH THE BIOLOGICAL MATERIAL COULD HAVE CAUSED THE PIECES TO BECOME STUCK. NORMAL MOVEMENT OF THE GUIDEWIRE IS AWAY FROM THE SHARPENED BEVEL AND WILL NOT DAMAGE THE WIRE. IF THE GUIDEWIRE DIRECTION IS REVERSED, THE GUIDEWIRE IS THEN PULLED AGAINST THE SHARPENED EDGE OF THE NEEDLE BEVEL AND CAN CAUSE SHEARING DAMAGE OF THE WIRE AND/OR CAUSE THE WIRE TO BECOME STUCK WITHIN THE NEEDLE. AN EXAMINATION OF THE WIRE STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. A BATCH HISTORY REVIEW (BHR) OF REFN3095 SHOWED 1 OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS BATCH NUMBER. THE COMPLAINTS FOR THIS BATCH NUMBER HAVE BEEN REPORTED FROM THE SAME FACILITY IN CHINA. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00782 | 2023-03-12 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE NURSE THAT THE PATIENT¿S WIFE IS SAYING THE LINE IS HARD TO FLUSH AND THEY ARE EXPERIENCING SIGNIFICANT RESISTANCE WHEN FLUSHING THE LINE. NURSE HAD NOT ACCESSED THE LINE AT THE TIME OF OUR PHONE CALL, SO ADVICE LINE NURSE ADVISED VISITING NURSE TO ACCESS THE LINE AND FLUSH IT, IN ORDER TO ACCURATELY REPORT THE PATENCY OF THE LINE. PATIENT¿S WIFE ALSO MENTIONED THAT FLUSHING THE LINE WAS VERY STIFF A FEW NIGHTS AGO BUT SEEMED TO BE GETTING BETTER BY YESTERDAY. NURSE UNSURE IF THEY WILL CONNECT PATIENT TONIGHT. NURSE TRIED FLUSHING THE LINE AND FLUSHED WELL WITH SLIGHT EXPECTED RESISTANCE OF PICC LINE- LOCKED THE LINE. THERE WAS NO PATIENT HARM REPORTED, NO MEDICAL INTERVENTION REQUIRED, AND NO HARM TO CLINICIAN. (B)(6) 2023 ADDITIONAL INFORMATION PROVIDED. IT WAS REPORTED BY THE CUSTOMER "EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. OCCLUSION WAS NOT RESOLVED, AIR BUBBLES IN THE LINE, HAD TO STOP INFUSION. NO PATIENT HARM. AFTER ADDITIONAL FOLLOW UP, IT IS UNKNOW IF THE PICC LINE WAS REMOVED." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01664 | 2023-03-12 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167L20 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT (SUSPECTED PINHOLE). ON (B)(6) CATHETER FELL OUT AFTER SURGERY IN ASYMMETRICALLY INFLATED STATE IN WARD . ON (B)(6) WHEN RECOVERED THE SAMPLE, THE BALLOON WAS ALREADY DEFLATED. THERE MIGHT BE A PINHOLE TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00779 | 2023-03-12 | 2023-02-05 | D, e, a, t, h | Manufacturer report | POWERPICC CATHETER W/SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED "ON (B)(6) 2023 CALLED TO ASSIST PATIENT WITH TPA IN PROBLEMATIC DUAL LUMEN PICC LOT# REGV0221 WHICH HAD RECEIVED MULTIPLE DOSES OF TPA SINCE INSERTION ON (B)(6) 2023. CHEST AND ARM XRAY SHOWED KINK/LOOP IN ARM. PICC TIP HAD BEEN READ SHORT ON CHEST XRAY AS EARLY AS (B)(6) 2023. LINE PULLED, VASCULAR STUDIES ORDERED. PATIENT WITH THROMBUS IN IJ, SUBCLAVIAN, AXILLA AND BASILIC. PIV PLACED. PATIENT TX TO ICU, CONTRALATERAL IJ PLACED." IT WAS STATED THE "PATIENT PASSED AWAY." | THE DATE THE PATIENT PASSED AWAY WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN DATE OF DEATH IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) OF REGV0211 SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS LOT NUMBER. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00786 | 2023-03-12 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE SENSOR IS INTERMITTENTLY GIVING INCORRECT READINGS. BD SALES REPRESENTATIVE HAS WORKED WITH THEM EXTENSIVELY AND FEELS THE SENSOR NEEDS TO BE REPLACED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF SENSOR IS INTERMITTENTLY GIVING INCORRECT READINGS WAS UNCONFIRMED. THE SENSOR WAS CONNECTED TO A TEST SR8, CALIBRATED, AND FUNCTIONS PROPERLY WITH NO INCORRECT READINGS. THE ROOT CAUSE OF THE REPORTED FAILURE IS INCONCLUSIVE AS THE REPORTED ISSUE COULD NOT BE REPRODUCED DURING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00774 | 2023-03-12 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6903190 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY HAD A LACERATION IN ONE OF THE CATHETER PATHWAYS AND AIR WAS ALLEGEDLY NOTICED IN THE CATHETER. IT WAS FURTHER REPORTED THAT BLOOD HAD ALLEGEDLY CAME OUT AT THE SITE OF THE CATHETER THAT WAS FOUND LACERATED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00784 | 2023-03-12 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC SV 4F DL BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3274118J |  | IT WAS REPORTED THAT LEAKAGE WAS FOUND AROUND THE SUTURE WING. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. A BATCH HISTORY REVIEW (BHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING BATCH NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01651 | 2023-03-12 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0166L28 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01659 | 2023-03-12 | 2023-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT WAS PLACED WITH FOLEY CATHETER AFTER THE OPERATION ON THE (B)(6) 2023. AT 4:20 ON THE (B)(6), THE PATIENT COMPLAINED THAT THE CATHETER FELL OUT. UPON CHECKING, IT WAS FOUND THAT THE FRONT TIP OF BALLOON ON THE FOLEY CATHETER HAD BURST, WHICH CAUSED THE CATHETER TO SLIP OUT. NO BLEEDING OCCURRED. THE PATIENT CONTINUED TO BE OBSERVED, AND THE PATIENT WAS ABLE TO URINATE ON THEIR OWN WITHOUT COMPLAINING OF DISCOMFORT. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 03MAR2023, CLARIFIED THAT TIP BURST REFERRED TO BALLOON BURST AND BREAK ALSO REFERRED TO BALLOON BURST IN DEVICE FAILURE PERFORMANCE IN THE ATTACHED SOURCE DOCUMENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATER ALLERGY". SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00788 | 2023-03-13 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT THE APPROXIMATELY ONE YEAR THREE MONTHS AND THIRTEEN DAYS POST PORT PLACEMENT PROCEDURE, WHEN THE PORT SYSTEM WAS FLUSHED WITH SALINE SOLUTION, SWELLING WAS FOUND AROUND THE PORT BODY. IT WAS FURTHER REPORTED THAT, LEAKAGE WAS ALLEGEDLY FOUND. REPORTEDLY, A CRACK WAS ALLEGEDLY FOUND ON THE CATHETER. THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
MW5115719 | 2023-03-13 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | POWERGLIDE | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | F120108PT-DOT PG PRO 20G 10CM RT |  |  | PATIENT ADMITTED TO HOSPITAL FOR RIGHT LOWER EXTREMITY PAIN AND WOUND. PATIENT STARTED ON IV ANTIBIOTICS AT TIME OF ADMISSION. ON (B)(6) 2023, PATIENT HAD TO HAVE A MIDLINE REPLACED. VASCULAR ACCESS DEVICE (VAD) TEAM MEMBER, REPORTED TO PATIENT'S ROOM AT 1530. MIDLINE PLACEMENT ATTEMPTED; HOWEVER, THE CATHETER WAS NOT ABLE TO BE ADVANCED PAST 5CM. PROCEDURE TERMINATED AND LINE PULLED. VAD TEAM MEMBER ASSESSED CATHETER LINE AND NOTED APPROXIMATELY 1.5CM OF LINE WAS NOT PRESENT AND POTENTIALLY "SHEARED-OFF" INTO THE PATIENT'S TISSUE. VAD TEAM MEMBER IMMEDIATELY NOTIFIED PRIMARY TEAM PHYSICIAN AND IMAGING ORDERED. PERIPHERAL IV PLACED INSTEAD OF MIDLINE AT THIS TIME. X-RAY PERFORMED AND CONFIRMED A 1.5CM LINEAR RADIOPAQUE DENSITY ALONG THE MEDIAL DISTAL PORTION OF THE UPPER ARM. ULTRASOUND STUDY PERFORMED AND DEMONSTRATED A CATHETER SHAPED FOREIGN BODY POSSIBLY WITHIN THE SUPERFICIAL VESSEL. CARDIOTHORACIC VASCULAR TEAM CONSULTED AND RECOMMENDED SURGICAL REMOVAL OF FOREIGN BODY. PATIENT REFUSED SURGERY FOR REMOVAL. RISKS EXPLAINED BY CONSULT TEAM AND PRIMARY TEAM OF LEAVING FOREIGN BODY IN PLACE. THESE RISKS INCLUDED SECONDARY INFECTION AND WOUND CREATION. PATIENT CONTINUED TO REFUSE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00787 | 2023-03-13 | 2023-01-26 | I, n, j, u, r, y | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808000 | 1808000 |  | IT WAS REPORTED THAT POST PORT PLACEMENT, THE PATIENT WAS DIAGNOSED WITH INFECTION WHICH REQUIRED ANTIBIOTIC INFUSION & AN ORAL ANTIBIOTIC. THE PATIENT'S CURRENT STATUS WAS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 11/2023). | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01671 | 2023-03-13 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166L28 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01682 | 2023-03-13 | 2023-02-22 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT AN INFECTION FROM THE PUREWICK AND HAD STOPPED USING IT. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "MAINTENANCE: REPLACE THE PUREWICK¿ FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICK¿ FEMALE EXTERNAL CATHETER WARNINGS: DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS." TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICK¿ FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICK¿ FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. "PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE". EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICK¿ FEMALE EXTERNAL CATHETER. " RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01672 | 2023-03-13 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE LEAKS OCCURRED IN THE FOLEY CATHETER AFTER SIX DAYS OF USE. THE USER NORMALLY USED THIS FOR A MONTH. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 20FEB2023, STATED THAT THE LOT NUMBER OF THE DEFECTIVE CATHETER WAS MYEY1619 ( BATCH# MYEY1619). THE PATIENT HAD ALREADY DISPOSED THE OTHER PACKAGING, SO THE LOT NUMBERS WERE UNKNOWN (BATCH# UNK). CLARIFIED THAT THE PACKAGE NORMALLY LASTS A MONTH, IF IT DID NOT LEAK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01674 | 2023-03-13 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT AFTER ONE DAY OF CATHETER PLACEMENT TO THE PATIENT, THE BALLOON RUPTURED. NO MISSING FRAGMENTS WERE FOUND. WATER LEAKED FROM THE BALLOON AREA AND NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. THE CATHETER WAS NOT PULLED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01678 | 2023-03-13 | 2023-02-20 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER LF HC | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THEY TOLD THE PATIENT AT THE HOSPITAL NOT TO USE THE PUREWICK FEMALE EXTERNAL CATHETER ANY LONGER BECAUSE IT WAS CAUSING URINARY TRACT INFECTIONS. THE PATIENT HAD BEEN USING THIS PRODUCT FOR MORE THAN 90DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. PER ADDITIONAL INFORMATION RECEIVED FROM LIBERATOR ON (B)(6) 2023,THE REPRESENTATIVE STATED THAT THE DOCTOR NO LONGER WANTED THE PATIENT TO USE THE THESE PUREWICK FEMALE EXTERNAL CATHETER, BECAUSE THEY THOUGHT IT WAS CONTRIBUTION TO THE PATIENT CONSTANT URINARY TRACT INFECTIONS, THE REPRESENTATIVE REQUESTED TO RETURN 3 BOXES OF UNOPENED WICKS AND 1 BOX OF AN UNOPENED ACCESSORY KIT. HE PATIENT HAD BEEN USING THIS PRODUCT FOR MORE THAN 90DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | N | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00789 | 2023-03-13 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808000 | 1808000 |  | IT WAS REPORTED THAT POST PORT PLACEMENT, THE PATIENT CONTINUED TO EXPERIENCE SWELLING, BRUISING AND PAIN . THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 07/2023). DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01684 | 2023-03-13 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 37303 | 37303 |  | IT WAS REPORTED THAT THE MALE EXTERNAL CATHETER WERE DIFFICULT TO REMOVE AS THE ADHESIVE WOULD NOT LET GO AND IT CAUSED SEVERE IRRITATION WHEN MALE EXTERNAL CATHETER REMOVED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: VERIFY CORRECT SIZE PRIOR TO USE. TRIM PUBIC HAIR IF NECESSARY. WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. UNROLL THE CATHETER OVER PENIS. GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: ENSURE DRAINAGE BAG IS EMPTY. DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER.¿ THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01685 | 2023-03-13 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 35303 | 35303 |  | IT WAS REPORTED THAT THE MALE EXTERNAL CATHETERS WERE DIFFICULT TO REMOVE AS THE ADHESIVE WOULD NOT LET GO AND CAUSED SEVERE IRRITATION WHEN IT REMOVED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "MECHANICAL FAILURE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: 1) VERIFY CORRECT SIZE PRIOR TO USE. 2) TRIM PUBIC HAIR IF NECESSARY. 3) WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. 4) OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. 5) PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. 6) UNROLL THE CATHETER OVER PENIS. 7) GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. 8) CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: 1) ENSURE DRAINAGE BAG IS EMPTY. 2) DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. 3) GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER.¿ | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01689 | 2023-03-13 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT A FEW OF THE INTERMITTENT CATHETERS DID NOT HAVE SALINE WATER AND IT WERE NOT USABLE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INCORRECT PROCESS PARAMETERS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED ISSUE. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01693 | 2023-03-13 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER HAD RUPTURED DURING PLACEMENT. THE CATHETER WAS FOUND TO BE TORN WHILE INDWELLING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01695 | 2023-03-13 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L22 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01694 | 2023-03-13 | 2022-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L24 |  | IT WAS REPORTED THAT THE CUSTOMER FOUND WATERMARKS AND SOME HOLES IN THE PRODUCT. NO OTHER INFORMATION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01696 | 2023-03-13 | 2023-02-27 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT URINARY TRACT INFECTION AND NO LONGER USING PUREWICK, AND PATIENT WAS NOT SURE THAT IF IT WAS FROM PUREWICK FEMALE EXTERNAL CATHETER. IT WAS NOTED THAT THE PATIENT HAD BEEN USING LESS THAN 90 DAYS. IT WAS UNKNOWN IF THE PUREWICK FEMALE EXTERNAL CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE OR MATERIAL SURFACE IS ROUGH, ABRASIVE OR UNCOMFORTABLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS: TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER. HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE. EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01670 | 2023-03-13 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® RUBBER UTILITY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 277712 | 277712 |  | IT WAS REPORTED THAT THE PATIENT FOUND A WHITE POWDER LIKE SUBSTANCE ON ONE INTERMITTENT CATHETER AND DID NOT WANT TO TRY TO USE. THE REPRESENTATIVE ADVISED TO REPLACE THAT CATHETER AND TO HOLD ON TO IT AS THE MANUFACTURER MAY WANT TO HAVE IT. PER CUSTOMER VIA PHONE ON (B)(6) 2022, IT WAS REPORTED THAT THERE WAS WHITE SUBSTANCE OR POWDER ON THE CATHETER. PER SAMPLE EVALUATION ON (B)(6) 2023, IT WAS NOTED THAT THERE WERE NO EYELETS ON THE INTERMITTENT CATHETER DURING EVALUATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01687 | 2023-03-13 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 35302 | 35302 |  | IT WAS REPORTED THAT THE MALE EXTERNAL CATHETERS WERE DIFFICULT TO REMOVE AS THE ADHESIVE WOULD NOT LET GO AND IT CAUSED SEVERE IRRITATION WHEN IT REMOVED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "VISCOMETER FAILURE OR MECHANICAL FAILURE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: 1) VERIFY CORRECT SIZE PRIOR TO USE. 2) TRIM PUBIC HAIR IF NECESSARY. 3) WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. 4) OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. 5) PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. 6) UNROLL THE CATHETER OVER PENIS. 7) GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. 8) CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: 1) ENSURE DRAINAGE BAG IS EMPTY. 2) DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. 3) GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER.¿ | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00828 | 2023-03-14 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT W/BIOPATCH AND PROBE COVER | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108BDP |  | IT WAS REPORTED BY THE CUSTOMER "MIDLINE IN PLACE AND WORKING FINE FOR A FEW DAYS. RN CALLED FOR A CONSULT TO SEE IF DEVICE WAS LEAKING. PICC NURSE VISUALLY INSPECTED THE DEVICE AND IDENTIFIED A SMALL HOLE THAT LEAKED WHEN THE DEVICE WAS FLUSHED.¿ ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: "MIDLINE PLACED IN LEFT BASILIC VEIN ON (B)(6) 2023, CUT BACK TO 10CM. NO ISSUES, NO OBVIOUS LEAKS. ON (B)(6) 2023, CALLED TO ASSESS MIDLINE FOR LEAKING. WHEN AT BEDSIDE DRESSING IS VISIBLY WET UNDER THE TEGADERM. MIDLINE FLUSHES WITH EASE, VISUALIZED WITH ULTRASOUND IN VEIN, NO S/S OF INFILTRATION. I REMOVED THE DRESSING, FLUSHED THE MIDLINE AND OBSERVED SMALL AMOUNT OF FLUSH ESCAPING THE MIDLINE NEAR THE INSERTION SITE, I COULD NOT VISUAL WHERE THIS WAS COMING FROM EXACTLY, BUT EVERY TIME THE MIDLINE WAS FLUSHED THIS HAPPENED. ON (B)(6) 2023 MIDLINE WAS REMOVED AND REPLACED WITH NO ISSUES. AFTER REMOVING I INSPECTED THE LINE FOR HOLES AND CANNOT OBSERVE EXACTLY WHERE THIS IS COMING FROM. THERE WAS NO PATIENT HARM." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00826 | 2023-03-14 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED THAT THE EXTENSION OF THE DEVICE SWELLS DURING INFUSION. FOLLOWING THE ALTERED STATE OF THE EXTENSION, THE DEVICE WAS REPAIRED WITH A DEDICATED KIT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF SWELLING OF THE TUBING DURING INFUSION IS CONFIRMED. TWO 4 FR SINGLE LUMEN GROSHONG NXT CLEARVUE CATHETERS WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLES. BOTH OF THE RETURNED CATHETERS WERE OBSERVED TO TERMINATE PROXIMAL TO THE VALVE AND THE DISTAL ENDS OF THE CATHETER WERE NOT RETURNED. ONE CATHETER WAS OBSERVED TO BE SHORTER THAN THE OTHER. TENSILE WEAKNESS WAS OBSERVED IN THE EXTENSION LEG TUBING JUST DISTAL THE WHITE OVERSLEEVE BETWEEN THE LUER HUB AND THE EXTENSION LEG IN BOTH SAMPLES, BUT A NOTICEABLY GREATER AMOUNT OF TENSILE WEAKNESS WAS OBSERVED IN THE SHORTER CATHETER SAMPLE. BOTH SAMPLES WERE FLUSHED WITH A 12 ML SYRINGE OF WATER AND WERE FOUND PATENT TO INFUSION AND ASPIRATION. WHEN THE CATHETERS WERE PRESSURIZED, THE EXTENSION LEG OF THE SHORTER CATHETER SAMPLE WAS OBSERVED TO BALLOON JUST DISTAL TO THE WHITE OVERSLEEVE. A MICROSCOPIC OBSERVATION REVEALED THE SURFACE OF THE AREA WHICH BALLOONED WHILE UNDER PRESSURE IN THE EXTENSION TUBING OF THE SHORTER CATHETER SAMPLE WAS SLIGHTLY DULLER IN LUSTER, AND THIS AREA WAS OBSERVED TO NECK WHEN THE TUBING WAS STRETCHED LONGITUDINALLY. THE LOCATION AND PHYSICAL CHARACTERISTICS OF THE DAMAGE OBSERVED IN THE RETURNED SAMPLE ARE INDICATIVE OF BALLOONING CAUSED BY OVER-PRESSURIZATION OF THE CATHETER, LIKELY A RESULT OF BUILDUP OF BIOLOGICAL MATERIAL WITHIN THE CATHETER, AND/OR STRETCHING OF THE EXTENSION LEG TUBING TO THE POINT OF CREATING ELASTIC WEAKNESS. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00829 | 2023-03-14 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK TIP LOCATION SYSTEM (TLS) STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9174108D |  | IT WAS REPORTED BY THE CUSTOMER THAT "BLOOD RETURN NOTED ON FLOW SHEETS FOR MULTIPLE DAYS. NO TPA ADMINISTRATIONS OR POWER INJECTIONS NOTED IN CHART. PATIENT REQUIRED NO ADDITIONAL TREATMENT OR INTERVENTION ONCE LINE REMOVED. UPON REMOVAL, TEAR IN EXTERNAL PORTION OF THE LINE BETWEEN 35-40CM MARKINGS. THE TEAR SEEMS TO BE 5CM LONG." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A SPLIT IN THE CATHETER IS CONFIRMED, BUT A ROOT CAUSE COULD NOT BE ESTABLISHED. ONE 4 FR S/L POWERPICC WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED BLOOD USE RESIDUES, AND A NEEDLELESS INJECTION CAP AND A STATLOCK STABILIZATION DEVICE WAS RETURNED ATTACHED TO THE DEVICE. A LARGE TEAR IN THE CATHETER WAS OBSERVED JUST PROXIMAL TO THE 34 CM DEPTH MARKER THAT RUNS JUST DISTAL OF THE 38 CM DEPTH MARKER. THE SPLIT WAS LONGITUDINALLY ALIGNED AND APPROXIMATELY 4 CM LONG. A MICROSCOPIC OBSERVATION REVEALED THE FRACTURE SITE TO BE ALONG AN EXTRUSION FEATURE, AND THE FRACTURE SURFACE APPEARED SMOOTH WITHOUT AN IDENTIFIABLE INITIATION POINT. THE WALL THICKNESS WAS MEASURED WHERE THE EXTRUSION FEATURE WAS IDENTIFIED AND WAS WITHIN DRAWING SPECIFICATIONS. BASED ON THE STATE OF THE LARGE SPLIT, A CAUSE COULD NOT BE IDENTIFIED FOR THE INITIATION OF THE SPLIT. POSSIBLE CAUSES OF THIS FAILURE MAY BE SHARP INSTRUMENT DAMAGE DURING REMOVAL, INTERACTIONS WITH BIOLOGICAL FEATURES WITHIN THE DYNAMIC VASCULATURE OF THE PATIENT, OR OTHER UNIDENTIFIED CAUSES. THE COMPLAINT OF A CATHETER SPLIT IS CONFIRMED; HOWEVER, AN EXACT CAUSE COULD NOT BE DETERMINED DUE TO THE NATURE OF THE COMPLAINT. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00791 | 2023-03-14 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT TI L/P 6 CF INT WSP ATT SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8716000 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, IT WAS DISCOVERED DURING PURCHASE ORDER PROCESSING THAT THE IMPLANTED POWER PORT WAS ALLEGEDLY EXPIRED AT THE TIME OF USE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 08/2022. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00831 | 2023-03-14 | 2023-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD MIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED, "A NURSE WAS ATTEMPTING TO INSERT A MIDLINE CATHETER INTO A PATIENT'S ARM. SHE WAS HAVING TROUBLE ADVANCING THE CATHETER AND SOMETHING DIDN'T FEEL RIGHT. SHE AND ANOTHER NURSE WHO WAS STANDING NEARBY HEARD AN AUDIBLE UNUSUAL CLICK. SHE DECIDED TO REMOVE THE CATHETER AND MAKE ANOTHER FRESH ATTEMPT. WHEN SHE WAS WITHDRAWING THE CATHETER SHE FELT RESISTANCE TRYING TO REMOVE IT FROM THE PATIENT'S SKIN. A STRONG PULL REMOVED IT AND IT WAS NOTED THAT THE GUIDEWIRE HAD EXITED THE SIDE OF THE CATHETER ABOUT ONE INCH FROM THE TIP WAS COILED INTO A SMALL KNOT." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00806 | 2023-03-14 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LOW ARTIFACT POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK LOW ARTIFACT POWER PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT WAS PLANNED TO BE REMOVED. IT WAS FURTHER REPORTED THAT THE CAP AND CATHETER ALLEGEDLY BROKE OFF THE PORT AND REMAINED IN THE PATIENT WHILE REMOVING. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00825 | 2023-03-14 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THE PATIENT HAD A PICC CATHETER PLACED ON (B)(6) AND WHEN THE NURSE GAVE MAINTENANCE ON (B)(6) SECONDS AFTER FLUSHING THE CATHETER, FLUID WAS FOUND TO BE OOZING FROM THE PUNCTURE SITE, AND A SMALL AMOUNT OF THE CATHETER WAS IMMEDIATELY PULLED OUT, WHILE FLUSHING THE CATHETER AND PULLING IT OUT, A PARTIAL BREAK WAS FOUND AT 4 CM FROM THE PUNCTURE POINT WHEN IT WAS PULLED OUT, AND THE LENGTH OF THE BREAK WAS THREE-QUARTERS OF THE CIRCUMFERENCE OF THE CATHETER, WHICH WAS TRIMMED AND THEN ROUTINELY MAINTAINED, AND WAS PULLED OUT ON FEBRUARY 5 AFTER COMPLETING CHEMOTHERAPY. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00830 | 2023-03-14 | 2023-02-22 | I, n, j, u, r, y | Manufacturer report | BARD POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER, "WE RECENTLY HAD AN INCIDENT RELATED TO THE POWERPICC WHERE ALL OR A PORTION OF THE GUIDEWIRE TIP BROKE OFF DURING INSERTION, UNFORTUNATELY THE INCIDENT RESULTED IN PATIENT HARM." NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00823 | 2023-03-14 | 2022-11-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER "WHILE PICC GROSHONG 4FR 60 CM WITH MICROINTRODUCER KIT WAS BEING PLACED IN THE PATIENT, IT STARTED PRESENTING A LEAKAGE IN NUMBER 45 AS A 'SQUIRT'". | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00808 | 2023-03-14 | 2023-02-25 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL BASIC TRAY W/ SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT GROSHONG PICC WAS PLACED ON 08 FEB AND DAILY USED FOR NUTRITION THERAPY. IT WAS DISCOVERED DISCONNECTION ON 25 FEB. PATIENT WAS SENT TO OPERATION ROOM FOR REMOVING THE PICC WHICH WAS ENTERED THE RIGHT HEART. ADDITIONAL INFORMATION RECEIVED 3/3/2023: ACCORDING TO USER¿S FEEDBACK, THE CATHETER COULD PASS/ENTER THE CONNECTOR ASSEMBLY SMOOTHLY WITHOUT ANY ISSUE IN THE PAST, BUT NOW, THEY NEED TO ROTATE/SCREW THE CATHETER TO PASS/ENTER THE CONNECTOR ASSEMBLY. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00817 | 2023-03-14 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER 4F RADIOLOGY BASIC TRAY SL 70 CM NITINOL GUIDEWIRE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3194355 |  | IT WAS REPORTED BY THE CUSTOMER THAT "WE HAD AN ISSUE LAST NIGHT WITH A PICC LINE. THE DOCTOR WAS ATTEMPTING TO REMOVE THE WIRE AFTER THE PICC CATHETER WAS INSERTED AND THE WIRE BECAME STUCK. ONCE HE WAS FINALLY ABLE TO REMOVE IT THE WIRE CAME OUT FRAYED. THERE WAS NO PATIENT HARM. THE WIRE DID NOT BREAK. THE EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING SITUATION." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00169 | 2023-03-14 | 2023-01-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U4150425RX | U4150425RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE SUPERFICIAL FEMORAL ARTERY VIA AN IPSILATERAL ANTEGRADE APPROACH, THERE WAS STRONG RESISTANCE FELT AND THE DEVICE ALLEGEDLY FAILED TO BE ADVANCED OVER THE GUIDEWIRE. IT WAS FURTHER REPORTED THAT THE SHAFT OF THE CATHETER WAS ALLEGEDLY KINKED AND WIRE LUMEN CRUSHED. THE BALLOON AND THE GUIDEWIRE WERE REMOVED TOGETHER AS ONE UNIT. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED GUIDEWIRE ADVANCEMENT FAILURE AND CATHETER LUMEN DEFORMATION COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. EXPIRY DATE: 10/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00170 | 2023-03-14 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SIDEKICK | SUPPORT CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | SD110ATJ |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THERE WAS RESISTANCE FELT UPON INSERTION AND THE CATHETER WAS ALLEGEDLY UNABLE TO BE ADVANCED THROUGH THE SHEATH OVER THE GUIDEWIRE. IT WAS FURTHER REPORTED THAT THE TIP OF THE CATHETER WAS FOUND PEELED. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE SIDEKICK SUPPORT CATHETER THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE SIDEKICK SUPPORT CATHETER ARE IDENTIFIED IN COMMON DEVICE NAME AND PMA/510K. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRY DATE: 09/2023. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00790 | 2023-03-14 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6903190 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY HAD A LACERATION IN ONE OF THE CATHETER PATHWAYS AND AIR WAS ALLEGEDLY NOTICED IN THE CATHETER. IT WAS FURTHER REPORTED THAT BLOOD HAD ALLEGEDLY CAME OUT AT THE SITE OF THE CATHETER THAT WAS FOUND LACERATED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00171 | 2023-03-14 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQF75104 | CQF75104 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE FOR MID CEPHALIC VEIN LESION, THE PTA BALLOON ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT WHEN THE BALLOON WAS ATTEMPTED TO REMOVE FROM THE PATIENT, THE BALLOON BECAME STUCK IN THE SHEATH. REPORTEDLY, THE BALLOON AND SHEATH WERE REMOVED AS ONE OVER THE WIRE. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 08/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00792 | 2023-03-14 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 7707540 | 7707540J |  | IT WAS REPORTED THAT ELEVEN YEARS EIGHT MONTH SEVEN DAYS POST A PORT PLACEMENT VIA THE LEFT SUBCLAVIAN VEIN FOR CHEMOTHERAPY, THE CATHETER WAS ALLEGEDLY DAMAGED. IT WAS FURTHER REPORTED THAT DURING CT EXAMINATION, IT WAS REVEALED THAT SALINE SOLUTION WAS LEAKED FROM TWO POINTS WITHIN THE SUBCUTANEOUS TUNNEL. THE CATHETER WAS REMOVED PERCUTANEOUSLY AND THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN COMMON DEVICE NAME AND PMA/510K. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00793 | 2023-03-14 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0602680CE |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE AREA WHERE THE CATHETER AND CATH-LOCK ATTACHES WAS ALLEGEDLY SNAPPED AND BROKE OFF COMPLETELY WHEN THE PORT WAS POSITIONED FOR PORT INSERTION. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE M.R.I IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE M.R.I IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F PRODUCTS ARE IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 12/2024. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01699 | 2023-03-14 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE FOLEY CATHETER BALLOON DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01700 | 2023-03-14 | 2023-02-20 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT AT THE TIME OF REMOVAL, THE WATER COULD NOT BE DRAINED AT ALL, AND DESPITE ATTEMPTS TO REMOVE THE WATER BY CUTTING THE VALVE AND INSERTING THE GUIDEWIRE, IT COULD NOT BE DRAINED AT ALL ,AND FINALLY IT WAS REMOVED BY PERCUTANEOUS PUNCTURE AND FORCED RUPTURE BY A NEARBY UROLOGIST. ALSO STATED THAT THE BALLOON WAS NOT DEFLATED SUFFICIENTLY. THE CATHETER WAS REMOVED BY FUNNEL CUT OR OTHER METHOD AND USUALLY BALLOON CAN BE DEFLATED WITHIN AROUND 15 SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01704 | 2023-03-14 | 2022-09-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® URETHRAL CATHETER TRAY, CLEAR PLASTIC CATHETER | URETHRAL CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | 802214 | 802214 |  | IT WAS REPORTED THAT THE NURSE CARING FOR THE PATIENT ATTEMPTED TO PERFORM A STRAIGHT CATHETERIZATION TO OBTAIN A URINE SPECIMEN. THE URETHRAL CATHETER FOLDED OVER ON ITSELF. THE CUSTOMER TOOK A NEW, CLEAN, CATHETER AND TRIED TO PERFORM ON A MANIKIN, AND THE SAME THING HAPPENED AND THE DEVICE DID NOT DO WHAT IT WAS SUPPOSED TO DO. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01679 | 2023-03-14 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER COUDE TIP | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0168L16 | 0168L16 |  | IT WAS REPORTED THAT THERE WAS A POSSIBLE DEFECTIVE FOLEY CATHETER OR POSSIBLE RECALL. AND MENTIONED THAT THE PRODUCT NUMBER 0168L16 WAS MISSING A CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01712 | 2023-03-14 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HYDROGEL-COATED FOLEY CATHETER KIT | FOLEY CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 |  | 730120 |  | IT WAS REPORTED THAT A PHYSICIAN IN AN ONLINE SURVEY STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE OF UNCLEAR INFORMATION, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO AQUAFIL BIOCATH FOLEY CATHETER, FEMALE LENGTH, 5ML BALLOON, FRENCH SIZE 20. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO INCORRECT TRANSLATION OR INADEQUATE INSTRUCTIONS. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER, OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
16536710 | 2023-03-14 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | PUREWICK MALE EXTERNAL CATHETER KIT | COLLECTOR, URINE, POWERED, NON INDWELLING CATHETER | C. R. BARD, INC. | PWM030 | PWM030 |  | MANUFACTURER INSTRUCTIONS DOES NOT ADDRESS THE NEED TO DISCONNECT THE PUREWICK FROM SUCTION. THERE ARE INSTANCES WHERE THIS DEVICE WOULD NEED TO BE DISCONNECTED FROM SUCTION IN A REAL LIFE SITUATION SUCH AS THERAPY, AMBULATION, TRANSPORTING TO AND FROM IMAGING. IN THESE CASES, WE CANNOT LOOK AT THE MANUFACTURER INSTRUCTIONS FOR ANY GUIDANCE FOR WHAT TO DO. IF WE JUST DISCONNECT THE TUBING, IT CAN LEAD TO THE URINE COMING OUT OF THE DEVICE IF THE PATIENT VOIDS WHICH IS A POSSIBILITY SINCE THESE CAN BE USED FOR INCONTINENCE MANAGEMENT. THIS LEADS TO OUR NURSES TRYING TO GET CREATIVE TO SOLVE THE PROBLEM OF DRIBBLING OR FLOW OF URINE. THE MOST LOGICAL IDEA BEING, PLUG THE PORT SO URINE DOESN¿T COME OUT AND HOOK BACK UP TO SUCTION UPON RETURN. WOULD LIKE TO BE ABLE TO UNHOOK PUREWICK FROM SUCTION TO MOVE PATIENT WITH THERAPY. CURRENTLY THERE ARE NO GUIDELINES IN THE MANUFACTURER INSTRUCTIONS ON HOW TO DO THAT. NO WAY TO CAP OR PLUG THE PUREWICK SO EITHER THE URINE LEAKS OUT THE BOTTOM OF THE PUREWICK WHILE THE PATIENT IS UP OR WE NEED TO TAKE OFF AND PUT A NEW PUREWICK ON AFTER WHICH COULD CAUSE SKIN ISSUES AND EXCESSIVE USE OF SUPPLIES LEADING TO INCREASED COST & WASTE. |  | N | Y |  |  | 2023-03
1018233-2023-01716 | 2023-03-14 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HYDROGEL-COATED FOLEY CATHETER KIT | FOLEY CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 |  | 730118 |  | IT WAS REPORTED THAT A PHYSICIAN IN AN ONLINE SURVEY STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE OF UNCLEAR INFORMATION , INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO AQUAFIL BIOCATH FOLEY CATHETER, FEMALE LENGTH, 5ML BALLOON, FRENCH SIZE 18. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO INCORRECT TRANSLATION OR INADEQUATE INSTRUCTIONS. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01721 | 2023-03-14 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT WHILE ATTEMPTING TO INFLATE THE BALLOON AFTER THE CATHETER PLACEMENT TO THE PATIENT IN WARD, SINCE A POPPING SOUND WAS HEARD, THE PROCEDURE WAS STOPPED. ON CHECKING THE CATHETER, FOUND THAT THE BALLOON HAD BURST. PER FOLLOW-UP RECEIVED VIA IBC ON (B)(6)2023, STATED THAT THERE WERE NO MISSING PIECES OBSERVED TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01724 | 2023-03-14 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | MAGIC 3 GO INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT NEEDED TO EXPLAIN THE FLIP UP AND DOWN AFTER THEY BURST IT GOT MORE LUBED FROM THE INTERMITTENT CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01727 | 2023-03-14 | 2023-02-21 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION. THE REPRESENTATIVE ADVISED THE PATIENT OF PROPER CLEANING PROCESS. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. PER FOLLOW UP VIA PHONE ON 08MAR2023, IT WAS REPORTED THAT THE CUSTOMER DID SEEK MEDICAL ATTENTION AND THE URINARY TRACT INFECTION WAS NOT CONFIRMED TO BE CAUSED BY THE WICKS. CUSTOMER WAS NOT CURRENTLY WICKS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | N | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00804 | 2023-03-14 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | IT WAS REPORTED THAT BURRS WERE FOUND WITH THE STYLET WITHIN THE POWER CATHETER DURING INSERTION. IT WAS UNEVEN. FOREIGN MATTERS WERE ADHERED TO THE SURFACE OF THE GUIDE WIRE. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00805 | 2023-03-14 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | IT WAS REPORTED THAT BURRS WERE FOUND WITH THE STYLET WITHIN THE POWER CATHETER DURING INSERTION. IT WAS UNEVEN. FOREIGN MATTERS WERE ADHERED TO THE SURFACE OF THE GUIDE WIRE. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00807 | 2023-03-14 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D2 |  | IT WAS REPORTED THAT IR PLACED A NEW DEVICE ON 1/17/23 AND WHEN THEY TOOK THAT OUT AT (B)(6)2023 THE PICC WAS KINKED IN 2 PLACES. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00809 | 2023-03-14 | 2023-01-25 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE PATIENT UNDERWENT ROUTINE OUTPATIENT PICC MAINTENANCE, WITH 5CM EXPOSED AND 41CM BUILT-IN. THERE WAS NO LOCAL ERYTHEMA PRESENT. WHEN FLUSHING AND SEALING THE TUBE, THERE WAS NO RETURN BLOOD PUMPING, AND SUBCUTANEOUS SWELLING ABOVE THE PUNCTURE BY GENTLY PUSHING THE SALINE. CLINICIAN SUSPECTED CATHETER LEAKAGE. WHEN GENTLY PULLING THE CATHETER TO FIND THE LEAKAGE, THE CATHETER BROKE OFF AND CAME OUT 1CM INSIDE THE PUNCTURE OPENING. THE BUILT-IN PART WAS NOT SEEN, IMMEDIATELY PRESSED ABOVE THE PUNCTURE POINT, THE FILM SHOWED THE CATHETER IN THE PULMONARY ARTERY. THE CATHETER WAS IMMEDIATELY PRESSED OVER THE PUNCTURE SITE AND THE RADIOGRAPH SHOWED THAT THE CATHETER WAS IN THE PULMONARY ARTERY. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00810 | 2023-03-14 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT CLINICIAN PLANNED TO PLACE ¿GROSHONG NXT CLEARVUE 4F PICC ¿ IN ONE OF HER PATIENTS ON (B)(6) 2023 AROUND 6.00PM BUT WHEN SHE OPENED THE PACKET SHE FOUND THERE WAS A KINK IN CATHETER ALONG WITH STYLET. SHE CHECK THE PATENCY OF CATHETER BY PUSHING SALINE INTO THE CATHETER AND SHE FOUND THERE WAS A LEAKAGE FROM KINK POINT AS SHOWN IN THE IMAGE. THEY HAD TO CANCEL THE PICC LINE INSERTION TODAY AND THEY HAVE GIVEN TREATMENT WITH PIV THIS TIME AND PLANNED PICC LINE INSERTION DURING NEXT SCHEDULED CYCLE TO INSERT AND GIVE FURTHER TREATMENT. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00812 | 2023-03-14 | 2023-01-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REPLACEMENT CONNECTOR FOR 4F GROSHONG NXT CLEARVUE PICC W/EXTENSION LEG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7812400 |  | IT WAS REPORTED "THE PATIENT WAS ADMITTED TO HOSPITAL ON (B)(6) AFTER HALF A MONTH OF CHEMOTHERAPY AFTER RECTAL TUMOR SURGERY. THE PICC CATHETER WAS PLACED IN THE RIGHT UPPER LIMB ON THAT DAY, AND THE PROCESS WAS SMOOTH. TODAY, THE PICC CATHETER TIP WAS BLOCKED DURING TUBE FLUSHING, AND THE TIP OF THE PICC CATHETER WAS REPLACED." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00813 | 2023-03-14 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (4F) (SINGLE-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8194118 |  | IT WAS REPORTED "THE PATIENT WAS IMPLANTED WITH A PICC CATHETER UNDER THE GUIDANCE OF B-ULTRASOUND ON (B)(6) 2022. DURING THE OPERATION, STARTING FROM (B)(6) 2023, A SMALL AMOUNT OF SEALING FLUID WAS FOUND AT THE PUNCTURE SITE DURING THE SEALING PROCESS, BUT CHEMOTHERAPY COULD BE CARRIED OUT NORMALLY. IN FEBRUARY, THE EXUDATION WAS MORE OBVIOUS THAN BEFORE. ON FEBRUARY 14, 2023, THE CATHETER WAS EXTUBATED IN ADVANCE, AND IT WAS FOUND THAT THERE WAS A LITTLE DAMAGE ABOUT 20CM AT THE FRONT END OF THE CATHETER." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00815 | 2023-03-14 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA & PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D2 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THE COMPLAINT OF THE CUSTOMER WAS THAT THE RUBBER STOPPER AT THE END OF THE STYLET T-LOCK ASSEMBLY FOR PICC PLACEMENT IS LEAKING WHEREAS IT NEVER WAS BEFORE. HE ALSO STATED THAT IT LOOKED LIKE A NEW STYLET STOPPER." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK FROM THE T-LOCK SEPTUM WAS INCONCLUSIVE DUE TO THE SAMPLE CONDITION. A SINGLE PHOTOGRAPH OF THE IMPLICATED T-LOCK ASSEMBLY AND STYLET WAS RETURNED FOR EVALUATION. THE T-LOCK ASSEMBLY AND STYLET HAD BEEN REMOVED FROM THE CATHETER. THE STYLET WAS SEEN TRAVERSING THROUGH THE YELLOW SEPTUM OF THE T-LOCK ASSEMBLY, AS PER DESIGN. NO LIQUID WAS SEEN EMANATING FROM THE SEPTUM OF THE T-LOCK DEVICE. HOWEVER, IT DID NOT APPEAR THAT THE DEVICE WAS ACTIVELY BEING INFUSED WHEN THE PHOTOGRAPH WAS TAKEN. WITHOUT THE PHYSICAL SAMPLE, FUNCTIONAL TESTING AND MICROSCOPIC EXAMINATION OF THE SEPTUM COULD NOT BE CONDUCTED. AS THE SUBMITTED PHOTOGRAPH DID NOT CONTAIN ENOUGH INFORMATION TO CONFIRM THE EVENT OR IDENTIFY A SPECIFIC ROOT CAUSE, THIS COMPLAINT WILL BE RECORDED AS INCONCLUSIVE AT THIS TIME. POSSIBLE CONTRIBUTING FACTORS COULD INCLUDE A DAMAGED SEPTUM OR OVER-PRESSURIZATION. IF THE PHYSICAL SAMPLE IS RECEIVED AT A LATER DATE, THIS INVESTIGATION WILL BE UPDATED WITH THE RELEVANT INFORMATION. THE REPORT OF A LEAK HAS BEEN DOCUMENTED AND INCLUDED IN COMPLAINT TRENDING. THIS TYPE OF COMPLAINT WILL CONTINUE TO BE MONITORED AS PART OF ONGOING EFFORTS TO IDENTIFY POTENTIAL MANUFACTURING/SUPPLIER RELATED ISSUES. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00822 | 2023-03-14 | 2023-02-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG 4F SINGLE-LUMEN PICC (60CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7717405 |  | IT WAS REPORTED THAT AFTER UNPACKING AND PRE-FLUSHING DURING PLACEMENT OF THE CATHETER, THE HEAD NURSE FOUND THAT THE CATHETER WAS RUPTURED. DUE TO CONCERN ABOUT THE QUALITY OF THE CATHETER, A NEW CATHETER WAS USED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00845 | 2023-03-15 | 2023-02-22 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE CATHETER WITH SHERLOCK 3CG(TPS) STYLET(4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000516 |  | IT WAS REPORTED "PICC LINE SNAPPED TODAY. MAY HAVE BEEN LIKE THAT FOR DAYS, JUST DIDN'T SEE UNTIL TODAY. WAS BULGING ABOVE INSERTION SITE ON FLUSHED SO REMOVED IT AND ONLY A FEW CMS CAME OUT SO HAD TO GO TO HOSPITAL FOR CARDIAC MONITORING AND RETRIEVAL." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00844 | 2023-03-15 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "WE RECENTLY HAD AN INCIDENT RELATED TO THE POWERPICC WHERE ALL OR A PORTION OF THE GUIDEWIRE TIP BROKE OFF DURING INSERTION." NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00835 | 2023-03-15 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD NEONATAL PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THEY NOTICED BLOOD INSIDE THE PICC LINE CATHETER AND IT WOULD NOT FLUSH. IT WAS STATED THE PICC LINE HAD TO BE REMOVED AND REPLACED. DEVICE WAS A SILASTIC NEONATAL PICC. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00838 | 2023-03-15 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED "THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT: REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. THE EVENT INVOLVED A TL POWERPICC PLACED IN THE UPPER ARM." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01757 | 2023-03-15 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175814 | 175814 |  | IT WAS REPORTED THAT THE BALLOON ON THE FOLEY CATHETER DID NOT INFLATE SYMMETRICALLY DURING TRAINING. DURING TRAINING FOR THE BARD FOLEY TRAYS IN BARCELONA, A HEALTHCARE PERSONNEL POINTED OUT THAT THE BALLOON HAD NOT INFLATED SYMMETRICALLY. ALSO STATED THAT IT APPEARS THAT THE BALLOON WOULD NOT SIT IN THE PATIENT¿S BLADDER TO ENSURE CORRECT DRAINAGE. THE CATHETER HAD BEEN INSERTED INTO AN ANATOMICAL MODEL . AS PER ADDITIONAL INFORMATION RECEIVED ON 22FEB2023, STATED THAT WHEN THE USER WAS PACKAGING THE CATHETER, THEY NOTICED THAT THE TIP WAS NOT SYMMETRICAL AND THERE WERE SOME JAGGED EDGES STICKING OUT THE TOP. | THE REPORTED EVENT IS CONFIRMED MANUFACTURING RELATED AS THE FLASH PRESENT ON THE DEVICE WAS LARGER THAN THE SPECIFICATION RANGE. VISUAL EVALUATION NOTED RECEIVED 1 FOLEY TRAY IN OPENED PACKAGING. VISUAL INSPECTION NOTED JAGGED PIECES OF EXTRA SILICON EXTENDING OUT FROM THE TIP. VISUAL EVALUATION NOTED RECEIVED 1 FOLEY TRAY IN OPENED PACKAGING. VISUAL INSPECTION NOTED JAGGED PIECES OF EXTRA SILICON EXTENDING OUT FROM THE TIP. VISUAL INSPECTION NOTED THAT THE TIP OF THE CATHETER WAS JAGGED WITH PIECES OF EXTRA SILICON EXTENDING OUT FROM THE TIP. THE PIECE WAS MEASURED IN LENGTH TO BE 0.1060 IN ON THE TIP AND 0.2080 IN ON THE SIDE OF THE TIP. A LARGE, JAGGED PIECE OF SILICON WAS MEASURED TO BE 0.0735 IN. ALSO, RECEIVED 3 PHOTO SAMPLES EACH SHOWING THE END OF CATHETER WITH A FLASH PRESENT. EACH PHOTO SHOWS THE FLASH FROM A DIFFERENT ANGLE. BASED ON PHYSICAL SAMPLES RECEIVED THE PRODUCT DOES NOT MEET SPECIFICATIONS ACCORDING TO WHICH STATES "FLASH SHOULD NOT BE GREATER THAN 0.015¿ IF FLASH IS DETECTED ON INSPECTION, MEASURE WITH VISION SYSTEM 20X." ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE IMPROPER CUTTING METHOD OR WORN KNIFE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE LABELLING REVIEW IS NOT REQUIRED AS LABELLING WOULD NOT HAVE PREVENTED THE REPORTED EVENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01756 | 2023-03-15 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175814 | 175814 |  | IT WAS REPORTED THAT THE BALLOON ON THE FOLEY CATHETER DID NOT INFLATE SYMMETRICALLY DURING TRAINING. DURING TRAINING FOR THE BARD FOLEY TRAYS IN BARCELONA, A HEALTHCARE PERSONNEL POINTED OUT THAT THE BALLOON HAD NOT INFLATED SYMMETRICALLY. ALSO STATED THAT IT APPEARS THAT THE BALLOON WOULD NOT SIT IN THE PATIENT¿S BLADDER TO ENSURE CORRECT DRAINAGE. THE CATHETER HAD BEEN INSERTED INTO AN ANATOMICAL MODEL. AS PER ADDITIONAL INFORMATION RECEIVED ON 22FEB2023, STATED THAT WHEN THE USER WAS PACKAGING THE CATHETER, THEY NOTICED THAT THE TIP WAS NOT SYMMETRICAL AND THERE WERE SOME JAGGED EDGES STICKING OUT THE TOP. AS PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 01MAR2023, CONFIRMED THAT THE CATHETER IN THE IMAGE IS THE ONE THAT WAS USED ON ANATOMICAL MODEL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00832 | 2023-03-15 | 2022-12-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT DURING A CATHETER PLACEMENT PROCEDURE, BLOOD RETURN COULD NOT BE CONFIRMED. THERE WAS NO REPORTED PATIENT INJURY. (B)(6) 2023 - THE RETURNED DEVICE EXHIBITED MULTIPLE LONGITUDINAL SPLITS. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF BLOOD RETURN NOT OCCURRING IN A GROSHONG CATHETER IS INCONCLUSIVE DUE TO THE NATURE OF THE RETURNED DEVICE. ONE 4 FR NXT CLEARVUE GROSHONG CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED BLOOD USE RESIDUES PRESENT THROUGHOUT THE ENTIRE CATHETER. NO STYLET WAS RETURNED IN THE CATHETER, AND THE CATHETER HAD BEEN TERMINATED ON THE PROXIMAL END AROUND 43.5 CM (BETWEEN THE 43 CM DEPTH MARKER AND THE 44 CM DEPTH MARKER). A MICROSCOPIC OBSERVATION REVEALED THE MEASUREMENTS OF THE DISTAL VALVE OF THE CATHETER TO BE WITHIN SPECIFICATIONS PER THE APPROPRIATE PART DRAWING. UPON FURTHER MICROSCOPIC EVALUATION OF THE CATHETER, MULTIPLE LONGITUDINAL SPLITS WERE OBSERVED AT THE PROXIMAL END OF THE CATHETER OF ABOUT 1 CM LONG EACH. THE SPLITS APPEARED TO MOSTLY HAVE LONGITUDINAL STRIATIONS AT THE FRACTURE SURFACE, ALTHOUGH THE DISTAL ENDS OF THE SPLITS BEGAN TO HAVE MORE GRANULAR FRACTURE SURFACES. IN AN ATTEMPT TO OBSERVE THE INSIDE OF THE CATHETER, ABOUT 1.5 CM OF THE PROXIMAL END OF THE CATHETER WAS CUT OFF FROM THE REST OF THE CATHETER AND CUT IN HALF SUBSEQUENTLY. THE INSIDE OF THE CATHETER APPEARED TO POTENTIALLY HAVE BRAIDED IMPRESSIONS SIMILAR TO STYLET DAMAGE; HOWEVER, IT IS UNCLEAR WHETHER THIS DAMAGE OCCURRED IN AN ATTEMPT TO DISASSEMBLE THE ORIGINAL DEVICE AFTER IT WAS ASSEMBLED WITH THE TWO-PIECE GROSHONG CONNECTOR. BEFORE THE CATHETER WAS CUT TO BE OBSERVED MICROSCOPICALLY, A FUNCTIONAL EVALUATION OF THE CATHETER INVOLVED USING WATER IN A 12 ML SYRINGE AND ATTACHING THE SYRINGE TO A NON-COMPLAINANT TWO-PIECE GROSHONG PROXIMAL CONNECTOR PIECE. AN ATTEMPT TO FLUSH THE CATHETER BY ATTACHING IT TO THE METAL CANNULA OF THE PROXIMAL CONNECTOR PROVED THE CATHETER TO BE PATENT TO INFUSION; HOWEVER, LEAKING WAS OBSERVED AT SPLITS OCCURRING AT THE PROXIMAL END OF THE CATHETER. ATTEMPTS TO ASPIRATE WATER FROM A SMALL BEAKER FAILED, BUT AIR WAS ABLE TO BE ASPIRATED INTO THE SYRINGE DUE TO THE SPLITS AT THE PROXIMAL END OF THE CATHETER. ABOUT 1.5 CM OF THE FRACTURED SECTION OF THE PROXIMAL END OF THE CATHETER WAS CUT OFF TO ATTEMPT WATER INFUSION AND ASPIRATION AND TO OBSERVE THE SPLITS MICROSCOPICALLY. ATTEMPTS TO ASPIRATE WATER USING THE SAME 12 ML SYRINGE AND NON-COMPLAINANT GROSHONG PROXIMAL CONNECTOR PROVED SUCCESSFUL AFTER THE CATHETER WAS TRIMMED AND THE SPLITS WERE NO LONGER PRESENT ON THE PROXIMAL END OF THE CATHETER. THE CATHETER WAS ALSO PATENT TO INFUSION. THE COMPLAINT OF A GROSHONG CATHETER UNABLE TO HAVE BLOOD RETURN IS INCONCLUSIVE DUE TO THE SPLITS OBSERVED AT THE PROXIMAL END OF THE CATHETER. IT CANNOT BE PROVEN WHETHER THE SPLITS PREVENTED ASPIRATION OF THE DEVICE, OR WHETHER THE SPLITS OCCURRED DURING DISASSEMBLY OF THE DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01750 | 2023-03-15 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166L14 |  | IT WAS REPORTED THAT A PHYSICIAN IN AN ONLINE SURVEY STATED NO, THEY DID NOT AGREE THAT THE URINARY DRAINAGE ACHIEVED FROM USE OF THIS PRODUCT BECAUSE IT WAS NOT MEANT FOR URINE DRAINAGE IN RELATION TO BIOCATH FOLEY CATHETER, STANDARD LENGTH, 30ML BALLOON, FRENCH SIZE 14. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01752 | 2023-03-15 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129416M |  | IT WAS REPORTED THAT INJECTING STERILE WATER, THE CATHETER BALLOON ASYMMETRICALLY INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01749 | 2023-03-15 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HYDROGEL-COATED FOLEY CATHETER KIT | FOLEY CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 |  | 730118 |  | IT WAS REPORTED THAT A PHYSICIAN IN AN ONLINE SURVEY STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE OF UNCLEAR INFORMATION , INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO AQUAFIL BIOCATH FOLEY CATHETER, FEMALE LENGTH, 5ML BALLOON, FRENCH SIZE 18. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
16544929 | 2023-03-15 | 2022-09-09 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | WHEN INTRODUCING THE CATHETER, NURSE MET RESISTANCE. INTRODUCER INSPECTED AND FOUND TO HAVE SLIGHT JAGGED EDGES ON EXTERNAL CANNULA. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01739 | 2023-03-15 | 2023-02-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THERE WERE MULTIPLE CUTS IN THE FOLEY CATHETER SHAFT AND THE CATHETER HAD FALLEN OUT. ALSO STATED THAT IT WAS MOST LIKELY CAUSED BY THE PATIENT. ALSO STATED THAT THE EVENT OCCURRED 4 DAYS AFTER PLACEMENT AND THERE WAS NO BALLOON RUPTURE OR TEAR OCCURRED. IT WAS ALSO NOTED THAT THE WATER LEAKED FROM THE SHAFT AND THE CATHETER WAS PULLED RARELY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01738 | 2023-03-15 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT ONCE AFTER THE URINARY CATHETER WAS IN PLACE, A LITTLE URINE WAS SEEN IN THE COLLECTION BAG. THERE WAS A LEAKAGE OF URINE OBSERVED IN THE PATIENT'S PROTECTION. WHEN THE BALLOON WAS CHECKED IT CONTAINED 10CC OF PPI WATER. THE NURSE THEN PROCEEDED TO REMOVE THE URINARY CATHETER HENCE CREATING THE RISK OF A BLADDER GLOBE. THEN A NEW URINARY CATHETER WAS PUT IN AND SO THE RISK OF INFECTION WAS ALSO INCREASED. THE CARERS ADVISED THAT THIS WAS THE 3RD TIME THEY HAD OBSERVED THIS TYPE OF INCIDENT WITH THIS REFERENCE FOR WHICH NO REPORT HAD BEEN MADE. THE BATCH WAS NO LONGER AVAILABLE IN THE DEPARTMENT AND THE CONCERNED CATHETER FOR THE INCIDENT WAS ALSO NOT KEPT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01747 | 2023-03-15 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT A PHYSICIAN IN AN ONLINE SURVEY STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE OF INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL IN RELATION TO BIOCATH FOLEY CATHETER, STANDARD LENGTH, 5ML BALLOON, FRENCH SIZE NOT SPECIFIED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01736 | 2023-03-15 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE URINE LEAKED FROM THE BIFURCATION OF THE CATHETER. THE CATHETER WAS INSERTED TO THE PATIENT BEFORE SURGERY AND DURING SURGERY, URINE LEAKAGE WAS FOUND IN THE BIFURCATION AREA OF THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01735 | 2023-03-15 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE URINE LEAKED FROM THE INFLATION LUMEN IN A LARGER AMOUNT ON (B)(6) AND IN A SMALLER AMOUNTS ON (B)(6). ON (B)(6), URINE LEAKAGE HAS OCCURRED. ON (B)(6) WHEN THE CATHETER WAS REMOVED, AIR CAME INTO THE SYRINGE VIGOROUSLY. ATTEMPTING SEVERAL TIMES TO REMOVE WATER IN THE BALLOON, IT WAS DIFFICULT TO REMOVE. AFTER ABOUT THE FOURTH ATTEMPT OF REMOVAL, PULLING THE SYRINGE PLUNGER SLOWLY, 6 ML OF FIXATION WATER WAS REMOVED. HOWEVER, THE BALLOON REMAINED SLIGHTLY INFLATED AFTER THE CATHETER REMOVAL.PER FOLLOW UP INFORMATION RECEIVED ON 03MAR2023, STATED THAT ONE DEVICE (ONE PATIENT) EXPERIENCED URINE LEAKAGE 4 TIMES DURING PLACEMENT TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01732 | 2023-03-15 | 2023-02-22 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT YEAST INFECTION FROM USE OF PUREWICK. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR YEAST INFECTION . PER ADDITIONAL INFORMATION RECEIVED ON 03MAR2023, THE PUREWICK FEMALE EXTERNAL CATHETER WAS WORKING GREAT BUT THIRD TIME THE WICK FALLS TO ABSORB ANY URINE. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01733 | 2023-03-15 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE URINE LEAKED FROM THE INFLATION LUMEN IN A LARGER AMOUNT ON (B)(6) AND IN A SMALLER AMOUNTS ON (B)(6). ON (B)(6), URINE LEAKAGE HAS OCCURRED . ON (B)(6), WHEN THE CATHETER WAS REMOVED, AIR CAME INTO THE SYRINGE VIGOROUSLY. ATTEMPTING SEVERAL TIMES TO REMOVE WATER IN THE BALLOON, IT WAS DIFFICULT TO REMOVE. AFTER ABOUT THE FOURTH ATTEMPT OF REMOVAL, PULLING THE SYRINGE PLUNGER SLOWLY, 6 ML OF FIXATION WATER WAS REMOVED. HOWEVER, THE BALLOON REMAINED SLIGHTLY INFLATED AFTER THE CATHETER REMOVAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00836 | 2023-03-15 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED "THIS PATIENT WAS ADMITTED TO HOSPITAL FOR CHEMOTHERAPY OF ADENOCARCINOMA IN THE RIGHT LUNG AND HAD THE THREE-WAY VALVE 4F CATHETER (7655405) PLACED UNDER ULTRASOUND ON (B)(6) 2022. 39 CM OF THE CATHETER WAS INSERTED INTERNALLY AND ANOTHER 5 CM WAS EXPOSED EXTERNALLY. DURING CATHETER MAINTENANCE IN THE AFTERNOON OF (B)(6) 2023, A SAND HOLE WAS NOTED AND LIQUID LEAKED. THIS PATIENT WAS DISCHARGED FROM HOSPITAL AND WILL RETURN TO HOSPITAL FOR RE-EXAMINATION 21 DAYS LATER, WHEN THE CATHETER WAS PLANNED TO BE REPLACED WITH A NEW ONE." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01730 | 2023-03-15 | 2023-02-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI20 | 0165SI20 |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS ATTEMPTED TO REPLACE WITH A NEW ONE 18 DAYS AFTER PLACEMENT, ALL THE STERILE WATER IN THE BALLOON COULD NOT BE REMOVED, AND THE CATHETER WAS PULLED AS IT WAS AND REMOVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01728 | 2023-03-15 | 2023-01-05 | I, n, j, u, r, y | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER HAD A TRANSURETHRAL RESECTION OF THE PROSTATE (TURP) SURGERY, AND THE DESIGN OF THE CATHETER WAS CAUSING THEM ISSUES. THE CUSTOMER DID NOT GIVE MEDICAL INTERVENTION INFORMATION. THE DATE OF THE EVENT WAS JANUARY 5TH. PER FOLLOW-UP INFORMATION RECEIVED ON 22FEB2023, STATED THAT THE CUSTOMER WAS USING A 3-WAY FOLEY AND HAD A SPONTANEOUS ERECTION. DURING THE ERECTION, THE CUSTOMER HAD A TEAR TO THEIR URETHRA. THE CUSTOMER DEVELOPED A STRICTURE AND WAS HAVING A PROCEDURE TO RECTIFY IT. THEY SUGGESTED THAT 2¿3 INCHES NEED TO BE ADDED TO THE CATHETER TO PREVENT THIS ISSUE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | N | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01766 | 2023-03-15 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119116 | 119116 |  | IT WAS REPORTED THAT IN THE PACKAGING OF 119116 CATHETER, IT WAS MENTIONED AS SILICONE ELASTOMER FOLEY CATHETER. WHEN THE USER LOOKED AT THE PHOTO SAMPLE, THE USER THOUGHT IT WAS A 100 PERCENT SILICONE FOLEY CATHETER. THE USER ASKED TO CONFIRM WHICH ONE IT WAS SILICONE ELASTOMER OR 100 PERCENT SILICONE AND IT WAS CONFIRMED THAT 119116 WAS A BARDEX ALL-SILICONE TEMPERATURE-SENSING 400-SERIES FOLEY CATHETER. THE USER ASKED ABOUT FIXING THIS LABEL ON THE COVER OF 119116 AS THE SHAFT STATED IT WAS A 100% SILICONE ELASTOMER. THE USER WAS NOT SURE HOW SOME OUS LABELS DIFFER FROM US. THEY FOUND IT ODD THAT IT STATED LATEX-FREE AND SILICONE-ELASTOMER AND THOUGHT IT COULD HAD DIFFERENT MEANINGS IN DIFFERENT COUNTRIES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01737 | 2023-03-15 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT UPON INJECTING STERILE WATER AFTER THE CATHETER INSERTION TO THE PATIENT, THE FOLEY CATHETER BALLOON BURST. ONLY THE CATHETER WAS RECOVERED. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 14MAR2023, STATED THAT THERE WERE NO MISSING PIECES WERE CONFIRMED TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00843 | 2023-03-15 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PROVENA POWERPICC SOLO 4F DL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000610A |  | IT WAS REPORTED BY THE CUSTOMER THAT "I PLACED A PICC EARLIER TODAY AND WHEN IT CAME TIME TO REMOVE THE PEEL-AWAY PORTION OF THE MICROINTRODUCER, ONE OF THE ORANGE TABS SEPARATED FROM THE SHEATH, WHICH RESULTED IN HAVING TO REMOVE THE REST OF THE SHEATH MANUALLY. LUCKILY, THERE WAS ENOUGH OF THE SHEATH TO GRAB AND I WAS ABLE TO REMOVE THE ENTIRE SHEATH WITHOUT FURTHER INCIDENT. A STATLOCK WAS USED FOR SECUREMENT." IT WAS STATED THERE WAS NO URGENT/LIFE THREATENING MEDICAL SITUATION INVOLVING THE EVENT. THERE WAS NO NEGATIVE IMPACT TO THE PATIENT. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00865 | 2023-03-16 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 8CF INT. W/SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808060 |  | IT WAS REPORTED THAT DURING PORT PLACEMENT PROCEDURE, THE DILATOR SLEEVE ALLEGEDLY TORN AND WAS NOT ABLE TO DILATE THROUGH THE GUIDE. THERE WAS NO REPORTED PATIENT INJURY | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 06/2024) THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00849 | 2023-03-16 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT TI L/P 6 CF INT WOSP ATT SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8716001 |  | IT WAS REPORTED THAT DURING PREPARATION OF A PORT PLACEMENT PROCEDURE, THE TIP WAS ALLEGEDLY FOUND TO BE BROKEN WHEN THEY OPENED THE SET. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 07/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00850 | 2023-03-16 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT SOME TIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN UNDER THE LEFT CLAVICLE WHILE INFUSION AND THE USE WAS SUSPENDED. IT WAS FURTHER REPORTED THAT SEVEN MONTHS AND NINE DAYS POST PORT PLACEMENT, THE PORT WAS REMOVED, AND UPON INSPECTION A BREAK WAS ALLEGEDLY FOUND AT 15 CM OF CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 08/2022). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00864 | 2023-03-16 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT SOMETIMES POST PORT PLACEMENT, THE PORT SYSTEM ALLEGEDLY COULD NOT BE INFUSED AND THEREFORE A COMPUTED TOMOGRAPHY EXAMINATION AS WELL AS CHEMOTHERAPY COULD NOT BE PERFORMED. IT WAS FURTHER REPORTED THAT THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE POWERPORT MRI IMPLANTABLE PORT, ONE CATH-LOCK AND ONE GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, GROSS VISUAL, MICROSCOPIC, TACTILE EVALUATION AND FUNCTIONAL TESTS WERE PERFORMED. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE TO INFUSE ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3 H11: B3, B5, H6 (METHOD, RESULT, CONCLUSION) H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01767 | 2023-03-16 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165PL08 |  | IT WAS REPORTED THAT A PHYSICIAN IN AN ONLINE SURVEY STATED NO. THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE OF UNCLEAR INFORMATION. INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO BARD BIOCATH HYDROGEL COATED LATEX PEDIATRIC FOLEY CATHETER, 3ML, FRENCH SIZE 8. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01794 | 2023-03-16 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316 | 119316 |  | IT WAS REPORTED THAT STERILE WATER LEAKED BEFORE PATIENT USE AND POSSIBILITY OF LUMEN TO LUMEN LEAK. THE EVENT OCCURRED DURING PRETEST AND THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED. WATER LEAKED FROM THE VALVE MENTIONED AS BACKFLOW. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE COULD BE LUMEN COMPROMISED BY A PUNCTURE OR CUT. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE SOLUTION IMMEDIATELY WENT INTO THE DRAINAGE LUMEN AT BIFURCATION AND CREATED LUMEN TO LUMEN LEAK. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿WARNINGS 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY INTENDED USE & EFFECT- EFFICACY: THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. DIRECTIONS FOR USE: 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDO-ELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR. (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01797 | 2023-03-16 | 2022-12-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRETEST BEFORE INSERTION, STERILIZED WATER WAS ABLE TO BE INJECTED BUT COULD NOT BE DRAINED AND ONLY ABOUT 4 CC OF WATER COULD BE DRAWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01795 | 2023-03-16 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 5323 | 35303 |  | IT WAS REPORTED THAT THE MALE EXTERNAL CATHETERS WERE ROLLING UP AND CAUSING LEAKING. ALSO STATED THAT THEY WERE HAVING PROBLEMS TAKING THE SHEATHS OFF AS THEY WERE SO ADHESIVE AND WERE TOO STICKY. THE PATIENT STATED THAT THIS HAD RESULTED IN BLEEDING. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "VISCOMETER FAILURE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "FITTING THE SHEATH THE SHEATH IS A STRAPLESS PENILE SHEATH: IT HAS AN ADHESIVE COATING INSIDE TO ENSURE A SAFE AND SIMPLE PROCEDURE FOR FITTING. IF NECESSARY TRIM PUBIC HAIR. CLEAN AND DRY PUBIC AREA. REMOVE SHEATH FROM ITS WRAPPING. PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CUP OF THE SHEATH. IF UNCIRCUMCISED, DO NOT RETRACT THE FORESKIN BUT ALLOW IT TO REMAIN OVER THE GLANS. AVOID CONTACT BETWEEN THE ADHESIVE AND THE GLANS. SLOWLY UNROLL THE SHEATH ALONG THE SHAFT OF THE PENIS, THEN GENTLY SQUEEZE THE SHEATH AROUND THE PENIS TO ENSURE EVEN ADHESION. TRY TO AVOID LEAVING A ROLLED `COLLAR´ OF SHEATH AROUND THE BASE OF THE PENIS. FINALLY, CONNECT THE URINE COLLECTION BAG. ALWAYS CHECK THAT THE CONNECTIONS ARE SECURE BEFORE USE. WEAR TIME WILL VARY FROM USER TO USER. IT IS RECOMMENDED THAT A SHEATH BE CHANGED EVERY 24 HOURS FOR REASONS OF HYGIENE. TO REMOVE: THE SHEATH MAY BE REMOVED BY GENTLY ROLLING IT O THE PENIS. AVOID SIMPLY PULLING THE SHEATH O. REMOVING THE SHEATH WHILST HAVING A BATH OR SHOWER MAY INCREASE COMFORT. DO NOT USE ON IRRITATED OR COMPROMISED SKIN." | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01791 | 2023-03-16 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT STATED WHEN THEY FOLD THE MAGIC3 HYDROPHILIC MALE INTERMITTENT CATHETER IT KINKED UP WHEN THEY WERE NOT SUPPOSED TO DO SO. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01789 | 2023-03-16 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175814 | 175814 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS USED FOR TRAINING DEMO, ONCE THE FOLEY WAS PLACED IN AN ANATOMICAL MODEL AND THE BALLOON INFLATED, THE BALLOON INFLATED ASYMMETRICALLY AND WOULD NOT DEFLATE FOR REMOVAL. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED 2 PHOTO SAMPLES. FIRST PHOTO SAMPLE FOCUSES IN ON INFLATED ASYMMETRICAL BALLOON WITH SOLUTION. THE SECOND PHOTO SAMPLE SHOWS OVERVIEW OF CATHETER. BASED ON SAMPLE RECEIVED PRODUCT DOES NOT MEET SPECIFICATIONS WHICH STATES "BALLOON TESTING FOR SYMMETRY MUST NOT EXCEED A RATIO OF 70:30 WHEN MEASURED FROM THE SIDE OF THE SHAFT." ALTHOUGH AN EXACT ROOT CAUSE COULD NOT DETERMINED A POTENTIAL ROOT CAUSE COULD BE SHAFT STRENGTH (MATERIALS OF SHAFT IS TOO SOFT AND TORSION OCCURS - NOT ENOUGH SUPPORT). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® EZ-LOK® SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK (FIG. 1A) OR SLIP TIP SYRINGES. OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. SWAB SURFACE OF BARD® EZ-LOK® SAMPLING PORT WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND REMOVE SYRINGE FROM SAMPLE PORT. UNKINK TUBING IF NECESSARY AND TRANSFER URINE SPECIMEN INTO SPECIMEN CUP OR FOLLOW HOSPITAL PROTOCOL. DISCARD SYRINGE ACCORDING TO HOSPITAL PROTOCOL. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB. STERILE UNLESS PACKAGE IS OPENED OR DAMAGED, EXCEPT FOR ANY INDIVIDUALLY PACKAGED COMPONENTS WITHIN THE TRAY WHICH ARE NOT LABELED AS STERILE. THESE COMPONENTS ARE NOT TERMINALLY STERILIZED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. PROPER TECHNIQUES FOR URINARY CATHETER INSERTION: PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION; INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT; USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE; DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED; MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED. CATHETER-TUBING JUNCTIONS: MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING; KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES; EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT; ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE; LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED. WASH HANDS AND DON CLEAN GLOVES: USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERIURETHRAL AREA; EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT; REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL; PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN. NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE: DON STERILE GLOVES; POSITION FENESTRATED DRAPE ON PATIENT; REMOVE FOLEY CATHETER FROM WRAP AND LUBRICATE CATHETER; PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND; WHEN CATHETER TIP HAS ENTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE. INSERT CATHETER TWO MORE INCHES AND INFLATE CATHETER BALLOON. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK. SECURE THE FOLEY CATHETER TO THE PATIENT USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED (SEE STATLOCK® FOLEY STABILIZATION DEVICE IFU). NOTE: PLEASE MAKE SURE PATIENT IS APPROPRIATE FOR USE OF STATLOCK® STABILIZATION DEVICE POSITION HANGER ON BED RAIL AT THE FOOT OF THE BED. NOTE: EXERCISE CARE TO KEEP BAG OFF THE FLOOR. USE GREEN SHEETING CLIP TO SECURE DRAINAGE TUBE TO THE SHEET. MAKE SURE TUBE IS NOT KINKED. INDICATE TIME AND DATE OF CATHETER INSERTION ON PROVIDED LABELS. PLACE DESIGNATED LABELS ON PATIENT CHART AND DRAINAGE SYSTEM. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL. INFLATE CATHETER BALLOON USING ENTIRE 10CC OF STERILE WATER PROVIDED IN THE PREFILLED SYRINGE. NOTE: USE OF LESS THAN 10CC CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00863 | 2023-03-16 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT AFTER SUCCESSFUL ENTRY OF THE INTRODUCER NEEDLE, THE GUIDE WIRE WAS INTRODUCED. DURING REMOVAL OF THE CORE, THE DISTAL END OF THE GUIDE WIRE GOT STUCK. ANOTHER ATTEMPT WAS MADE AND THE NEEDLE WAS REMOVED SUCCESSFULLY. THE GUIDEWIRE WAS ALMOST REMOVED TOTALLY WHEN THE "CORE" WAS REMOVED. THE INTRODUCER SHEATH COULD NOT PASS OVER THE GUIDE WIRE. EVENTUALLY, THE DISTAL END OF THE GUIDE WIRE WAS CUT OFF WITH STERILE SCISSORS BEFORE INTRODUCTION OF THE INTRODUCER SHEATH TO COMPLETE THE PROCEDURE. IT WAS DIFFICULT TO CUT OFF THE DISTAL END OF THE GUIDEWIRE LEADING TO INCREASED WORK TIME. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00862 | 2023-03-16 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT AFTER SUCCESSFUL ENTRY OF THE INTRODUCER NEEDLE, THE GUIDE WIRE WAS INTRODUCED. DURING REMOVAL OF THE CORE, THE DISTAL END OF THE GUIDE WIRE GOT STUCK. ANOTHER ATTEMPT WAS MADE AND THE NEEDLE WAS REMOVED SUCCESSFULLY. THE GUIDEWIRE WAS ALMOST REMOVED TOTALLY WHEN THE "CORE" WAS REMOVED. THE INTRODUCER SHEATH COULD NOT PASS OVER THE GUIDE WIRE. EVENTUALLY, THE DISTAL END OF THE GUIDE WIRE WAS CUT OFF WITH STERILE SCISSORS BEFORE INTRODUCTION OF THE INTRODUCER SHEATH TO COMPLETE THE PROCEDURE. IT WAS DIFFICULT TO CUT OFF THE DISTAL END OF THE GUIDEWIRE LEADING TO INCREASED WORK TIME. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01788 | 2023-03-16 | 2023-02-23 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE MAGIC 3 INTERMITTENT CATHETER WAS NOT WORKING, AND IT MIGHT CAUSE URINARY TRACT INFECTION. THE REPRESENTATIVE SENT SAMPLES OF OTHER CATHETERS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. PER FOLLOW UP VIA PHONE 13MAR2023, CUSTOMER EXPERIENCED UTI AND WAS TREATED WITH ANTIBIOTICS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿COLLAPSED DRAINAGE LUMEN DUE TO WALL THICKNESS AND DUROMETER". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE HYDROPHILIC COATED SILICONE INSTRUCTIONS FOR USE INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." | Y | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01780 | 2023-03-16 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER WAS PAINFUL AND COULD NOT GET IN. PATIENT HAD AN ENLARGED PROSTATE. NO MEDICAL INTERVENTION WAS REPORTED. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR IS NOT REPORTABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01783 | 2023-03-16 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS TORN. AS PER ADDITIONAL INFORMATION RECEIVED ON 01MAR2023, IT WAS STATED THAT AFTER INSERTING THE CATHETER, STERILIZED WATER WAS INJECTED, AND WHEN THE CATHETER WAS PULLED TO FIX IT TO THE BLADDER NECK, IT WAS REMOVED WITHOUT ANY RESISTANCE, AND THE BALLOON PART WAS TORN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "WALL THICKNESS TOO THIN". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿WARNINGS: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. APPLICABLE PATIENTS: ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: METHOD FOR USE: DO NOT REUSE. DO NOT RE-STERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. MATERIAL: BALLOON CATHETER: SILICONE; SILVER COATING. THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO 22 EVERY 2FR: BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER; ACCESSORIES: CLOSED DRAINAGE BAG; THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS. SYRINGE PREFILLED WITH STERILE WATER. WATER SOLUBLE LUBRICANT; ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION; TWEEZERS; GAUZE PADS; WATERPROOF SHEET; COTTON BALLS; GLOVES; STATLOCK FOLEY. INTENDED USE & EFFECT- EFFICACY: THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. DIRECTIONS FOR USE: METHOD OF USE: CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. PRECAUTIONS FOR USE: TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT: OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING: CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. WHEN ENDO-ELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR. WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING: REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING: DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.); BALLOON-RUPTURE METHOD. WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. NON-RUPTURE METHOD: ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. BALLOON-RUPTURE METHOD: INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. MRI SAFETY: BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. MALFUNCTION AND ADVERSE EVENTS: MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE; DIFFICULTY OR FAILURE TO REMOVE THE DEVICE; OCCLUSION OF CATHETER INNER LUMENS; ENCRUSTATION; ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE; DEVICE DAMAGE DUE TO INAPPROPRIATE USE; FAILURE TO MEASURE TEMPERATURE; IMPROPER TEMPERATURE INDICATION. ADVERSE EVENTS: URINARY-TRACT INFECTION; HEMORRHAGE, HEMATURIA; ALLERGY REACTION TO THE DEVICE; CALCULUS FORMATION; EDEMA; PAIN; DISCOMFORT; INJURY OF BLADDER OR URETHRAL; URETHRITIS, URINARY INCONTINENCE; RETAINED BALLOON FRAGMENTS. STORAGE METHOD AND EXPIRATION DATE: STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01777 | 2023-03-16 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 35302 | 35302 |  | IT WAS REPORTED THAT ALL OF THE MALE EXTERNAL CATHETERS WERE DIFFICULT TO REMOVE AS THE ADHESIVE WOULD NOT LET GO AND CAUSED SEVERE IRRITATION WHEN IT REMOVED. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "VISCOMETER FAILURE OR MECHANICAL FAILURE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: 1) VERIFY CORRECT SIZE PRIOR TO USE. 2) TRIM PUBIC HAIR IF NECESSARY. 3) WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. 4) OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. 5) PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. 6) UNROLL THE CATHETER OVER PENIS. 7) GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. 8) CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: 1) ENSURE DRAINAGE BAG IS EMPTY. 2) DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. 3) GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER.¿ | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00872 | 2023-03-16 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 3F SL SOLO MAX BARRIER PLUS KIT, BIOPATCH, PROBE COVER & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1193108D5 |  | IT WAS REPORTED VIA MS&S: "CALLER STATES HER MOM HAS A POWERPICC PROVENA SOLO. SHE STATES WHEN SHE OR THE NURSE FLUSH THE PICC THEY HEAR A WHISTLING SOUND AROUND THE ROUND PORTION OF THE CATHETER HUB. CALLER STATES THERE IS NO SIGN OF LIQUID LEAKING WHEN THEY FLUSH ONLY THE SOUND OF AIR ESCAPING." MS&S RESPONSE: "INFORMED CALLER THAT THEY SHOULD DISCONTINUE USE OF THE CATHETER AND HAVE IT EVALUATED BY A DOCTOR AS SOON AS POSSIBLE. THE RISK OF BLEEDING AND INFECTION IF THERE IS A HOLE IN THE CATHETER IS POSSIBLE." NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01778 | 2023-03-16 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53612G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE ALL NOT LUBRICATED ENOUGH OUT THE PACKET, AND THEREFORE IT CAUSED SEVERE DISCOMFORT FOR THE PATIENT, MAKING THEM SORE. STATED THAT THEY FELT QUITE DRY, AS THOUGH THEY HAD BEEN LEFT OUT TOO LONG. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW-UP INFORMATION RECEIVED VIA IBC ON 03MAR2023, STATED THAT NO MEDICAL INTERVENTION WAS NEEDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS. " | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01775 | 2023-03-16 | 2023-02-23 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE PATIENT HAD TWO URINARY TRACT INFECTIONS WITHIN THE LAST 12 MONTH FROM USING MAGIC 3 GO INTERMITTENT CATHETER. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿PATIENT SENSITIVITY TO MATERIALS.¿ THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. INSTRUCTIONS FOR USE: REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE; OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE; IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE; POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN; REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM); WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET; WASH HANDS." THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01776 | 2023-03-16 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175814 | 175814 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS USED FOR TRAINING DEMO, ONCE THE FOLEY WAS PLACED IN AN ANATOMICAL MODEL AND THE BALLOON INFLATED, THE BALLOON INFLATED ASYMMETRICALLY AND WOULD NOT DEFLATE FOR REMOVAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01772 | 2023-03-16 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | INTS14 | INTS14 |  | IT WAS REPORTED THAT THERE WERE NO HOLES AT THE TIP TO DRAIN BLADDER AND THE PACKAGING WITH LOT NUMBER WAS NOT SAVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01773 | 2023-03-16 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165PL08 |  | IT WAS REPORTED THAT A PHYSICIAN IN AN ONLINE SURVEY STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE OF UNCLEAR INFORMATION, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO BARD BIOCATH HYDROGEL COATED LATEX PEDIATRIC FOLEY CATHETER, 3ML, FRENCH SIZE 8. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO INCORRECT TRANSLATION OR INADEQUATE INSTRUCTIONS. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01771 | 2023-03-16 | 2023-01-05 | I, n, j, u, r, y | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER HAD A FOLEY CATHETER INSERTED, WHILE IT WAS INSERTED, THEY HAD A SPONTANEOUS ERECTION, AND THEIR URETHRA WAS TORN OR RIPPED. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW UP VIA PHONE ON 22FEB2023, IT WAS REPORTED THAT THE PATIENT WAS USING A 3 WAY FOLEY AND HAD A SPONTANEOUS ERECTION. DURING THE ERECTION, THEY HAD A TEAR TO HIS URETHRA, DEVELOPED A STRICTURE AND HAD A PROCEDURE TO RECTIFY IT. THE PATIENT SUGGESTED THAT 2-3 INCHES BE ADDED TO THE CATHETER TO PREVENT THIS ISSUE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01770 | 2023-03-16 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTERMITTENT ROCHESTER CATHETERS | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT CUSTOMER CURRENTLY USING THE INTERMITTENT CATHETER AND USED 3 TO 4 TIMES PER DAY. STATED THAT EVERY 4TH CATHETER HAD A RIDGE DOWN THE LENGTH OF THE CATHETER AND COULD FEEL THE RIDGE WHEN APPLYING LUBRICANT. WHEN THEY USED, THE CATHETER CAUSED TEARING, BLEEDING, AND EXTREME DISCOMFORT. CUSTOMER THROWING AWAY ABOUT 25 PERCENT OF PRODUCT RECEIVED AND WANTED SUGGESTION ON A DIFFERENT CATHETER MORE COMFORTABLE. REPRESENTATIVE SUGGESTED MAGIC 3 COUDE AND BUC16C. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PLC FAILURE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. | N | Y | LAY USER/PATIENT |  | 2023-03
16551870 | 2023-03-16 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | UNKNOWN |  |  | WHILE ATTEMPTING PICC INSERTION, ADVANCED PICC THROUGH INTRODUCER RESISTANCE WAS MET. PULLED CATHETER BACK AND CHECKED INTEGRITY OF CATHETER AND 3CG WIRE. THE WIRE WAS BENT AT ALMOST 90 DEGREE ANGLE ABOUT 4CM FROM THE TIP. WIRE HAD TO BE REPLACED. MANUFACTURER RESPONSE FOR PICC, PICC (PER SITE REPORTER) [REDACTED DATE] SENT INITIAL CONTACT FOR IDENTIFIERS. RETRIEVED THE SAMPLE; EMAILED REP. TO REQUEST AN RMA. |  | N | Y |  |  | 2023-03
3006260740-2023-00848 | 2023-03-16 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6903230 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CLIP ON THE CATHETER WAS ALLEGEDLY RUPTURED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00847 | 2023-03-16 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6903190 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CLIP ON THE CATHETER WAS ALLEGEDLY RUPTURED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01774 | 2023-03-16 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HYDROGEL-COATED FOLEY CATHETER KIT | FOLEY CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 |  | 730120 |  | IT WAS REPORTED THAT A PHYSICIAN IN AN ONLINE SURVEY STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE OF UNCLEAR INFORMATION, INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROFESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO AQUAFIL BIOCATH FOLEY CATHETER, FEMALE LENGTH, 5ML BALLOON, FRENCH SIZE 20. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00891 | 2023-03-17 | 2023-02-28 | I, n, j, u, r, y | Manufacturer report | MIDLINE CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING AROUND PIV SITE. DEVICE INVOLVED IS A POWER MIDLINE IN THE UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00896 | 2023-03-17 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | NA | 5603200J |  | IT WAS REPORTED THAT THE PATIENT HAD A POWER TRIALYSIS IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, DURING CRRT, THE CATHETER WAS OCCLUDED AFTER 21 HOURS OF USE. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF CATHETER OCCLUSION IS CONFIRMED. ONE 20 CM POWERTRIALYSIS DIALYSIS CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED ABUNDANT BLOOD USE RESIDUES THROUGHOUT THE DEVICE. FUNCTIONAL TESTING INCLUDED FLUSHING WATER INTO EACH LUMEN USING A 12 ML SYRINGE TO TEST THE ABILITY TO INFUSE INTO THE POWERTRIALYSIS CATHETER. THE ARTERIAL AND POWER INJECTION LUMENS WERE EACH PARTIALLY OCCLUDED BY BLOOD PRODUCTS INITIALLY. AFTER DRIED BLOOD PRODUCTS WERE CLEARED FROM THE CATHETER, EACH LUMEN PROVED PATENT TO INFUSION. THE VENOUS LUMEN WAS UNREMARKABLE. A MICROSCOPIC OBSERVATION REVEALED BLOOD PRODUCTS DRIED INSIDE EACH LUMEN. THE COMPLAINT OF AN OCCLUDED CATHETER IS CONFIRMED AS BLOOD PRODUCT ACCUMULATION WAS A POTENTIAL CAUSE OF THE PARTIAL OCCLUSION OF THE DEVICE. ONCE THE DEVICE WAS FLUSHED WITH WATER AND SOME BLOOD PRODUCT WAS REMOVED, THE LUMENS PROVED TO BE PATENT TO INFUSION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00895 | 2023-03-17 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602833 | 0602833 |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT, A HOLE WAS FORMED ON THE CATHETER IN THE CATHETER LOCK. IT WAS FURTHER REPORTED THAT THE PATIENT ALLEGEDLY EXPERIENCED SYMPTOMS OF PAIN. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00870 | 2023-03-17 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PRISTINE | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PRISTINE |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE TIP ALLEGEDLY CATCHES AT THE CAVAL JUNCTION. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00890 | 2023-03-17 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD SL POWERPICC PROVENA CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT: FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. THERE WAS NO REDNESS OR SWELLING AT THE SITE. EVENT INVOLVED A SL POWERPICC PROVENA PLACED IN THE UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00886 | 2023-03-17 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER BASIC KIT 4F SL 20CM | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | P6154118 |  | CUSTOMER REPORTED MEDICATION (CYMEVAN) EXTRAVASATED FROM THE MIDLINE AND REPORTS THE SEVERITY OF THE EVENT AS "MAJOR." NO OTHER INFORMATION WAS PROVIDED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00871 | 2023-03-17 | 2022-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS CT POWER-INJECTABLE IMPLANTABLE PORT, CHRONOFLEX, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 7480000 | 7480000 |  | IT WAS REPORTED THAT TEN MONTHS AND FOUR DAYS POST PORT PLACEMENT, IT WAS CONFIRMED IN MAGNETIC RESONANCE IMAGING THAT THE PORT HAD ALLEGEDLY MOVED. IT WAS FURTHER REPORTED THAT THE PORT ALLEGEDLY HAD SHIFTED A LITTLE AS IT WAS NOT ACCESSED IN THE SAME SPOT BUT WOULD TYPICALLY ACCESSED IT WITH THE NEEDLE. FURTHERMORE, THE PATIENT ALLEGEDLY NOTICED SOME PAIN AND TUGGING SENSATION. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING INTERVENTION OR MEDICATION TO TREAT PAIN AND TUGGING SENSATION. THERE WAS NO REPORTED PATIENT INJURY. | H10: ADDITIONAL INFORMATION WAS RECEIVED, THERE WAS NO PROBLEM WITH THE PORT AND THE PORT IS WORKING AS INTENDED. HENCE THIS RECORD WILL NOT BE CONSIDERED AS A COMPLAINT. HOWEVER, SINCE AN INITIAL MDR WAS SUBMITTED, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 04/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00876 | 2023-03-17 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5903190 | 5903190 |  | IT WAS REPORTED THAT APPROXIMATELY EIGHT MONTHS POST DIALYSIS CATHETER PLACEMENT PROCEDURE, THE DIALYSIS CATHETER ALLEGEDLY HAD A CRACK ON THE Y SECTION. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE 19CM GLIDEPATH D/L CATHETER WAS RETURNED FOR EVALUATION. GROSS VISUAL, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED FRACTURE ISSUE AS NO FRACTURES WERE NOTED AND NO LEAKS WERE OBSERVED AS WATER EXITED THE DISTAL END OF THE CATHETER SUCCESSFULLY. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 09/2022). H11: H6 (METHOD, RESULT, CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00878 | 2023-03-17 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0605420CE |  | IT WAS REPORTED THAT DURING PORT PLACEMENT PROCEDURE, THE PLASTIC ROD COMING OUT OF THE CHAMBER ALLEGEDLY BROKE INTO SEVERAL PIECES. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE BARDPORT MRI IMPLANTABLE PORT WAS RECEIVED FOR EVALUATION AND TWO ELECTRONIC PHOTOS WERE PROVIDED FOR REVIEW. VISUAL, MICROSCOPIC VISUAL AND FUNCTIONAL TESTS WERE PERFORMED. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE AND THE IDENTIFIED MATERIAL SEPARATION ISSUE AS THE PORT STEM WAS NOTED TO HAVE A FRACTURE AND A PORTION OF THE PORT STEM WAS NOTED TO BE MISSING. THE PHOTO REVIEW ALSO CONFIRMS THE SAME. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 11/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00179 | 2023-03-17 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE TIP OF THE CATHETER ALLEGEDLY SEPERATED AFTER TWO HUNDRED AND FORTY SECONDS OF ACTIVATION. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN COMMON DEVICE NAME AND PMA/510K. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 03/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01800 | 2023-03-17 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 998804 | 998804 |  | IT WAS REPORTED THAT WHEN INFLATING THE BALLOON DILATION CATHETER DURING OPERATION, THE BALLOON FAILED TO BE INFLATED UNDER SCOPE FOLLOWING WATER INJECTION BY THE ASSISTANT USING A PRESSURE PUMP. THE BALLOON WAS JUDGED TO BE DAMAGED AND THEN REMOVED. A NEW KIT WAS UNPACKED AND USED, CAUSING NO ADVERSE IMPACT ON THE PATIENT. PER FOLLOW UP INFORMATION RECEIVED VIA IBC ON 03MAR2023, STATED THE BALLOON COULD NOT BE INFLATED WITH LIQUID INJECTED AND THERE WAS NO BURST. ALSO STATED THAT THE BALLOON WAS INSERTED INTO THE PATIENT, BUT ONCE THE FAILURE OF INFLATION WAS FOUND, THE BALLOON WAS REMOVED FROM THE PATIENT. AND NO HARM OR INJURY CAUSED ON PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01803 | 2023-03-17 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE HYDROPHILIC INTERMITTENT SILICONE CATHETERS WAS NOT LUBRICATED AND CAUSED INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. PER CUSTOMER VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE CUSTOMER HAD AN INFECTION PRIOR TO USING THE CATHETER. CUSTOMER STATED THAT HE RETURNED THE CATHETERS WITH NO LUBE AND WAS IN TURN SENT THE CORRECT CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01807 | 2023-03-17 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT A 92 YEAR OLD PATIENT HAD URINARY RETENTION. THE DOCTOR DETAINED THE PATIENT ON FEBRUARY 22ND AND PLACED A CATHETER AT THE UROLOGY OUTPATIENT DEPARTMENT. THE PATIENT RETURNED HOME TO A SHORT STAY DESTINATION. AT AROUND 9:00 AM ON FEBRUARY 23RD, THE CATHETER WAS PULLED OUT AT THE SHORT-STAY SITE. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. SAMPLE WAS EVALUATED AND FOUND TEAR ON SHAFT OF CATHETER THAT ALLOW WATER TO LEAK OUT. THE TEAR LOOKS LIKE IT HAD BEEN EXPOSED TO SHARP OBJECT. THERE WAS SCRATCH OBSERVED ALONG SHAFT NEAR THE TEAR. THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE (EXAMPLE: USER RELATED/OPERATOR ERROR/ROUGH HANDLING OR EXPOSE TO SHARP OBJECT/MECHANICAL FORCE). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01805 | 2023-03-17 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE TARGET TEMPERATURE WAS 36C, HOWEVER PATIENT TEMPERATURE WAS 34.5C WITH FOLEY AND WATER TEMPERATURE WAS ONLY 30C ON THE ARCTIC SUN DEVICE. THE FLOW RATE WAS 2.8LPM. IT WAS NOTED THAT THE WEIGHT OF THE PATIENT WAS 62KG WITH MEDIUM PADS IN PLACE. THE EVENT LOG SHOWED ALERT 50 (PATIENT TEMPERATURE 1 ERRATIC) AND ALARM 15 (UNABLE TO OBTAIN A STABLE PATIENT TEMPERATURE). FLEXED TEMPERATURE IN CABLE AND PATIENT TEMPERATURE REMAINED STABLE. ESOPHAGEAL PROBE TO BEDSIDE MONITOR CORRELATING WITH 35C. TREND INDICATOR SHOWED ONE ARROW UP. THE WATER TEMPERATURE WAS 38C BY THE END OF THE CALL. MIS RECOMMENDED NURSE TO CHECK PROTOCOL TO SEE WHAT EXTERNAL INTERVENTIONS MIGHT BE ABLE TO TAKE LIKE WARM BLANKETS, BAIR HUGGER, COVER HEAD, HANDS AND FEET. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "SENSOR WIRE TOO WEAK / THERMISTOR WIRE TOO WEAK". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD WAS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01806 | 2023-03-17 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS REMOVED SPONTANEOUSLY IN THE WARD AFTER PLACEMENT. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. POTENTIAL ROOT CAUSE FOR THIS POTENTIAL FAILURE MODE COULD BE DUE TO THIN RUBBERIZE THICKNESS/MECHANICAL FAILURE/OPERATOR ERROR OR BUBBLE VOID ON RUBBERIZE LAYER THAT MAY LEAD TO LUMEN-TO-LUMEN LEAK. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: BARD® SILVER TSC TRAY SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01810 | 2023-03-17 | 2023-01-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT THE PATIENT'S FOLEY CATHETER WAS ORDERED TO BE REMOVED. WHEN ATTEMPTING TO REMOVE THE CATHETER, RESISTANCE WAS MET AND WAS UNABLE TO BE REMOVED. ONCE REMOVED, THE BALLOON OF THE CATHETER APPEARED TO BE ROLLED EVEN THOUGH THE BALLOON WAS EMPTIED PRIOR TO REMOVAL. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INCORRECT BALLOON DESIGN (BALLOON WALL THICKNESS EXCESSIVE)". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "WARNINGS: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: METHOD FOR USE: DO NOT REUSE; DO NOT RE-STERILIZE. THIS DEVICE CONTAINS 10% POVIDINE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRACT MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. DIRECTIONS FOR USE: CLEAN THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
16564559 | 2023-03-17 | 2023-01-20 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | PowerPICC® Catheters | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 1174108D | 1174108D |  | REFERENCE NUMBER GIVEN TO SUPERVISOR. STEEL NEEDLE ON PICC WOULD INSERT INTO PATIENT. NEEDLE WAS PERFECTLY CENTERED IN VEIN. NEEDLE WOULD NOT GIVE BLOOD RETURN. PATIENT HAD TO BE RE-ACCESSED WITH SHORTER NEEDED LOWER, ACCESS WELL, DIFFICULT PASS BUT WITH PATIENT MOVING WOULD PASS. PICC LINE WOULD NOT PROVIDE READING ON MACHINE. WITHOUT READING PICC LINE COULD NOT BE TRIMMED IF NEEDED WITH CURRENT PATIENT POSITION. X-RAY TAKEN AND PATIENT NEEDED TO GO TO OPERATING ROOM AS PICC WAS TOO FAR OUT FOR DISCHARGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01812 | 2023-03-17 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE LEAD WIRE WAS CUT AFTER EVERY OPERATION, BUT THIS TIME STERILIZED WATER LEAKED OUT AND THE FOLEY CATHETER WAS SPONTANEOUSLY PULLED OUT. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 10MAR2023, IT WAS STATED THAT THEY DID NOT USE THE TEMPERATURE SENSOR ANYMORE, SO THEY CUT IT. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. A POTENTIAL ROOT CAUSE COULD BE LUMEN COMPROMISED BY A PUNCTURE OR CUT. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED THAT THE TEMPERATURE WIRE WAS CUT. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE SOLUTION IMMEDIATELY WENT INTO THE DRAINAGE LUMEN AT BIFURCATION AND CREATED LUMEN TO LUMEN LEAK. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿WARNINGS 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY INTENDED USE & EFFECT- EFFICACY: THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. DIRECTIONS FOR USE: 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDO-ELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR. (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01823 | 2023-03-17 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT 8 DAYS AFTER PLACEMENT TO THE PATIENT, THE CATHETER SPONTANEOUSLY FELL OFF AND ON CHECKING IT, THE BALLOON WAS FOUND BURST. CATHETER WAS NOT PULLED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01826 | 2023-03-17 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT FOLEY CATHETER BALLOON FAILED TO DRAIN DURING THE TEST. THE CUSTOMER CHANGE TO A NEW ONE. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE THIN RUBBERIZE WALL THAT CAUSING COLLAPSE/PINCH INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT¿ AND MIGHT BE CONTRIBUTED BY USER RELATED, EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00883 | 2023-03-17 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TPS STYLET FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | CK000377 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN THE PICC WAS INSERTED THE P WAVE WAS INTERMITTENT, THE PATIENT WAS SENT FOR A CHEST X-RAY TO CONFIRM LOCATION. THE TIP OF THE PICC WAS KINKED A 1CM KINK. THE IR CONSULTANT RECOMMENDED THAT THE PICC LINE BE PULLED BACK A CENTIMETER THEN HAVE ANOTHER CHEST X RAY. THE KINK IN THE PICC WAS STILL EVIDENT AND A LARGER KINK THEY STATE 2 CM. THE LINE WAS IN THE CORRECT PLACE SO WAS LEFT IN SITU AND THE PATIENT WAS BOOKED TO RETURN TO X-RAY ON THE MONDAY TO HAVE THE KINK STRAIGHTENED BY THE CONSULTANT. THE PATIENT WAS TRANSFERRED TO THE HOSPITAL OVER THE WEEKEND AND DETAILS OF THE PICC LINE WERE SENT WITH THE PATIENT. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00881 | 2023-03-17 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 18G X 2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1182250 |  | IT WAS REPORTED BY THE CUSTOMER "CLINICIAN REPORTS ULTRASOUND VISUALIZATION OF A RETAINED FOREIGN BODY AT THE SITE OF A PREVIOUS 18G ACCUCATH ACE INSERTION RIGHT UPPER ARM LOCATION. PATIENT HAD 'RIPPED IV OUT' ON (B)(6) 2023 PER FACILITY STAFF WITH NO REPORTED ABNORMAL FINDING AT THAT TIME. PATIENT TO BE ASSESSED BY CARE TEAM TO DETERMINE PLAN FOR FOREIGN BODY REMOVAL." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00882 | 2023-03-17 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD DL TRIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT: REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. THE EVENT INVOLVED A DL POWERTRIALYSIS CATHETER PLACED IN THE CHEST. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01832 | 2023-03-17 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY PROBE WAS READING ERRATIC TEMPERATURES, SO THEY STOPPED THERAPY AND WAS READING PATIENT TEMPERATURE AT 18C . NURSE HAD AN ESOPHAGEAL AND A RECTAL THAT WERE BOTH CORRELATING AT 34.2C. MIS HAD NURSE TO FLEX TEMPERATURE IN CABLE AND PATIENT TEMPERATURE WAS ERRATIC. NURSE GOT ANOTHER TEMPERATURE IN CABLE FROM ANOTHER DEVICE AND THE SAME ISSUE. THEY REPLACED THE FOLEY AND AFTER A FEW MINUTES THE TEMPERATURES CORRELATING. MIS ADVISED NURSE TO HOLD ONTO THE BARD FOLEY FOR INVESTIGATION. PER FOLLOW UP INFORMATION RECEIVED VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT MALE PATIENT DID COMPLETE THERAPY AND THERE WAS NO IMPACT. NURSE COULD NOT REMEMBER THE AGE AND WEIGHT OF THE PATIENT. NURSE ADVISED BY MIS TO CHANGE THE FOLEY AND THAT RESOLVED THE ISSUE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "SENSOR WIRE TOO WEAK / THERMISTOR WIRE TOO WEAK". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW THAT CANNOT BE COMPLETED DUE TO NO LOT NUMBER. THE PRODUCT CATALOG NUMBER AND LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01834 | 2023-03-17 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616 | 53616 |  | IT WAS REPORTED THAT THE PATIENT OPENED ABOUT SIX INTERNMENT CATHETERS, AND THEY DID NOT HAVE LUBRICANT IN THE PACKETS. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "AIR IN LINES". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. THE CATHETER BECOMES SLIPPERY WHEN WETTED WITH WATER, ELIMINATING THE NEED FOR A SEPARATE LUBRICANT. FOR YOUR ADDED CONVENIENCE, THIS CATHETER IS PACKAGED WITH ITS OWN STERILE WATER. SIMPLY RELEASE THE WATER FROM ITS FOIL PACKET AND THEN TIP THE UN-OPENED CATHETER PACKAGE END-TO-END. THE CATHETER ACTS LIKE A MAGNET TO ATTRACT THE WATER AND ACTIVATE ITS SLIPPERY COATING. FOLLOW THESE STEPS FOR BEST RESULTS. RELEASE THE WATER: PRIOR TO OPENING THE SEALED CATHETER POUCH: APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. WET THE CATHETER: HOLD PACKAGE WITH PRINTED SIDE UP. TIP PACKAGE END-TO-END THREE TO SIX TIMES TO WET CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. USE THE CATHETER: PEEL BACK PACKAGE TO EXPOSE FUNNEL END OF CATHETER. IF DESIRED, USE THE SELF-ADHESIVE TAPE ON THE PACKAGE TO TEMPORARILY ATTACH THE POUCH TO ANY DRY VERTICAL SURFACE. REMOVE CATHETER AND USE ACCORDING TO PHYSICIAN¿S INSTRUCTIONS. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE, DO NOT RESTERILIZE. MADE OF SILICONE ELASTOMER." THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01840 | 2023-03-17 | 2023-02-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165812 | 165812 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON FOLDED BACK ON ITSELF WHEN IT WAS REMOVED, RESULTING IN DIFFICULT REMOVAL. IT CREATED A RING THAT ADDED TO THE CATHETER CIRCUMFERENCE. THIS WAS THE SECOND TIME THIS HAD HAPPENED ON THIS UNIT. THE FIRST-TIME USER ALMOST COULD NOT GET THE CATHETER OUT OF THE LITTLE BOY. PER SAMPLE RECEIVED, THE MATERIAL NUMBER FOR THE RETURN CATHETER IS 165812 WHICH DOES NOT MATCH THE MATERIAL NUMBER IN THIS COMPLAINT. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE BALLOON MATERIAL DOES NOT SHRINK QUICKLY ENOUGH. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: [WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00884 | 2023-03-17 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC 4F SINGLE LUMEN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED THAT THE CUSTOMER PLACED A 4 FRENCH SINGLE LUMEN PICC. PATIENT WENT HOME AND CAME BACK BECAUSE THEIR PICC WASN'T WORKING. WHEN THEY TOOK IT OUT, THE IR DEPARTMENT SAID THE CATHETER WAS KINKED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01843 | 2023-03-17 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® SILASTIC® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 33620 | 33620 |  | IT WAS REPORTED THE FOLEY CATHETER BALLOON DEFLATED AND CAME OUT OF THE URETHRA WHEN THE DOCTOR RETURNED AFTER THEY PLACED IT FOR THE FIRST TIME ON THE PATIENT. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. AS IT WAS UNKNOWN HOW THE INFLATION CAP/VALVE CAME LOOSE AND DISCONNECTED, THIS EVENT WOULD BE CONFIRMED AND THE CAUSE WAS UNKNOWN. ALTHOUGH A SPECIFIC CAUSE COULD NOT BE DETERMINED, A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE, "IMPROPER TOLERANCING (FIT BETWEEN THE CHECK VALVE/CAP ASSEMBLY AND INFLATION FUNNEL)". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER." "VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE." "STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01842 | 2023-03-17 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | MAGIC 3 COUDE INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE LUBE WAS NOT AS GOOD AND SMOOTH COMING OUT WHILE USING THE INTERMITTENT CATHETER. PER FOLLOW UP VIA PHONE ON 08MARCH2023, THE CUSTOMER HAD NO ISSUES WITH INSERTION, BUT UPON REMOVING THE CATHETER THAT IT APPEARED THE LUBRICANT FADES BECAUSE THE WAS MET WITH A LITTLE RESISTANCE. THE PATIENT STATED THAT THE MAGIC 3 GO STILL HAD LUBRICANT UPON REMOVING AND WOULD LIKE THIS PRODUCT TO BE THE SAME ALSO STATED THAT IT IS NOT A MAJOR ISSUE AND JUST WANTED TO REPORT IT FOR FUTURE REFERENCE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR, VISCOMETER FAILURE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00888 | 2023-03-17 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER, SHERLOCK 3CG, 5F MBT, MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1295108D |  | IT WAS REPORTED BY THE CUSTOMER THAT "WHILE PUTTING ON STERILE GOWN FOR PICC PROCEDURE, CUSTOMER NOTICED A TEAR (OR RIP) IN THE SEAM BETWEEN THE GOWN SLEEVE AND THE CUFF." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01821 | 2023-03-17 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX TEMP-SENSING CATHETER | UNKNOWN TEMPERATURE SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 129416M | UNKNOWN |  | IT WAS REPORTED THAT THERE WAS INACCURATE TEMPERATURE READINGS, OR NO READINGS AT ALL. THE REPRESENTATIVE SUGGESTED TO CHANGE THE INTERFACE CABLE AND THAT DID NOT HELP. THEY ALSO TRIED AN ESOPHAGEAL PROBE AND GOT A TEMPERATURE READING AND SO THEY FELT THAT THE FOLEY TEMP CATHETERS WERE NOT WORKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00889 | 2023-03-17 | 2023-02-28 | I, n, j, u, r, y | Manufacturer report | BARD DL PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT: REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. THE EVENT INVOLVED A DL POWERPICC PLACED IN THE CHEST. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00875 | 2023-03-17 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6CF INT W SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806060 |  | IT WAS REPORTED THAT PRIOR TO A PORT PLACEMENT PROCEDURE IN RIGHT SUBCLAVIAN VEIN, THE PUNCTURE TIP OF THE PUNCTURE SHEATH WITH THE DILATOR WAS ALLEGEDLY DAMAGED AND HAD BURRS. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 10/2023). DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00874 | 2023-03-17 | 2023-02-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE INFUSION ALLEGEDLY LEAKED FROM THE PORT AROUND THE ACCESS NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00877 | 2023-03-17 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, WHEN BLOOD RETURN WAS CONFIRMED USING THE NON-CORING STRAIGHT NEEDLE, AIR WAS ALLEGEDLY ASPIRATED AND BLOOD RETURN COULD NOT BE ALLEGEDLY CONFIRMED. IT WAS FURTHER REPORTED THAT WHEN THE NON-CORING ANGLE NEEDLE WAS USED, BLOOD RETURN COULD BE CONFIRMED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00873 | 2023-03-17 | 2023-02-02 | D, e, a, t, h | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT WAS ALLEGEDLY UNABLE TO OBTAIN A BLOOD RETURN. IT WAS FURTHER REPORTED THAT DURING A FOLLOW-UP, THE PATIENT EXPIRED DUE TO PANCREATIC CANCER, EMBOLISM, AND CANCER SPREADING TO THE LIVER CAUSING LIVER FAILURE. REPORTEDLY, THE PORT REMAINS IMPLANTED IN THE PATIENT. THE PATIENTS CAUSE OF DEATH IN RELATION TO THE DEVICE IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01850 | 2023-03-20 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WERE NO ABNORMALITIES IN THE EXPANSION OF THE BALLOON BEFORE INSERTION, BUT AFTER INSERTION INTO THE PATIENT, THE BALLOON WAS DAMAGED AND SPONTANEOUS WITHDRAWAL OCCURRED FOUR TIMES. ALL THE EVENT WAS CONFIRMED WITH THE SAME LOT AND THE REMAINING ONE WAS COLLECTED WITHOUT BEING USED. THE USER THEN USED 29030J14 DUE TO THE LACK OF (B)(4). SIMILARLY, THERE WAS A REPORT DRAFTED SEPARATELY FOR SPONTANEOUS WITHDRAWAL DUE TO BALLOON DAMAGE. PER INVESTIGATOR NOTIFICATION RECEIVED VIA EMAIL ON 12MAR2023 STATED THAT PER SAMPLES RECEIVED. 8 SAMPLES WERE RETURNED. 4 SAMPLES NOT IN PACKAGING APPEARED TO BE USED WITH MATERIAL NUMBER (B)(4) AND MANUFACTURING LOT NUMBER NGGS2164. 4 SAMPLES IN UNOPENED PACKAGING WITH MATERIAL NUMBER (B)(4), 3 WITH BATCH NUMBER NGGT3722 , AND 1 WITH BATCH NUMBER NGGS3548. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01848 | 2023-03-20 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT ALTHOUGH IT WAS CONFIRMED THAT THERE WAS NO ABNORMALITY IN THE BALLOON BEFORE INSERTION, THE BALLOON WAS DAMAGED AFTER INSERTION AND WAS REMOVED SPONTANEOUSLY. A SIMILAR EVENT HAD ALSO OCCURRED WITH THE SAME DEVICE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "HIGH MODULUS SILICONE/ WALL THICKNESS TOO THIN". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "(4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.]". (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] IT ALSO STATES "[DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED". IT ALSO STATES "STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED . | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01845 | 2023-03-20 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT HAD BEEN USING MAGIC3 CATHETERS, BUT RECENTLY THEY NOTICED A CHANGE IN THE LUBRICATION. THEY STATED THE PREVIOUS PACKAGING HAD A WHITE ADHESIVE TAB, AND THEY LIKED HOW MUCH LUBE THERE WAS ON THE CATHETER. THE NEW PACKAGING HAD A BLUE ADHESIVE AND DID NOT SEEM TO HAVE ENOUGH LUBRICATION, AND THEY WERE ALL PRODUCTS 53614G. BD REPRESENTATIVE CONFIRMED THAT THE PRODUCT (53614G LOT: JUGW0050) WAS STILL IN DATE AND WOULD EXPIRE IN 2027. THE CUSTOMER BELIEVED BARD HAD CHANGED SOMETHING WITH THE ADHESIVE AND LUBRICANT. USED DOCUMENT RM7601719 AND CONFIRMED THE ADHESIVE ON 53614G WAS BLUE, BUT DID NOT SEE A RECENT CHANGE IN COLOR. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01855 | 2023-03-20 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POP-ON¿ SELF-ADHERING MALE EXTERNAL CATHETER | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 32302 | 32302 |  | IT WAS REPORTED THAT THE ADHESIVE MADE IT DIFFICULT TO APPLY AND REMOVE THE MALE EXTERNAL CATHETER . THE PATIENT¿S HAIR GOT CAUGHT IN THE ADHESIVE CAUSING PAIN. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: VERIFY CORRECT SIZE PRIOR TO USE. TRIM PUBIC HAIR IF NECESSARY. WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. UNROLL THE CATHETER OVER PENIS. GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: ENSURE DRAINAGE BAG IS EMPTY. DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER.¿ DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01860 | 2023-03-20 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316 | 119316 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON FELL OFF IMMEDIATELY AFTER THE PHYSICIAN INSERTED IT TO THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
16572058 | 2023-03-20 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERGLIDE PRO 20GA, 10CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | F220108PT | F220108PT |  | ATTEMPTED MIDLINE CATHETER INSERTION USING ULTRASOUND GUIDANCE IN THE LEFT CEPHALIC VEIN. WHEN ATTEMPTING TO DEPLOY GUIDEWIRE THEN CATHETER, I MET RESISTANCE, WHICH WAS ALSO HURTING THE PATIENT. SO I STOPPED, PULLED CATHETER BACK, PULLED GUIDEWIRE BACK, THEN REMOVED NEEDLE FROM HIS ARM. INSPECTED THE NEEDLE/CATHETER/GUIDEWIRE AND FOUND THAT PART OF THE CATHETER BROKE OFF FROM THE NEEDLE. I USED THE ULTRASOUND AND WAS ABLE TO VISUALIZE THE CATHETER PARTIALLY IN THE CEPHALIC VEIN. MANUFACTURER RESPONSE FOR POWERGLIDE PRO MIDLINE CATHETER, (BRAND NOT PROVIDED) (PER SITE REPORTER). THEY ARE GOING TO INVESTIGATE TO SEE IF IT WAS A MANUFACTURING DEFECT OR USER ERROR. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01878 | 2023-03-20 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 37302 | 37302 |  | IT WAS REPORTED THAT THE ADHESIVE MADE IT DIFFICULT TO APPLY AND REMOVE THE MALE EXTERNAL CATHETER. THE PATIENT¿S HAIR GOT CAUGHT IN THE ADHESIVE CAUSING PAIN. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "VISCOMETER FAILURE OR MECHANICAL FAILURE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: 1) VERIFY CORRECT SIZE PRIOR TO USE. 2) TRIM PUBIC HAIR IF NECESSARY. 3) WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. 4) OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. 5) PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. 6) UNROLL THE CATHETER OVER PENIS. 7) GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. 8) CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: 1) ENSURE DRAINAGE BAG IS EMPTY. 2) DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. 3) GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER.¿ | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00897 | 2023-03-20 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1606062 | 1676300 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE ACCESS NEEDLE ALLEGEDLY CRACKED. IT WAS FURTHER REPORTED THAT THE NEEDLE ALLEGEDLY HAD SUCTION PROBLEM AND THE EDGES ARE TOO SHARP. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01869 | 2023-03-20 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE JELLY LEAKED AND THE JELLY PACKAGING WAS DAMAGED WHEN THE PRODUCT WAS TAKEN OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01849 | 2023-03-20 | 2023-02-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167L20 |  | IT WAS REPORTED THAT ON THE NEXT DAY AFTER REPLACEMENT ON FEBRUARY 24 THE FOLEY CATHETER SPONTANEOUSLY FELL OFF. CHECKING THE CATHETER BALLOON THERE WAS A CRACK FOUND ON IT. THE PATIENT DID NOT APPEAR TO HAVE PULLED ON THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. EVALUATION FOUND IRREGULAR BALLOON BURST WITHOUT MISSING PIECE. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX; PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES]: BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO26EVERY 2FR; BALLOON CATHETER: ACCESSORIES: SYRINGE PREFILLED WITH STERILE WATER; WATER-SOLUBLE LUBRICANT; (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.). [INTENDED USE & EFFECT- EFFICACY]: THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE]: METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. PRECAUTIONS FOR USE: WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL.SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS]: PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS): EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.); BALLOON-RUPTURE METHOD: WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. NON-RUPTURE METHOD: ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. BALLOON-RUPTURE METHOD: INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. MALFUNCTION AND ADVERSE EVENTS: MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE; DIFFICULTY OR FAILURE TO REMOVE THE DEVICE; OCCLUSION OF CATHETER INNER LUMENS; ENCRUSTATION; ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE; DEVICE DAMAGE DUE TO INAPPROPRIATE USE. ADVERSE EVENTS: URINARY-TRACT INFECTION; HEMORRHAGE, HEMATURIA; ALLERGY REACTION TO THE DEVICE; CALCULUS FORMATION; EDEMA; PAIN; DISCOMFORT; INJURY OF BLADDER OR URETHRAL; URETHRITIS, URINARY INCONTINENCE; RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]: STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01866 | 2023-03-20 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175814 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS TOO LONG, THE PATIENT SAID THEY WOULD HAVE TO BE A COUPLE OF FEET TALLER FOR THE ITEM TO FIT PROPERLY. ALSO, THE PATIENT STATED THAT THE TUBING GOES LONGER THAN USUAL AND ONLY ONE CATHETER HAD THIS ISSUE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿INSPECTION RESULTS (MEASUREMENT, TESTING DATA) NOT VERIFIED BY IQA ASSIGNEE ERROR OF INSPECTOR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE SILICONE. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LAWS AND REGULATIONS.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01861 | 2023-03-20 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER LEAKED FROM THE FOLEY CATHETER BALLOON AREA. | THE REPORTED EVENT WAS CONFIRMED HOWEVER THE CAUSE WAS UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND SOLUTION IMMEDIATELY LEAKED OUT OF A 0.013" PINHOLE FOUND ON THE BALLOON. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00898 | 2023-03-20 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, WHEN THE VESSEL WAS PUNCTURED AFTER CONNECTING A SYRINGE TO THE HUB OF THE INTRODUCER NEEDLE, BLOOD COULD NOT BE ASPIRATED, AND AIR WAS ALLEGEDLY ASPIRATED. IT WAS FURTHER REPORTED THAT WHEN THE HUB OF THE INTRODUCER NEEDLE WAS CHECKED, MULTIPLE CRACKS WERE FOUND ON THE HUB OF THE INTRODUCER NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE INTRODUCER NEEDLE WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A CRACK WAS NOTED ON THE INTRODUCER NEEDLE HUB. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED SUCTION AND THE AIR/ GAS IN DEVICE ISSUES AS THE EXACT CIRCUMSTANCE AT THE TIME OF THE EVENT REPORTED WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00181 | 2023-03-20 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U357574 | U357574 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE DEVICE ALLEGEDLY BROKEN. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 03/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01882 | 2023-03-20 | 2023-02-23 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129616 |  | IT WAS REPORTED THAT 3 DAYS AFTER PLACEMENT, HAVING ATTEMPTED TO REMOVE THE BALLOON BY INSERTING A SYRINGE INTO THE INFLATION VALVE BUT WAS UNABLE TO REMOVE THE WATER. CONSULTED A DOCTOR, AND REMOVED THE CATHETER BY THE BALLOON BURSTING METHOD. THE USER INDICATED THAT THE CATHETER TIP FELT HARDER THAN THE PRODUCT HE NORMALLY USES. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE POOR SAMPLE CONDITION. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. OBSERVED THE BALLOON ALREADY DAMAGE DUE TO USER PUNCH THE BALLOON. SAMPLE WAS EVALUATED AND INFLATION EYE ARE MEET INTERNAL SPECIFICATION. DUE TO POOR SAMPLE CONDITION THE COMPLETE INVESTIGATION CANNOT BE PERFORMED. A POTENTIAL ROOT CAUSE OF THIS FAILURE MODE IS COULD BE THIN RUBBERIZE WALL THAT COULD RESULTS IN LATEX DIP SPEED OUT TOO FAST AND ALSO MIGHT BE USER RELATED (OVER ASPIRATED, INCORRECT SYRINGE)/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00923 | 2023-03-21 | 2023-02-02 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED "THE CATHETER WAS INSERTED INTO THE PATIENT IN (B)(6) 2022, AND THE ROUTINE CHEST RADIOGRAPH EXAMINATION BEFORE TREATMENT ON (B)(6) 2023, SHOWED THAT THE CATHETER WAS BROKEN ABOUT 2CM AWAY FROM THE PUNCTURE POINT, AND THE BROKEN CATHETER WAS LOCATED IN THE PULMONARY ARTERY. THE BROKEN CATHETER WAS REMOVED BY INTERVENTIONAL OPERATION THE NEXT DAY, DURING WHICH THE PATIENT AND HIS FAMILY HAD NO DISSATISFACTION." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BROKEN CATHETER IS CONFIRMED AND WAS DETERMINED TO BE RELATED TO FLEXURE FATIGUE OF THE CATHETER. ONE 4 FR S/L GROSHONG NXT CLEARVUE CATHETER WITH THE TWO-PIECE CONNECTOR ATTACHED WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUES THROUGHOUT THE CATHETER AND A COMPLETE BREAK AT THE 34 CM DEPTH MARKER. THE CATHETER WAS RETURNED IN TWO SEGMENTS: ONE FRAGMENT OF CATHETER FROM THE 34 CM DEPTH MARKER TO THE DISTAL VALVE, AND ONE SECTION OF CATHETER ATTACHED TO THE TWO-PIECE CONNECTOR ENDING ON THE DISTAL END AT THE 34 CM DEPTH MARKER. A MICROSCOPIC OBSERVATION REVEALED A SPLIT IN THE CATHETER BETWEEN THE 33 CM AND 34 CM DEPTH MARKERS IN ADDITION THE BREAK FOUND AT THE 34 CM DEPTH MARKER. THE BREAK EXHIBITED A SMOOTH AND ROUNDED FRACTURE SURFACE ON ONE HALF OF THE BREAK, AND A MORE GRANULAR FRACTURE SURFACE ON THE OTHER HALF OF THE BREAK. THE SPLIT APPEARED TO HAVE A VERY ROUNDED, GLOSSY FRACTURE SURFACE THAT WAS CIRCUMFERENTIALLY ALIGNED. THE CORNERS OF THE SPLIT APPEARED TO BE VERTICALLY ALIGNED WITH THE LENGTH OF THE CATHETER. THE ROUNDED AND SHINY EDGES OF BOTH THE SPLIT AND THE BREAK ARE LIKELY DUE TO FLEXURAL FATIGUE OF THE DEVICE. THE COMPLAINT OF A BROKEN CATHETER WITHIN THE PATIENT IS CONFIRMED. MICROSCOPIC OBSERVATIONS OF THE BREAK AND THE SPLIT REVEAL THE DEVICE TO BE WORN FROM RUBBING OF THE CATHETER AGAINST ITSELF IN A KINKED POSITION, WHICH IN TURN LEADS TO FLEXURAL FATIGUE OF THE SILICONE AND ULTIMATELY A SPLIT AND/OR A BREAK IN THE CATHETER. THE LENGTH OF TIME THE CATHETER REMAINED INSIDE THE PATIENT COULD ALSO BE A CONTRIBUTING FACTOR. THE LOCATION OF THE DAMAGE SUGGESTED THAT CATHETER SECUREMENT TECHNIQUES MAY HAVE CONTRIBUTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00926 | 2023-03-21 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY THE CUSTOMER, "WHILE ATTEMPTING TO INSERT AN ACCUCATH ACE THIS MORNING (LOT # REGZ1729), I WAS ABLE TO GET INTO THE VEIN NO PROBLEM VIA ULTRASOUND, ADVANCED THE GUIDE WIRE SMOOTHLY, ADVANCED THE CATHETER BUT IT WOULDN'T ADVANCE SO I WAS TRYING TO RETRACT THE GUIDE WIRE AND LOOK WITH ULTRASOUND AGAIN, BUT THE GUIDEWIRE WOULDN'T RETRACT. I WAS MEETING EXTREME RESISTANCE AND DID NOT WANT TO FORCE IT. AFTER WAITING A FEW MOMENTS, I TRIED AGAIN. I WAS ABLE TO RETRACT IT HALFWAY BUT NO FURTHER. I PRESSED THE BUTTON TO HOPEFULLY RETRACT THE NEEDLE AND GUIDE WIRE AND IT DIDN'T DO ANYTHING. AFTER A FEW ATTEMPTS, IT RETRACTED THE NEEDLE, AND THE GUIDEWIRE WAS BENT." 03/09/2023 ADDITIONAL INFORMATION PROVIDED: PATIENT IS BEING TREATED FOR ELEVATED FERRITIN-THERAPEUTIC PHLEBOTOMY. PATIENT DOES HAVE A LOT OF SCAR TISSUE FROM CHEMOTHERAPY TREATMENTS. IT WAS STATED THERE WAS NO PATIENT HARM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DIFFICULTY ENCOUNTERED WHILE INSERTING AN ACCUCATH CATHETER WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. A SINGLE PHOTOGRAPH OF THE IMPLICATED 20G ACCUCATH ACE DEPLOYMENT DEVICE WAS RETURNED FOR EVALUATION. THE CATHETER HAD BEEN PREVIOUSLY DEPLOYED, AND THE NEEDLE HAD BEEN RETRACTED BACK INTO THE DEPLOYMENT DEVICE. IT WAS REPORTED THAT DIFFICULTY WAS ENCOUNTERED ADVANCING THE CATHETER. HOWEVER, THE CAUSE FOR THE INITIAL INSERTION DIFFICULTY OF THE CATHETER COULD NOT BE DETERMINED FROM THE RETURNED PHOTOGRAPH. POSSIBLE CONTRIBUTING FACTORS COULD INCLUDE A DAMAGED CATHETER TIP PREVENTING A SMOOTH CATHETER INSERTION OR INSERTION TECHNIQUE. THE DISTAL SECTION OF THE GUIDEWIRE WAS SEEN PROTRUDING FROM THE DEPLOYMENT DEVICE, AS PER DESIGN. THE GUIDEWIRE APPEARED TO BE INTACT WITH THE COILED TIP PRESENT ON THE WIRE. IT COULD NOT BE DETERMINED FROM THE PHOTOGRAPH IF THE DISTAL TIP OF THE GUIDEWIRE CONTAINED ANY DAMAGE OR DEFORMATION. A BEND WAS SEEN IN THE WIRE, APPROXIMATELY HALF-WAY BETWEEN THE DISTAL END OF THE DEPLOYMENT DEVICE AND THE DISTAL END OF THE WIRE. THE KINK IN THE WIRE MAY BE INDICATIVE OF A DIFFICULT PLACEMENT/REMOVAL. POSSIBLE CONTRIBUTING FACTORS COULD INCLUDE THE INTERACTION BETWEEN THE GUIDEWIRE AND THE NEEDLE OR DAMAGE TO THE GUIDEWIRE PREVENTING IT FROM PROPERLY RETRACTING. THE IFU WARNS, ¿DO NOT FORCE OR RETRACT THE GUIDEWIRE. RETRACTING THE GUIDEWIRE MAY INCREASE THE RISK OF GUIDEWIRE DAMAGE. IF THE GUIDEWIRE MUST BE RETRACTED, REMOVE THE ENTIRE DEVICE TO PREVENT THE NEEDLE FROM DAMAGING OF SHEARING THE GUIDEWIRE.¿ IT WAS ALSO REPORTED THAT THE USER HAD DIFFICULTY RETRACTING THE NEEDLE. POSSIBLE CONTRIBUTING FACTORS COULD INCLUDE THE INTERACTION BETWEEN THE GUIDEWIRE AND THE NEEDLE OR A BENT NEEDLE. THE IFU WARNS, ¿DO NOT BEND THE NEEDLE BEFORE OR DURING USE AS THIS MAY AFFECT PROPER NEEDLE RETRACTION.¿ ALTHOUGH THE MECHANISM OF DAMAGE COULD NOT BE DETERMINED, IT IS LIKELY THAT THE DIFFICULT CATHETER INSERTION, DIFFICULTY RETRACTING THE GUIDEWIRE, AND DIFFICULTY ENGAGING THE SAFETY WERE CASCADING EVENTS AND WERE LIKELY RELATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00901 | 2023-03-21 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY FIVE YEARS AND TWO MONTHS POST PORT PLACEMENT VIA THE LEFT SUBCLAVIAN VEIN, THE ANTICANCER AGENT ALLEGEDLY LEAKED NEAR THE PORT BODY. IT WAS FURTHER REPORTED THAT WHEN COMPUTED TOMOGRAPHY EXAMINATION WAS PERFORMED, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN SUBCUTANEOUSLY. REPORTEDLY THE PORT BODY AND THE DISTAL CATHETER SEGMENT WERE REMOVED AND ANOTHER NEW PORT SYSTEM WAS REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 03/2020). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00906 | 2023-03-21 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE THE CATHETER ALLEGEDLY FRACTURED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRY DATE: 04/2024. SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00910 | 2023-03-21 | 2023-01-26 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1806060 | 1806060 |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS AND NINE MONTHS POST PORT PLACEMENT, THE TIP OF THE CATHETER ALLEGEDLY BROKE OFF IN THE SUPERIOR VENA CAVA DURING PORT REMOVAL PROCEDURE. IT WAS FURTHER REPORTED THAT UPON FLUOROSCOPY EXAMINATION IT WAS CONFIRMED THAT A FRAGMENT OF THE CATHETER WAS ALLEGEDLY RETAINED IN THE INTRAVASCULAR SPACE. REPORTEDLY THE RETAINED FRAGMENT WAS REMOVED BY PERCUTANEOUS ENDOVASCULAR RETRIEVAL PROCEDURE. THE PORT WAS REMOVED. THE CURRENT STATUS OF THE PATIENT IS STABLE. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00908 | 2023-03-21 | 2022-12-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT LOW-PROFILE IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0655870 | 0655870 |  | IT WAS REPORTED THAT APPROXIMATELY ONE MONTH EIGHT DAYS POST PORT PLACEMENT, A HOLE WAS ALLEGEDLY FOUND IN THE CATHETER. IT WAS FURTHER REPORTED THAT THE DRUG HAD ALSO FLOWED INTO THE TISSUE. REPORTEDLY THE PORT WAS REMOVED AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 08/2025. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00909 | 2023-03-21 | 2023-02-08 | I, n, j, u, r, y | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 7707540 | 7707540 |  | IT WAS REPORTED THAT FOUR MONTHS AND FOUR DAYS POST PORT PLACEMENT, THE CATHETER ALLEGEDLY BROKE AND MIGRATED TO HEART. IT WAS FURTHER REPORTED THAT TWO SEGMENTS OF THE CATHETER WERE REMOVED AND ONE SEGMENT OF THE CATHETER HAS NOT BEEN REMOVED. REPORTEDLY, THE PATIENT HAD A CARDIORESPIRATORY ARREST AND WAS RESUSCITATED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 01/2024), H6 (METHOD). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00924 | 2023-03-21 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC FT CATHETER WITH SHERLOCK 3CG TPS STYLET (5F) (DL) (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000671 |  | IT WAS REPORTED BY HEALTHCARE PROFESSIONAL, "NURSE HAS A PATIENT WITH POWERPICC FT 5F DL CK000671 REGV3387. SHE STATES IT IS CLOTTED AND THEY WANT TO USE ALTEPLASE AND NEED THE PV OF THE CATHETER. NURSE WAS INFORMED THE TOTAL LENGTH OF THIS CATHETER PRIOR TO CUTDOWN IS 55CM WITH PV OF 0.56 ML FOR EACH LUMEN. INFORMED THAT SHE WILL NEED TO FIND OUT WHAT THE CATHETER WAS CUTDOWN TO AND CALCULATE THE CURRENT PV." | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01891 | 2023-03-21 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN THE FOLEY CATHETER BALLOON WAS PRE-TESTED BEFORE USE ON THE PATIENT, IT WAS DIFFICULT TO INFLATE. NEXT, THE STERILE WATER WAS ALSO DIFFICULT TO REMOVE WHEN ATTEMPTING TO REMOVE IT. THEN, ANOTHER CATHETER WAS USED. | THE REPORTED EVENT IS CONFIRMED MANUFACTURING RELATED AS EXTRA SILICONE WAS PRESENT BLOCKING MAJORITY OF THE LUMEN. VISUAL EVALUATION NOTED RECEIVED A 3-WAY TEMP SENSING CATHETER WITH CUT PORTION OF INLET TUBING. THERE APPEARED TO BE A SMALL BUBBLE OF EXTRA SILICON BLOCKING THE MAJORITY OF THE LUMEN. PIN GAUGES 0.024 AND 0.040 WERE INSERTED FROM BOTH DIRECTIONS AND COULD NOT GO THROUGH THE LUMEN. ALSO RECEIVED 3 PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWS OVERVIEW OF CATHETER. SECOND PHOTO SAMPLE SHOWS SLIGHTLY INFLATED BALLOON WITH SOLUTION. THIRD PHOTO SAMPLE THAT SHOWS WRITTEN DOCUMENT IN JAPANESE. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE: LUMEN COLLAPSES UNDER THE PRESSURE OF THE BALLOON; PINCHED LUMEN; KINKED LUMEN. THE PRODUCT WAS USED FOR PATIENT DIAGNOSTIC OR TREATMENT. THE PRODUCT WAS INFLUENCED BY THE REPORTED EVENT. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. LABELLING REVIEW IS NOT REQUIRED AS LABELLING WOULD NOT HAVE PREVENTED THE REPORTED EVENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00922 | 2023-03-21 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY THE CUSTOMER THAT "ON ACCUCATH INSERTION ATTEMPTED TO DEPLOY GUIDEWIRE AND MET RESISTANCE. RETRACTED WIRE BACK AND REPOSITIONED ACCUCATH . ATTEMPTED TO RE ADVANCE GUIDEWIRE AND UNABLE. NOTED ACCUCATH WIRE EXITING THRU THE INSTAFLASH HOLE ON NEEDLE WHEN ACCUCATH REMOVED FROMPATIENT AFTER NOT BEING ABLE TO DEPLOY WIRE SECOND TIME." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00921 | 2023-03-21 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO (TLS) STYLET (5F) (TRIPLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S9395108D |  | IT WAS REPORTED BY THE CUSTOMER THAT THE "POWERPICC SOLO PROVENA TRIPLE LUMEN PICCS ARE KINKING." THEY STATE THIS IS NOT HAPPENING ON THE NON-SOLO TRIPLE LUMEN PROVENA PICCS. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01846 | 2023-03-21 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS FOUND TO BE DISLODGED DURING DIAPER CHANGE ON (B)(6) 2023. THE MICROPORE TAPE SECURING THE TUBING WAS INTACT AND THE BALLOON WAS FULLY DEFLATED. 10 ML OF WATER WAS INJECTED INTO THE BALLOON PORT AND WAS NOTICED TO BE LEAKING. THE USER REPLACE IT WITH A NEW CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE COULD BE DUE TO BUBBLE VOID ON RUBBERIZE LAYER THAT MAY LEAD TO ISSUE LUMEN TO LUMEN LEAKAGE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00920 | 2023-03-21 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THAT WHEN THE NURSE WOULD CHANGE THE DRESSING AND FLUSH THE PICC LINE, IT LEAKS AT THE WINGS. ADDITIONAL INFORMATION RECEIVED 03/13/2023: CUSTOMER REPORTS THE CATHETER WAS SECURED WITH A STATLOCK AND WAS REMOVED. NO NEW CATHETER WAS PLACED, AND NO PATIENT HARM REPORTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK NEAR THE SUTURE WINGS WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR S/L POWERPICC SOLO. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED JUST DISTAL OF THE MOLDED SUTURE WINGS. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR ¿ 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER ¿ OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING) AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00918 | 2023-03-21 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THAT THERE WAS LEAKAGE FROM THE WING ON FEBRUARY 23RD AFTER GIVEN TREATMENT WITH CHEMO (EPIRUBICIN AND CYKLOFOSFAMID). ADDITIONAL INFORMATION RECEIVED: "THE REPORTED LEAK WAS DISCOVERED DURING TREATMENT. THE PATIENT SAW FLUID UNDER THE DRESSING (THE LEAK WAS EXTERNAL). THE PATIENT HAD RECEIVED THE ENTIRE DAY'S TREATMENT AND WAS BOOKED FOR A NEW PICC WITHIN TREATMENT INTERVALS. CUSTOMER STATES IT WAS DIFFICULT TO DETERMINE THE IMPACT ON THE SKIN BECAUSE THE PATIENT HAD REACTED TO THE DRESSING OVER THE PICC. THE PICC WAS REMOVED IN ONE PIECE. THE PICC HAD BEEN SECURED WITH STATLOCK. WE LOOKED AND MEASURED THE CATHETER BEFORE IT WAS PULLED AND FLUSHED THE CATHETER AFTER IT WAS PULLED AND SAW THAT IT WAS LEAKING." | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00917 | 2023-03-21 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | CUSTOMER REPORTED THAT WHEN UNPACKING THE POWERPICC, THERE WAS BLACK FOREIGN MATTER IN THE PACKAGE OF THE CATHETER. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01922 | 2023-03-21 | 2022-10-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DAVOL® FEMALE CATH KIT | FEMALE CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 0035380 | 0035380 |  | IT WAS REPORTED THAT CATHETER IN THE FEMALE STRAIGHT CATHETER KIT WAS VERY LOOSE INSIDE OF CAP. CATHETER USED TO BE HELD IN PLACE INSIDE CAP WHICH WOULD ALLOW IT TO BE PULLED FORWARD BUT WOULD PREVENT IT FROM SLIDING BACK IN TUBE OR FALLING OUT. STATED THAT THE DANGER WAS THAT THE CATHETER COULD BE ADVANCED TO FAR FORWARD INTO THE URETHRA AND BECOME DIFFICULT TO REMOVE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INLET PORT OF THE CAP / CATHETER (FR.) THICKNESS OUT OF SPECIFICATION". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "OPEN PACKAGE AND REMOVE PLASTIC WALLET. OPEN PLASTIC WALLET AND DON GLOVES. PULL CATHETER OUT OF TUBE TO DESIRED LENGTH. LAY TUBE IN STERILE FIELD. OPEN LUBRICANT AND LUBRICATE CATHETER. OPEN SWAB PACKET. CLEANSE VAGINAL AREA. PROCEED WITH CATHETERIZATION. PULL CATHETER OUT OF TOP; TIGHTEN COVER AND DEPRESS BLUE SPOUT. FILL OUT LABEL, PLACE ON CENTRIFUGE TUBE. SEND TO LAB IN NORMAL MANNER. IMPORTANT: USE PLASTIC WALLET AS STERILE FIELD. PULL CATHETER OUT OF CENTRIFUGE TUBE TO THE PROPER LENGTH IMMEDIATELY AFTER DONNING GLOVES. NOTE: IF URINE DOES NOT FLOW FREELY INTO THE TUBE, THE CAP MAY NEED TO BE LOOSENED SLIGHTLY." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01923 | 2023-03-21 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT 1.5 ML OF THE STERILE WATER WAS UNABLE TO REMOVE ON THE FOLEY CATHETER REMOVAL. THE CATHETER WAS SUCCESSFULLY REMOVED, NO BLEEDING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01921 | 2023-03-21 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED THAT THE PATIENT COMPLAINED THAT THE LUBRICANT WAS WATERY AND LEAKY MORE THAN IT HAD BEEN BEFORE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿FOR UROLOGICAL USE ONLY. INTENDED FOR USE BY PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICES ARE PASSED TO THE URINARY BLADDER VIA THE URETHRA. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. PRIOR TO OPENING THE SEALED CATHETER POUCH, APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. WET THE CATHETER BY HOLDING THE PACKAGE WITH THE PRINTED SIDE UP AND TIP THE PACKAGE END-TO-END THREE TO SIX TIMES TO WET THE CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. PEEL OPEN THE PACK AT THE FUNNEL/HANDLE END JUST ENOUGH TO EXPOSE 5 CM (2¿) OF THE CATHETER, OR FOR PRODUCTS WITH AN INSERTION SLEEVE, EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. USING THE CATHETER . WITH INSERTION SLEEVE (GRIPPER): HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. WITHOUT INSERTION SLEEVE (GRIPPER): HOLD THE HANDLE/FUNNEL END AND REMOVE THE CATHETER FROM THE PACKAGING. GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA AND ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER.¿ DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01920 | 2023-03-21 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT WAS INDWELLED WITH A 22F THREE-LUMEN BALLOON CATHETER AFTER THE OPERATION ON FEBRUARY 24. AFTER THE OPERATION, THEY RETURNED TO THE ROOM AND CONTINUED TO IRRIGATE THE BLADDER, AND THE FLUSHING FLUID WAS CLARIFIED. AT 19:10, IT WAS FOUND THAT THE PATIENT'S IRRIGATION FLUID WAS NOT SMOOTH, AND THE PATIENT'S URINARY CATHETER WAS CHECKED. IT WAS FOUND THAT THE URINARY CATHETER HAD FALLEN OFF AND THE BALLOON WAS RUPTURED. THEN A NEW URINARY CATHETER WAS USED, AND THE BLADDER IRRIGATION CONTINUED. AS PER THE DESCRIPTION OF EVENT CAUSE ANALYSIS, THE RUPTURE OF THE URINARY CATHETER BALLOON MIGHT BE A PRODUCT QUALITY PROBLEM. ALSO STATED THAT THERE WAS AN INCREASED RISK OF URINARY TRACT INFECTION IN PATIENTS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01919 | 2023-03-21 | 2023-03-01 | I, n, j, u, r, y | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MAGIC3 GO MALE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE PATIENT GOT A REALLY BAD URINARY TRACT INFECTION FROM THE MAGIC 3 GO MALE INTERMITTENT CATHETERS AND ALSO STATED THAT THE PATIENT ENDED UP IN THE EMERGENCY ROOM. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " . | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01917 | 2023-03-21 | 2023-02-28 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER COUDE TIP | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0168L14 | 0168L14 |  | IT WAS REPORTED THAT THE CATHETER GOT BROKEN. ABOUT 25CM OF THE DISTAL PART OF THE CATHETER WAS RETRIEVED, BUT THE CATHETER TIP HAD REMAINED IN THE PATIENT BLADDER WHICH WAS RECOVERED LATER. PER ADDITIONAL INFORMATION RECEIVED VIA INVESTIGATORS NOTIFICATION ON 07MAR2023, STATED THAT THE CATHETER FRAGMENT LEFT BEHIND IN THE PATIENT'S BLADDER WAS SPONTANEOUSLY DISCHARGED FROM THE PATIENT BODY ON MARCH 2. THE SALES REPRESENTATIVE RECOVERED IT ON MARCH 7.THE PATIENT WAS MALE. THE PATIENT WAS HOSPITALIZED IN A PSYCHIATRIC WARD AND HE MAY HAVE SELF-EXTRACTED THE CATHETER. PER FOLLOW UP INFORMATION RECEIVED FROM IBC ON 08MAR2023, CATHETER BREAKAGE. | THE REPORTED EVENT IS CONFIRMED, CAUSE UNKNOWN. THE USED TWO WAY FOLEY CATHETER WAS RETURNED WITHOUT THE ORIGINAL PACKAGING. PHOTO SAMPLES WERE SUBMITTED SHOWING THE CATHETER BROKEN INTO TWO SEPARATE PIECES. VISUAL EVALUATION OF THE PHYSICAL SAMPLE NOTED BOTH BROKEN PIECES CONNECTED PERFECTLY WITH EACH OTHER, CONFIRMING NO PIECES MISSING FROM THE CATHETER AS A RESULT OF THE BREAK. AS IT IS UNCLEAR IF THE PATIENT SELF-EXTRACTED THE CATHETER, THIS EVENT WILL BE CONFIRMED CAUSE UNKNOWN. AS THE EVENT STATES THE PATIENT MAY HAVE SELF EXTRACTED THE CATHETER, POTENTIAL ROOT CAUSES FOR THIS EVENT COULD BE, "USER (PATIENT) MEDICATED, DELIRIOUS, CONFUSED, OR RECKLESS" OR "INADEQUATE DESIGN". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.]" "TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00919 | 2023-03-21 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK PICC PLUS CRESCENT TRICOT FIX-POST | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | PIC0222 |  | IT WAS REPORTED THAT THERE WAS A TUG AND A SNAP SOUND WAS HEARD, AND PICC CAME OFF WITH THE STATLOCK WHEN THE CHILD¿S MOTHER WAS CARRYING CHILD SEATED BY THE BEDSIDE WHILE SHE STOOD UP HUGGING THE CHILD. EVEN THOUGH WE ARE CONSIDERING FACTORS LIKE THE ¿TUGGING¿ OF THE LINE DUE TO A CHANGE IN POSITION, WE WOULD LIKE TO CHECK WITH VENDORS ON HOW SECURE THE STATLOCK IS TO WITHSTAND ANY TUGS OR PULLS. PATIENT STATUS/ INTERVENTION: PATIENT WAS INSERTED WITH A NEW PICC AND STATLOCK. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01909 | 2023-03-21 | 2023-02-28 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETER SUCTIONED FLESH OFF THE PATIENT'S LABIA AND IT CAUSED BLEEDING AND SORES. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE COULD BE DUE TO "DIMENSIONS NOT SPECIFIED CORRECTLY". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "WARNINGS: TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS." "RECOMMENDATIONS: PRIOR TO CONNECTING THE PUREWICKTM FEMALE EXTERNAL CATHETER TO HOSPITAL WALL SUCTION TUBING, VERIFY SUCTION FUNCTION BY COVERING THE OPEN END OF THE SUCTION TUBING WITH ONE HAND AND OBSERVING THE PRESSURE DIAL. IF THE PRESSURE DOES NOT INCREASE WHEN THE LINE IS COVERED, VERIFY THAT THE TUBING IS SECURED, CONNECTED, AND NOT KINKED. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS." "SETUP: CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE DRYDOCTM VACUUM STATION, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE DRYDOCTM VACUUM STATION USER GUIDE FOR FURTHER INFORMATION. USING STANDARD SUCTION TUBING, CONNECT THE PUREWICKTM FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER." BAW2456229 REV 0 WAS REVIEWED FOR THIS INVESTIGATION AND IT STATES: "SETUP: CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE PUREWICK¿ URINE COLLECTION SYSTEM, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE PUREWICK¿ URINE COLLECTION SYSTEM USER GUIDE FOR FURTHER INFORMATION. USING STANDARD SUCTION TUBING, CONNECT THE PUREWICK¿ FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER. WARNINGS: TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICK¿ FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. RECOMMENDATIONS: PRIOR TO CONNECTING THE PUREWICK¿ FEMALE EXTERNAL CATHETER TO HOSPITAL WALL SUCTION TUBING, VERIFY SUCTION FUNCTION BY COVERING THE OPEN END OF THE SUCTION TUBING WITH ONE HAND AND OBSERVING THE PRESSURE DIAL. IF THE PRESSURE DOES NOT INCREASE WHEN THE LINE IS COVERED, VERIFY THAT THE TUBING IS SECURED, CONNECTED, AND NOT KINKED. ENSURE THE PUREWICK¿ FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICK¿ FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICK¿ FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICK¿ FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICK¿ FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-01910 | 2023-03-21 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE 5 HOURS PATIENT WAS INTO REWARMING AND THE ARCTIC SUN DEVICE ALERTED ABOUT TEMPERATURE BEING UNSTABLE AND THERAPY STOPPED. THEY DISCONNECTED AND RECONNECTED CABLES AND PADS. RESTARTED THERAPY AND IT APPEARS TO BE NORMAL NOW. NURSE NOTED THAT THEY ALSO TURNED OFF WARMER TO SEE IF THAT WAS CREATING AN ISSUE. NURSE NOTED THAT AXILLARY SHOWS PATIENT STABLE NOW AT 97.4F BUT TEMPERATURE JUMPED FROM 93.9F TO 96.4F FROM 19:00-20:00. THE PATIENT TEMPERATURE WAS 35.7C, TARGET TEMPERATURE WAS 36.5C, WATER TEMPERATURE WAS 33.1 AND WATER FLOW RATE WAS 0.9 L/M. NEONATAL PAD IN USE. TREND SHOWS 1 BAR TRENDING DOWNWARD. THEY WERE USING SKIN PROBE 35.7C TO RUN THERAPY AND SECONDARY PROBE IS ESOPHAGEAL PROBE SHOWED 33.9C. NOTED THAT NURSE ENCOURAGE USE OF THE ESOPHAGEAL PROBE OR ANOTHER TRUE CORE TEMPERATURE PROBE OVER A SKIN PROBE. SKIN PROBE WAS REGISTERING SURFACE TEMPERATURE WHICH SHOWS A PRETTY SIGNIFICANT GAP BETWEEN THE SKIN AND THE ESOPHAGEAL READING. MIS EXPLAINED RADIANT WARMER OF WARM BLANKETS COULD AFFECT THE TEMPERATURE READING AND THERAPY SINCE THEY WERE UTILIZING A SKIN PROBE. MIS ENCOURAGED NURSE TO DISCUSS RUNNING THERAPY OFF OF ESOPHAGEAL PROBE AND TURNING RADIANT WARMER BACK ON. NURSE NOTED RADIANT WARMER WOULD ASSIST AND WOULD NOT AFFECT ACCURATE TEMPERATURE READING IF THEY WERE USING A CORE TEMPERATURE PROBE. NURSE VERIFIED THEY DO HAVE BABY "TACOED" IN WITH PADS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "SENSOR WIRE TOO WEAK / THERMISTOR WIRE TOO WEAK". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01894 | 2023-03-21 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT THE CUSTOMER HAD AN ISSUE WITH THE 16FR FOLEY KIT. THE ISSUE OCCURRED WHEN THEY TRIED TO REMOVE THE CATHETER, THEY DEFLATED THE BALLOON AND IT APPEARED THE CATHETER ROLLED UP UPON ITSELF. ONE PATIENT WAS MALE AND THEY BROUGHT IN A PHYSICIAN TO TRY AND REMOVE THE CATHETER. AFTER MUCH TIME THEY CUT THE CATHETER AT THE Y WHERE THE INFLATION PORT RESIDED AND THEY WERE ABLE TO REMOVE. THE SECOND ONE WAS A FEMALE BUT SINCE THEY EXPERIENCED THE FIRST ONE, THEY DID THE SAME AND IT CAME OUT. IT WAS NOTED THAT THEIR TEAM AT MH USED THE 14FR KIT AND HAD NO ISSUES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01896 | 2023-03-21 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT THE CUSTOMER HAD AN ISSUE WITH THE 16FR FOLEY KIT. THE ISSUE OCCURRED WHEN THEY TRIED TO REMOVE THE CATHETER THEY DEFLATED THE BALLOON AND IT APPEARED THE CATHETER ROLLED UP UPON ITSELF. ONE PATIENT WAS MALE AND THEY BROUGHT IN A PHYSICIAN TO TRY AND REMOVE THE CATHETER. AFTER MUCH TIME THEY CUT THE CATHETER AT THE Y WHERE THE INFLATION PORT RESIDED AND THEY WERE ABLE TO REMOVE. THE SECOND ONE WAS A FEMALE BUT SINCE THEY EXPERIENCED THE FIRST ONE, THEY DID THE SAME AND IT CAME OUT. IT WAS NOTED THAT THEIR TEAM AT MH USED THE 14FR KIT AND HAD NO ISSUES. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. THE DEVICE HAD NOT MET SPECIFICATIONS. THE PRODUCT WAS USED FOR PATIENT TREATMENT. THE PRODUCT CAUSED THE REPORTED FAILURE. VISUAL: RECEIVED 3 PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWS OUTER PACKAGING OF CATHETER. SECOND AND THIRD PHOTO SAMPLES SHOW END OF CATHETER WITH BALLOON CUFFING AROUND THE END OF CATHETER. THEREFORE, PRODUCT DOES NOT MEET SPECIFICATIONS WHICH STATES "BALLOON MUST NOT CUFF AFTER DEFLATION." A POTENTIAL ROOT CAUSE COULD BE BALLOON MATERIAL DOES NOT SHRINK QUICKLY ENOUGH. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INVESTIGATION IS CONCLUDED, AND NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED DIRECTIONS FOR USE WASH HANDS AND DON CLEAN GLOVES EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERIURETHRAL AREA 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL 5. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP 6. DON STERILE GLOVES 7. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE 8. POSITION FENESTRATED DRAPE ON PATIENT 9. SATURATE 3 FOAM SWAB STICKS IN POVIDONE IODINE 10.ATTACH THE WATER FILLED SYRINGE TO THE INFLATION PORT NOTE: IT IS NOT NECESSARY TO PRE-TEST THE FOLEY CATHETER BALLOON 11. REMOVE FOLEY CATHETER FROM WRAP AND LUBRICATE CATHETER 12. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA INORA MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD 13. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND A. WHEN CATHETER TIP HAS NTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE B. INSERT CATHETER TWO MORE INCHES AND INFLATE CATHETER BALLOON 14. INFLATE CATHETER BALLOON USING ENTIRE 10CC OF STERILE WATER PROVIDED IN THE PREFILLED SYRINGE NOTE: USE OF LESS THAN 10CC CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON 15. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK 16. SECURE THE FOLEY CATHETER TO THE PATIENT USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED (SEE STATLOCK® FOLEY STABILIZATION DEVICE IFU) NOTE: PLEASE MAKE SURE PATIENT IS APPROPRIATE FOR USE OF STATLOCK® STABILIZATION DEVICE 17. POSITION HANGER ON BED RAIL AT THE FOOT OF THE BED NOTE: EXERCISE CARE TO KEEP BAG OFF THE FLOOR 18. USE GREEN SHEETING CLIP TO SECURE DRAINAGE TUBE TO THE SHEET. MAKE SURE TUBE IS NOT KINKED 19. INDICATE TIME AND DATE OF CATHETER INSERTION ON PROVIDED LABELS. PLACE DESIGNATED LABELS ON PATIENT CHART AND DRAINAGE SYSTEM 20. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL FOLEY CATHETER REMOVAL 1. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE 2. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY 3. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION 4. IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL 5. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT¿. CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01898 | 2023-03-21 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT UNABLE TO DEFLATE NATURALLY DURING PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01900 | 2023-03-21 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS UNABLE TO DEFLATE NATURALLY DURING PRETEST. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT S POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE COLLAPSED LUMEN AND BLOCKED LUMEN. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. ATTEMPTED TO DEFLATE BALLOON PASSIVELY AND SOLUTION COULD NOT BE WITHDRAWN. SOLUTION WAS WITHDRAWN BY ASPIRATION SUCCESSFULLY. DISSECTED BALLOON AND FOUND THAT THE NOTCH WAS PERFORATED COMPLETELY. DISSECTED INFLATION FUNNEL AND PIN GAUGES 0.039¿ AND BELOW COULD FREELY PASS THROUGH LUMEN. 0.040¿ - 0.045¿ TOOK SOME EFFORT TO GET THROUGH AND 0.046¿ AND ABOVE CANNOT PASS THROUGH. BASED ON PHYSICAL SAMPLE RECEIVED THE PRODUCT DOES NOT MEET WHICH STATES "SHAFT MUST NOT BE DAMAGED OR PINCHED." THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01895 | 2023-03-21 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53812G | 53812G |  | IT WAS REPORTED THAT MULTIPLE CATHETERS WITH BENT TIPS WERE MIXED IN THE 53812G BOX. SO PATIENT BOUGHT ONE ACTUAL PRODUCT TO THE HOSPITAL AND IT WAS UNKNOWN WHETHER IT WAS A DEFECTIVE PRODUCT. | THE REPORTED EVENT IS CONFIRMED - CAUSE UNKNOWN. VISUAL INSPECTION NOTED RECEIVED 1 MAGIC 3 GO IN OPENING PACKAGING. NO SHARPS OR FLASHES WERE PRESENT. CATHETER WAS BENT. THEREFORE, PRODUCT DOES NOT MEET SPECIFICATION. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED, A POTENTIAL ROOT CAUSE COULD BE OPERATOR ERROR. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, USE THE ADHESIVE LABEL(S) TO HANG THE PACKAGE ON A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. CLEAN THE OPENING TO THE URETHRA- AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6" FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01904 | 2023-03-21 | 2023-02-28 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK URINE COLLECTION SYSTEM WAS NOT SUCTIONING, AND PATIENT GOT A RASH AND URINARY TRACT INFECTION . THE PATIENT HAD BEEN USING THIS PRODUCT FOR LESS THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. AS PER CUSTOMER VIA PHONE ON 15MAR2023, IT WAS REPORTED THAT THE CUSTOMER SOUGHT MEDICAL INTERVENTION AND WAS PRESCRIBED ANTIBIOTICS. THE CUSTOMER STATED THE PER HER MEDICAL PROVIDER THE UTI WAS CAUSED FROM USING THE WICKS. CUSTOMER IS NO LONGER USING THE DEVICE AND UTI HAS CLEARED WITH MEDICATION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE OR MATERIAL SURFACE IS ROUGH, ABRASIVE OR UNCOMFORTABLE ". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS: ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: ¿ AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER ¿ HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE ¿ EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ¿ EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01905 | 2023-03-21 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119216M | 119216M |  | IT WAS REPORTED THAT THE URETHRAL CATHETER WAS REMOVED DUE TO COMPLAINTS OF BLADDER RETENTION. THE BULB WAS DEFLATED WHICH HAD ONLY 3 CC OF A CLEAR YELLOW COLOR FLUID. CATHETER BULB WAS RECHECKED AND RE-INFLATED TO RULE OUT MALFUNCTION AND IT WAS NOTED THAT THE BULB DID HAVE A LEAK. IT WAS ALSO STATED THAT THE PATIENT WAS ABLE TO URINATE ON HER OWN AND WAS NOT RECATHETERIZED. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01906 | 2023-03-21 | 2023-02-28 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT CONTINUES TO GET THE URINARY TRACT INFECTIONS. PATIENT USED THE PUREWICK FEMALE EXTERNAL CATHETERS FOR 4 MONTHS IN THE HOSPITAL AND HAD 2 URINARY TRACT INFECTIONS THERE. ONCE THE PATIENT STARTED USING THE PUREWICK AT HOME AND THE PATIENT HAD 2 MORE. PATIENT'S DOCTOR TOLD THE PATIENT THAT THEY WERE FROM THE WICKS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED AND THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01907 | 2023-03-21 | 2023-02-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT CONTINUED TO GET THE URINARY TRACT INFECTIONS. PATIENT USED THE PUREWICK FEMALE EXTERNAL CATHETERS FOR 4 MONTHS IN THE HOSPITAL AND HAD 2 URINARY TRACT INFECTIONS THERE. ONCE THE PATIENT STARTED USING THE PUREWICK AT HOME AND THE PATIENT HAD 2 MORE URINARY TRACT INFECTIONS. PATIENT'S DOCTOR TOLD THE PATIENT THAT THEY WERE FROM THE WICKS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED AND THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD WAS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01902 | 2023-03-21 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE-SENSING FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER IN THE BALLOON DID NOT DRAIN DURING PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01937 | 2023-03-22 | 2023-03-01 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE TUBING CONNECTION TOO LOOSE AND ELBOW CONNECTOR WAS NOT HOLDING THE TUBING IN PLACE CORRECTLY. THE PATIENT HAD BEEN USING THE PUREWICK PRODUCT FOR LESS THAN 90 DAYS. IT WAS NOTED THAT THE PUREWICK ACCESSORY REPLACEMENT KIT WAS REPLACED ON (B)(6) 2022. PER ADDITIONAL INFORMATION RECEIVED FROM LIBERATOR ON (B)(6) 2023, IT WAS STATED THAT THE PUREWICK URINE COLLECTION SYSTEM NOT WORKING WELL AND GOT "URINARY TRACT INFECTION". IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS: DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS." "PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE". " RECOMMENDATIONS: ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS". THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01949 | 2023-03-22 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT 14 DAYS AFTER PLACEMENT, THE CATHETER WAS SPONTANEOUSLY FELL OFF AND BALLOON DEFLATION WAS OBSERVED. THERE WAS NO BALLOON RUPTURE OR TEAR. IT WAS UNKNOWN IF WATER LEAKED BECAUSE THE EVENT WAS CATHETER FELL OUT AND CUSTOMER DID NOT TRY TO REPRODUCE THE EVENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01939 | 2023-03-22 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN THE FOLEY CATHETER BALLOON WAS PRE-TESTED BEFORE USE ON THE PATIENT, IT WAS DIFFICULT TO INFLATE. NEXT, THE STERILE WATER WAS ALSO DIFFICULT TO REMOVE WHEN ATTEMPTING TO REMOVE IT. THEN, ANOTHER CATHETER WAS USED. PER FOLLOW UP INFORMATION RECEIVED FROM IBC ON 22MAR2023,THE DEVICE WAS NOT USED ON THE PATIENT AND THE EVENT OCCURRED DURING A PRETEST. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED A ALSO RECEIVED 3 PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWS OVERVIEW OF CATHETER. SECOND PHOTO SAMPLE SHOWS SLIGHTLY INFLATED BALLOON. THIRD PHOTO SAMPLE SHOWS DOCUMENT IN JAPANESE. BASED ON PHYSICAL SAMPLE RECEIVED PRODUCT DOES NOT MEET SPECIFICATIONS 3-WAY TEMP SENSING CATHETER WITH CUT PORTION OF INLET TUBING. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE DIFFICULT WITHDRAWAL OF BALLOON THROUGH URINARY TRACT PATIENT IMPACT, URETHRAL TISSUE DAMAGE / IRRITATION OR TEMPORARY BLEEDING. THE PRODUCT WAS USED FOR PATIENT DIAGNOSTIC OR TREATMENT. THE PRODUCT WAS INFLUENCED BY THE REPORTED EVENT. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00952 | 2023-03-22 | 2023-03-01 | I, n, j, u, r, y | Manufacturer report | BARD TRIALYSIS HD CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. THE EVENT INVOLVED A DL POWERTRIALYSIS PLACED IN THE CHEST. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01943 | 2023-03-22 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT INJECTING STERILE WATER DURING A PRETEST TO INFLATE THE BALLOON, THE WATER DID NOT ENTER THE CATHETER AND LEAKED FROM THE INFLATION VALVE. THE ACTUAL SAMPLE WAS DISCARDED BY THE CUSTOMER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00954 | 2023-03-22 | 2023-02-27 | I, n, j, u, r, y | Manufacturer report | POWERPICC® CATHETER, 4F SL, 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000887 |  | IT WAS REPORTED BY THE CUSTOMER "DURING PICC PROCEDURE GUIDEWIRE WAS RETAINED INTO THE PATIENT." NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: APPROXIMATELY 20CM OF WIRE EXTERNAL TO VENIPUNCTURE SITE. WIRE NOT MAINTAINED WHILE PICC INSERTED DILATED, WIRE RETAINED INTO PATIENT. TOURNIQUET PLACED ON AFFECTED ARM, RAPID RESPONSE CALLED, PATIENT HAD RETRIEVAL IN IR THROUGH FEMORAL VEIN. PATIENT STABLE THROUGHOUT. CATHETER WAS NOT INSERTED AS THE PROCEDURE WAS ABORTED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01954 | 2023-03-22 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT'S FOLEY WAS PULLED, WHEN PULLING THE SALINE FROM THE BULB THEY NOTICED IT LOOKED LIKE URINE. ONCE THE FOLEY WAS REMOVED, THEY ATTACHED A NEW 10ML FLUSH TO THE PORT FOR THE BALLOON. THERE APPEARED TO BE A LEAK AT SOMEWHERE WITH THE BALLOON BECAUSE THERE WAS A TINY BIT OF MOISTURE WHERE THEY HAD DRIED THE FOLEY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00963 | 2023-03-22 | 2023-03-04 | I, n, j, u, r, y | Manufacturer report | FULL KIT 20G X 10 CM WITH GUARDIVA AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120108PT |  | BD REPRESENTATIVE PROVIDED THE FOLLOWING INFORMATION: "NOTIFIED THIS MORNING BY THE ACCOUNT THAT A POWERGLIDE CATHETER BROKE OFF IN A PATIENTS ARM. AT THIS TIME I AM AWAITING FOR THE PICC NURSE TO CALL ME TO DESCRIBE THE EVENT. THIS NURSE WAS THE PLACER OF THE DEVICE. WAS PROVIDED A SMALL BIT OF INFORMATION FROM THE NURSE AS SHE STATED THAT AT THIS TIME THE CATHETER IS STILL IN THE PATIENTS ARM AND THE PATIENT IS REFUSING REMOVAL. AN X-RAY WAS CONFIRMED THAT THERE WAS A FOREIGN BODY IN THE LEFT CEPHALIC VEIN. BASED OF THE REMAINING CATHETER THAT WAS REMOVED FROM THE PATIENT, IT APPEARS THAT ABOUT 1.3CM IS REMAINING IN THE PATIENTS UPPER EXTREMITY. " ADDITIONAL INFORMATION RECEIVED 3/8/2023: THE CLINICIAN PROVIDED THE FOLLOWING DETAILS TO THE BD REPRESENTATIVE: RN REPORTED THAT SHE ASSESSED THE L. CEPHALIC VEIN AND A 20G WAS SUITABLE FOR THE PLACEMENT. REPORTED VESSEL WAS COMPRESSIBLE AND ABOUT 1CM DEEP. STATED WHEN SHE LAUNCHED THE GUIDEWIRE THE WIRE WENT IN SMOOTHLY AND THEN ADVANCED HER CATHETER WHICH WENT SMOOTHLY UP UNTIL ABOUT HALFWAY WHEN SHE MET SLIGHT RESISTANCE. REPORTED STOPPING THE PROCEDURE DUE TO THE RESISTANCE AND PULLED THE WIRE BACK AND THEN THE DEVICE WITHOUT DIFFICULTY. STATED THAT AFTER SHE REMOVED THE DEVICE SHE NOTICED THE CATHETER LOOKING SHORTER AND THAT THE PURPLE REINFORCED TIP WAS NOT AT THE TOP OF THE CATHETER. AT THAT TIME SHE ASSESSED THE SITE AND STATED THE PATIENT REPORTED ZERO PAIN. X-RAY WAS THEN ORDERED THAT CONFIRMED FOREIGN BODY IN THE L. CEPHALIC VEIN. PATIENT REFUSED REMOVAL OF FOREIGN BODY. UNKNOWN AT THIS TIME OF THE PATIENT STATUS. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01953 | 2023-03-22 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123614A |  | IT WAS REPORTED THAT CATHETER LEAKING AND DEFLATING DURING PRODUCT USE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00942 | 2023-03-22 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED VIA MS&S "CALLER STATES PATIENT'S POWERPICC GOT PARTIALLY PULLED OUT AND IT IS NO LONGER IN THE SVC. SHE STATES THE PATIENT IS GOING HOME AND THE PICC LINE WAS LEFT IN FOR IV ZOSYN. SHE STATES IT IS POSITIONED SIMILAR TO A MIDLINE AND IS SITTING IN THE "LATERAL BORDER OF THE LEFT SECOND RIB JUNCTION OF THE AXILLARY AND SUBCLAVIAN VEIN. CALLER INFORMS THAT AN X-RAY WAS PERFORMED TO DETERMINE PLACEMENT. CALLER WANTS TO KNOW IF THEY CAN KEEP IT IN THIS AREA AND ADMINISTER ZOSYN. CALLER STATES THE PATIENT IS GOING TO BE DISCHARGED HOME. CALLER STATES SHE DOES NOT HAVE FURTHER DETAILS OF HOW IT BECAME MISPOSITIONED." BD MEDICAL SERVICES EXPLAINED "BD CANNOT RECOMMEND USING A POWERPICC WHERE THE TIP DOES NOT RESIDE IN THE LOWER ONE-THIRD OF THE SVC. POWERPICCS ARE DESIGNED TO DWELL CENTRALLY AND AS CURRENTLY POSITIONED, THIS PICC IS A PERIPHERAL LINE. USING THIS DEVICE IN THIS MANNER WOULD BE CONSIDERED OFF-LABEL USE OF THE PICC." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00957 | 2023-03-22 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7274118 |  | THE VASCULAR ACCESS REPORTS THAT DURING THE CATHETER PLACEMENT PROCEDURE, UPON REMOVING THE STYLET IT'S GETTING STUCK TO THE CATHETER AND DEFORMS THE CATHETER. NO OTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION RECEIVED ON 10.APR.2023: ¿ THE MATERIALS WERE NOT USED ON THE PATIENT/THERE WAS NO IMPACT ON THE PATIENT. __ ADDITIONAL INFORMATION RECEIVED 17.MAR.2023 EVIDENCED DURING DEVICE VERIFICATION; STYLET STICKING TO THE DEVICE AND ITS RESISTANCE DURING PURGE, INSERTION IS ABANDONED AND A NEW ONE IS REQUESTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00964 | 2023-03-22 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT THERE WERE ISSUES WITH THE T-LOCK CONNECTOR ON THE PICC LINE. FLUID LEAKING OUT WHEN FLUSHED AND AIR PULLING IN WHEN ATTEMPTING TO ASPIRATE FOR BLOOD. 03/15/2023 ADDITIONAL INFORMATION PROVIDED THERE WERE NOT ANY PATIENT COMPLICATIONS, NO DELAY OF TREATMENT. PICC LINE INSERTION WAS THE PROCEDURE BEING DONE, LIDOCAINE WAS USED FOR LOCAL ANESTHETIC. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00967 | 2023-03-22 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3175335 |  | IT WAS REPORTED BY THE CUSTOMER THERE WAS LEAKING AT THE HUB CONNECTION BETWEEN PICC AND CONNECTOR. ADDITIONAL INFORMATION RECEIVED 3/15/2023: CUSTOMER REPORTED 2 PICCS WERE USED WITH MINIMAL PROBLEM ASIDE FROM THE LEAKING AT THE HUB THEN "LIKELY REMOVED AFTER TREATMENT WAS COMPLETE." THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00966 | 2023-03-22 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3175335 |  | IT WAS REPORTED BY THE CUSTOMER THERE WAS LEAKING AT THE HUB CONNECTION BETWEEN PICC AND CONNECTOR. ADDITIONAL INFORMATION RECEIVED 3/15/2023: CUSTOMER REPORTED 2 PICCS WERE USED WITH MINIMAL PROBLEM ASIDE FROM THE LEAKING AT THE HUB THEN "LIKELY REMOVED AFTER TREATMENT WAS COMPLETE." THIS REPORT ADDRESSES THE FIRST DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01938 | 2023-03-22 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT COULD NOT GET THE MAGIC 3 GO MALE HYDROPHILIC INTERMITTENT CATHETER OPEN. PATIENT WAS DROPPING THINGS ON THE FLOOR AND SAID THAT THEY COULD NOT EVEN GET THEM OPEN. PATIENT WOULD TAKE THE BARD CATHETER THAT THEY PREFER. THE PATIENT STATED THAT THE BLUE TAB WAS NOT BIG ENOUGH TO BE A GRIP TO HOLD THE CATHETER STURDY. THE PATIENT ONLY HAD ROUNDED COUNTERS, NOT EDGED ONES AND THE ADVISED THAT THE CATHETER WAS A HYDRO CATHETER AND ADVISED THE PATIENT ON USAGE OF THE HYDRO AND ALSO ADVISED THAT THE LIBERATOR WOULD SEND A DIFFERENT CATHETER SINCE THEY COULD NOT OPEN THE ONES THAT THEY HAD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00955 | 2023-03-22 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F DUAL-LUMEN POLYURETHANE CATHETER RADIOLOGY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275135 |  | IT WAS REPORTED BY THE CUSTOMER THAT "DOCTOR INSERTED 5F DOUBLE LUMEN PICC. WE COULD NOT REMOVE THE WIRE. I BELIEVE IT WAS WAS FROM THE PURPLE PORT BUT I AM NOT 100% CERTAIN ANYMORE." NO PATIENT OR USER IMPACT NOTED. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00951 | 2023-03-22 | 2023-02-22 | I, n, j, u, r, y | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT SEVEN MONTHS AND TEN DAYS POST PORT PLACEMENT PROCEDURE VIA THE SUBCLAVIAN VEIN FOR CHEMOTHERAPY, THE BLOOD RETURN COULD NOT BE CONFIRMED DURING USING THE PORT SYSTEM FOR CHEMOTHERAPY. IT WAS FURTHER REPORTED THAT SUBCUTANEOUS LEAK OF THE ANTICANCER AGENT OCCURRED. FURTHERMORE, AN X-RAY EXAMINATION REVEALED THAT THE CATHETER WAS BROKEN AND THE DISTAL CATHETER SEGMENT MIGRATED INTO THE VENTRICLE. THE PATIENT FELT DISCOMFORT AND PAIN, AND SWELLING DUE TO SUBCUTANEOUS LEAK OF THE ANTICANCER AGENT. REPORTEDLY, THE PORT BODY AND THE DISTAL CATHETER SEGMENT WAS REMOVED FROM THE PATIENT AND ANOTHER DEVICE WAS PLACED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE X-PORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WERE RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE, AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION AND IDENTIFIED DENT ISSUE AS A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED ON THE DISTAL END OF THE CATHETER ATTACHED THAT WAS ELLIPTICAL IN SHAPE. THE SURFACE WAS NOTED TO BE GRANULAR WITH RESIDUE THROUGHOUT. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MIGRATION, FLUID LEAK AND SUCTION ISSUE, AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED, AND THE REPORTED EVENT COULD NOT BE REPRODUCED IN THE LAB. THE CHARACTERISTICS OF THE BREAK ARE INDICATIVE OF PINCH-OFF. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01925 | 2023-03-22 | 2022-12-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI12 | 0165SI12 |  | IT WAS REPORTED THAT STERILE WATER LEAKED DURING PRETEST. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS EVALUATED AND OBSERVED THERE WAS A PINHOLE AT THE BOTTOM PART OF SAC COLLAR THAT ALLOWED WATER TO LEAK OUT. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE NON-UNIFORM LATEX FILLING AT CATHETER SAC COLLAR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE(S) THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS. PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER. WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. CORRECTION: D, F, H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01932 | 2023-03-22 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT URINE LEAKED FROM NEAR THE JOINT BETWEEN THE FUNNEL AND THE SHAFT OF THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00931 | 2023-03-22 | 2022-10-24 | D, e, a, t, h | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED POST PORT PLACEMENT, THE HEALTHCARE PROFESSIONAL ALLEGEDLY WAS UNABLE TO OBTAIN A BLOOD RETURN. IT WAS FURTHER REPORTED THAT DURING FLUSHING, SWELLING WAS ALLEGEDLY IDENTIFIED TWO CENTIMETERS OVER THE PORT SITE. REPORTEDLY, THE HEALTHCARE PROFESSIONAL SCHEDULED A PORT REMOVAL PROCEDURE. THE PATIENT REPORTEDLY EXPIRED PRIOR TO THE PORT REMOVAL PROCEDURE. THE PATIENTS CAUSE OF DEATH IS UNKNOWN. | DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO AND IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00937 | 2023-03-22 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT APPROXIMATELY FIVE YEARS POST PORT PLACEMENT PROCEDURE, THE BLOOD ALLEGEDLY BACKED UP WHEN ANTIBIOTIC WAS ADMINISTERED. IT WAS FURTHER REPORTED THAT THE DEVICE ALLEGEDLY COAGULATED AND THEREFORE DIFFICULT TO BE FLUSHED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00928 | 2023-03-22 | 2022-04-13 | I, n, j, u, r, y | Manufacturer report | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602680 | 0602680 |  | IT WAS REPORTED THAT APPROXIMATELY ONE MONTH POST PORT PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY FRACTURED WAS FOUND UPON TOMOGRAPHY EXAMINATION. IT WAS FURTHER REPORTED THAT FRACTURED CATHETER ALLEGEDLY MIGRATED TO THE LATERAL SEGMENT OF THE LEFT LOBE OF THE LIVER. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 12/2024). DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00933 | 2023-03-22 | 2023-02-22 | D, e, a, t, h | Manufacturer report | PRISTINE | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PRISTINE |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY DISSECTED INTO THE ILIAC VEIN, CAUSING THE PATIENT TO BLEED TO DEATH. REPORTEDLY, THE HEALTHCARE PROFESSIONAL WAS UNABLE TO DETERMINE IF THE CAUSE OF DEATH WAS FROM THE CATHETER OR HOW THE CATHETER WAS PLACED. PATIENT REPORTEDLY EXPIRED. | H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT PERFORMED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS OR IMAGES WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED ILIAC VEIN DISSECTION, BLEEDING, AND PATIENT DEATH COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3, H6 (METHOD) H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01933 | 2023-03-22 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT ON THE NEXT DAY AFTER USE, URINE LEAKED FROM THE FOLEY CATHETER. ALSO STATED THAT THE DETAIL LOCATION OF THE URINE LEAKAGE WAS UNKNOWN. PER NOTIFICATION RECEIVED FROM INVESTIGATOR ON (B)(6)2023, IT WAS REPORTED THAT THE INVESTIGATION PERTAINING TO THE TEMPERATURE WIRE UNWINDING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00940 | 2023-03-22 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT STC PPICC, 5FR TL, | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1385118Q |  | CUSTOMER REPORTED THAT THE GREY SECTION OF THE INTRODUCER WAS COILING BACK WHEN ATTEMPTING TO INSERT. SKIN DERMATOTOMY PERFORMED X 3 WITHOUT SUCCESS DUE TO COILING. ALTERNATE MICROINTRODUCER UTILIZED. THIS CAUSED PATIENT DISCOMFORT AND A PROLONGED PROCEDURE. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00939 | 2023-03-22 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT STC PPICC, 5FR TL, | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1385118Q |  | CUSTOMER REPORTED THAT THE GREY SECTION OF THE INTRODUCER WAS COILING BACK WHEN ATTEMPTING TO INSERT. SKIN DERMATOTOMY PERFORMED X 3 WITHOUT SUCCESS DUE TO COILING. ALTERNATE MICROINTRODUCER UTILIZED. THIS CAUSED PATIENT DISCOMFORT AND A PROLONGED PROCEDURE. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00938 | 2023-03-22 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT STC PPICC, 5FR TL, | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1385118Q |  | CUSTOMER REPORTED THAT THE GREY SECTION OF THE INTRODUCER WAS COILING BACK WHEN ATTEMPTING TO INSERT. SKIN DERMATOTOMY PERFORMED X 3 WITHOUT SUCCESS DUE TO COILING. ALTERNATE MICROINTRODUCER UTILIZED. THIS CAUSED PATIENT DISCOMFORT AND A PROLONGED PROCEDURE. IT WAS REPORTED THIS OCCURRED WITH THREE DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00947 | 2023-03-22 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC, IR, 135 CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174335J |  | IT WAS REPORTED THAT THE PATIENT HAD A POWERPICC IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, AIR LEAKAGE WAS SUSPECTED. THE DEVICE WAS REMOVED FROM THE PATIENT¿S BODY. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE, COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAK IN THE CATHETER WAS UNCONFIRMED AS THE COMPLAINT COULD NOT BE REPRODUCED. THE PRODUCT RETURNED FOR EVALUATION WAS A 4FR SL POWERPICC CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A LEAK WAS NOT OBSERVED. NO LEAKS WERE OBSERVED DURING PRESSURIZATION OF THE SAMPLE. THE INVESTIGATION FINDINGS DID NOT PRODUCE A RESULT BASED ON THE DESCRIPTION OF THE EVENT, THEREFORE BOTH CONFIRMATION OF THE REPORTED EVENT AND IDENTIFICATION OF A ROOT CAUSE(S) WERE NOT ESTABLISHED. CONSEQUENTLY, THIS COMPLAINT IS UNCONFIRMED AT THIS TIME. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00944 | 2023-03-22 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THE CUSTOMER NOTICED SOME OF THE WIRES INSIDE THE PICCS HAVE BEEN KINKED AND DIFFICULT TO RE-ADVANCE THE WIRE AFTER TRIMMING." ADDITIONAL INFORMATION RECEIVED 03/15/2023: "KINK NOTED IN PICC WIRE AT THE END OF THE PROCEDURE WHEN WIRE WAS REMOVED. SOME DIFFICULTY WAS NOTED WHEN RE-ADVANCING THE WIRE AFTER TRIMMING." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00946 | 2023-03-22 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2295108 |  | IT WAS REPORTED BY THE CUSTOMER THAT "CUSTOMER NOTICED SOME OF THE WIRES INSIDE THE PICCS HAVE BEEN KINKED AND DIFFICULT TO RE-ADVANCE THE WIRE AFTER TRIMMING." THERE AS NO PATIENT HARM REPORTED PATIENT IS BEING TREATED FOR RT BREAST CA . CONCOMITANT THERAPY IS UNKNOWN. PATIENT DOES NOT HAVE AN INFECTIONS DISEASE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00950 | 2023-03-22 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT STC PPICC, 5FR TL, | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1385118Q |  | IT WAS REPORTED BY THE CUSTOMER THAT "THERE WAS A KINK REPORTED IN A TRIPLE LUMEN POWERPICC BY THE VAT TEAM. THE KINK OCCURRED AT 6CM MARK UNDER THE PATIENT'S SKIN. PICC WAS PLACED IN PATIENT ON (B)(6) 2023 AND REMOVED ON (B)(6) 2023 DUE TO KINK. VAT TEAM HAD TO REPLACE LINE IN OPPOSITE ARM, WHICH LED TO EXTRA UNNECESSARY POKE AND PAIN TO THE PATIENT. STAFF ATTEMPTED TPA X 2 ON (B)(6) AND (B)(6) PRIOR TO PERFORMING X-RAY TO UNCOVER PICC WAS KINKED. UNNECESSARY STICKS TO PATIENT (REPLACING LINE), UNNECESSARY MEDICATIONS GIVEN, SECOND INVASIVE PROCEDURE, PATIENT ANXIETY AND WORRY REGARDING SOMETHING HAPPENING TO NEW LINE." ADDITIONAL INFORMATION RECEIVED ON 5/23/2023: CATHETER PLACED RIGHT UPPER BASILIC VEIN. CATHETER TRIMMED TO 42 CM WITH 3CM LEFT EXTERNAL. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00948 | 2023-03-22 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD STATLOCK | DEVICE, INTRAVASCULAR, CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | CLINICIAN REPORTED SHE HAS NOTICED RECENTLY THAT A FEW PATIENTS HAD A PICC THAT PARTIALLY CAME OUT AFTER THE NURSE HAD CHANGED THE BANDAGE AT THE PATIENTS' HOME. THE CUSTOMER STATED "SOMETIMES NURSES WHO COME TO THE HOME ARE NOT VERY EXPERIENCED IN TAKING CARE OF PICC'S AND THAT IS ONE OF THE REASONS THAT IT SOMETIMES GOES WRONG. THOSE LITTLE HOLES IN WHICH YOU HAVE TO PUT THE BLUE PINS, THEY ARE REALLY SMALL. YOU HAVE TO PUT IT TOGETHER WEARING GLOVES, THAT'S NOT EASY! YOU ALSO HAVE TO PUSH VERY FIRMLY TO CLOSE THE STATLOCK AND IT IS HARD NOT TO MOVE THE CATHETER WHILE DOING SO." THE CUSTOMER REPORTED THIS OCCURRED WITH A FEW PATIENTS HOWEVER THE EXACT QUANTITY INVOLVED WAS NOT PROVIDED TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01927 | 2023-03-22 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | LATEX TEMP-SENSING CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT NURSE TRYING TO START THERAPY ON THE ARCTIC SUN DEVICE AND GETTING AN ALARM 14 (PATIENT TEMPERATURE 1 PROBE OUT OF RANGE). NURSE CONFIRMED THE PROBE GOT A READING ON THE BEDSIDE MONITOR. NURSE WAS ABLE TO REMOVE THE CABLE FROM OUTLET CONTROL TEMPERATURE (T2), MOVE TO OUTLET MONITOR TEMPERATURE (T1) AND START THERAPY. PER FOLLOW UP INFORMATION RECEIVED VIA PHONE ON 10MAR2023, IT WAS REPORTED THAT THERE WAS NO IMPACT TO THE MALE PATIENT AND THERAPY WAS COMPLETED. MIS DID TROUBLESHOOTING AND ADVISED NURSE TO SWITCH THE TEMPERATURE PROBE. THIS RESOLVED THE ISSUE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "SENSOR WIRE TOO WEAK / THERMISTOR WIRE TOO WEAK". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW THAT CANNOT BE COMPLETED DUE TO NO LOT NUMBER. THE PRODUCT CATALOG NUMBER AND LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01929 | 2023-03-22 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT STERILIZED WATER DID NOT ENTER AND LEAKED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00972 | 2023-03-23 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT MRI ISP, 8FR. GROSHONG, INT W SP, ATTACHABLE SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808560 |  | IT WAS REPORTED THAT DURING A PORT REMOVAL PROCEDURE, UPON X-RAY EXAMINATION THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN. IT WAS FURTHER REPORTED THAT THE DEVICE WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE SAMPLE WAS NOT RETURNED FOR EVALUATION. THREE ELECTRONIC PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, AS THE CATHETER WAS NOTED TO BE FRACTURED INTO THREE SEGMENTS. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00976 | 2023-03-23 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT MRI ISP, IMPLANTABLE PORT, GROSHONG SINGLE LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT ONE YEAR ONE MONTHS AND THIRTEEN DAYS POST PORT PLACEMENT VIA RIGHT INTERNAL JUGULAR VEIN, THE CATHETER ALLEGEDLY FRACTURED AT APPROXIMATELY 4 CM ABOVE PORT CONNECTION. IT WAS FURTHER REPORTED THAT PATIENT EXPERIENCED TENDERNESS OVER RIGHT NECK FEW DAYS AFTER LAST CYCLE CHEMO AND HYPERPIGMENTATION WAS NOTED ALONG THE CATHETER. REPORTEDLY, SLIGHT PALPABLE PUFFINESS WAS NOTED ADJACENT TO LINE AROUND RIGHT CLAVICLE AFTER 250ML NS INFUSION. THE BROKEN CATHETER WAS REMOVED AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS AND IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 02/2023). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00187 | 2023-03-23 | 2023-02-23 | I, n, j, u, r, y | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80126 | ATG80126 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, BALLOON ALLEGEDLY GOT STUCK IN THE PATIENT. IT WAS FURTHER REPORTED THAT BALLOON ALLEGEDLY DISCONNECTED FROM THE CATHETER. SINCE IT WAS ATTACHED TO THE FISTULA THE CATHETER WAS CUT TO PULL IT THRU. REPORTEDLY THE BALLOON WAS RUPTURED AND CATHETER WAS STILL ON THE WIRE. DETACHED SEGMENT WAS REMOVED USING SNARE. FEMORAL ACCESS WAS GAINED TO PULL THE BALLOON AND CATHETER. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ATLAS GOLD PTA DILATATION CATHETER DETACHED IN THREE SEGMENTS WAS RECEIVED FOR EVALUATION. SEGMENT ONE CONSISTED OF THE Y-BODY AND STRAIN RELIEF CONNECTED TO A 3-WAY STOP COCK. SEGMENT TWO CONSISTED OF THE OUTER CATHETER WITH A COMPLETE CIRCUMFERENTIAL BREAK. SEGMENT THREE CONSISTED OF THE INNER GUIDEWIRE LUMEN STILL ATTACHED TO THE BALLOON WITH THE PROXIMAL END FRAYED. NO FUNCTIONAL TESTING WAS PERFORMED DUE TO THE CONDITION OF THE SAMPLE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED DETACHMENT AS THE BALLOON WAS RETURNED DETACHED IN THREE SEGMENTS. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED BALLOON RUPTURE AND DEFLATION PROBLEM AS NO FUNCTIONAL TESTING COULD BE PERFORMED DUE TO THE CONDITION OF THE SAMPLE. THE INVESTIGATION IS ALSO INCONCLUSIVE FOR THE REPORTED REMOVAL DIFFICULTY AS THE CONDITIONS FOR USE COULD NOT BE REPLICATED IN THE LABORATORY. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED BALLOON RUPTURE, DETACHMENT, REMOVAL DIFFICULTY AND DEFLATION PROBLEM COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01985 | 2023-03-23 | 2023-02-27 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT AT HOME, THERE WAS NO OUTFLOW OF URINE FROM THE CATHETER, SO WHEN IT WAS ATTEMPTED TO REMOVE, IT COULD NOT BE REMOVED, SO AT KOKURA MEMORIAL HOSPITAL, THE PENIS WAS PUNCTURED AND THE BALLOON PORTION WAS RUPTURED AND REMOVED. THE BALLOON WAS TRIED TO BE DEFLATED BY PERCUTANEOUS BALLOON RUPTURE. THE POSSIBILITY OF INDWELLING IN THE URETHRA WAS SUSPECTED BECAUSE NO URINE OUTFLOW HAS BEEN CONFIRMED AND THE PENIS HAS BEEN PUNCTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00975 | 2023-03-23 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE TIP SIDE OF THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00186 | 2023-03-23 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG120164 | ATG120164 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE FOR OCCLUSION OF THE LEFT BRACHIOCEPHALIC VEIN, THE BALLOON ALLEGEDLY BURST. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 06/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00974 | 2023-03-23 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT DURING PREPARATION OF A PORT PLACEMENT, WHEN PRIMING OF THE PORT AND PUNCTURING THE SEPTUM WITH THE NON-CORING NEEDLE, THE SEPTUM WAS ALLEGEDLY CRACKED. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONTACT. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE X-PORT IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC VISUAL AND FUNCTIONAL EVALUATIONS WERE PERFORMED ON THE RETURNED DEVICE. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED AT THE DISTAL END OF THE CATHETER ATTACHED AND THE EDGES OF THE COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED TO BE EVEN. THE SURFACE WAS NOTED TO BE GLOSSY WITH STRIATIONS THROUGHOUT. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED MATERIAL SPLIT ISSUE AS A SPLIT WAS NOTED ON THE PORT SEPTUM ALONG WITH MINOR ACCESS POINTS. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 01/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00978 | 2023-03-23 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D2 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN SETTING UP FOR PICC PLACEMENT, CLINICIAN NOTED THAT RUBBER STOPPER THAT 3CG GUIDEWIRE GOES THROUGH WAS NOTICEABLY SHORTER (SMALLER) THAN WHAT WAS PREVIOUSLY SEEN IN THE KIT. PICC WAS TRIMMED AND FLUSHED, AND CLINICIAN NOTED A DROP OR TWO OF SALINE COME OUT THROUGH STOPPER. THEY DIDN'T THINK IT WAS TOO BAD, AND PROCEEDED WITH INSERTING PICC LIKE NORMAL. WHILE INSERTING PICC, THEY FLUSHED A FEW TIMES WHILE ADVANCING AND NOTED DROPS OF SALINE COMING OUT THROUGH STOPPER, WHEN NOTED PICC WAS IN POSITION PER SHERLOCK, ASPIRATED TO ENSURE GOOD BLOOD FLOW AND NOTED AIR WAS COMING IN THROUGH RUBBER STOPPER. CLINICIAN STOPPED, AND ENSURED T-CONNECTION WAS TIGHT AND ATTEMPTED AGAIN, NOTICING SAME ISSUE. THEY DID NOT ATTEMPT TO FLUSH AS A "BUNCH OF AIR" HAD BEEN INTRODUCED, REMOVED T-CONNECTOR AND GUIDEWIRE AND WAS ABLE TO ASPIRATE AND FLUSH, THEN PLACED A NEEDLELESS CONNECTOR ONTO PORT WITHOUT FURTHER ISSUES. CUSTOMER REPORTS THE PICC WAS OTHERWISE INTACT, SECURED WITH STATLOCK AND A SECUREPORT IV DRESSING. PATIENT HAD PICC INSERTED FOR LONG-TERM ANTIBIOTICS AND SENT HOME. CUSTOMER STATES NO INJURY OCCURRED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01970 | 2023-03-23 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® URETHRAL CATHETER TRAY, RED RUBBER CATHETER | URETHRAL CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | 802115 | 802115 |  | IT WAS REPORTED THAT "FOLEY" TRAY WAS MISSING BETADINE SWABS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00968 | 2023-03-23 | 2022-12-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | UNK EQUISTREAM SPLIT TIP |  | IT WAS REPORTED THAT APPROXIMATELY SEVEN MONTHS POST DIALYSIS CATHETER PLACEMENT, CATHETER ALLEGEDLY HAD A HOLE IN THE ARTERIAL PATH NEXT TO THE CATHETER'S WHITE PART. IT WAS FURTHER REPORTED THAT CATHETER WAS EXPLANTED. REPORTEDLY ANOTHER CATHETER WAS SCHEDULED TO BE PLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01935 | 2023-03-23 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | URETHRAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 4A6142 | 4A5144 |  | IT WAS REPORTED THAT THE PATIENT AFTER TRYING 6 TOUCHLESS PLUS UNISEX PRE-LUBRICATED URETHRAL CATHETER KITS SAID IT DID NOT HAVE A HOLE IN THE TUBE THAT ALLOWED IT TO DRAIN. THEY HAD 100 AND USED 6. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-00973 | 2023-03-23 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT FOUR MONTHS AND TWENTY-FOUR DAYS POST PORT PLACEMENT, THE PORT ALLEGEDLY COULDN¿T GET BLOOD REFLEX AND LEAKED NEAR HEART AREA. IT WAS FURTHER REPORTED THAT, THE DEVICE WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE POWERPORT IMPLANTABLE PORT ATTACHED TO CATHETER WAS RECEIVED FOR EVALUATION AND ONE X-RAY IMAGE WAS PROVIDED FOR REVIEW. VISUAL, MICROSCOPIC, TACTILE EVALUATION AND FUNCTIONAL TESTS WERE PERFORMED. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED LEAK ISSUE AS THE CONTRAST WAS NOTED TO BE LEAKING PRIOR TO THE TIP. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED SUCTION ISSUE AS THE EXACT CIRCUMSTANCE AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00977 | 2023-03-23 | 2023-01-11 | I, n, j, u, r, y | Manufacturer report | POWERPORT MRI, 9.6 FR. SILICONE, INTER. KIT W/SUTURE PLUG, S/L | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1859600 | 1859600 |  | IT WAS REPORTED THAT TWO MONTHS AND FIFTEEN DAYS POST PORT PLACEMENT, ONLY A PORTION OF THE PORT ALLEGEDLY CAME OUT DURING THE PORT REMOVAL. IT WAS FURTHER REPORTED THAT THE REMAINING REMNANT OF THE CATHETER WAS SEVERED INTERNALLY AND LEFT IN PATIENT. REPORTEDLY, PATIENT HAD TO BE SCHEDULED FOR SURGERY FOR A FOREIGN BODY REMOVAL. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURES AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED ISSUE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01974 | 2023-03-23 | 2023-02-28 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT EXPERIENCED URINARY TRACT INFECTIONS (UTIS) WHILE USING THE PUREWICK FEMALE EXTERNAL CATHETER. THE HOSPITAL NURSE ADVISED THE PATIENT TO USE A REGULAR CATHETER INSTEAD. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. PER FOLLOW UP VIA PHONE ON 15-MAR-2023, THE PATIENT¿S DAUGHTER REPORTED THAT THE PATIENT HAD PASSED AWAY ON (B)(6) 2023. NO ADDITIONAL INFORMATION WAS PROVIDED. PER CLINICAL FOLLOW UP VIA PHONE ON 20-MAR-2023, THE PATIENT¿S DAUGHTER STATED THAT OVER TIME THE PUREWICK DEVICE GOT LOUDER, AND THERE SEEMED TO BE LESS SUCTION. A REPLACEMENT DEVICE WAS OBTAINED. SHE CONFIRMED THE PATIENT HAD DEVELOPED A UTI, AND IT WAS TREATED WITH ANTIBIOTICS. THE PATIENT¿S DAUGHTER REPORTED THAT THE 89-YEAR-OLD FEMALE PATIENT WAS BEDRIDDEN AND PASSED AWAY ON (B)(6) 2023 DUE TO CONGESTIVE HEART FAILURE (CHF). SHE ADDITIONALLY STATED THE PATIENT¿S DEATH WAS UNRELATED TO THE PUREWICK DEVICE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-01967 | 2023-03-23 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THERE WAS A SMALL CRACK AT THE TIP OF THE TEMPERATURE SENSOR CORD AT THE TIP OF THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01963 | 2023-03-23 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175814 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED INTO A PATIENT AND THEN FELL OUT AFTER THE BALLOON WAS INFLATED WITHIN 1 MINUTE OF INSERTION. UPON EXAMINATION, A SMALL HOLE WAS DISCOVERED IN THE CATHETER BALLOON. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 02MAR2023, IT WAS STATED THAT NO HARM OCCURRED TO THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01973 | 2023-03-23 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE DRAINAGE LUMEN. DURING PRETEST TO CHECK THE BALLOON, UPON INJECTING SALINE SOLUTION TO THE CATHETER, THE BALLOON WAS INFLATED, AND AT THE SAME TIME, WATER LEAKING FROM THE DRAINAGE LUMEN WAS OBSERVED. PER FOLLOW UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, THERE WERE NO OBVIOUS VISIBLE DEFECTS TO THE RETURNED SAMPLE. IT WAS WATER LEAK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00992 | 2023-03-24 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD MIDLINE CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT: FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. EVENT INVOLVED A 20G MIDLINE PLACED IN THE UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02020 | 2023-03-24 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165814 | 165814 |  | IT WAS REPORTED THAT THE SILICONE CATHETER LEAKED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02019 | 2023-03-24 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 73022L | 73022L |  | IT WAS REPORTED THAT THERE WAS A HOLE JUST DISTAL TO THE THREE-WAY PORT OF THE FOLEY CATHETER, SO THE IRRIGANTS WERE SQUIRTING OUT. STATED THAT THE SAMPLE WAS NOT AVAILABLE SINCE IT WAS WITH THE PATIENT, AND IT WAS DISCARDED. AS PER FOLLOW UP VIA EMAIL RECEIVED ON 03FEB2023, IT WAS STATED THAT THEY HAD THREE INCIDENTS AT THE HOSPITAL WHERE 22FRENCH SILICON 3 WAYS CATHETERS WERE PLACED. ONCE THE CBI WAS STARTED IT WAS NOTICED THAT WATER WAS GUSHING OUT OF THE SIDES DUE TO A HOLE. IN TERMS OF IMPACT, THE HOLE IN THE CATHETER WOULD BREAK STERILITY. IT WAS ALSO STATED THAT IT WAS PAINFUL FOR PATIENTS TO HAVE CATHETERS REPLACED ESPECIALLY ONE AS LARGE AS A 22FRENCH CATHETERS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON 03MAR2023, IT WAS REPORTED THAT THE ANOTHER CATHETER WITH A HOLE WAS FOUND TODAY. THEY CURRENTLY HAD 8 AND A HALF BOXES FULL OF 22FR SILICON 3 WAY CATHETERS AND THEY WERE NO LONGER FEEL COMFORTABLE USING THESE CATHETERS. THE HOLE WAS SMALL AND ON THE THIRD PORT SECTION OF THE TUBE AND VERY DISTAL AND JUST PROXIMAL TO THE BIFURCATION AND ONLY SEE THE HOLE ONCE THE CBI IRRIGANT WAS TURNED ON AND IT SQUIRTS OUT. IN TERMS OF PATIENT CARE, EXCHANGING THE CATHETER WAS PAINFUL AND INCREASES THEIR RISK OF DEVELOPING A URINARY TRACT INFECTION. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00982 | 2023-03-24 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5913190 | 5913190 |  | IT WAS REPORTED THROUGH RESULTS OF A CLINICAL TRIAL THAT FOUR MONTHS AND EIGHTEEN DAYS POST INDEX PROCEDURE COMPLETION, THE SUBJECT HAD ADVERSE EVENT OF CLOTTED CATHETER. IT WAS FURTHER REPORTED THAT, THE ADVERSE EVENT RELATIONSHIP TO STUDY DEVICE WAS POSSIBLY RELATED AND THUS THE ACTION WAS TAKEN TO REMOVE THE DEVICE. REPORTEDLY, THE CATHETER WAS REMOVED. THE PROCEDURE WAS COMPLETED EXCHANGING ANOTHER NEW CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 02/2023). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00983 | 2023-03-24 | 2023-03-02 | I, n, j, u, r, y | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED BY THE CUSTOMER "ON MARCH 2ND BY MORNING, DURING PICC CATHETER PLACEMENT PROCEDURE, THE GUIDEWIRE USED FOR THE PUNCTURE UNRAVELED. THE PUNCTURE WAS PERFORMED, DURING THE INTRODUCTION OF THE GUIDE WIRE IT PRESENTED A RESISTANCE, WHEN ATTEMPTING TO REMOVE IT THE GUIDEWIRE GOT STUCK. THE NEEDLE WAS REMOVED, BUT WE HAD TOO MUCH DIFFICULT TO REMOVE THE GUIDEWIRE ENTIRELY. AN INTENSIVE PROFESSIONAL WAS CALLED WHICH WAS ABLE TO REMOVE IT, BUT THEN IT WAS OBSERVED THAT THE GUIDE WAS UNRAVELED. THE VASCULAR PROFESSIONAL WAS CALLED TO VERIFY IF THERE IS ANY REMAINING MATERIAL IN PATIENT." | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00986 | 2023-03-24 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER, 5F DL, 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000590A |  | IT WAS REPORTED BY THE CUSTOMER VIA PHONE, THAT THERE WAS A NOTICEABLE RED FOREIGN MATTER LOCATED INSIDE THE POWERPICC DEVICE. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00987 | 2023-03-24 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (18G) (8CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | F118080T |  | IT WAS REPORTED BY THE CUSTOMER "UPON PULLING BACK ON DEVICE HOUSING TO SAFETY WITH RIGHT HAND WHILE HOLDING WINGS STEADY WITH LEFT HAND THE NEEDLE SEPARATED FROM THE HOUSING." THERE WAS NO IMPACT TO THE PATIENT. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00989 | 2023-03-24 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY (4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4174118 |  | IT WAS REPORTED, "THE CATHETER SEEMS TO BE FINE, PRESENTS FLOW AND RE-FLOW WHEN TESTED WITH SALINE SOLUTION. WHEN SET IS CONNECTED TO ADMINISTER MEDICATION/SALINE THERAPY, IT PRESENTS A LEAKAGE IN THE CONNECTION BETWEEN THE CATHETER'S HUB WITH THE MICROCLAVE. THE MICROCLAVE IS REMOVED AND THE LEAKAGE PERSISTS. THE SET IS EXCHANGED, AND THE LEAKAGE KEEPS OCCURRING IN THE CONNECTION BETWEEN THE SET WITH THE CATHETER'S HUB IN ALL TESTS PERFORMED. IT'S REQUIRED TO CHANGE THE CATHETER." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00990 | 2023-03-24 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F TL MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1395108D4 |  | IT WAS REPORTED BY THE CUSTOMER THAT "PLACER NOTED LEAKING OF FLUID AND BLOOD FROM THE YELLOW STOPPER END WHERE THE STYLET PASSES THROUGH TO THE PICC. PLACER ALSO NOTED AIR ENTERING LINE THROUGH THE SAME AREA WHEN DRAWING THROUGH. THIS ISSUE WAS NOTICED DURING INSERTION OF OUR TL SOLO PICC. THERE WAS NO PATIENT IMPACT. NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00994 | 2023-03-24 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERGLIDE MIDLINE CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | CUSTOMER REPORTED THE POWERGLIDE WAS OCCLUDED AND WAS NOT ABLE TO BE FLUSHED. WHEN INCLUDED STATLOCK WAS REMOVED, A KINK WAS FOUND AT THE INSERTION SITE. REPLACEMENT WITH STANDARD PIVC STATLOCK RELIEVED THE KINK AND ALLOWED FOR CONTINUED USE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00993 | 2023-03-24 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERGLIDE MIDLINE CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER, POWERGLIDE OCCLUDED AND WAS NOT ABLE TO BE FLUSHED. WHEN INCLUDED STATLOCK WAS REMOVED, A KINK WAS FOUND AT THE INSERTION SITE. REPLACEMENT WITH STANDARD PIVC STATLOCK RELIEVED THE KINK AND ALLOWED FOR CONTINUED USE. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00995 | 2023-03-24 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERGLIDE MIDLINE CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | CUSTOMER REPORTED THE POWERGLIDE WAS OCCLUDED AND WAS NOT ABLE TO BE FLUSHED. WHEN INCLUDED STATLOCK WAS REMOVED, A KINK WAS FOUND AT THE INSERTION SITE. REPLACEMENT WITH STANDARD PIVC STATLOCK RELIEVED THE KINK AND ALLOWED FOR CONTINUED USE. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02025 | 2023-03-24 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119416M | 119416M |  | IT WAS REPORTED THAT TEMPERATURE SENSING FOLEY CATHETER NOT GETTING AN ACCURATE TEMPERATURE. CUSTOMER HAD TWO INSTANCES, ONE IN LAST WEEK AND TWO MORE IN THIS WEEK OF THIS SO 3 TOTAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00988 | 2023-03-24 | 2023-02-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D3 |  | IT WAS REPORTED BY THE CUSTOMER, "REGARDING PICC INSERTION KITS THAT RECENTLY HAVE TRANSITIONED TO NEWER STYLE KITS WITH A DIFFERENT EXTENSION SET WITH THE WIRE EXTENDING THROUGH THE PORT. THE NEW STYLE T-PORT HAS A MORE CONCAVE STYLE IN WHICH THE WIRE FEEDS THROUGH. THE INSERTION OF SINGLE LUMEN LOT REGX3255 AIR ENTERED THE CATHETER AT THE T-PORT INTO THE PICC AND PULLED BACK AIR AS THEY RETRACTED THE T-PORT SYRINGE. THE PICC WAS IMMEDIATELY REMOVED OR CLAMPED AND MANAGED WITHOUT PATIENT HARM. THE SL PICC WAS ABORTED, AND A NEW KIT USED TO PLACE THE LINE." | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02016 | 2023-03-24 | 2023-03-04 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD KINK IN THE LINE AND WOKE UP WET AT LAST NIGHT WHILE USING THE PUREWICK FEMALE EXTERNAL CATHETER. IT WAS STATED THAT THE FIRST NIGHT THE PATIENT USED IT WORKED GREAT. THE PATIENT KNEW LAST NIGHT HAD IT IN THE RIGHT PLACE AND JUST NEEDED BE SURE THE LINE WAS CLEAR THE FIRST MORNING BUT THE PATIENT WOKE UP A LITTLE WET. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. PER ADDITIONAL INFORMATION RECEIVED ON 04MAR2023, THE PATIENT HAD URINARY TRACT INFECTION AND THE HOSPITAL THOUGHT IT WAS BECAUSE OF THE PRODUCT. THE REPRESENTATIVE ADVISED ON THE PROPER CLEANING AND ADVISED TO REACH OUT TO THE DOCTOR FOR HELP. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. IT WAS UNKNOWN IF THE PUREWICK FEMALE EXTERNAL CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN . | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION. MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE OR MATERIAL SURFACE IS ROUGH, ABRASIVE OR UNCOMFORTABLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE SETUP: 1. CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE DRYDOCTM VACUUM STATION, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE DRYDOCTM VACUUM STATION USER GUIDE FOR FURTHER INFORMATION. 2. USING STANDARD SUCTION TUBING, CONNECT THE PUREWICKTM FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER. PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. REMOVAL: 5. TO REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER, FULLY SEPARATE THE LEGS, GLUTEUS, AND LABIA. TO AVOID POTENTIAL SKIN INJURY UPON REMOVAL, GENTLY PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER DIRECTLY OUTWARD. ENSURE SUCTION IS MAINTAINED WHILE REMOVING THE PUREWICKTM FEMALE EXTERNAL CATHETER. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. MAINTENANCE: 6. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. INDICATION FOR USE: THE PUREWICKTM FEMALE EXTERNAL CATHETER IS INTENDED FOR NON-INVASIVE URINE OUTPUT MANAGEMENT IN FEMALE PATIENTS. CONTRAINDICATIONS: PATIENTS WITH URINARY RETENTION WARNINGS: DO NOT USE THE PUREWICKTM FEMALE EXTERNAL CATHETER WITH BEDPAN OR ANY MATERIAL THAT DOES NOT ALLOW FOR SUFFICIENT AIRFLOW. TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER. HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON DO NOT USE BARRIER CREAM ON THE PERINEUM WHEN USING THE PUREWICKTM FEMALE EXTERNAL CATHETER. BARRIER CREAM MAY IMPEDE SUCTION. PROCEED WITH CAUTION IN PATIENTS WHO HAVE UNDERGONE RECENT SURGERY OF THE EXTERNAL UROGENITAL TRACT. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. MAINTAIN SUCTION UNTIL THE PUREWICKTM FEMALE EXTERNAL CATHETER IS FULLY REMOVED FROM THE PATIENT TO AVOID URINE BACKFLOW. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. PRIOR TO CONNECTING THE PUREWICKTM FEMALE EXTERNAL CATHETER TO HOSPITAL WALL SUCTION TUBING, VERIFY SUCTION FUNCTION BY COVERING THE OPEN END OF THE SUCTION TUBING WITH ONE HAND AND OBSERVING THE PRESSURE DIAL. IF THE PRESSURE DOES NOT INCREASE WHEN THE LINE IS COVERED, VERIFY THAT THE TUBING IS SECURED, CONNECTED, AND NOT KINKED. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." | Y | N | LAY USER/PATIENT |  | 2023-03
2020394-2023-00188 | 2023-03-24 | 2022-11-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER IQ ULTRASONIC | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | XCTO146 | XCTO146 |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE IN LEFT FEMORAL ARTERY VIA ANTIGRADE APPROACH, THE CATHETER COULD NOT ALLEGEDLY CROSS AND WAS STUCK TO THE WIRE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE CROSSER IQ CATHETER WAS RECEIVED AND EVALUATED. MATERIAL SEPARATION WAS NOTED WITH TIP STILL CONNECTED TO INNER CORE WIRE. UNKNOWN MATERIAL WAS ALSO NOTED ON THE DISTAL TIP. UPON FUNCTIONAL TESTING, AN ATTEMPT WAS MADE TO REMOVE THE RETURNED GUIDE-WIRE AND WAS UNSUCCESSFUL. THEREFORE, BASED ON THE FINDINGS, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED MATERIAL SEPARATION ISSUE. THE INVESTIGATION IS ALSO CONFIRMED FOR THE REPORTED DEVICE-DEVICE INCOMPATIBILITY ISSUE AS AN UNKNOWN MATERIAL WAS NOTED ON THE TIP AND THE RETURNED GUIDEWIRE WAS STUCK DURING FUNCTIONAL TESTING. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE TO ADVANCE ISSUE AS NO FURTHER FUNCTIONAL TESTING WAS PERFORMED. A DEFINITIVE ROOT CAUSE FOR THE REPORTED FAILURE TO ADVANCE, DEVICE-DEVICE INCOMPATIBILITY AND THE IDENTIFIED MATERIAL SEPARATION ISSUES COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRY DATE: 07/2023). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00979 | 2023-03-24 | 2022-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5913230 | 5913230 |  | IT WAS REPORTED THROUGH RESULTS OF A CLINICAL TRIAL THAT ELEVEN MONTHS AND SEVENTEEN DAYS POST INDEX PROCEDURE COMPLETION, THE SUBJECT HAD ADVERSE EVENT OF CLOTTED CATHETER. IT WAS FURTHER REPORTED THAT, THE ADVERSE EVENT RELATIONSHIP TO STUDY DEVICE WAS POSSIBLY RELATED AND THUS THE ACTION WAS TAKEN TO REMOVE THE DEVICE. REPORTEDLY, THE CATHETER WAS REMOVED. THE PROCEDURE WAS COMPLETED EXCHANGING ANOTHER NEW CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 06/2022. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00980 | 2023-03-24 | 2020-05-12 | I, n, j, u, r, y | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5913190 | 5913190 |  | IT WAS REPORTED THROUGH RESULTS OF A CLINICAL TRIAL THAT NINE MONTHS AND TWENTY-SIX DAYS POST INDEX PROCEDURE COMPLETION, THE SUBJECT HAD ADVERSE EVENT OBSTRUCTED CATHETER WITH FIBRIN SHEATH FORMATION. IT WAS FURTHER REPORTED THAT THE ADVERSE EVENT RELATIONSHIP TO STUDY DEVICE WAS POSSIBLY RELATED WHICH REQUIRES ADDITIONAL SURGICAL INTERVENTION AND SO ACTION WAS TAKEN TO REMOVE THE DEVICE. REPORTEDLY, THE CATHETER WAS REMOVED. THE PROCEDURE WAS COMPLETED EXCHANGING ANOTHER NEW CATHETER. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 11/2020). | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01987 | 2023-03-24 | 2023-02-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTERMITTENT ROCHESTER CATHETERS | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER HAD BEEN SUPPLIED WITH THE MAGIC3 HYDROPHILIC COATED SILICONE CATHETER. THEY READ AND UNDERSTOOD THE INSTRUCTIONS BUT SEEM TO STILL HAVE PROBLEMS WHEN THEY GO TO REMOVE IT. IT DID NOT SEEM TO HAVE THE LUBRICOSITY OF THE RED RUBBER ONES THAT THEY APPLY GEL TO. CUSTOMER QUESTIONED IF THEY COULD APPLY MORE LUBRICANT TO THE DEVICE, THEY FULLY UNDERSTOOD THE IMPLICATIONS OF THE ADDED CHANCES OF INFECTION BY TOUCHING MORE OBJECTS BUT WOULD STILL LIKE TO TRY AND USE THE CATHETER BECAUSE THEY HAD A LARGE SUPPLY AT THIS POINT. THEY WERE WORRIED THAT THE COATING THAT WAS ALREADY ON THERE MIGHT BE INCOMPATIBLE WITH ANOTHER LUBRICANT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND LOT NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00193 | 2023-03-24 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE CATHETER ALLEGEDLY MADE AN ABNORMAL SOUND AND FAILED TO VIBRATE. IT WAS FURTHER REPORTED THAT THE COREWIRE WAS FOUND EXTENDED. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01994 | 2023-03-24 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER FELL OUT OF THE PATIENT AFTER 2 DAYS OF PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01998 | 2023-03-24 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119416M | 119416M |  | IT WAS REPORTED THAT TEMPERATURE SENSING FOLEY CATHETER NOT GETTING AN ACCURATE TEMPERATURE. CUSTOMER HAD TWO INSTANCES, ONE IN LAST WEEK AND TWO MORE IN THIS WEEK OF THIS SO 3 TOTAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01990 | 2023-03-24 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER IN THE FOLEY CATHETER BALLOON WAS UNABLE TO BE REMOVED. THE CATHETER WAS NOT USED ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01995 | 2023-03-24 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51616 | 51616 |  | IT WAS REPORTED THAT THERE WAS NOT ENOUGH LUBRICANT ON THE MAGIC 3 FEMALE INTERMITTENT CATHETER. PATIENT HAD TO GO AND BUY THE LUBE TO USE WITH THE CATHETER. PER CLARIFICATION MAIL RECEIVED FROM LIBERATOR ON 22MAR2023, STATED THAT THE PATIENT WAS A MAN AND HE USES FEMALE-LENGTH INTERMITTENT CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-02003 | 2023-03-24 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS DIFFICULT TO INFLATE. ACCORDING TO THE ATTACHED DOCUMENT, UPON INJECTING WATER TO THE CATHETER DURING THE PRE-TEST, THE BALLOON WAS DIFFICULT TO INFLATE. WATER LEAKAGE FROM THE BALLOON WAS NOT OBSERVED. ALSO STATED THAT SWELLING OF BALLOON WITH WATER TO REPLACE BUT WHEN IT POURED THE NATIONAL WATER THEY HEARD A SOUND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01988 | 2023-03-24 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT AT HOME, THERE WAS NO OUTFLOW OF URINE FROM THE CATHETER, SO WHEN IT WAS ATTEMPTED TO REMOVE, IT COULD NOT BE REMOVED, SO AT (B)(6) HOSPITAL, THE PENIS WAS PUNCTURED AND THE BALLOON PORTION WAS RUPTURED AND REMOVED. THE BALLOON WAS TRIED TO BE DEFLATED BY PERCUTANEOUS BALLOON RUPTURE. THE POSSIBILITY OF INDWELLING IN THE URETHRA WAS SUSPECTED BECAUSE NO URINE OUTFLOW HAS BEEN CONFIRMED AND THE PENIS HAS BEEN PUNCTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02008 | 2023-03-24 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER CAME OUT WHEN THE PATIENT WAS MOVED TO THE OPERATING ROOM AND PLACED ON THE OPERATING TABLE AFTER THE CATHETER WAS PLACED IN THE WARD. WHEN THE WATER WAS ADDED AGAIN, THE WATER FLOWED INTO THE LUMEN. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. BASED ON THE EVALUATION, A TEAR WAS FOUND ON THE RUBBERIZED LAYER CAUSING INTERLUMEN LEAKAGE. THE REPORTED EVENT IS CONFIRMED AS MANUFACTURING RELATED ISSUE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON UNIFORM RUBBERIZED THICKNESS ETC CAUSING TORN RUBBERIZE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00981 | 2023-03-24 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT APPROXIMATELY ONE MONTH AND ELEVEN DAYS POST PORT PLACEMENT PROCEDURE, THE COMPUTED TOMOGRAPHIC IMAGE EXAMINATION REVEALED THAT THE DEVICE ALLEGEDLY HAD A SUBCUTANEOUS LEAK NEAR THE PORT BODY. IT WAS FURTHER REPORTED THAT, THE DEVICE DOES NOT HAVE ANY DAMAGE. REPORTEDLY, THE PORT WAS STILL IMPLANTED. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 09/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-01992 | 2023-03-24 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER LEAKED FROM THE VICINITY OF THE INFLATION VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00996 | 2023-03-25 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY THE CUSTOMER "POWERPICCS LEAKING AT THE RUBBER AREA AT THE PROXIMAL END. THEY WERE TO RECEIVE A DIFFERENT LOT NUMBER BUT INSTEAD RECEIVED THE SAME LOT NUMBER [NEW PICC] AND WHEN THEY TRIED TO INSERT AGAIN, THEY HAD THE SAME LEAKING." ADDITIONAL INFORMATION RECEIVED ON 3/6/2023: IT WAS REPORTED BY THE CUSTOMER "ON FRIDAY, MARCH 3RD, PRIOR TO INSERTION, WHEN FLUSHING THE PICC WITH SALINE, IT WAS DISCOVERED THAT THE PICC HAD THE SAME DEFECT WHERE SALINE LEAKED FROM THE RUBBER STOPPER. OPENED A SECOND PICC WITH THE SAME LOT NUMBER AND FLUSHED IT WITH SALINE AND AGAIN, SALINE LEAKED OUT THROUGH THE RUBBER STOPPER ALONG THE INSERTION WIRE. THIS REPORT ADDRESSES THE FIRST DEVICE. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00997 | 2023-03-25 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED BY THE CUSTOMER "POWERPICCS LEAKING AT THE RUBBER AREA AT THE PROXIMAL END. THEY WERE TO RECEIVE A DIFFERENT LOT NUMBER BUT INSTEAD RECEIVED THE SAME LOT NUMBER [NEW PICC] AND WHEN THEY TRIED TO INSERT AGAIN, THEY HAD THE SAME LEAKING." ADDITIONAL INFORMATION RECEIVED ON 3/6/2023: IT WAS REPORTED BY THE CUSTOMER "ON FRIDAY, MARCH 3RD, PRIOR TO INSERTION, WHEN FLUSHING THE PICC WITH SALINE, IT WAS DISCOVERED THAT THE PICC HAD THE SAME DEFECT WHERE SALINE LEAKED FROM THE RUBBER STOPPER. OPENED A SECOND PICC WITH THE SAME LOT NUMBER AND FLUSHED IT WITH SALINE AND AGAIN, SALINE LEAKED OUT THROUGH THE RUBBER STOPPER ALONG THE INSERTION WIRE. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01001 | 2023-03-26 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC, 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER, "REGARDING PICC INSERTION KITS THAT RECENTLY HAVE TRANSITIONED TO NEWER STYLE KITS WITH A DIFFERENT EXTENSION SET WITH THE WIRE EXTENDING THROUGH THE PORT. THE NEW STYLE T-PORT HAS A MORE CONCAVE STYLE IN WHICH THE WIRE FEEDS THROUGH. THE PICC WAS IMMEDIATELY REMOVED OR CLAMPED AND MANAGED WITHOUT PATIENT HARM." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01000 | 2023-03-26 | 2022-12-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (5F) DL (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8295118 |  | IT WAS REPORTED THAT "THIS PATIENT NEEDS TO HAVE A PICC CATHETER PLACED. DURING PLACEMENT, THE CUSTOMER FOUND THAT THE CATHETER¿S RADIAN WAS NOT EVENT, WHICH LED TO DIFFICULTY IN CATHETER INTRODUCTION. THE CATHETER WAS THEREFORE REPLACED WITH ANOTHER ONE, WHICH WAS INSERTED SUCCESSFULLY." 03/20/2023 - THE RETURNED STYLET EXHIBITED BREAKS AND BENDS IN THE COIL WIRE. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE ANALYSIS, PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED STYLET IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE 5 FR DUAL LUMEN POWERPICC SOLO CATHETER WITH A COILED STIFFENING STYLET AND T-LOCK INSERTED WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLE. A BEND WAS OBSERVED IN THE STYLET WITHIN THE CATHETER AT THE 24 CM DEPTH MARKER. THE STYLET WAS REMOVED FROM THE CATHETER AND THE BEND WAS FOUND TO BE APPROXIMATELY 13.5 CM PROXIMAL TO THE DISTAL TIP OF THE STYLET. NO BREAK WAS FOUND IN THE CORE WIRE OF THE RETURNED STYLET DURING TACTILE EVALUATION. A MICROSCOPIC OBSERVATION REVEALED AT LEAST TWO BREAKS IN THE COILED WIRE OF THE STYLET AT THE LOCATION OF THE BEND. THE BREAK SITES IN THE COILED WIRE WERE OBSERVED TO BE ANGLED AND SOMEWHAT UNEVEN. SOME OF THE COILS NEAR THE BREAK SITES WERE OBSERVED TO BE BENT AND DEFORMED. STRIATIONS OBSERVED ON THE SURFACE OF THE DISTAL END OF THE CATHETER INDICATED THAT IT HAD BEEN TRIMMED WITH A SHARP INSTRUMENT SUCH AS SCISSORS OR A SCALPEL AT THE 42 CM DEPTH MARKER. THE WELD TIP OF THE STYLET WAS OBSERVED TO BE PRESENT AND INTACT. WHILE THE EXACT CAUSE OF THE DAMAGE OBSERVED IN THE RETURNED STYLET COULD NOT BE DETERMINED, POSSIBLE CAUSES INCLUDE CONTACT WITH A SHARP INSTRUMENT, ACCIDENTAL STYLET DAMAGE DURING TRIMMING OF THE CATHETER, AND CONTACT WITH A METALLIC INSTRUMENT SUCH AS HEMOSTATS. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01006 | 2023-03-26 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT ON (B)(6) PATIENT RETURNS FROM ENDOSCOPY AFTER NASOGASTRIC PROBE PLACEMENT, PICC DEVICE WITH FILM WET, PRESENTING RESISTANCE AND WITH A LEAKAGE OF SALINE SOLUTION IN INSERTION SITE. THE PHYSICIAN HAS BEEN MADE AWARE. THE PICC CATHETER WAS REMOVED, AND IT WAS NOTED IT HAS A HOLE IN INSERTION - RUPTURE OF CATHETER'S LUMEN. ON MARCH 6 ANOTHER PICC CATHETER WAS PLACED. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-00998 | 2023-03-26 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD3 |  | IT WAS REPORTED BY THE CUSTOMER, "REGARDING PICC INSERTION KITS THAT RECENTLY HAVE TRANSITIONED TO NEWER STYLE KITS WITH A DIFFERENT EXTENSION SET WITH THE WIRE EXTENDING THROUGH THE PORT. THE NEW STYLE T-PORT HAS A MORE CONCAVE STYLE IN WHICH THE WIRE FEEDS THROUGH. AIR ENTERED THE CATHETER AT THE T-PORT INTO THE PICC AND PULLED BACK AIR AS THEY RETRACTED THE T-PORT SYRINGE. THE PICC WAS IMMEDIATELY REMOVED OR CLAMPED AND MANAGED WITHOUT PATIENT HARM." | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | OTHER |  | 2023-03
3006260740-2023-01005 | 2023-03-26 | 2023-01-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER THAT DURING THE PICC PLACEMENT PROCEDURE, IT WAS NOTED THAT THE CATHETER IS DAMAGED, BLOOD IS REFLOWING, AND THE VALVE IS NOT EFFECTIVE. PROCEDURE WAS COMPLETED BY REMOVING THE CATHETER AND PLACEMENT OF A NEW ONE. NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02056 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 | FOL0100 | FOL0100 |  | IT WAS REPORTED THAT THE CUSTOMER HAD BEEN HAVING ISSUES GETTING THE COVIDIEN SILICONE FOLEYS TO FIT IN THE STATLOCK FOR THE PAST YEAR. ALSO STATED IF THEY FLIP THE FOLEY OVER, THE WHITE CIRCULAR PART PUSHED THE PATIENT AND WAS UNCOMFORTABLE. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-01017 | 2023-03-27 | 2023-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602833 | 0602833 |  | IT WAS REPORTED THAT THREE MONTHS AND TWENTY-THREE DAYS POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY FOUND TO BE CRACKED NEAR THE PORT SYSTEM. IT WAS FURTHER REPORTED THAT THE SUBCUTANEOUS LEAKAGE HAD ALLEGEDLY OCCURRED FROM THE CATHETER. REPORTEDLY, THE CATHETER WAS REMOVED. THE PROCEDURE WAS COMPLETED REPLACING ANOTHER CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 12/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01014 | 2023-03-27 | 2022-12-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118J |  | IT WAS REPORTED THAT IT WAS DIFFICULT TO INSERT THE GUIDEWIRE AND DIFFICULT TO REMOVE IT. THERE WAS NO REPORTED PATIENT INJURY. ON 03/21/2023: IT WAS FOUND DURING AN INVESTIGATION THAT A SMALL SECTION OF THE COILED WIRE OF THE GUIDEWIRE WAS UNRAVELED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE ANALYSIS, PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GUIDEWIRE IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE 0.018 IN. NITINOL GUIDEWIRE WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED SOME USE RESIDUE ON THE RETURNED SAMPLE. THE DISTAL END OF THE GUIDEWIRE WAS OBSERVED TO BE SLIGHTLY CURVED. TACTILE EVALUATION OF THE GUIDEWIRE CONFIRMED THE CORE WIRE WAS NOT BROKEN. A MICROSCOPIC OBSERVATION REVEALED A SMALL SECTION OF THE COILED WIRE OF THE GUIDEWIRE WAS UNRAVELED NEAR THE ADHESIVE FILLET AT ITS PROXIMAL END, AND THE ADHESIVE FILLET WAS OBSERVED TO BE SLIGHTLY DAMAGED. IT COULD NOT BE DETERMINED WHEN OR WHERE THE RETURNED GUIDEWIRE WAS DAMAGED FROM THE EVIDENCE OBSERVED ON THE RETURNED SAMPLE. POSSIBLE CONTRIBUTING FACTORS INCLUDE DAMAGE FROM A METALLIC INSTRUMENT AND EXCESSIVE FORCE DURING HANDLING, MANIPULATION, OR USE OF THE GUIDEWIRE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02057 | 2023-03-27 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT IT WAS DIFFICULT TO DEFLATE FOLEY CATHETER DURING PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02059 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT SOME OF THE INTERMITTENT CATHETERS WERE CLOGGED AND DEFECTIVE. | THE REPORTED EVENT IS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO INSUFFICIENT EYELET AREA BLOCKED EYE OR INNER DIAMETER OF DRAINAGE EYE SMALL ENOUGH TO ALLOW FOR OCCLUSION DURING USE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: INSTRUCTIONS FOR USE. INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DONT REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-02061 | 2023-03-27 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT CONNECTING THE SYRINGE WITH THE INFLATION VALVE TO INJECT STERILE WATER, THE WATER LEAKED FROM THE INFLATION VALVE. IT WAS FURTHER REPORTED THAT THE STERILE WATER WAS CLOUDY . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02064 | 2023-03-27 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE DURING A PRETEST, UPON INJECTING WATER INTO THE CATHETER, WATER LEAKED IN THE BLUE PLASTIC BAG. IT IS UNCLEAR WHETHER THAT LEAKAGE IS FROM THE BALLOON OR THE CATHETER. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED. PER ADDITIONAL INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, STATED THAT THE LEAK LOCATION WAS NOT CONFIRMED. HOWEVER, CHECKING THE RETURNED SAMPLE, THE DRAINAGE EYE PART WAS TORN. THE LEAK MAY HAVE OCCURRED FROM THE TEAR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02065 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123618A | 123618A |  | IT WAS REPORTED THAT THE 2 WAY FOLEY CATHETER WERE DEFECTIVE, AND THE NURSE STATED THE WATER WAS GOING IN BUT NONE COMING OUT. THE PATIENT HAD HAD URINE LEAKING ON THEM SINCE THE PATIENT STATED USED THESE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-01011 | 2023-03-27 | 2021-02-28 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1808060 |  | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT ELEVEN DAYS POST A PORT PLACEMENT, THE PORT ALLEGEDLY DID NOT FUNCTION AND THE PATIENT ALLEGEDLY EXPERIENCED HEMATOMA, ARM PAIN AND FEVER DUE TO THE PORT. IT WAS FURTHER REPORTED THAT X-RAY EXAMINATION AND LAB TESTS WERE PERFORMED WHICH CONFIRMED A POTENTIAL BLOOD STREAM INFECTION. REPORTEDLY, THE PATIENT WAS ADMITTED TO THE HOSPITAL; HOWEVER, THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. THE MEDICAL RECORD ALLEGES THAT A BODY IMPLANT PORT WAS IMPLANTED TO A PATIENT FOR THE PURPOSE OF BLOOD DRAW AND IMMUNOTHERAPY DELIVERY. APPROXIMATELY A YEAR AND ONE MONTH LATER THE OLD PORT WAS REMOVED THROUGH UNKNOWN MAIN APPROACH YOUR NEW BARD POWER PORT ISP MRI HAS BEEN IMPLANTED AFTER BEING DIAGNOSED WITH THE CROHN'S DISEASE. THE PORT WAS LAYING PERFECTLY FLAT ON THE CHEST, AND IT WAS EASILY PALPABLE. THE PATIENT WAS RECOVERY AND IN STABLE CONDITION WITHOUT NO COMPLICATION. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE AS NO OBJECTIVE EVIDENCE FOR THE REPORTED DEFICIENCY WITH THE PORT IN THE SUBMITTED MEDICAL RECORD REVIEW. ADDITIONALLY, NO OBJECTIVE EVIDENCE OF INFECTION AND HEMATOMA COULDN'T BE CONFIRMED. FURTHERMORE, THE CLINICAL CONDITION ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 05/2021), G3, H6 (METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02072 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175812 |  | IT WAS REPORTED THAT A PATIENT'S RECENT CATHETER HAD BURST INSIDE THEIR BLADDER (BATCH# NGGT2360). ALSO STATED THAT SEVERAL CATHETER BALLOON WERE DEFLATED INSIDE THEIR BLADDER PREVIOUSLY (BATCH# NGFX3156 AND BATCH# NGFV3848). THE CUSTOMER WANTED TO MAKE THE REPRESENTATIVE AWARE IN CASE OF A PRODUCTION OR BATCH ERROR. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 23MAR2023, STATED THAT THE FOLEY CATHETER BALLOON OF BATCH # NGGT2360 WAS SPLIT, CATHETER BALLOON OF BATCH # NGFX3156 WAS DEFLATED AND CATHETER BALLOON OF BATCH # NGFV3848 WAS ALSO DEFLATED. THE BALLOON APPEARED TO BE CLEAN SPLIT. THE CUSTOMER WAS UNSURE IF ANY MATERIAL WAS LEFT IN BLADDER, SO THEY ESCALATED TO GENERAL PRACTITIONER WHO DIDN¿T FEEL APPROPRIATE FOR A ULTRASOUND SCAN TO CONFIRM THIS UNLESS THE PATIENT BECOMES SYMPTOMATIC. ON (B)(6) 2023, THE PREVIOUS FOLEY CATHETER BALLOON INSERTED WITH BATCH # NGGT2360 WAS SPLIT, ON (B)(6) 2023 THE BALLOON DEFLATED FOR WHICH THEY WERE UNSURE OF LOT NUMBER AS THIS WAS NOT DOCUMENTED (BATCH# UNK) WHEN IT WAS INSERTED AND ON (B)(6) 2022, THE CATHETER BALLOON WITH BATCH # NGFV3848 DEFLATED. THEY WERE UNSURE IF BALLOON WITH BATCH NGGT2360 WAS EITHER DEFLATED OR SPLIT FIRST AND IT WAS UNABLE TO DETERMINE. UPON ARRIVAL THIS CATHETER WAS ALREADY OUT DUE TO FALLING OUT AND ALREADY HAD SPLIT AS PER THE PICTURE PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02070 | 2023-03-27 | 2023-01-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE-SENSING FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE DEVICE WAS USED IN THE OPERATING ROOM AND THE WATER LEAKED FROM THE BALLOON. THERE WAS NO BALLOON RUPTURE OR TEAR. THE LOCATION OF LEAK IS UNKNOWN BECAUSE THE EVENT WAS CATHETER FELL OUT. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02073 | 2023-03-27 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | MAGIC 3 INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT PATIENT AND CAREGIVER STATED THAT THE MAGIC 3 INTERMITTENT CATHETER WAS DIFFICULT TO INSERT SO THESE WERE NOT WORKING FOR THE PATIENT. PATIENT STATED THESE WERE STRAIGHT CATHETERS. THE PATIENT WANTED COUDE MAGIC 3 INTERMITTENT CATHETER 50614. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-02078 | 2023-03-27 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT THE CUSTOMER HAD AN ISSUE WITH A FOLEY CATHETER LAST WEEK. STATED THAT WHEN ONE OF THEIR NURSES DEFLATED THE FOLEY BALLOON, A RIDGE FORMED NEAR THE TIP OF THE CATHETER. UNFORTUNATELY, THEY DID NOT HAVE THE FOLEY TRAY PACKAGE THAT WAS USED. HOWEVER, THEIR OPERATING ROOM (OR) CLINICAL MANAGER WAS ABLE TO DUPLICATE THE PROBLEM WITH ANOTHER TRAY. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED ONE PHOTO SAMPLE THAT SHOWS END OF CATHETER WITH BALLOON MUSHROOMED. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE INCORRECT BALLOON DESIGN. DHR REVIEW IS NOT REQUIRED AS NO LOT NUMBER WAS REPORTED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: DESCRIPTION: BARD® ALL-SILICONE FOLEY CATHETERS AID IN CLINICAL MANAGEMENT BY ALLOWING FOR CONTINUOUS DRAINAGE AND MEASUREMENT OF URINE. INDWELLING URINARY CATHETERS OR FOLEY CATHETERS CAN BE USED IN PATIENTS OF ALL AGES. UP TO 10 CH/FR IS GENERALLY CONSIDERED PEDIATRIC SIZING, HOWEVER, THE APPROPRIATE SIZE SHOULD BE DETERMINED BY THE HEALTHCARE PROFESSIONAL. INTENDED USE: FOR UROLOGICAL USE ONLY. INDICATIONS: BARD® ALL-SILICONE FOLEY CATHETERS ARE INDICATED FOR ANY CLINICAL CONDITION REQUIRING DRAINAGE AND/OR COLLECTION AND/OR MEASUREMENT OF URINE. GENERALLY, DRAINAGE IS ACCOMPLISHED BY INSERTING THE CATHETER THROUGH THE URETHRA AND INTO THE BLADDER. CONTRAINDICATION: NO KNOWN CONTRAINDICATIONS CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. DO NOT USE DEVICE IF PACKAGE IS OPENED OR DAMAGED. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS, OR DEATH OF THE PATIENT. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. WARNINGS: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. IT MAY CAUSE BALLOON TO BURST AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE, AND FEDERAL LAWS AND REGULATIONS. USERS AND/OR PATIENTS WITHIN THE EUROPEAN UNION, SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND THE COMPETENT AUTHORITY OF THE MEMBER STATE IN WHICH THE USER AND/ OR PATIENT IS ESTABLISHED. USERS OUTSIDE OF THE EUROPEAN UNION SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/ OR PATIENT IS ESTABLISHED. STORE CATHETERS AT ROOM TEMPERATURE AND AWAY FROM DIRECT EXPOSURE TO SUNLIGHT PREFERABLY IN ORIGINAL PACKAGING. INSTRUCTIONS FOR USE: VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. FOLLOW YOUR FACILITY PROTOCOL FOR CATHETER STABILIZATION AND URINE COLLECTION. INSERTION: WASH HANDS. USE PROPER ASEPTIC TECHNIQUE TO PREPARE THE PATIENT FOR CATHETER INSERTION. REMOVE THE CATHETER FROM WRAP AND LUBRICATE THE CATHETER. IF HOSPITAL POLICY PERMITS, IT IS POSSIBLE TO INJECT LUBRICANT INTO THE URETHRA. CATHETERIZE THE PATIENT USING DOMINANT HAND. WHEN CATHETER TIP HAS ENTERED BLADDER, URINE WILL FLOW THROUGH THE CATHETER. FOR FEMALE, INSERT CATHETER 2 MORE INCHES (APPROXIMATELY 5 CM). FOR MALE, INSERT CATHETER ALL THE WAY TO BIFURCATION. INFLATE THE BALLOON WITH PRE-FILLED WATER SYRINGE IF INCLUDED. IF PRE-FILLED SYRINGE IS NOT INCLUDED, THEN USE A SLIP TIP/LUER LOCK SYRINGE TO FILL CATHETER BALLOON WITH STERILE WATER. DO NOT USE NEEDLE. DO NOT EXCEED RECOMMENDED CAPACITIES AS STATED ON THE INFLATION LUMEN OF THE CATHETER. GENTLY PULL THE CATHETER UNTIL THE CATHETER BALLOON IS SNUG AGAINST THE BLADDER NECK REMOVAL: WASH HANDS. GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION, IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF BALLOON DOES NOT DEFLATE, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. AFTER BALLOON IS FULLY DEFLATED, REMOVE CATHETER, AND DISPOSE OF IN ACCORDANCE WITH HOSPITAL POLICY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
MW5116005 | 2023-03-27 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARDEX COUDE SILVER HYDROGEL CATHETER, BARDEX 22FR 5CC | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | 1758SI22 | 1758SI22 |  | SILICONE 22FR CATHETER IN PATIENT AND CATHETER WAS NOT DRAINING. MD ATTEMPTED TO REMOVE CATHETER AFTER BALLOON DEFLATION BUT WAS UNABLE. MD CUT BALLOON FILLING PORT AND BALLOON STAYED INFLATED. MD REMOVED CATHETER WITH INFLATED BALLOON. MD WENT BACK IN WITH CYSTOSCOPE TO ASSURE HOMEOSTATIS AND THEN REPLACED NEW CATHETER WITHOUT DIFFICULTY. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
MW5116007 | 2023-03-27 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | COUDE CATHETER 12 FR. 5CC RIBBED BALLOON | CATHETER, COUDE | C. R. BARD, INC. |  | 0102L12 |  | PATIENT UNDERWENT L4-L5 LAMINECTOMY. A 12-FR. FOLEY COUDE CATHETER WAS PLACED AND BALLOON INFLATED PRIOR TO PROCEDURE WITHOUT DIFFICULTY. POST OPERATIVELY, THE RN DEFLATED BALLOON TIP WITH EMPTY SYRINGE. WHEN UNABLE TO REMOVE ANY MORE FLUID FROM TIP RN ATTEMPTED TO PULL CATHETER OUT BUT RESISTANCE MET CAUSING PATIENT DISCOMFORT. CHARGE RN TRIED TO TROUBLE SHOOT BY TRYING TO SEE IF ANY MORE FLUID COULD BE REMOVED FROM TIP. NOTHING CAME OUT. FOLEY LEFT IN PLACE. DURING SHIFT, NIGHT RN AND DAY RN MADE ANOTHER ATTEMPT TO PULL FOLEY AND THIS TIME IT CAME OUT. TIP OF CATHETER HAD BALLOON STILL SLIGHTLY INFLATED. PATIENT IN STABLE CONDITION WITH NO INJURY. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01010 | 2023-03-27 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA RIGHT INTERNAL JUGULAR VEIN, RESISTANCE WAS MET WHEN FLUSHING THE CATHETER AND ASPIRATING THE HEPARINIZED WATER IN VITRO. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT INVESTIGATION SUMMARY: ONE POWERPORT IMPLANTABLE PORT, ONE TUNNELER, ONE 6.5 PEEL-APART SHEATH AND VESSEL DILATOR, ONE VEIN PICK, ONE STRAIGHT NON-CORING NEEDLE, ONE CATH-LOCK AND ONE FLUSHING CONNECTOR LOADED TO A CATHETER WERE RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, FUNCTIONAL AND TACTILE EVALUATIONS WERE PERFORMED. NO LEAKS WERE OBSERVED THROUGHOUT THE INFUSION TEST. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO FLUSH AND SUCTION ISSUES AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 12/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02068 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE-SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, THE CATHETER BALLOON WAS UNEVENLY INFLATED. IT WAS FURTHER REPORTED THAT AFTER INJECTING STERILE WATER INTO THE CATHETER, THE WATER WAS DIFFICULT TO DRAIN BACK INTO THE SYRINGE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02055 | 2023-03-27 | 2023-03-06 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD LEAKING. THE PATIENT WAS CURRENTLY IN THE HOSPITAL DUE TO URINARY TRACT INFECTION FROM SITTING IN THE URINE. THE REPRESENTATIVE WENT OVER PLACEMENT AND OVERFLOW VALVE. THE PATIENT HAD BEEN USING THIS PRODUCT FOR LESS THAN 90 DAYS. IT WAS UNKNOWN IF THE PUREWICK FEMALE EXTERNAL CATHETER CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE " DIMENSIONS NOT SPECIFIED CORRECTLY AND/OR IMPROPER MATERIALS USED ASSEMBLY COLLAPSES UNDER VACUUM". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. WARNINGS: DO NOT USE THE PUREWICK¿ FEMALE EXTERNAL CATHETER WITH BEDPAN OR ANY MATERIAL THAT DOES NOT ALLOW FOR SUFFICIENT AIRFLOW. NEVER INSERT THE PUREWICK¿ FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICK¿ FEMALE EXTERNAL CATHETER HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE DO NOT USE BARRIER CREAM ON THE PERINEUM WHEN USING THE PUREWICK¿ FEMALE EXTERNAL CATHETER. BARRIER CREAM MAY IMPEDE SUCTION. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICK¿ FEMALE EXTERNAL CATHETER. MAINTAIN SUCTION UNTIL THE PUREWICK¿ FEMALE EXTERNAL CATHETER IS FULLY REMOVED FROM THE PATIENT TO AVOID URINE BACKFLOW. RECOMMENDATIONS: ¿ ENSURE THE PUREWICK¿ FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICK¿ FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICK¿ FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICK¿ FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICK¿ FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICK¿ FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS. | Y | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-02054 | 2023-03-27 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT AFTER THE CATHETER INSERTION. UPON CHECKING THE CATHETER, THE BALLOON RUPTURE WAS FOUND. NO MATERIALS POSSIBLY CAME IN CONTACT TO THE BALLOON. THE CATHETER WAS NOT PULLED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02050 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON BURST. THERE WAS NO PROBLEM WHEN PRE-TESTED BEFORE INSERTION, BUT THE OUTFLOW OF URINE DURING SURGERY WORSENED, SO WHEN THE USER CHECKED IT, THE CATHETER WITH THE RUPTURED BALLOON CAME OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01008 | 2023-03-27 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN, WHITE, 4.2F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 0601610 | 0601610 |  | IT WAS REPORTED THAT DURING PREPARATION OF A CATHETER REPAIR PROCEDURE, THE GLUE WAS ALLEGEDLY HARDENED. THERE WAS NO PATIENT CONTACT. | A VOLUNTARY RECALL HAS BEEN INITIATED FOR CATHETER REPAIR KITS WHICH ARE PRODUCT CATALOG/LOT NUMBER SPECIFIC. REPORTEDLY THE CATHETER REPAIR KITS HAVE HARDENED/COAGULATED ADHESIVE. HARDENED OR COAGULATED ADHESIVE HAS THE POTENTIAL TO CAUSE DELAYS WHILST AN ALTERNATIVE REPAIR KIT, ADHESIVE OR REPLACEMENT CATHETER IS OBTAINED; THEREFORE, POTENTIALLY PROLONGING SURGERY OR NECESSITATING EXCHANGE. TO DATE THERE HAVE BEEN NO ADVERSE EVENTS WORLDWIDE REPORTED RELATED TO THIS ISSUE. AS RESULT OF THE FIELD ACTION, THIS EVENT IS BEING REPORTED AS A MALFUNCTION REPORTABLE EVENT. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRY DATE: 12/2024. SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01009 | 2023-03-27 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN, WHITE, 4.2F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 0601610 | 0601610 |  | IT WAS REPORTED THAT DURING PREPARATION OF A CATHETER REPAIR PROCEDURE, THE GLUE WAS ALLEGEDLY HARDENED. THERE WAS NO PATIENT CONTACT. | A VOLUNTARY RECALL HAS BEEN INITIATED FOR CATHETER REPAIR KITS WHICH ARE PRODUCT CATALOG/LOT NUMBER SPECIFIC. REPORTEDLY THE CATHETER REPAIR KITS HAVE HARDENED/COAGULATED ADHESIVE. HARDENED OR COAGULATED ADHESIVE HAS THE POTENTIAL TO CAUSE DELAYS WHILST AN ALTERNATIVE REPAIR KIT, ADHESIVE OR REPLACEMENT CATHETER IS OBTAINED; THEREFORE, POTENTIALLY PROLONGING SURGERY OR NECESSITATING EXCHANGE. TO DATE THERE HAVE BEEN NO ADVERSE EVENTS WORLDWIDE REPORTED RELATED TO THIS ISSUE. AS RESULT OF THE FIELD ACTION, THIS EVENT IS BEING REPORTED AS A MALFUNCTION REPORTABLE EVENT. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRY DATE: 12/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02052 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HEMATURIA FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 2551H20 |  | IT WAS REPORTED THAT AFTER 2 DAYS OF PLACEMENT, THE CATHETER BALLOON WHICH WAS REPLACED AFTER THE BALLOON BURST HAPPENED A DAY AFTER TURP PROCEDURE HAD ALSO BURST AGAIN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02033 | 2023-03-27 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE WATER WAS DIFFICULT TO BE INJECTED. DURING THE TEST, WATER WAS INJECTED TO EXPAND THE FOLEY CATHETER BALLOON, BUT IT WAS NOT POSSIBLE. THE USER TRIED TO CONNECT THE VALVE AND INJECT IT, BUT THEY COULD NOT DO IT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02036 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI20 | 0165SI20 |  | IT WAS REPORTED THAT URINE LEAKED DUE TO THE ASYMMETRICAL FOLEY CATHETER BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02037 | 2023-03-27 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE CATHETER WAS PLACED TO THE PATIENT IN THE OPERATING ROOM, AND ON THE NEXT DAY, IT FELL OUT OF THE PATIENT IN THE WARD. THE BALLOON WAS DRAINED. THE EVENT OCCURRED ONE DAY AFTER THE PLACEMENT. NO BALLOON RUPTURE OR TEAR WAS OBSERVED ON THE FOLEY CATHETER. THE LEAK LOCATION WAS UNKNOWN BECAUSE THE EVENT WAS CATHETER FELL OUT AND THE USER DID NOT TRY TO REPRODUCE THE EVENT. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02041 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123618A | 123618A |  | IT WAS REPORTED THAT THE 2 WAY FOLEY CATHETER WERE DEFECTIVE, AND THE NURSE STATED THE WATER WAS GOING IN BUT NONE COMING OUT. THE PATIENT HAD URINE LEAKING ON THEM SINCE THE PATIENT STATED USED THESE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. NOT EVALUATED REASON: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
2020394-2023-00200 | 2023-03-27 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U357544 | U357544 |  | IT WAS REPORTED THAT DURING PREPARATION FOR AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY LEAKED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: ADDITIONAL INFORMATION RECEIVED, THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 12/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02038 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED THAT THE PATIENT COMPLAINED THAT THE LUBRICANT WAS WATERY AND LEAKY MORE THAN IT HAD BEEN BEFORE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE, "INCORRECT PROCESS PARAMETERS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-02040 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI20 | 0165SI20 |  | IT WAS REPORTED THAT URINE LEAKED DUE TO THE ASYMMETRIC FOLEY CATHETER BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01012 | 2023-03-27 | 2023-02-28 | I, n, j, u, r, y | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | A710962 | A710962 |  | IT WAS REPORTED THAT APPROXIMATELY TEN MONTHS POST PORT PLACEMENT PROCEDURE, TISSUE ABOVE THE PORT HAD ALLEGEDLY BROKEN DOWN. IT WAS FURTHER REPORTED THAT THE PORT HAD ALLEGEDLY ERODED TO THE POINT IT WAS EXPOSED AND THE DEVICE ALLEGEDLY LEAKED. REPORTEDLY PATIENT HAD DEVELOPED INFECTION AND BRUISING IN THE TISSUE ABOVE THE PORT. THE PORT WAS REMOVED. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERFLOW IMPLANTABLE PORT ATTACHED TO A CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. A COMPLETE DIAGONAL BREAK WAS NOTED AT THE DISTAL END OF THE CATHETER ATTACHED AND THE EDGES OF THE COMPLETE DIAGONAL BREAK WERE NOTED TO BE UNEVEN WITH THE SURFACE NOTED TO BE GLOSSY WITH STRIATIONS AND RESIDUE THROUGHOUT THE SURFACE. UPON INFUSION, A LEAK WAS ALSO OBSERVED FROM THE FUNNEL. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FLUID LEAK ISSUE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED CATHETER EXPULSION AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT ARE UNKNOWN. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRY DATE: 11/2022). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02046 | 2023-03-27 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HEMATURIA FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 2551H20 |  | IT WAS REPORTED THAT THE CATHETER BALLOON BURST. ON THE NEXT DAY AFTER THE TURP PROCEDURE, WHEN THE USER ATTEMPTED TO REMOVE THE FIXATION WATER FROM THE BALLOON IN THE WARD, THE BALLOON BURST WAS FOUND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02042 | 2023-03-27 | 2022-09-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT A PATIENT'S RECENT CATHETER HAD BURST INSIDE THEIR BLADDER (BATCH# NGGT2360). ALSO STATED THAT SEVERAL CATHETER BALLOON WERE DEFLATED INSIDE THEIR BLADDER PREVIOUSLY (BATCH# NGFX3156 AND BATCH# NGFV3848). THE CUSTOMER WANTED TO MAKE THE REPRESENTATIVE AWARE IN CASE OF A PRODUCTION OR BATCH ERROR. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 23MAR2023, STATED THAT THE BATCH # OF NGT2360 WAS SPLIT, BATCH # OF NGFX3156 WAS DEFLATED AND BATCH # OF NGFV3848 WAS DEFLATED. THE BALLOON APPEARED TO BE CLEAN SPLIT. THE CUSTOMER WAS UNSURE IF ANY MATERIAL LEFT IN BLADDER, ESCALATED TO GENERAL PRACTITIONER WHO DIDN¿T FEEL APPROPRIATE FOR A ULTRASOUND SCAN TO CONFIRM THIS UNLESS HE BECOMES SYMPTOMATIC. ON (B)(6) 2023, THE PREVIOUS CATHETER INSERTED WITH BATCH # NGGT2360 SPLIT, ON (B)(6) 2023 THE BALLOON DEFLATED FOR WHICH THEY WERE UNSURE OF LOT NUMBER ON THIS ONE AS NOT DOCUMENTED WHEN IT WAS INSERTED AND ON (B)(6) 2022, THE CATHETER BALLOON WITH NGFV3848 DEFLATED. THEY WERE UNSURE IF BALLOON EITHER DEFLATED OR SPLIT FIRST AND WAS UNABLE TO DETERMINE. UPON ARRIVAL THE CATHETER WAS ALREADY OUT DUE TO FALLING OUT AND ALREADY HAD SPIT AS PER THE PICTURE PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02047 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT 24 OF THE INTERMITTENT CATHETERS PACKAGES WERE EMPTY (NO WATER INSIDE). | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED ISSUE. DEVICE EVALUATED BY MFR: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-02051 | 2023-03-27 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT UPON OPENING THE PACKAGE, THE CATHETER TIP WAS FOUND KINK AND THE CATHETER WAS UNABLE TO BE INSERTED. THE CATHETER WAS NOT USED ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02045 | 2023-03-27 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT. STATED THAT THERE MIGHT BE A CRACK IN THE BALLOON. PER FOLLOW UP INFORMATION RECEIVED ON 14MAR2023, STATED THAT THE EVENT WAS NOT BALLOON BURST ISSUE. THERE WERE NO OBVIOUS VISIBLE DEFECTS TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01004 | 2023-03-28 | 2023-03-02 | I, n, j, u, r, y | Manufacturer report | GROSHONG 4F SINGLE-LUMEN PICC (60CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7717405 |  | IT WAS REPORTED THAT THIS PATIENT WAS DIAGNOSED WITH ACUTE PROMYELOCYTIC LEUKEMIA AND HAD A GROSHONGPICC CATHETER INSERTED BLINDLY FROM THE LEFT MEDIAN CUBITAL VEIN BY A QUALIFIED PICC PLACEMENT NURSE AT THE HEMATOLOGY DEPARTMENT ON (B)(6), 2022. 48 CM OF THE CATHETER WAS INSERTED INTERNALLY AND ANOTHER 7 CM WAS EXPOSED EXTERNALLY. THE CATHETER TIP WAS POSITIONED IN THE SUPERIOR VENA CAVA. THE PATIENT UNDERWENT CHEMOTHERAPY AND CATHETER MAINTENANCE PERIODICALLY ON THE DOCTOR¿S ORDER, WHEN NO DISCOMFORT SYMPTOMS WERE COMPLAINED OF, SUCH AS CHEST PAIN, CHEST DISTRESS, PALPITATION AND ANHELATION. ON (B)(6), 2023, THE PATIENT WENT BACK HOSPITAL FOR CHEMOTHERAPY AND CHEST CT SHOWED CORD-LIKE HIGH-DENSITY SHADOW IN THE VASCULAR LUMEN OF AN OUTER BASAL SEGMENT OF THE INFERIOR LOBE IN THE RIGHT LUNG. BASED ON THE CLINICAL SITUATION, RESIDUES OF BROKEN PICC CATHETER WERE CONSIDERED. THE CARDIOTHORACIC SURGERY DEPARTMENT AND THE VASCULAR INTERVENTION DEPARTMENT WERE CONSULTED IMMEDIATELY. THE PICC CATHETER WAS REMOVED ON THE DOCTOR¿S ORDER AND APPROXIMATELY 19.5 CM WAS MISSING FROM THE PROXIMAL CATHETER. DURING RETRIEVAL OF THE RESIDUAL PORTION, THE PATIENT DEVELOPED DISCOMFORTS SUCH AS PALPITATION AND ANHELATION. THE RETRIEVAL WAS CEASED THEREFORE AND FAILED EVENTUALLY. THE RESIDUAL PORTION IS STILL IN THE INFERIOR LOBE OF THE RIGHT LUNG OF THIS PATIENT. EXPERTS¿ CONSULTATION AND TREATMENT ARE AWAITING. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02102 | 2023-03-28 | 2023-03-08 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT STATED THAT PUREWICK FEMALE EXTERNAL CATHETER GAVE AN INFECTION. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CAUSED THE INFECTION AND UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. WARNINGS: TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-02103 | 2023-03-28 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT THE STAFF FOUND THE PATIENT¿S TEMPERATURE WAS UNDERSTATED BY 2 DEGREES. STATED THAT THE PATIENT COMPLAINED OF FEELING HOT, THEY DOUBLE CHECKED THEIR TEMPERATURE, AND FOUND THAT THE TEMPERATURE SENSING FOLEY CATHETER WAS INCORRECT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01018 | 2023-03-28 | 2023-01-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6903270 |  | IT WAS REPORTED THAT SOMETIME POST CATHETER PLACEMENT, THE PROXIMAL END OF CATHETER ALLEGEDLY KINKED. IT WAS FURTHER REPORTED THAT SURGICAL INTERVENTION REQUIRED MORE THAN ONCE IN SAME PATIENT. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 11/2023. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02100 | 2023-03-28 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE STERILIZED WATER THAT WAS INJECTED DURING PRETEST DID NOT COME OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01019 | 2023-03-28 | 2022-12-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6903270 |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT, THE PROXIMAL EXTENSION OF THE CATHETER ALLEGEDLY KINKED, WHERE IT'S NEEDED TO CLAMP IT AND HEALTH CARE PROVIDER INSTRUCTED TO CHANGE THE CLAMP. IT WAS FURTHER REPORTED THAT SURGICAL INTERVENTION WAS REQUIRED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 11/2023. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01021 | 2023-03-28 | 2022-11-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6903270 |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT, PHYSICIAN REPORTS THAT THE PROXIMAL END OF THE CATHETER WAS ALLEGEDLY KINKED. IT WAS FURTHER REPORTED THAT THE PATIENT HAD TO SUFFER A NEW SURGERY PROCEDURE TO IMPLANT THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 08/2023). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01020 | 2023-03-28 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602830 | 0602830 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, DEVICE ALLEGEDLY HAD FRACTURE IN THE CONNECTION. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 07/2026). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02105 | 2023-03-28 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT WHEN THE OPERATOR USED A THREE-LUMEN CATHETER TO INSERT IT TO THE PATIENT'S URETHRA, THE DRAINAGE WAS NOT SMOOTH AND THEN INSERTED IT TO THE PATIENT'S BLADDER. AFTER USING IT FOR A PERIOD OF TIME, THE BALLOON RUPTURED AND THE PATIENT FELT SEVERE PAIN. THE PATIENT WAS GIVEN PAIN RELIEF AND REPLACED THE SPARE STERILE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02099 | 2023-03-28 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ALL SILICONE CATHETER | UNKNOWN SILICONE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT THE CUSTOMER HAD AN ISSUE WITH A FOLEY CATHETER LAST WEEK. STATED THAT WHEN ONE OF THEIR NURSES DEFLATED THE FOLEY BALLOON, A RIDGE FORMED NEAR THE TIP OF THE CATHETER. UNFORTUNATELY, THEY DID NOT HAVE THE FOLEY TRAY PACKAGE THAT WAS USED. HOWEVER, THEIR OPERATING ROOM (OR) CLINICAL MANAGER WAS ABLE TO DUPLICATE THE PROBLEM WITH ANOTHER TRAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02092 | 2023-03-28 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® URETHRAL CATHETER | URETHRAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 802410 |  | IT WAS REPORTED THAT CUSTOMER STATED INTERMITTENT CATHETER WAS INCORRECT AND THERE WAS DISCREPANCY ON THE DIAMETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-02096 | 2023-03-28 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER THAT ENTERED THE BALLOON DURING THE PRETEST COULD NOT BE DRAINED. WHEN THE SYRINGE WAS PULLED WHILE PUSHING THE CUFF BY HAND, THE STERILIZED WATER WAS FINALLY RECOVERED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02095 | 2023-03-28 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT DURING THE PRETEST, STERILE WATER LEAKED FROM AROUND THE VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02094 | 2023-03-28 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE CATHETER WAS PLACED IN AN OPERATION, AND IT CAME OUT SEVERAL DAYS AFTER RETURNING TO THE WARD. THE BALLOON WAS DEFLATED AND THERE WAS NO STERILE WATER. THERE WAS NO TEAR OR RUPTURE OBSERVED ON THE BALLOON. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6), 2022, IT WAS STATED THAT DUE TO WATER LEAK THE CATHETER FELL OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02088 | 2023-03-28 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT ON (B)(6) 2023, AT AROUND 10 AM, THE INSERTION OF THE URINARY CATHETER WAS PERFORMED DURING THE OPERATION. ON (B)(6) 2023, AT 1:20 AM THE CATHETER WAS FOUND TO HAVE FELL OUT FROM THE BODY. THERE WAS NO LIQUID IN THE AIR BAG, AND A SLOW AIR LEAK WAS FOUND WHEN IT WAS INFLATED AGAIN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02093 | 2023-03-28 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT THE BALLOON CUFF WAS TORN AND WAS NOT USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
16627124 | 2023-03-28 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® LUBRICATH® FOLEY CATHETER | CATHETER, COUDE | C. R. BARD, INC. | A899916 | A899916 |  | UPON PLACING A URINARY CATHETER PRIOR TO A SCHEDULE C-SECTION, I NOTICED A BLUE FLECK OF WHAT APPEARED TO BE PLASTIC IN THE COLLECTION TUBE WHILE CHECKING FOR URINE FOR CORRECT PLACEMENT. A BLUE PIECE OF SOMETHING SEMI-TRANSPARENT WAS IMMEDIATELY FLOATING IN THE URINE. IT HAD A DARKER BLUE LINE DOWN THE MIDDLE. IT WAS ABOUT THE SIZE OF AN ERASER TIP. I POINTED IT OUT TO THE NURSE HELPING ME PLACE THE CATHETER AS SOON AS I SAW IT. THERE WERE NO DIFFICULTIES INSERTING IT, SO IT WAS NOT AN OUTSIDE CONTAMINANT. I TOOK PHOTOS OF THE PACKAGING AND I TOOK A PHOTO OF THE PLASTIC LATER THAT DAY WHEN I EMPTIED THE CATHETER DURING RECOVERY. I ASKED THE PATIENT IF IT LOOKED FAMILIAR TO HER IN CASE IT WAS SOMETHING THAT COULD BE EXPLAINED. I TOLD A FEW NURSES AT THE DESK AND THEY GUESSED THAT IT WAS PROBABLY PART OF THE PLASTIC SHEATH THAT THE CATHETER TIP IS COVERED WITH PRIOR TO PLACEMENT. THIS MADE SENSE BUT I AM NOT 100% SURE. EITHER WAY THIS COULD POTENTIALLY HARM A PATIENT OR CLOG A CATHETER OR POTENTIALLY AFFECT TEST RESULTS. IT NEEDS TO BE INVESTIGATED. |  | N | Y |  |  | 2023-03
1018233-2023-02087 | 2023-03-28 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119416M | 119416M |  | IT WAS REPORTED THAT UPON INSERTION OF FOLEY, NURSE WAS UNABLE TO INFLATE BALLOON, ABLE TO INSERT 2 ML AND THE INFLATION PORT ENLARGED AT THE HUB. ALSO STATED THAT PROPER INSERTION TECHNIQUE WAS USED, URINE RETURNED AND CATHETER WAS INSERTED ALL THE WAY INTO THE HUB. NURSE REMOVED THE CATHETER BALLOON STILL WOULD NOT INFLATE. NEW FOLEY WAS PLACED AND THE SENSOR DID NOT WORK. THIS HAPPENED ON (B)(6) 2023. THEN AGAIN SAME ISSUES WERE NOTED WITH THE PROBES, CORDS SWITCHED AND CONNECTIONS CHECKED, BUT DID NOT FIX. IT WAS STATED THAT THE MACHINE STOPPING THERAPY BECAUSE IT CANNOT OBTAIN AN ACCURATE TEMPERATURE FROM THE BARD TEMP SENSING FOLEY. FOLEY WAS TRANSMITTING TO THE MONITOR, HOWEVER IT WAS UNABLE TO TRANSMIT TO THE ARCTIC SUN MACHINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02083 | 2023-03-28 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER WAS PLACED ON (B)(6) EVENING AND THE CATHETER CAME OUT AND WAS PULLED OUT AROUND 4:00 AM ON FEBRUARY 28. ON (B)(6), WHEN WATER WAS INJECTED INTO THE CATHETER THAT HAD BEEN PULLED OUT, THE BALLOON WAS INFLATED WITHOUT ANY PROBLEMS. AFTER THAT, WHEN IT WAS LEFT AS IT WAS, IT WAS IN A WITHERED STATE ON (B)(6). IT WAS INFLATED ON (B)(6). THERE WAS NO TEAR OR RUPTURE OBSERVED ON THE BALLOON AND THE WATER LEAKED FROM THE DRAINAGE EYE/ LEAKAGE TO INLET TUBE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02081 | 2023-03-28 | 2023-03-07 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THESE PUREWICK FEMALE EXTERNAL CATHETERS CAUSED URINARY TRACT INFECTION, BUT THE PATIENT THOUGHT IT WAS BECAUSE THEY WERE NOT CHANGING IT VERY OFTEN. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. PER CLARIFICATION MAIL RECEIVED ON 09MAR2023, STATED THAT IT SEEMS LIKE THE PATIENT WAS WEARING THE WICK LONGER THAN RECOMMENDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-02080 | 2023-03-28 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER THAT ENTERED THE BALLOON COULD NOT BE DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02082 | 2023-03-28 | 2023-03-07 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WAS NOT USING THE PUREWICK FEMALE EXTERNAL CATHETER BECAUSE THEY KEPT GETTING YEAST INFECTION. THE PATIENT HAD BEEN USING THIS PRODUCT FOR MORE THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE YEAST INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-02091 | 2023-03-28 | 2023-02-16 | I, n, j, u, r, y | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT WHEN THE OPERATOR USED A THREE-LUMEN CATHETER TO INSERT IT TO THE PATIENT'S URETHRA, THE DRAINAGE WAS NOT SMOOTH AND THEN INSERTED IT TO THE PATIENT'S BLADDER. AFTER USING IT FOR A PERIOD OF TIME, THE BALLOON RUPTURED AND THE PATIENT FELT SEVERE PAIN. THE PATIENT WAS GIVEN PAIN RELIEF AND REPLACED THE SPARE STERILE CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
16627219 | 2023-03-28 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BROVIAC 4.2 F SINGLE-LUMEN CV CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 0600520 | 0600520 |  | DURING ROUTINE LINE ASSESSMENT, BLOOD WAS NOTED TO BE ON THE PATIENT'S GOWN. BLEEDING WAS NOTED TO BE ORIGINATING FROM A PART OF THE LINE THAT IS MORE DISTAL FROM THE PATIENT. IT WAS DISCOVERED THAT THERE WAS A CRACK ON THE LINE AND WAS DUE TO CLAMPING THE LINE INAPPROPRIATELY. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01022 | 2023-03-28 | 2022-11-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUISTREAM SPLIT TIP | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6903270 |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT PROCEDURE, THE PROXIMAL END OF THE CATHETER WAS ALLEGEDLY KINKED. IT WAS FURTHER REPORTED THAT NEW SURGICAL INTERVENTION WAS REQUIRED FOR CATHETER IMPLANT. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 11/2023. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02108 | 2023-03-28 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS UNABLE TO BE DEFLATED. USER HAD ALREADY CUT THE INFLATION ARM, BUT NO FLUID CAME FROM THE BALLOON. REPRESENTATIVE DISCUSSED UTILIZING A STYLET OR GUIDEWIRE. REPRESENTATIVE ADVISED THAT MORE INVASIVE TECHNIQUE WOULD BE FOR UROLOGIST TO DEFLATE THE BALLOON WITH A TRANSURETHRAL OR TRANSVAGINAL BALLOON PUNCTURE AND CYSTOSCOPY. REPRESENTATIVE ALSO DISCUSSED MULTIPLE NURSING TEXTBOOKS HAVE SUGGESTED PROCEDURES AND THAT WERE ALWAYS DEFAULT TO HOSPITAL PROTOCOL AND TO CONSULT UROLOGIST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02107 | 2023-03-28 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53812G | 53812G |  | IT WAS REPORTED THAT THESE MAGIC 3 GO MALE CATHETERS WERE NO LUBRICATED ENOUGH. NO ADDITIONAL INFORMATION PROVIDED. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-01046 | 2023-03-28 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT "LINE PLACED ON (B)(6) 2023 AT A FACILITY. ON 2/28/23 NOTED TO BE LEAKING NEAR HUB. UPON FURTHER INSPECTION 21G LUMEN NOTED TO BE PARTIALLY DETACHED FROM HUB. THIS WAS ONLY EVIDENT WHEN HORIZONTAL TENSION PLACED ON LUMEN. PICC LINE DISCONTINUED DUE TO MALFUNCTION. PICC LINE WAS DISCONTINUED AND PATIENT RECEIVED MULTIPLE PIV ATTEMPTS DUE TO DIFFICULT ACCESS. ALSO PLACED RISK FOR INFECTION." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01045 | 2023-03-28 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT W/GUARDIVA AND PROBE COVER | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108DGP |  | IT WAS REPORTED VIA MS&S: CALLER STATES SHE IS WORKING WITH A PATIENT THAT HAS POWERPICC S4153108DGP (LOT NUMBER REGY0370) AND CAN FLUSH INTO THE PICC EASILY BUT CANNOT GET A BLOOD RETURN. CALLER ASKING IF WE CAN PROVIDE CATHFLO TO HELP WITH THE OCCLUSION AND HOW TO ADMINISTER CATHFLO. MS&S RESPONSE: EXPLAINED THAT A FIBRIN TAIL CAN FORM NEAR THE TOP OF THE CATHETER. WHEN THE PICC IS FLUSHED WITH SALINE THE FIBRIN TAIL IS NOT AN ISSUE. WHEN ASPIRATED FOR A BLOOD RETURN THE FIBRIN TAIL CAN BE DRAWN INTO THE CATHETER TIP OCCLUDING THE LUMEN. CONFIRMED THE CATHFLO CAN BE USED TO HELP ALLEVIATE OCCLUSIONS CAUSED BY FIBRIN TAILS. EXPLAINED SHE WOULD NEED TO FOLLOW THE INSTRUCTIONS FOR CATHFLO USAGE. EXPLAINED THAT BD DOES NOT PROVIDE CATHFLO TO CUSTOMERS. SUGGESTED SHE CALL THE DOCTOR AND DISCUSS FURTHER. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01047 | 2023-03-28 | 2023-01-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH DIALYSIS CATHETER 12F X 12.5CM (PRE-CURVED) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5554120 |  | IT WAS REPORTED THAT THIS PATIENT HAD THE SHORT DIALYSIS CATHETER PLACED ON (B)(6) 2023. AFTER THE PENETRATION WAS COMPLETED, THE GUIDE WIRE COULD NOT BE REMOVED NORMALLY. A NEW CATHETER WAS INSERTED IN THIS PATIENT EVENTUALLY. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01043 | 2023-03-28 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC HF CATHETER W/ SHERLOCK 3CG (TPS) STYLET 5F MAX BARRIER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1385108QD |  | IT WAS REPORTED BY THE CUSTOMER, DILATOR WAS BLOCKED AND WOULDN'T LET WIRE PASS THROUGH. 03/20/2023 ADDITIONAL INFORMATION PROVIDED EASILY ADVANCED DILATOR OVER THE WIRE BUT MET RESISTANCE WHEN ATTEMPTING TO PASS DILATOR INTO THE SKIN AT INSERTION SITE DUE TO RESISTANCE. UPON INSPECTION THE DILATOR WAS NOTED TO HAVE IRREGULAR TIP WHICH WOULD NOT ALLOW FURTHER ADVANCEMENT OF DILATOR. ADDITIONAL DILATOR OBTAINED AND PICC LINE PLACEMENT COMPLETED WITHOUT FURTHER PROBLEMS. PROCEDURE BEING PERFORMED WAS A PICC PLACEMENT. CATHETER SECUREMENT DEVICE SORBAVIEW CONTOUR DRESSING, MASTISOL, AND SECUREPORT IV. PROCEDURE WAS COMPLETED AFTER OBTAINING A NEW KIT. NO PATIENT HARM OTHER THEN DISCOMFORT, ADDITIONAL LIDOCAINE WAS USED AT THE INSERTION SITE. PATIENT WAS BEING TREATED FOR AN INFECTION. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01039 | 2023-03-28 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE PROXIMAL END OF THE INTRODUCER SHEATH WAS CRACKED AND A NEW ONE WAS UNPACKED AND USED. NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01040 | 2023-03-28 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D2 |  | IT WAS REPORTED "WE EXPERIENCED LEAKING FROM AROUND THE WIRE AS FLUSH WAS DONE THROUGH THE SIDE PORT EVEN AFTER THE HUB CONNECTION WAS CHECKED. THE FLUID IS COMING DIRECTLY FROM THE WIRE PUNCTURE THROUGH THE RUBBER INJECTION SITE." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01038 | 2023-03-28 | 2023-03-10 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE PICC CATHETER WAS INSERTED IN THIS PATIENT AT THE END OF LAST DECEMBER. EVIDENT LIQUID LEAKAGE WAS OBSERVED DURING THE CHEMOTHERAPY IN EARLY MARCH. THE CATHETER WAS REMOVED AND TWO EVIDENT LEAKING SITES WERE NOTED NEAR THE INSERTION SITE. EXTRAVASATION OF CHEMOTHERAPY DRUGS OCCURRED IN THIS PATIENT, CAUSING OBVIOUS ULCERATION ABOVE THE ELBOW AND BENEATH THE INSERTION SITE. THE CUSTOMER WAS UNSATISFIED WITH THE DIFFICULT TREATMENT AND ADVERSE IMPACT. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01035 | 2023-03-28 | 2023-02-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED "THE PATIENT WAS INFECTED AFTER A NECROTIC METATARSAL AMPUTATION OF THE PROXIMAL FOURTH METATARSAL BONE OF THE LEFT FOOT. ACCORDING TO THE DOCTOR'S ADVICE, IN ORDER TO REDUCE LOCAL INFECTION AND INFLAMMATION CAUSED BY REPEATED PUNCTURE, THE PICC CATHETER WAS INSERTED ON (B)(6) 2023, AND THE OPERATION WAS NORMAL UNTIL (B)(6) 2023. DURING THIS PERIOD, THE PATIENT WAS TREATED WITH PICC. AT 23:50 ON (B)(6) 2023, THE PICC CATHETER WAS FOUND TO HAVE SPONTANEOUSLY BROKEN, AND THE NURSE ON DUTY IMMEDIATELY CHECKED BY THE BED, IMMEDIATELY REMOVED THE BROKEN RESIDUAL CATHETER AT 23:55 ON (B)(6) 2023, FOLLOWED BY LOCAL DRESSING CHANGE AND PRESSURE DRESSING. THE NEXT DAY, THE PATIENT WAS FOUND TO HAVE NO OBVIOUS LOCAL BLOOD INFILTRATION, WHICH DID NOT CAUSE ADVERSE EFFECTS ON THE PATIENT." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01036 | 2023-03-28 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT AFTER INSERTION OF THE CATHETER, A PINHOLE WAS FOUND ON THE CATHETER. THE DEVICE WAS REMOVED AND ANOTHER NEW CATHETER WAS REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01033 | 2023-03-28 | 2023-03-13 | I, n, j, u, r, y | Manufacturer report | BARD POWERMIDLINE CATHETER BASIC KIT (3F) (SINGLE-LUMEN) (20CM) | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P6153118 |  | IT WAS REPORTED THAT A PATIENT HAS EXPERIENCED MILD TO SEVER ALLERGIC REACTIONS TO MIDLINE INSERTIONS APPROX 4 DAYS POST INSERTION. SYMPTOMS INCLUDE SKIN RASH. THESE EVENTS OCCURRED OVER THE PAST 6 MONTHS. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01042 | 2023-03-28 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7957505J |  | IT WAS REPORTED THAT RESISTANCE WAS FELT WHEN THE MICROINTRODUCER WITH SHEATH WAS INSERTED. WHEN THE MICROINTRODUCER WITH SHEATH WAS CHECKED, THE SHEATH WAS PARTLY TURNED UP. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01030 | 2023-03-28 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 20G X 10 CM WITH PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F220108PTX |  | IT WAS REPORTED BY THE CUSTOMER: "ATTEMPTED MIDLINE CATHETER INSERTION USING ULTRASOUND GUIDANCE IN THE LET CEPHALIC VEIN. WHEN ATTEMPTING TO DEPLOY GUIDEWIRE THEN CATHETER, I MET RESISTANCE, WHICH WAS ALSO HURTING THE PATIENT. SO I STOPPED, PULLED CATHETER BACK, PULLED GUIDEWIRE BACK, THEN REMOVED NEEDLE FROM HIS ARM. INSPECTED THE NEEDLE/CATHETER/GUIDEWIRE AND FOUND THAT PART OF THE CATHETER BROKE OFF FROM THE NEEDLE. I USED THE ULTRASOUND AND WAS ABLE TO VISUALIZE THE CATHETER PARTIALLY IN THE CEPHALIC VEIN. PUTTING A MIDLINE IN FOR ANTIBIOTIC ADMINISTRATION. I DID CONDUCT A BRIEF CHART REVIEW. THAT PATIENT UNDERWENT IR PROCEDURE ON 3/8 FOR RETRIEVAL OF THE CATHETER FRAGMENT. THIS PROCEDURE WAS SUCCESSFUL WITHOUT COMPLICATION. PATIENT HAS SINCE DISCHARGED HOME WITH HOME HEALTH. NO RECOMMENDATION FOR FURTHER FOLLOW-UP BY IR. " | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02109 | 2023-03-28 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | MAGIC 3 COUDE INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE NOT LUBRICATED ENOUGH. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR, VISCOMETER FAILURE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-03
3006260740-2023-01031 | 2023-03-28 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD TRIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED "THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. THE EVENT INVOLVED A TL POWERTRIALYSIS PLACED IN THE NECK." NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION AND SPECIFIC PATIENT DETAILS WERE NOT DISCLOSED AS PART OF THE SURVEY. AS SUCH, ADDITIONAL INFORMATION AND THE SUBJECT SAMPLE CANNOT BE OBTAINED.THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01024 | 2023-03-28 | 2023-03-09 | I, n, j, u, r, y | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY THE CUSTOMER THAT "CUSTOMER ATTEMPTED TO REMOVE ACCUCATH IV AND WIRE WAS TEMPORARILY LEFT IN THE PATIENT. WIRE WAS SUCCESSFULLY REMOVED.". | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01023 | 2023-03-28 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D |  | IT WAS REPORTED BY THE CUSTOMER THAT "CUSTOMER IS SAYING THAT PART OF THE INTRODUCER BROKE OFF. STOPPED PROCEDURE. HAD TO DO SAFETY CHECKS. WILL NEED TO RE-INSERT PICC." ADDITIONAL EVENT INFORMATION RECEIVED: "PICC LINE WAS BEING PLACED IN PATIENT, DILATOR (INTACT) THREADED OVER GUIDEWIRE AND ENTERED VESSEL WITH NO RESISTANCE. PICC ATTEMPTED TO BE THREADED BUT WAS UNABLE TO PASS SHOULDER. INNER CANNULA OF DILATOR WAS RE-INSERTED INTO PATIENT AND THEN REMOVED WITH OUTER SHEATH AS ONE PIECE. AT THIS TIME, IT WAS NOTED THAT TIP OF INNER DILATOR CANNULA WAS NOT INTACT. RN LOOKED ON STERILE TRAY AND DRAPE AND DID NOT SEE PIECE. PT HAD TO STAY IN FACILITY ONE EXTRA DAY FOR WORK UP TO FIND POTENTIAL PIECE OF CATHETER. PATIENT HAD X-RAY AND CT SCAN AS WELL. THE LINE WAS UNABLE TO BE PLACED." | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01025 | 2023-03-28 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED "THERE WAS A HOLE IN PICCLINE. THE PICC LINE THAT LEAKED FROM THE INSERTION SITE DURING FLUSHING BEFORE CYTOTHERAPY. THE PICC LINE WAS DISMANTLED AND SAW A HOLE ABOUT 1.2 CM INSIDE THE INSERTION SITE. THE PICC LINE HAS BEEN IN PLACE FOR A LONG TIME AND THE CLINICIAN "HOPES THAT IS WHAT COULD BE THE CAUSE". SET OCTOBER 2021 AND WITHDRAWN MARCH 2023. IT WAS ANCHORED WITH SECURACATH. PATIENT RECEIVED A SINGLE CYCLE WITH FOLFOX (CYT: OXALIPLATIN + FLUOROURACIL) WHEN IT WAS NEWLY ADDED IN OCTOBER 2021 AND THEN UNTIL IT WAS DISCONTINUED IN 230306 IMMUNOTHERAPY WAS GIVEN: NIVOLUMAB." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02110 | 2023-03-28 | 2023-03-09 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | MAGIC 3 INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT THE PATIENT THOUGHT THAT THEY GOT URINARY TRACT INFECTIONS BECAUSE OF USING THESE MAGIC 3 GO HYDROPHILIC MALE INTERMITTENT CATHETER. REPRESENTATIVE ADVISED THAT GETTING URINARY TRACT INFECTIONS WAS NORMAL WHEN USING THE CATHETERS, BUT THE PATIENT SAID THEY WERE USING THE MAGIC 3 GO CATHETERS AND THE PATIENT NEVER GOT URINARY TRACT INFECTION IN 2 YEARS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY.". THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
3006260740-2023-01027 | 2023-03-28 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET (4F) SL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000819 |  | IT WAS REPORTED BY THE CUSTOMER, "BROWN DISCOLORATION AFTER USING SCISSORS. THE DISCOLORATION ON THE RING DISAPPEARED WITHIN A COUPLE HOURS. WHILE TRIMMING STATSEAL DISC BROWN DISCOLORATION WAS NOTED TO COME OFF SCISSORS. NEW RING AND SCISSORS OBTAINED BUT PICC HAD ALREADY BEEN PLACED AND SCISSORS USED TO TRIM CATHETER. STATLOCK WAS USED TO SECURE PICC. NO ISSUES AS OF 3/8/2023 WILL CONTINUE TO MONITOR. THERE WAS NO CONCOMITANT THERAPY AS EVENT OCCURRED WITH PICC PLACEMENT." | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01029 | 2023-03-28 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED "THE PATIENT WAS HOSPITALIZED WITH THE DIAGNOSIS OF 'MALIGNANT BREAST TUMOR', AND PERFORMED PICC CATHETERIZATION ACCORDING TO THE DOCTOR'S ADVICE ON (B)(6) 2023. WHEN THE NURSE ON DUTY PREPARED MATERIALS FOR PUNCTURE, IT WAS FOUND THAT THE PUNCTURE SHEATH OF THE CENTRAL VENOUS CATHETER KIT FOR PERIPHERAL INTUBATION HAD A TEAR, AND IMMEDIATELY REPLACED THE CENTRAL VENOUS CATHETER KIT AND ACCESSORIES FOR PERIPHERAL INTUBATION FOR THE PATIENT, BUT THERE WERE STILL CRACKS IN THE PUNCTURE SHEATH AFTER THE NEW REPLACEMENT. CAUSING PATIENTS TO QUESTION. THE NURSE ON DUTY IMMEDIATELY CHANGED AGAIN AND EXPLAINED THE SITUATION TO THE PATIENT." THIS REPORT ADDRESSES THE SECOND CATHETER KIT USED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01028 | 2023-03-28 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED "THE PATIENT WAS HOSPITALIZED WITH THE DIAGNOSIS OF 'MALIGNANT BREAST TUMOR', AND PERFORMED PICC CATHETERIZATION ACCORDING TO THE DOCTOR'S ADVICE ON (B)(6) 2023. WHEN THE NURSE ON DUTY PREPARED MATERIALS FOR PUNCTURE, IT WAS FOUND THAT THE PUNCTURE SHEATH OF THE CENTRAL VENOUS CATHETER KIT FOR PERIPHERAL INTUBATION HAD A TEAR, AND IMMEDIATELY REPLACED THE CENTRAL VENOUS CATHETER KIT AND ACCESSORIES FOR PERIPHERAL INTUBATION FOR THE PATIENT, BUT THERE WERE STILL CRACKS IN THE PUNCTURE SHEATH AFTER THE NEW REPLACEMENT. CAUSING PATIENTS TO QUESTION. THE NURSE ON DUTY IMMEDIATELY CHANGED AGAIN AND EXPLAINED THE SITUATION TO THE PATIENT." THIS REPORT ADDRESSES THE FIRST CATHETER KIT USED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01026 | 2023-03-28 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PPICC 4F DL 3CG BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY MEDWATCH THAT "THE "SHERLOCK" DEVICE USED TO PLACE BEDSIDE PICC LINES ISN'T READING THE TIP OF THE PICC CORRECTLY. THIS MEANS THE PATIENT HAS TO HAVE A CHEST X-RAY TO DETERMINE THE PICC TIP." THE DEPARTMENT IS VASCULAR ACCESS. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02126 | 2023-03-29 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE VICINITY OF THE BIFURCATION AREA WHEN THE USER INSERTED A SYRINGE INTO THE INFLATION VALVE AND INJECTED STERILIZED WATER. THE EVENT OCCURRED JUST AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BIFURCATION AREA AND THE VALVE MENTIONED AS BACKFLOW. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01049 | 2023-03-29 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT PRIOR TO A PORT PLACEMENT PROCEDURE, A SLICE WAS ALLEGEDLY NOTICED IN THE PRODUCT WRAPPER. REPORTEDLY, STERILITY WAS COMPROMISED. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 10/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01053 | 2023-03-29 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/GUARDIVA, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD3 |  | IT WAS REPORTED BY THE CUSTOMER, "REGARDING PICC INSERTION KITS THAT RECENTLY HAVE TRANSITIONED TO NEWER STYLE KITS WITH A DIFFERENT EXTENSION SET WITH THE WIRE EXTENDING THROUGH THE PORT. THE NEW STYLE T-PORT HAS A MORE CONCAVE STYLE IN WHICH THE WIRE FEEDS THROUGH. AIR ENTERED THE CATHETER AT THE T-PORT INTO THE PICC AND PULLED BACK AIR AS THEY RETRACTED THE T-PORT SYRINGE. THE PICC WAS IMMEDIATELY REMOVED OR CLAMPED AND MANAGED WITHOUT PATIENT HARM." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02140 | 2023-03-29 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 7283 | 37302 |  | IT WAS REPORTED THAT THE CUSTOMER THOUGHT THAT THEY WERE DOING SOMETHING WRONG WHEN THEY WERE UNABLE TO CONNECT A BAG TO THE MALE EXTERNAL CATHETER. THE REPRESENTATIVE WENT TO HELP AND FOUND A FAULTY SHEATH. ONE OF THE MALE EXTERNAL CATHETER WAS FOUND TO BE HAVING THE OUTLET ON THE FUNNEL OF THE SHEATH CLOSED OFF. THE OUTLET SHOULD BE OPEN TO ALLOW URINE TO DRAIN. THEY CHECKED ALL OTHER SHEATHS IN THE BOX AND FOUND THAT OTHER SHEATHS WERE OKAY. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2022, IT WAS STATED THAT THE DEFECT FOUND WHILE BEING USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-02128 | 2023-03-29 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165L16 | 0165L16 |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEVERAL PRODUCT FAILURES SPECIFICALLY WITH THE BALLOON EASILY BREAKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02129 | 2023-03-29 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165L18 | 0165L18 |  | IT WAS REPORTED THAT THERE HAVE BEEN THREE LEAKS IN THE PAST FEW DAYS WITH THE FOLEY CATHETER TRAY. STATED THAT THE NURSES WHO REPORTED THIS SAID THEY THOUGHT THE URINE WAS LEAKING AROUND THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02127 | 2023-03-29 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L20 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS DIFFICULT TO REMOVE. THE ACTUAL PRODUCT ANALYSIS WAS REQUESTED IN CASES WHERE THE CATHETER WAS DIFFICULT TO REMOVE, NOT IN CASES WHERE IT WAS DIFFICULT TO REMOVE WATER. PER ADDITIONAL INFORMATION RECEIVED ON (B)(6)2023, STATED THAT CALCIUM-LIKE DEPOSITS WERE OBSERVED AROUND THE TIP. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02125 | 2023-03-29 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FIXATION WATER LEAKED FROM THE FUNNEL BIFURCATION OF THE FOLEY CATHETER. THE EVENT OCCURRED JUST AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BIFURCATION AREA. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01048 | 2023-03-29 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY BROKEN INTO TWO PIECES. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY HAD A LEAKAGE. REPORTEDLY, THE CATHETER WAS REMOVED, AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 05/2024. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02114 | 2023-03-29 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE USER FELT MORE RESISTANCE THAN USUAL WHEN INJECTING STERILE WATER FROM THE INFLATION VALVE OF THE FOLEY CATHETER DURING THE PRETEST. THE STERILE WATER DID NOT RETURN TO THE SYRINGE WHEN THE STERILE WATER WAS WITHDRAWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02113 | 2023-03-29 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI12 | 0165SI12 |  | IT WAS REPORTED THAT AFTER INDWELLING, THE CATHETER WAS PULLED OUT AND THE BALLOON PART WAS FOUND RUPTURED DURING THE PATIENT OBSERVATION. THE EVENT OCCURRED 6 HOURS AFTER PLACEMENT AND BALLOON TEAR/RUPTURE WAS OBSERVED. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02115 | 2023-03-29 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | MAGIC 3 INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THESE INTERMITTENT CATHETER WERE HARD TO INSERT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OUTSIDE DIAMETER IS TOO LARGE ROUGH OR IRREGULAR SHAPE PROTRUSIONS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER.". | N | Y | LAY USER/PATIENT |  | 2023-03
1018233-2023-02117 | 2023-03-29 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1715J12 |  | IT WAS REPORTED THAT THE TIP OF THE FOLEY CATHETER BALLOON WAS BENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02112 | 2023-03-29 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED, THAT THE WATER LEAKED FROM THE SHAFT OF THE FOLEY CATHETER UPON INJECTING STERILE WATER. AFTER, THE CATHETER PLACEMENT TO THE PATIENT IN THE EMERGENCY ROOM. THE EVENT OCCURRED JUST AFTER THE PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE SHAFT. IT WAS STATED, THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00208 | 2023-03-29 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQ5054 | CQ5054 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE BALLOON WAS ALLEGEDLY BURST DURING PRESSURIZATION WITH PRESSURE PUMP. IT WAS FURTHER REPORTED THAT THE BLOOD VESSEL WAS DAMAGED AND RECOVERED AFTER COMPRESSION TO STOP BLEEDING. REPORTEDLY, DAMAGED BALLOON WAS REMOVED AND NEW BALLOON WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRATION DATE: 01/2025). DEVICE EVALUATED BY MFR: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01050 | 2023-03-29 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER CLAMP WAS ALLEGEDLY BROKEN. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 05/2023. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02124 | 2023-03-29 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE USER FELT MORE RESISTANCE THAN USUAL WHEN INJECTING STERILE WATER FROM THE INFLATION VALVE OF THE FOLEY CATHETER DURING THE PRETEST. THE STERILE WATER DID NOT RETURN TO THE SYRINGE WHEN THE STERILE WATER WAS WITHDRAWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00207 | 2023-03-29 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A CALCIFIED TARGET LESION IN THE ANTERIOR TIBIAL ARTERY BELOW THE KNEE VIA AN IPSILATERAL APPROACH, THE TIP OF THE CATHETER ALLEGEDLY SEPARATED. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 09/2023). DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01065 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THEY PLACED 7 PICCS YESTERDAY AND ALL OF THE DLS HAD SMALL TO SCANT AMOUNTS OF LEAKING THROUGH THE STOPPER AROUND THE WIRE WHEN FLUSHED." THIS REPORT ADDRESSES THE SECOND DEVICE. | DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01067 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THEY PLACED 7 PICCS YESTERDAY AND ALL OF THE DLS HAD SMALL TO SCANT AMOUNTS OF LEAKING THROUGH THE STOPPER AROUND THE WIRE WHEN FLUSHED." THIS REPORT ADDRESSES THE FOURTH DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01066 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THEY PLACED 7 PICCS YESTERDAY AND ALL OF THE DLS HAD SMALL TO SCANT AMOUNTS OF LEAKING THROUGH THE STOPPER AROUND THE WIRE WHEN FLUSHED." THIS REPORT ADDRESSES THE THIRD DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01073 | 2023-03-30 | 2023-03-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD 4FR S/L PER-Q-CATH | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED THAT WHEN PREFLUSHING THE CATHETER PRIOR TO PLACEMENT, THE NURSE FOUND A SANDHOLE WITH THE CATHETER AND IMMEDIATELY CONTACTED US AND RETAINED THE PROBLEMATIC PRODUCT. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01060 | 2023-03-30 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | THE FOLLOWING WAS REPORTED: "PATIENT REFERRED TO FACILITY FOR PICC FRACTURE BY DISTRICT NURSING TEAM AS PICC LEAKING AT SITE WHEN FLUSHED WITH SALINE." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01074 | 2023-03-30 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH DIALYSIS CATHETER KIT 12F X 20CM | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5553200 |  | IT WAS REPORTED THAT WHEN THE GUIDE WIRE NEEDS TO BE PULLED OUT AFTER THE PATIENT'S PUNCTURE, THE GUIDE WIRE CANNOT BE PULLED OUT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01069 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THEY PLACED 7 PICCS YESTERDAY AND ALL OF THE DLS HAD SMALL TO SCANT AMOUNTS OF LEAKING THROUGH THE STOPPER AROUND THE WIRE WHEN FLUSHED." THIS REPORT ADDRESSES THE SIXTH DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01076 | 2023-03-30 | 2023-03-13 | I, n, j, u, r, y | Manufacturer report | FULL KIT 20G X 10 CM WITH GUARDIVA AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F120108PT |  | IT WAS REPORTED BY CUSTOMER THAT "POWERGLIDE TIP CATHETER DAMAGE- TIP BROKE OFF IN PATIENT. DEVICE WAS "LOOSE" OUT OF PACKAGE PER RN." IT WAS REPORTED BY THE CUSTOMER THAT "PT IS SUPER HARD STICK WITH DVT TO RUE AND L SUBCLAVIAN COMPLETELY OCCLUDED. HAD PIV IN R AC THAT INFILTRATED. ATTEMPTED LINE IN SMALL VEIN 6CM DISTAL TO L AC. VEIN IMMEDIATELY TIED INTO A LARGER VEIN BUT WOULDN'T HAVE HAD QUITE ENOUGH SPACE IF I STUCK WHERE IT GOT LARGER SO TRIED JUST BELOW IT. VEIN WAS APPROX 1/2 CM DEEP SO WENT IN W ALMOST A FLAT ANGLE. WAS A LITTLE OFF TO ONE SIDE SO HAD TO ADJUST ANGLE SL. LOOKED LIKE I WAS IN VEIN AND THOUGHT I HAD A LITTLE BLOOD RETURN IN CATHETER SO SLID THE GUIDEWIRE OFF. DIDN'T SEEM TO HAVE ANY RESISTANCE. WHEN I TRIED TO SLIDE OFF THE CATHETER, IT ONLY WENT ABOUT 5CM AND MET RESISTANCE. I PULLED IT BACK ONTO NEEDLE WITHOUT DIFFICULTY BUT WAS UNABLE TO DISENGAGE THE GUIDEWIRE TO PULL IT BACK. THE BUTTON WOULDN'T EVEN PUSH DOWN. SO, I SLID THE CATHETER BACK OFF AS FAR AS IT WOULD GO AND REMOVED THE HOUSING WITHOUT DIFFICULTY. NOTED THE GUIDEWIRE HAD A BEND TO IT. WAS UNABLE TO ASPIRATE ANY BLOOD SO SLOWLY REMOVED CATHETER WHILE ATTEMPTING TO ASPIRATE. NEVER GOT BLOOD RETURN AND GOT TO WHERE I THOUGHT END OF CATHETER WAS AND IT STUCK! WHEN I FLUSHED IT, IT LEAKED OUT OF INSERTION SITE AND UNDER THE SKIN SOME. TOURNIQUET WAS STILL ON TO TRY TO KEEP VEIN AT MAX SIZE. COULD PUSH THE CATHETER BACK IN ABOUT 1CM BUT AGAIN STUCK WHEN TRIED TO PULL IT OUT. TRIED JUST HOLDING GENTLE TRACTION ON CATHETER AND COUNTER TRACTION ON SKIN, TRIED LIGHTLY MASSAGING THE TISSUE AROUND INSERTION SITE. ROTATED CATHETER SLIGHTLY BACK AND FORTH. ON ULTRASOUND IT LOOKED LIKE CATHETER TIP WAS JUST IN VEIN BUT AT 5 O'CLOCK POSITION. CATH LAB NURSE CAME TO HELP. TRIED THE SAME THINGS I TRIED. EVENTUALLY WAS ABLE TO REMOVE IT. NEVER APPLIED FORCE TO IT. WHEN IT CAME OUT WE CHECKED THE TIP AND IT LOOKED SHEARED AND THEN WHEN COMPARED TO UNUSED CATHETER AND MEASURED, THERE WAS 1.5 CM MISSING. SENT PT TO ER WITH TOURNIQUET ON. THEY DID A DOPPLER AND DIDN'T THINK THERE WAS A FOREIGN BODY IN THE VEIN BUT SAID THERE APPEARED TO BE A THROMBUS NEAR INSERTION SITE. XRAY SAYS THERE IS 1.5CM FOREIGN BODY IN THE SOFT TISSUE. AND THERE WAS A THROMBUS TO RADIAL AND CEPHALIC VEINS AND " VASCULAR CALCIFICATIONS THROUGHOUT THE SOFT TISSUES". ON XRAY HIS ARM LOOKS LIKE IT HAS ROPES IN IT THERE IS SO MUCH CALCIFICATION AND THROMBUS. THE VEIN I WAS ATTEMPTING WAS EASILY COMPRESSIBLE. ER DOC TOLD HIM TO FOLLOW UP WITH HIS SURGEON FOR POSSIBLE REMOVAL." ADDITIONAL INFORMATION 03/21/2023 PATIENT DID NOT RECEIVE AN IV FOR LONG-TERM ANTIBIOTICS. UNKNOWN TREATMENT OR OUTCOME FOLLOWING ER. PATIENT WAS BEING TREATED FOR DIABETIC FOOT ULCER/LONG TERM IV ANTIBIOTICS. RELEVANT HISTORY: PT HAS HAD NUMEROUS PICCS BILATERALLY. HAD KNOWN DVTS IN R ARM. HAD KNOWN TOTAL OCCLUSION TO LEFT SUBCLAVIAN WITH ONLY COLLATERAL BLOOD FLOW. PT'S LAST PIV WAS IN R AC AND INFILTRATED WHICH RESULTED IN NEED FOR IV ACCESS ON THIS DATE. VENOUS DOPPLER STUDY OF LEFT ARM IN ER (POST ATTEMPTED INSERTION) SHOWED THROMBUS TO RADIAL AND CEPHALIC VEINS AND " VASCULAR CALCIFICATIONS THROUGHOUT THE SOFT TISSUES". PT IS NEAR END OF OPTIONS FOR IV ACCESS. HAS EJECTION FRACTION OF 9% PER WIFE. SURGEON REPORTED PT NOT A CANDIDATE FOR PORT A CATH PLACEMENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01070 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THEY PLACED 7 PICCS YESTERDAY AND ALL OF THE DLS HAD SMALL TO SCANT AMOUNTS OF LEAKING THROUGH THE STOPPER AROUND THE WIRE WHEN FLUSHED." THIS REPORT ADDRESSES THE SEVENTH DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01064 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THEY PLACED 7 PICCS YESTERDAY AND ALL OF THE DLS HAD SMALL TO SCANT AMOUNTS OF LEAKING THROUGH THE STOPPER AROUND THE WIRE WHEN FLUSHED." THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. INITIAL MEDWATCH REPORT WAS SUBMITTED, UPON FURTHER REVIEW IT WAS FOUND THAT THIS MEDWATCH REPORT 3006260740-2023-01064 IS A DUPLICATE FILE AND HAS BEEN VOIDED. THE ORIGINAL EVENT WAS SUBMITTED ON MEDWATCH REPORT 3006260740-2023-02260. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00212 | 2023-03-30 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U41502H15RX | U41502H15RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A HIGHLY CALCIFIED TARGET LESION IN THE ANTERIOR TIBIAL ARTERY VIA AN IPSILATERAL ANTEGRADE APPROACH, THE PTA BALLOON WAS ALLEGEDLY UNABLE TO CROSS THE LESION. IT WAS FURTHER REPORTED THE SHAFT OF THE CATHETER ALLEGEDLY BENT AND CONTRAST LEAKED FROM THE CATHETER LUMEN WHEN THE PHYSICIAN PUSHED AND TRIED TO MAKE THE CATHETER CROSS THE LESION. REPORTEDLY, THE FAILURE TO INFLATE WAS FOUND OUTSIDE THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ULTRAVERSE RX PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, MULTIPLE KINKS THROUGHOUT THE CATHETER SHAFT AND BUNCHING OF THE BALLOON WERE OBSERVED. THE SAMPLE WAS INADVERTENTLY SEVERED INTO TWO SEGMENTS AND THUS FUNCTIONAL TESTING COULDN'T BE PERFORMED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED MATERIAL DEFORMATION AS MULTIPLE KINKS THROUGHOUT THE CATHETER SHAFT AND BUNCHING OF THE BALLOON WERE OBSERVED. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED INABILITY TO CROSS LESION AND LEAK AS FUNCTIONAL TESTING COULD NOT BE PERFORMED. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED INABILITY TO CROSS LESION, MATERIAL DEFORMATION AND LEAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: (EXPIRATION DATE: 10/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
2020394-2023-00213 | 2023-03-30 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U41502H4RX | U41502H4RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A HIGHLY CALCIFIED TARGET LESION IN THE ANTERIOR TIBIAL ARTERY VIA AN IPSILATERAL ANTEGRADE APPROACH, THE PTA BALLOON ALLEGEDLY RUPTURED UNDER THE RATED BURST PRESSURE. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONLY THE OUTER PACKAGE WAS RETURNED AND THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. PER THE REPORTED EVENT DETAILS, THE TARGET LESION WAS HIGHLY CALCIFIED. THEREFORE, IT IS POSSIBLE THE INTERACTION BETWEEN THE LESION AND THE BALLOON CONTRIBUTED TO THE REPORTED BALLOON RUPTURE. HOWEVER, THE DEFINITIVE ROOT CAUSE FOR THE REPORTED BALLOON RUPTURE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 01/2025), G3. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01059 | 2023-03-30 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D |  | IT WAS REPORTED BY THE CUSTOMER ¿WHILE INSERTING PICC-AFTER ONE PASS AND MEETING RESISTANCE. 3CG WIRE WAS REMOVED FROM PATIENT AND TIP WAS NOTED TO BE BENT AT A 90-DEGREE ANGLE. NEW 3CG WIRE WAS ACQUIRED AND USED FOR COMPLETION OF PROCEDURE.¿ NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01058 | 2023-03-30 | 2023-02-28 | I, n, j, u, r, y | Manufacturer report | POWER PORT MRI ISP, 8 FR. GROSHONG, INT W SP, ATTACHABLE SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808560 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY HAD BROKEN. IT WAS FURTHER REPORTED THAT PATIENT ALLEGEDLY EXPERIENCED EXTRAVASATION AND ADDITIONAL SURGERY NEEDED TO TREAT PATIENT. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01072 | 2023-03-30 | 2023-01-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716000J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE GUIDEWIRE WAS ALLEGEDLY DIFFICULT TO INSERT THROUGH THE PUNCTURE NEEDLE. IT WAS FURTHER REPORTED THAT THE GUIDEWIRE WAS ALLEGEDLY STRETCHED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE INTRODUCER NEEDLE LOADED WITH A J-TIP GUIDEWIRE AND A GUIDEWIRE STRAIGHTENER IN A GUIDEWIRE HOOP WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC VISUAL, FUNCTIONAL AND DIMENSIONAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO INSERT, AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED AND THE SAMPLE EVALUATION RESULTS INDICATING DIFFICULTY UNDER LABORATORY CONDITIONS ARE NOT BY THEMSELVES SUFFICIENT TO CONFIRM THAT THIS EVENT DID NOT OCCUR UNDER CLINICAL CONDITIONS. HOWEVER, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED STRETCHED, IDENTIFIED FRACTURE, PHYSICAL RESISTANCE AND MATERIAL PROTRUSION ISSUES AS THE GUIDEWIRE WAS NOTED TO BE UNCOILED ON A DISTAL END PORTION AND THE UNCOILED GUIDEWIRE FELT STUCK WITHIN THE INTRODUCER NEEDLE. THE FLAT CORE WIRE WAS NOTED TO HAVE A COMPLETE BREAK AND WAS NOTED TO BE PROTRUDING THE UNCOILED GUIDEWIRE AREA. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 04/2023) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01062 | 2023-03-30 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6CF INT W SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806060 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE THE CATHETER ALLEGEDLY BENDS IN PLACEMENT. IT WAS FURTHER REPORTED THE CATHETER DOESN'T HAVE RETURN. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 04/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01071 | 2023-03-30 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6CF INT W SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806060 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY BENDS IN PLACEMENT. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 04/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01063 | 2023-03-30 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6CF INT W SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806060 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE THE CATHETER ALLEGEDLY BENDS IN PLACEMENT. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 04/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01068 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D3 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THEY PLACED 7 PICCS YESTERDAY AND ALL OF THE DLS HAD SMALL TO SCANT AMOUNTS OF LEAKING THROUGH THE STOPPER AROUND THE WIRE WHEN FLUSHED." THIS REPORT ADDRESSES THE FIFTH DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01078 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC SOLO 4FR S/L, 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 22194118 |  | IT WAS REPORTED THAT 4FR SL POWERPICC SOLO, SMALL LEAK OBSERVED APPROXIMATELY 5CM FROM INSERTION SITE, SOME CHEMOTHERAPY FROM THE INFUSER PUMP LEAKED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01061 | 2023-03-30 | 2022-12-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER "PATIENT ATTENDED FACILITY WITH LEAKING FROM SITE WHEN PICC WAS FLUSHED. PICC DISCARDED ONCE REMOVED." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01075 | 2023-03-30 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D |  | IT WAS REPORTED BY THE CUSTOMER ¿WHILE ATTEMPTING PICC INSERTION, ADVANCED PICC THROUGH INTRODUCER RESISTANCE WAS MET. PULLED CATHETER BACK AND CHECKED INTEGRITY OF CATHETER AND 3CG WIRE. THE WIRE WAS BENT AT ALMOST 90 DEGREE ANGLE ABOUT 4 CM FROM THE TIP. WIRE HAD TO BE REPLACED.¿ NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02145 | 2023-03-30 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119416M | 119416M |  | IT WAS REPORTED THAT UPON INSERTION OF FOLEY, NURSE WAS UNABLE TO INFLATE BALLOON, ABLE TO INSERT 2 ML AND THE INFLATION PORT ENLARGED AT THE HUB. ALSO STATED THAT PROPER INSERTION TECHNIQUE WAS USED, URINE RETURNED AND CATHETER WAS INSERTED ALL THE WAY INTO THE HUB. NURSE REMOVED THE CATHETER BALLOON STILL WOULD NOT INFLATE. NEW FOLEY WAS PLACED AND THE SENSOR DID NOT WORK. THIS HAPPENED ON (B)(6)2023. THEN AGAIN SAME ISSUES WERE NOTED WITH THE PROBES, CORDS SWITCHED AND CONNECTIONS CHECKED, BUT DID NOT FIX. IT WAS STATED THAT THE MACHINE STOPPING THERAPY BECAUSE IT CANNOT OBTAIN AN ACCURATE TEMPERATURE FROM THE BARD TEMP SENSING FOLEY. FOLEY WAS TRANSMITTING TO THE MONITOR, HOWEVER IT WAS UNABLE TO TRANSMIT TO THE ARCTIC SUN MACHINE . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02146 | 2023-03-30 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT ON THE NEXT DAY AFTER PLACEMENT, THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. IT WAS A OBSTETRICS AND GYNECOLOGY PATIENT AND THE PATIENT EXPERIENCED DYSURIA AFTER DELIVERY. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6)2022, IT WAS ALSO STATED THAT THE DYSURIA WAS UNRELATED TO SPONTANEOUS CATHETER REMOVAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02150 | 2023-03-30 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119110 | 119110 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DUE TO WATER LEAK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02148 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE CATHETER WAS INSERTED BEFORE OPERATION AND WHEN THE PATIENT MOVED THE BODY AFTER OPERATION, THE CATHETER FELL OFF OUT OF THE PATIENT SPONTANEOUSLY AND WAS FOUND THAT THE BALLOON HAS BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02153 | 2023-03-30 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE VICINITY OF THE BIFURCATION AREA WHEN THE USER INSERTED A SYRINGE INTO THE INFLATION VALVE AND INJECTED STERILIZED WATER. THE EVENT OCCURRED JUST AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BIFURCATION AREA AND THE VALVE MENTIONED AS BACKFLOW. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02154 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT IT WAS DIFFICULT TO DRAIN WATER FROM THE CATHETER AS THE SYRINGE STOPPED WORKING. THE USER STATED THAT A PRESSURE WAS FELT MORE THAN USUAL WHEN INJECTING WATER DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01077 | 2023-03-30 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC SV 3F SL,BASIC, CE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 66173118 |  | IT WAS REPORTED THAT SMALL VEIN 3FR PICC LEAKING UNDER SECURE-A-CATH SITE. INSERTED 8TH FEB AND REMOVED 10TH MARCH. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED. ONE 3 FR POWERPICC SV WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUES THROUGHOUT THE RETURNED DEVICE, AND DRIED ADHESIVE-LIKE MATERIAL WAS OBSERVED IN THE VICINITY OF THE SUGGESTED SPLIT. A TACTILE EVALUATION OF THE RETURNED CATHETER REVEALED THE ADHESIVE MATERIAL SPANNING ABOUT 3 CM LONG WAS HARDENED. FUNCTIONAL TESTING OF THE DEVICE INVOLVED PRESSURIZING WATER INTO THE COMPLAINANT CATHETER BY CLAMPING THE END OF THE CATHETER AND USING A 12 ML SYRINGE TO INFUSE WATER INTO THE NEEDLELESS INJECTION CAP. UPON PRESSURIZATION OF THE DEVICE, A LEAK WAS FOUND BETWEEN THE 9 CM AND 10 CM DEPTH MARKERS WHERE THE ADHESIVE WAS LOCATED. IT TOOK SIGNIFICANT FORCE TO IDENTIFY THE SPLIT DUE TO THE AMOUNT OF ADHESIVE IN THE VICINITY OF THE SPLIT. AN ATTEMPT TO WIPE THE ADHESIVE MATERIAL OFF OF THE CATHETER WITH A GERMICIDAL WIPE WAS UNSUCCESSFUL. A MICROSCOPIC OBSERVATION REVEALED DIVOTS IN THE ADHESIVE AROUND THE AREA OF THE LEAK, AND FLATTENING OF THE CATHETER WAS OBSERVED NEAR THE LEAK SITE; HOWEVER, DUE TO THE AMOUNT OF ADHESIVE IN THE VICINITY OF THE SPLIT, IT WAS DIFFICULT TO IDENTIFY THE FEATURES OF THE SPLIT. THE FLATTENING OF THE CATHETER SHAFT AT THE SPLIT SITE SUGGESTED THAT CATHETER COMPRESSION/KINKING LIKELY CAUSED/CONTRIBUTED TO THE OBSERVED CATHETER DAMAGE. POSSIBLE CONTRIBUTING FACTORS INCLUDE INSERTION, MAINTENANCE AND SECUREMENT TECHNIQUES. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02160 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT DURING PRETEST, THE WATER DID NOT COME OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02157 | 2023-03-30 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT 4 DAYS AFTER PLACEMENT, THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. THE BALLOON WAS DEFLATED, AND THE DIAPER WAS WET, AS THE FIXATION WATER HAD LEAKED OUT. THERE WAS NO TEAR/RUPTURE ON THE BALLOON. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER AND IT WAS INSERTED BY HAND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
16643339 | 2023-03-30 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP FOLEY TRAY SYSTEM BARD® LUBRI-SIL FOLEY CATHETER TRAY | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | A942216 | A942216 |  | NARRATIVE FROM STAFF: BARD FOLEY CATHETER NOTED TO HAVE LEAK ALONG SEAM OF THE CATHETER TUBE WHEN INSERTED IN PATIENT. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02164 | 2023-03-30 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50000000 | UNK |  | IT WAS REPORTED THAT UPON INSERTION OF FOLEY, NURSE WAS UNABLE TO INFLATE BALLOON, ABLE TO INSERT 2 ML AND THE INFLATION PORT ENLARGED AT THE HUB. ALSO STATED THAT PROPER INSERTION TECHNIQUE WAS USED, URINE RETURNED, AND CATHETER WAS INSERTED ALL THE WAY INTO THE HUB. NURSE REMOVED THE CATHETER BALLOON STILL WOULD NOT INFLATE. NEW FOLEY WAS PLACED AND THE SENSOR DID NOT WORK. THIS HAPPENED ON (B)(6) 2023. THEN AGAIN SAME ISSUES WERE NOTED WITH THE PROBES, CORDS SWITCHED, AND CONNECTIONS CHECKED, BUT DID NOT FIX. IT WAS STATED THAT THE MACHINE STOPPING THERAPY BECAUSE IT CANNOT OBTAIN AN ACCURATE TEMPERATURE FROM THE BARD TEMP SENSING FOLEY. FOLEY WAS TRANSMITTING TO THE MONITOR, HOWEVER IT WAS UNABLE TO TRANSMIT TO THE ARCTIC SUN MACHINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02152 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT IT WAS DIFFICULT TO DRAIN WATER FROM THE CATHETER AS THE SYRINGE STOPPED WORKING. THE USER STATED THAT A PRESSURE WAS FELT MORE THAN USUAL WHEN INJECTING WATER DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02141 | 2023-03-30 | 2023-03-08 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175814 |  | IT WAS REPORTED THAT THE FOLEY CATHETER DID NOT DRAIN QUICKLY ENOUGH OR INCOMPLETE IN QUANTITY. IT SEEMED TO BE OBSTRUCTED. WHEN USING DISPOSABLE CATHETERS, THE BLADDER DID EMPTY COMPLETELY. THE CONTROL WAS PERFORMED WITH A BLADDER SCAN. PREVIOUSLY THERE WERE NO COMPLAINTS BUT SINCE 2 WEEKS COMPLAINTS INCREASED. THE IMPLICATIONS WAS POSSIBLE OVERFILLING OF THE BLADDER DUE TO FASTER FILLING OF THE BLADDER THAN EMPTYING VIA THE CATHETER. THE RISKS WERE PATIENT COLLAPSE OR BLADDER RUPTURE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02165 | 2023-03-30 | 2023-02-21 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETERS WERE PAINFUL. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. MEDICAL INTERVENTION WAS UNKNOWN. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE PATIENT WAS CURRENTLY NOT USING THE PUREWICK IN ORDER TO CLEAR UP THE UTI. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-02166 | 2023-03-30 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WAS NO OUTFLOW OF URINE FROM FOLEY CATHETER INTO THE BAG DURING SURGERY. THE PATIENT HAD A SWOLLEN BLADDER AND HAD TO REPLACE THE CATHETER. NO MEDICAL INTERVENTION WAS REPORTED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 23MAR2022, IT WAS CONFIRMED THAT THE NO OUTFLOW OF URINE" ISSUE WAS WITH CATHETER. THE INFORMATION REGARDING THE NO OUTFLOW OF URINE FROM CATHETER CAUSED SWELLING IN BLADDER WEREN'T PROVIDED FROM CUSTOMER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02167 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT ON INJECTING WATER INTO THE CATHETER, THE WATER WAS DIFFICULT TO ENTER AND IT LEAKED FROM THE INFLATION VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01055 | 2023-03-30 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK TLS TECH 5F MAX | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9275108D |  | BD MEDICAL SERVICES RECEIVED THE FOLLOWING VIA PHONE CALL: DOCTOR STATES HE JUST PLACED A POWERPICC IN A PATIENT. DOCTOR STATES HE TRIMMED THE CATHETER WHILE THE WIRE WAS STILL IN. WHEN HE REMOVED THE WIRE, IT WAS STILL 50CM IN LENGTH SO NO WIRE WAS LEFT IN THE PATIENT HOWEVER IT APPEARS SOME OF A COATING WAS SHAVED OFF AND IS INSIDE THE CATHETER IN THE PATIENT. DOCTOR IS ASKING FOR THE MATERIAL OF THE COATING OR IF WE KNOW HOW FLOPPY IT IS TO KNOW IF HE WOULD BE ABLE TO REMOVE IT WITH A SNARE. BD MEDICAL SERVICES RESPONDED "INFORMED DOCTOR THE POWERPICC KIT CONTAINS A FLEXURA NITINOL GUIDEWIRE. THE MATERIAL IS NITINOL AND IS NOT COATED. INFORMED DOCTOR UNSURE WHAT "COATING" WOULD HAVE CAME OFF OR IF THE WIRE ITSELF WAS SHAVED OFF. INFORMED DOCTOR WE HAVE NO INFORMATION REGARDING A POTENTIAL SHAVING OF THE GUIDEWIRE IN A PATIENT AND REMOVING IT OR IF HE SHOULD REPLACE THE ENTIRE CATHETER. HOW TO PROCEED IS A CLINICAL DECISION. ADDITIONAL INFORMATION 3/19 DURING FOLLOW UP PHONE CALL WITH CUSTOMER: ¿"FIXED THE PROBLEM: 2 COMMENTS HAS STIFLER WIRING. FORGOT TO PULL THE WIRE BACK BEFORE CUTTING, CUTTING THE OUTER SHEATH OFF OF IT. THE WIRE TOO SHORT TO ACT AS LEAD WIRE. ADVANCED STIFLLER WIRE TO INSERT CATHETER AND CAUGHT DISTAL END OF END AND SHEATH BROKE OFF. SURGEON REMOVED THE END OF THE SHEATH." CATHETER WAS FLUSHED PRIOR TO REMOVING THE STYLET. NO ADVERSE EFFECTS TO THE PATIENT. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02178 | 2023-03-30 | 2023-02-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123614A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PLACED FOR ABOUT 23 HOURS, THE PATIENT HEARD A VIBRATING SOUND FROM THE ABDOMEN, AND THE URINARY CATHETER FELL OUT FROM THE URETHRAL OPENING. AS PER THE DESCRIPTION OF EVENT CAUSE ANALYSIS, STATED THAT THERE WAS SPONTANEOUS RUPTURE OF THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02173 | 2023-03-30 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 996081 | 996081 |  | IT WAS REPORTED THAT THE GAUGE BROKE OFF ON EAGLE INFLATION DEVICE FROM X-FORCE BALLOON KIT. | THE REPORTED EVENT IS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION AND FURTHER INVESTIGATION WAS NOT CONCLUSIVE. THOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, POTENTIAL ROOT CAUSES FOR THIS EVENT COULD BE, INADEQUATE SPECIFICATIONS OR INADEQUATE SUPPLIER. THE EXACT ROOT CAUSE COULD NOT BE IDENTIFIED. ALTHOUGH A ROOT CAUSE COULD NOT BE DEFINITIVELY IDENTIFIED, BASED ON THE RISK DOCUMENTATION REVIEW, A POTENTIAL ROOT CAUSE FOR THIS TYPE OF FAILURE COULD BE INAPPROPRIATE PACKAGE DESIGN. HOWEVER, THERE WAS INSUFFICIENT INFORMATION TO CONFIRM THIS POTENTIAL ROOT CAUSE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: BEFORE USE, INSPECT THE DEVICE TO VERIFY THAT NO DAMAGE HAS OCCURRED DURING SHIPPING AND HANDLING. UDI FORMAT UPDATED WITH AVAILABLE PRODUCT INFORMATION PER GUDID. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02177 | 2023-03-30 | 2023-03-09 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT NOW HAD BLISTERS ON THEIR VAGINAL AREA SINCE USING THE PUREWICK FEMALE EXTERNAL CATHETER AND URINARY TRACT INFECTION. THEY STOPPED USING THE PRODUCTS AND THE BLISTERS AND THE URINARY TRACT INFECTION HAD NOW GONE AWAY. THE PATIENT WAS GIVEN ANTIBIOTICS BY DOCTOR. THE PATIENT HAD BEEN USING THIS PRODUCT FOR MORE THAN 90 DAYS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
1018233-2023-02180 | 2023-03-30 | 2023-02-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BALLOON BIPOLAR TEMPORARY PACING ELECTRODE, 5FR 110CM | BALLOON FLOW-ASSISTED BIPOLAR ELECTRODE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 006173P | 006173P |  | IT WAS REPORTED THAT THE BALLOON OF THE TEMPORARY PACING ELECTRODE CATHETER WAS RUPTURED. STATED THAT THE USER WAS EXPOSED TO HAZARDOUS, BLOOD, OR BODILY FLUIDS. MEDICAL INTERVENTION WAS UNKNOWN. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 20MAR2022, IT WAS STATED THAT THERE WAS NO MISSING PIECES OF BALLOON INSIDE THE PATIENT. IT WAS NOTED THAT THERE WAS NO INJURY AND MEDICAL INTERVENTION WAS REQUIRED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02181 | 2023-03-30 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | A900216 | 119416 |  | IT WAS REPORTED THAT AFTER MOVING PATIENT TO A BED IN THE POST-ANESTHESIA CARE UNIT, TEMPERATURE SENSING FOLEY CATHETER SLID OUT OF PATIENT WHEN NURSE WAS MOVING FOLEY BAG. IT WAS OBSERVED THAT THE BALLOON WAS NOT INFLATED AT THE END OF THE CASE. THE FOLEY BALLOON WAS INFLATED AT BEGINNING OF CASE BECAUSE THERE WAS URINE RETURN. THEY CHECKED THE FOLEY BALLOON TO SEE IF IT HAD A LEAK AND FOUND THAT THE BALLOON WAS INTACT BUT THE INJECTION PORT TO INFLATE THE BALLOON WAS LEAKING SLOWLY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01054 | 2023-03-30 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | A710962 | A710962 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT ALLEGEDLY BLED THROUGH THE CATHETER. HOWEVER, THERE IS NO INFORMATION REGARDING TREATMENT PROVIDED FOR BLEEDING. THE PATIENT'S CURRENT STATUS WAS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 11/2022) THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02215 | 2023-03-31 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS PLACED IN THE OPERATING ROOM, THE CATHETER SPONTANEOUSLY FELL OFF IN THE WARD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01083 | 2023-03-31 | 2023-03-08 | I, n, j, u, r, y | Manufacturer report | BARD POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED THAT THERE WAS REDNESS OR SWELLING AT OR AROUND CENTRAL LINE INSERTION SITE AND THERE IS FLUID LEAKING FROM THE AROUND CENTRAL LINE INSERTION SITE. BD POWERPICC DOUBLE LUMEN IN CHEST. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02220 | 2023-03-31 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT THE FIXED WATER COULD NOT BE INJECTED EVEN THOUGH THE SYRINGE WAS CONNECTED AND INJECTED. WHEN THE USER TRIED WITH ANOTHER PRODUCT, WATER WAS ABLE TO BE INJECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02218 | 2023-03-31 | 2023-02-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT STERILE WATER LEAKS BETWEEN SYRINGE AND VALVE. THERE WAS NO STERILE WATER DURING INSERTION. AS PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 29MAR2023, NO DAMAGE DUE TO LEAK, IT'S NOT THAT THERE'S NO WATER, IT'S THAT THE WATER COULDN'T BE INJECTED. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. THIS INVESTIGATION DOES NOT REQUIRE A DHR REVIEW DUE TO A CLOSURE LETTER NOT REQUIRED FOR THIS COMPLAINT. LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02209 | 2023-03-31 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OFF ON THE DAY AFTER CATHETER PLACEMENT. THERE WAS NO TEAR/ RUPTURE ON THE BALLOON. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. THERE WAS NO PATIENT'S URINE STONE AND THE CATHETER WAS NOT PULLED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, WIRE PRESSING TECHNIQUE, STRIPPING TECHNIQUE ETC. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02224 | 2023-03-31 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE WATER LEAKED WHILE INJECTING STERILE WATER. THE LOCATION OF THE LEAK WAS UNKNOWN. THERE WAS NO BALLOON RUPTURE OR TEAR, AND THE WATER LEAKED FROM AN UNKNOWN LOCATION BECAUSE THE EVENT WAS CATHETER FELL OUT. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER, AND THE CATHETER WAS INSERTED BY HAND. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 16MAR2022, IT WAS STATED THAT THE CATHETER DID NOT FELL OUT OF THE PATIENT. THE CHECK MARK IN THE INQUIRY SHEET WAS MISTAKENLY ENTERED. IT WAS ALSO STATED THAT THE WHILE INFLATING THE BALLOON DID NOT FEEL ANY RESISTANCE. | THE REPORTED EVENT WAS CONFIRMED ¿ UNKNOWN CAUSE. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. BASED ON THE EVALUATION, THE CATHETER WAS INFLATED USING 10ML OF WATER AND OBSERVED WATER LEAKED FROM THE BALLOON. OBSERVED THE APPEARANCE OF THE BALLOON SAC HAS CHANGED COLOR AND OBSERVED SMALL TEAR ON SAC BODY. THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. THE REPORTED EVENT IS CONFIRMED AS UNKNOWN CAUSE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. CORRECTION: F,H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01087 | 2023-03-31 | 2023-01-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR (RADEL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770352 |  | IT WAS REPORTED CORD IS PULLING AWAY/BREAKING AT THE STRAIN RELIEF. (B)(6) 2023 IT WAS FOUND THAT THE USB CABLE IS CUT IN THE MIDDLE OF THE CABLE WITH WIRES SHOWING. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF DAMAGED STRAIN RELIEF WAS CONFIRMED. THE STRAIN RELIEF AND USB CABLE AT THE STRAIN RELIEF ARE NOT DAMAGED BUT THE USB CABLE IS CUT IN THE MIDDLE OF THE CABLE WITH WIRES SHOWING. THE ROOT CAUSE OF THE REPORTED FAILURE IS DUE TO USE OF THE DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01095 | 2023-03-31 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE SENSOR IS NOT PICKING UP THE PICC 75% OF THE TIME. ERRATIC AND INACCURATE READINGS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF THE SENSOR IS NOT PICKING UP THE PICC 75% OF THE TIME. ERRATIC AND INACCURATE READINGS WAS CONFIRMED. DURING TESTING, THE SENSOR PASSED THE ECG FUNCTIONALITY BUT FAILED THE MAGNET TRACKING FUNCTIONALITY. THE SENSOR INTERMITTENTLY DOES NOT DETECT THE MAGNET TRACKING. THE SENSOR WAS CONNECTED TO THE AUTOMATED CALIBRATION FIXTURE AND FAILED WITH 10 SENSORS NOT READING. THE ROOT CAUSE OF THE REPORTED FAILURE WAS IDENTIFIED AS AN INTERNAL FAILURE WITHIN THE MAGNET TRACKING PCB. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01081 | 2023-03-31 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716000J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE VIA THE SUBCLAVIAN VEIN, THE GUIDEWIRE ALLEGEDLY COULD NOT BE INSERTED. IT WAS FURTHER REPORTED THAT THE GUIDEWIRE WAS ALLEGEDLY STRETCHED. REPORTEDLY, THE NEEDLE AND THE GUIDEWIRE WAS REMOVED AS A SINGLE UNIT. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE INTRODUCER NEEDLE AND ONE J-TIP GUIDEWIRE LOADED INTO A GUIDEWIRE HOOP WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC VISUAL, FUNCTIONAL AND DIMENSIONAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO INSERT, AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED. HOWEVER, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED STRETCHED AND IDENTIFIED MATERIAL SEPARATION ISSUES AS COMPLETE BREAKS WERE NOTED TO BOTH THE ROUND AND FLAT CORE WIRES. THE DISTAL PORTION OF THE GUIDEWIRE WAS NOTED TO BE BENT AND UNCOILED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 06/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02223 | 2023-03-31 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT LEAKAGE OCCURRED DURING USE AFTER PLACEMENT OF THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02205 | 2023-03-31 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WERE NO ABNORMALITIES IN THE EXPANSION OF THE BALLOON BEFORE INSERTION, BUT AFTER INSERTION INTO THE PATIENT, THE BALLOON WAS DAMAGED AND SPONTANEOUS WITHDRAWAL OCCURRED FOUR TIMES. ALL THE EVENT WAS CONFIRMED WITH THE SAME LOT AND THE REMAINING ONE WAS COLLECTED WITHOUT BEING USED. THE USER THEN USED 29030J14 DUE TO THE LACK OF 119314. SIMILARLY, THERE WAS A REPORT DRAFTED SEPARATELY FOR SPONTANEOUS WITHDRAWAL DUE TO BALLOON DAMAGE. PER INVESTIGATOR NOTIFICATION RECEIVED VIA EMAIL ON 12MAR2023 STATED THAT PER SAMPLES RECEIVED. 8 SAMPLES WERE RETURNED. 4 SAMPLES NOT IN PACKAGING APPEARED TO BE USED WITH MATERIAL NUMBER 119314 AND MANUFACTURING LOT NUMBER NGGS2164. 4 SAMPLES IN UNOPENED PACKAGING WITH MATERIAL NUMBER 119314, 3 WITH BATCH NUMBER NGGT3722, AND 1 WITH BATCH NUMBER NGGS3548. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 15MAR2023, STATED THAT THERE WAS NO BALLOON RUPTURE. THE CATHETER HAD A BALLOON PINHOLE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3 UNOPENED 3-WAY TEMP SENSING CATHETERS. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE 3 CATHETER BALLOONS WERE INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. NO PINHOLE WAS NOTED ON THE BALLOONS. THE BALLOONS PASSIVELY DEFLATED WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02212 | 2023-03-31 | 2023-01-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT AFTER INJECTING STERILIZED WATER INTO THE BALLOON DURING PRETEST, WATER LEAKAGE FROM THE BALLOON WAS CONFIRMED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02203 | 2023-03-31 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT WHEN THE BALLOON WAS INFLATED DURING THE TEST BEFORE INSERTION, THE SYRINGE COULD BE INSERTED, BUT WATER LEAKED FROM NEAR THE VALVE. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. THE AFFECTED SYRINGE WAS MANUFACTURED BY HANSCENT. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER ¿ HANSCENT. THIS INVESTIGATION DOES NOT REQUIRE A DHR REVIEW DUE TO A CLOSURE LETTER NOT REQUIRED FOR THIS COMPLAINT. THE REPORTED EVENT IS UNDERSTOOD, CHARACTERIZED, AND CONFIRMED AS UNRELATED TO THE LABELING ISSUE. IT WAS CONFIRMED RELATED TO A SUPPLIER, THEREFORE LABELING REVIEW IS NOT REQUIRED FOR THIS REPORTED EVENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02196 | 2023-03-31 | 2023-03-12 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD A URINARY TRACT INFECTION THE FIRST WEEK AFTER USING THE PUREWICK FEMALE EXTERNAL CATHETER. CUSTOMER ALSO STATED THAT THE PUREWICK URINE COLLECTION SYSTEM WAS LOUD. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. WARNINGS: TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
16655004 | 2023-03-31 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARD STRAIGHT CATHETER | CATHETER, UROLOGICAL | C. R. BARD, INC. |  |  |  | WHEN ATTEMPTING TO PERFORM AN INTERMITTENT SELF-CATHETERIZATION (ISC), RN DISCOVERED STRAIGHT CATH TUBING HAD BECOME DISCONNECTED FROM BAG, BREAKING THE CLOSED SYSTEM. FORTUNATELY, THE CATHETER DID NOT REACH THE PATIENT. MANUFACTURER RESPONSE FOR STRAIGHT CATHETER TUBING, BARD PK7648761 (PER SITE REPORTER) IT WAS REPORTED TO THE BARD REP AND AWAITING FURTHER INSTRUCTIONS FROM THEM. |  | N | Y |  |  | 2023-03
1018233-2023-02190 | 2023-03-31 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT'S URINARY CATHETER FELL OUT, AND THE BALLOON FOR FIXATION HAS SHRUNK INTO AN OVAL SHAPE. WITH THE URINATION OBSERVATION AND COMMUNICATION WITH THE PATIENT AND FAMILY MEMBERS FOR PSYCHOLOGICAL COMFORT, THE PATIENT WAS ABLE TO URINATE BY THEMSELVES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02191 | 2023-03-31 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165814 | 165814 |  | IT WAS REPORTED THAT ANOTHER COMPLAINT WAS RECEIVED FROM THE PHARMACY ABOUT THE FOLEY CATHETER IN QUESTION. THE PATIENT REPORTEDLY HAD 4 MORE FAULTY BLADDER PROBE OF THE BARDEX BRAND IN THE PAST FEW WEEKS. THE PROBE WAS INSERTED IN THE MORNING AND CAME OUT BY ITSELF IN THE EVENING. SO, ACCORDING TO THE PHARMACY, THERE MUST BE SOMETHING DEFECTIVE ABOUT THE BALLOON OR THE VALVE TO KEEP THE BALLOON INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02194 | 2023-03-31 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THE BALLOON BURST DURING THE PRETEST. IT WAS MENTIONED THAT IT WAS NOT INFLATED IN BLUE VINYL. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE PINHOLE IN BALLOON OR CUFF. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING AND A SYRINGE. VISUAL INSPECTION NOTED BALLOON BURST MEASURING 0.1860" WITH NO MISSING PIECES NOTED. THIS IS OUT OF SPECIFICATION WHICH STATES, "BALLOON MUST NOT BE TORN". A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02192 | 2023-03-31 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT DURING A PRETEST, ON INJECTING STERILE WATER, THE WATER LEAKED FROM THE INFLATION VALVE. THERE WAS NO BALLOON TEAR/RUPTURE AND THE WATER LEAKED FROM THE VALVE MENTIONED AS BACKFLOW. THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. THE AFFECTED SYRINGE WAS MANUFACTURED BY HANSCENT. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER ¿ HANSCENT. THE DEFORMATION AT THE TIP OF THE SYRINGE IS THE ROOT CAUSE THAT CONTRIBUTED TO THE LEAK; HOWEVER, THE ROOT CAUSE OF THE DEFORMATION IS UNKNOWN. NO FAILURES WERE FOUND WITHIN THE SUPPLIER INVESTIGATIONS. BOTH FLEX AND HANSCENT SHOWED NO CLEAR EVIDENCE THAT THE FAILURES OCCURRED DURING THEIR PROCESSES. HANSCENT DID NOT FIND ANY ROOT CAUSE OF THE LEAKAGE. ALL PROCESSES WITHIN HANSCENT WERE UNDER CONTROL AND NO DEFORMATION AND LEAKAGE WERE FOUND IN PRODUCTION RETAIN SAMPLES, INVENTORY, AND SIMULATION PRODUCTION PRODUCTS, NO LEAKAGE RECORD IN DHR. AS BD WILL NO LONGER BE USING HANSCENT AND FLEX FOR THE SYRINGES, THIS INVESTIGATION WILL BE CONCLUDED AS CAUSE UNKNOWN. PRIME CARE WILL BECOME BD¿S QUALIFIED SUPPLIER. DHR REVIEW IS NOT REQUIRED. A LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02193 | 2023-03-31 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WERE NO ABNORMALITIES IN THE EXPANSION OF THE BALLOON BEFORE INSERTION, BUT AFTER INSERTION INTO THE PATIENT, THE BALLOON WAS DAMAGED AND SPONTANEOUS WITHDRAWAL OCCURRED FOUR TIMES. ALL THE EVENT WAS CONFIRMED WITH THE SAME LOT AND THE REMAINING ONE WAS COLLECTED WITHOUT BEING USED. THE USER THEN USED 29030J14 DUE TO THE LACK OF 119314. SIMILARLY, THERE WAS A REPORT DRAFTED SEPARATELY FOR SPONTANEOUS WITHDRAWAL DUE TO BALLOON DAMAGE. PER INVESTIGATOR NOTIFICATION RECEIVED VIA EMAIL ON 12MAR2023 STATED THAT PER SAMPLES RECEIVED. 8 SAMPLES WERE RETURNED. 4 SAMPLES NOT IN PACKAGING APPEARED TO BE USED WITH MATERIAL NUMBER 119314 AND MANUFACTURING LOT NUMBER NGGS2164. 4 SAMPLES IN UNOPENED PACKAGING WITH MATERIAL NUMBER 119314, 3 WITH BATCH NUMBER NGGT3722, AND 1 WITH BATCH NUMBER NGGS3548. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02189 | 2023-03-31 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT WHEN STERILE WATER WAS INJECTED TO INFLATE THE BALLOON AFTER INSERTING THE CATHETER IN THE OPERATION ROOM, STERILE WATER LEAKED OUT FROM THE VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02186 | 2023-03-31 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS USED DURING THE OPERATION FOR URINE DRAINAGE. IT WAS FOUND THAT THE BLADDER IRRIGATION LUMEN OF THE CATHETER WAS NARROW, NOT REACHING THE STANDARD DIAMETER OF THE CAVITY. THEN A NEW CATHETER WAS IMMEDIATELY REPLACED. THERE WERE NO ADVERSE EVENTS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02185 | 2023-03-31 | 2022-12-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD OR BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND THEIR MEDICAL APPLICATIONS BECAUSE SOMETIMES THE SIZE WAS NOT IDEAL IN RELATION TO BIOCATH FOLEY CATHETER PRECONNECTED TO A BARDIA 2L BEDSIDE BAG, FEMALE LENGTH, FRENCH SIZE 12. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT/ MISSING TRANSLATION; MISSING INSTRUCTIONS; VENDOR/PRINTER ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: THE IFU STATES, "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02206 | 2023-03-31 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WERE NO ABNORMALITIES IN THE EXPANSION OF THE BALLOON BEFORE INSERTION, BUT AFTER INSERTION INTO THE PATIENT, THE BALLOON WAS DAMAGED AND SPONTANEOUS WITHDRAWAL OCCURRED FOUR TIMES. ALL THE EVENT WAS CONFIRMED WITH THE SAME LOT AND THE REMAINING ONE WAS COLLECTED WITHOUT BEING USED. THE USER THEN USED (B)(4) DUE TO THE LACK OF 119314. SIMILARLY, THERE WAS A REPORT DRAFTED SEPARATELY FOR SPONTANEOUS WITHDRAWAL DUE TO BALLOON DAMAGE. PER INVESTIGATOR NOTIFICATION RECEIVED VIA EMAIL ON (B)(6) 2023 STATED THAT PER SAMPLES RECEIVED. 8 SAMPLES WERE RETURNED. 4 SAMPLES NOT IN PACKAGING APPEARED TO BE USED WITH MATERIAL NUMBER (B)(4) AND MANUFACTURING LOT NUMBER NGGS2164. 4 SAMPLES IN UNOPENED PACKAGING WITH MATERIAL NUMBER (B)(6) , 3 WITH BATCH NUMBER (B)(4), AND 1 WITH BATCH NUMBER (B)(4) . PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THERE WAS NO BALLOON RUPTURE. THE CATHETER HAD A BALLOON PINHOLE. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02182 | 2023-03-31 | 2023-03-07 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PUREWICK DISPOSABLE | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED, AND THE NURSES KEPT INSERTING THE PUREWICK FEMALE EXTERNAL CATHETER INTO THE PATIENT VAGINA THE WHOLE STAY THAT THE PATIENT USED THE WICKS ON HOSPITALIZATION THAT LASTED FOR 11 DAYS. THE PATIENT SUFFERED A LOT OF DISCOMFORT ALL THE TIME AND ACQUIRED A URINARY TRACT INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-03
3006260740-2023-01082 | 2023-03-31 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE DEVICE ALLEGEDLY HAD A CRACK IN THE CATHETER EXTENSION LINE BELOW THE HUB. IT WAS FURTHER REPORTED THAT BLOOD ALLEGEDLY LEAKED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
1018233-2023-02187 | 2023-03-31 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT WHEN STERILE WATER WAS INJECTED TO INFLATE THE BALLOON AFTER INSERTING THE CATHETER IN THE OPERATION ROOM, STERILE WATER LEAKED OUT FROM THE VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-03
3006260740-2023-01113 | 2023-04-01 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE 20G X 2.25IN | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | PCK00010 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE "ACCUCATH WIRE GOT SHEARED OFF INTO PATIENT- WHEN HIT SAFETY BUTTON, NEEDLE RETRACTED BUT GUIDEWIRE DID NOT." NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02227 | 2023-04-01 | 2023-03-09 | I, n, j, u, r, y | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER TURNED GREEN, BUT THE PATIENT URINE WAS NOT GREEN IT WAS YELLOW. LET THE PATIENT KNEW THEY MIGHT NEED TO CLEAN THE BAG WITH VINEGAR AND WATER OR SOAPY WATER. ALSO ADVISED PATIENT TO CONTACT DOCTOR'S OFFICE. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE PATIENT USED A DIFFERENT BRAND OF FOLEY CATHETERS. THEY HAD IT CHANGED ON (B)(6) 2023, BUT NOW HAS A UTI AND WAS PRESCRIBED ANTIBIOTICS BY DOCTOR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01110 | 2023-04-01 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S4153108D |  | IT WAS REPORTED BY THE CUSTOMER, "I¿M STILL NOTICING THIS PROBLEM WITH THE INTRODUCERS BASICALLY BUCKLING. WHEN I TRY TO ADVANCE IT, IT DOES GO SMOOTHLY AND IT'S LIKE THE GREY PART OF THE PEEL AWAY SHEATH GETS CAUGHT ON THE SKIN AND RESISTS ADVANCING, CAUSES THE PLASTIC TO CRUMPLE UP. IT'S LIKE THE MATERIAL ISN¿T STIFF ENOUGH, BUT ALSO THE TRANSITION TO THE TAPER WHITE INNER DILATOR IS NOT SMOOTH ENOUGH." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01109 | 2023-04-01 | 2023-03-09 | I, n, j, u, r, y | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 2174108 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE DEVICE WAS USED FOR VENOUS ACCESS FOR CHEMOTHERAPY INFUSION FOR NH LYMPHOMA. ON (B)(6) 2023 THE DEVICE WAS FOUND TO BE FUNCTIONAL, IN FACT, SAMPLING, WASHING AND DRESSING WAS PERFORMED. ON (B)(6) 2023 THE DRESSING WAS WET WITH BLOOD REFLUX IN THE SMALL TUBE, ON ASPIRATION BUBBLES IN THE TUBE, ON WASHING SF LEAKED FROM THE BASE OF THE CATHETER (BETWEEN CONE AND CATHETER); THE SYRINGE USED FOR WASHING IS A 10 CC LL PRE-FILLED SYRINGE). REMOVAL OF DM, AFTER BLOOD CULTURE FOR FEVER AND RAISE INDICATIONS OF INFLAMMATION. THE CATHETER WAS CUT OFF THE TIP AND SENT FOR BACTERIAL EXAMINATION. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01112 | 2023-04-01 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | CUSTOMER REPORTED "ON (B)(6), A PICC 4FR GROSHONG CATHETER WAS INSERTED, AND AT THE END, IT WAS OBSERVED THAT A DROP OF SERUM APPEARED BETWEEN THE TRANSPARENT PART AND THE CATHETER. SEROLIZED IN PUSH AND EXTRAVASATION OBSERVED AGAIN AT THE SAME SITE. DURING THE PASSAGE OF THE CATHETER, NOTHING THAT CALLED ATTENTION WAS OBSERVED. FOR PATIENT SAFETY, AS WE ADMINISTER CHEMOTHERAPY IN OUR SERVICE, I OPENED A NEW CATHETER AND USED A NEW REPAIR." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01108 | 2023-04-01 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1274108D2 |  | IT WAS REPORTED BY THE CUSTOMER WE EXPERIENCED LEAKING FROM AROUND THE WIRE AS FLUSH WAS DONE THROUGH THE SIDE PORT EVEN AFTER THE HUB CONNECTION WAS CHECKED. THE FLUID IS COMING DIRECTLY FROM THE WIRE PUNCTURE THROUGH THE RUBBER INJECTION SITE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01106 | 2023-04-01 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5 F POLYURETHANE CATHETER NURSING BASIC TRAY WITH MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 3175118 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE POWER PICC CATHETER DID NOT COME WITH THE MARKINGS EVERY 1CM, AT THE HEIGHT OF THE 32CM MARKING THERE IS NO MARKING, REAPPEARING AT THE END OF THE CATHETER. THERE IS NO SAMPLE, DUE TO THE CATHETER BEING INSERTED IN THE PATIENT. (B)(6), 2023: ADDITIONAL INFORMATION RECEIVED. MISSING CATHETER MARKINGS WAS NOTICED PRIOR TO USE. NO PATIENT OR CAREGIVER HARM. THERE WAS NO NEED FOR MEDICAL AND/OR SURGICAL INTERVENTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF MISSING DEPTH MARKINGS ON THE CATHETER WAS CONFIRMED AND APPEARS TO BE RELATED TO THE MANUFACTURING PROCESS. A SINGLE PHOTOGRAPH OF A PICC CATHETER WAS RETURNED FOR EVALUATION. THE PROXIMAL END OF THE DEVICE WAS NOT VISIBLE IN THE PHOTOGRAPH, SO THE CONFIGURATION COULD NOT BE VERIFIED, BUT THE SAMPLE APPEARED CONSISTENT WITH THE IMPLICATED 5FR S/L POWERPICC CATHETER. THE CATHETER SHAFT APPEARED TO BE MISSING DEPTH MARKINGS. THE PRINTING DISAPPEARED AT THE 33CM DEPTH MARKING AND THEN REAPPEARED AGAIN NEAR THE DISTAL END OF THE CATHETER. THE EXACT NUMBER OF DEPTH MARKINGS MISSING FROM THE CATHETER COULD NOT BE DETERMINED FROM THE PHOTOGRAPH BUT APPEARED TO BE APPROXIMATELY 19 CENTIMETERS WORTH OR PRINTING. A REVIEW OF THE CATHETER DRAWING CONFIRMED THAT THE CATHETER SHAFT SHOULD CONTAIN CONTINUOUS DEPTH MARKINGS FROM 0CM THROUGH THE 54CM MARKING DOT. AS THE MISSING DEPTH MARKINGS WERE NOTED PRIOR TO PLACEMENT OF THE DEVICE, THE MOST PROBABLE CAUSE IS DUE TO THE MANUFACTURING PRINTING PROCESS. BD IS WORKING CLOSELY WITH MANUFACTURING TO HELP PREVENT RECURRENCE OF THE REPORTED EVENT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01105 | 2023-04-01 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | IT WAS REPORTED BY THE CUSTOMER, 11:32 ON (B)(6) 2023, THE BARD 4F SINGLE-LUMEN PICC CATHETER WAS INSERTED. THE PROCESS WENT SMOOTHLY. THE PERIPHERALLY INTUBATED CENTRAL VENOUS CATHETER KIT AND THE CONNECTOR OF THE ACCESSORY WERE CONNECTED TO THE INFUSION SET WITHOUT LEAKAGE.(B)(6) AT 09:00, WHEN THE FLUID WAS INFUSED AGAIN, THE FLUID LEAKED FROM THE CONNECTOR AND THE PICC CONNECTING TUBE. REPLACE THE HEPARIN CAP, AND THE INFUSION PROCESS GOES SMOOTHLY. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01130 | 2023-04-03 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT APPROXIMATELY TWO YEARS AND NINE MONTHS POST PORT PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE DISTAL CATHETER SEGMENT AND THE PORT BODY WERE REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER IN TWO SEGMENTS WERE RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED AT THE DISTAL END OF THE CATHETER ATTACHED. THE EDGES OF THE BREAK WERE NOTED TO BE UNEVEN, AND THE SURFACE WAS NOTED TO BE GRANULAR. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND MATERIAL SEPARATION ISSUES. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01127 | 2023-04-03 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT MRI ISP, 8 FR GROSHONG, INT WO SP, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8808561 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE HEALTHCARE PROFESSIONAL WAS POORLY REPORTED OF THE PRESENCE OF A STYLET CATHETER INSIDE THE KIT. HOWEVER, PHYSICIAN NOTED THE LACK OF RADIOLOGICAL DETECTION OF THE STYLET INSIDE THE CATHETER. IT WAS FURTHER REPORTED THAT THE IMPLANTATION WAS CARRIED OUT UNDER SCOPING AND THE X-RAY REVEALED THAT THE STYLET WAS ALLEGEDLY NOT VISIBLE AFTER IMPLANTATION. REPORTEDLY, THE STYLET WAS REMOVED WITHOUT FURTHER COMPLICATIONS. THE PROCEDURE WAS COMPLETED WITHOUT ANY CONSEQUENCES FOR THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION, TWO PHOTOS WERE PROVIDED FOR REVIEW. THE PHOTO SHOWS THE PRODUCT MATERIAL#, FLUSHING HUB ATTACHED TO THE CATHETER IN WHICH A CATHLOCK IS ATTACHED IN THE MIDDLE OF THE CATHETER. ALSO, A WARNING TAG WITH "FLUSH CATHETER PRIOR TO USE" WAS NOTED HUNG IN THE CATHETER. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED IMPROPER PROCEDURE ISSUE AS THE PORT WAS IMPLANTED ALONG WITH THE STYLET. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED INADEQUATE INSTRUCTION FOR HEALTHCARE PROFESSIONAL AND NO DISPLAY ISSUES, AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01124 | 2023-04-03 | 2023-03-01 | I, n, j, u, r, y | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT THREE MONTHS AND THREE DAYS POST PORT PLACEMENT, AN X-RAY EXAMINATION WAS PERFORMED AND REVEALED THAT THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE DISTAL CATHETER SEGMENT MIGRATED INTO THE ATRIUM. THE PORT BODY AND THE DISTAL CATHETER SEGMENT WERE REMOVED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER IN TWO SEGMENTS WERE RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED WHICH APPEARED ELLIPTICAL IN SHAPE. THE EDGES OF THE BREAK WAS NOTED TO BE UNEVEN AND THE SURFACE WAS NOTED TO BE GRANULAR. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION AND THE IDENTIFIED DENT ISSUES. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MIGRATION ISSUE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01121 | 2023-04-03 | 2022-02-08 | I, n, j, u, r, y | Manufacturer report | POWER PORT MRI ISP, IMPLANTABLE PORT, GROSHONG SINGLE LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED SOMETIME POST PORT PLACEMENT, THE CATHETER OF THE DEVICE WAS ALLEGEDLY FOUND TO BE BROKEN AND SEPARATED. IT WAS FURTHER REPORTED THAT THE DISTAL CATHETER SEGMENT MIGRATED INTO THE RIGHT HEART. REPORTEDLY, ONLY 6 CM OF THE CATHETER WAS REMOVED AND THE DISTAL CATHETER SEGMENT REMAINED INSIDE THE RIGHT HEART. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRY DATE: 03/2023). | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02265 | 2023-04-03 | 2023-03-16 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETER CAUSED A URINARY TRACT INFECTION. THE PATIENT HAD BEEN USING THIS PRODUCT FOR LESS THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE SETUP: 1. CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE DRYDOCTM VACUUM STATION, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE DRYDOCTM VACUUM STATION USER GUIDE FOR FURTHER INFORMATION. 2. USING STANDARD SUCTION TUBING, CONNECT THE PUREWICKTM FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER. PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. REMOVAL: 5. TO REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER, FULLY SEPARATE THE LEGS, GLUTEUS, AND LABIA. TO AVOID POTENTIAL SKIN INJURY UPON REMOVAL, GENTLY PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER DIRECTLY OUTWARD. ENSURE SUCTION IS MAINTAINED WHILE REMOVING THE PUREWICKTM FEMALE EXTERNAL CATHETER. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. MAINTENANCE: 6. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. INDICATION FOR USE: THE PUREWICKTM FEMALE EXTERNAL CATHETER IS INTENDED FOR NON-INVASIVE URINE OUTPUT MANAGEMENT IN FEMALE PATIENTS. CONTRAINDICATIONS: PATIENTS WITH URINARY RETENTION. WARNINGS: DO NOT USE THE PUREWICKTM FEMALE EXTERNAL CATHETER WITH BEDPAN OR ANY MATERIAL THAT DOES NOT ALLOW FOR SUFFICIENT AIRFLOW. TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER. HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE. EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON. DO NOT USE BARRIER CREAM ON THE PERINEUM WHEN USING THE PUREWICKTM FEMALE EXTERNAL CATHETER. BARRIER CREAM MAY IMPEDE SUCTION. PROCEED WITH CAUTION IN PATIENTS WHO HAVE UNDERGONE RECENT SURGERY OF THE EXTERNAL UROGENITAL TRACT. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. MAINTAIN SUCTION UNTIL THE PUREWICKTM FEMALE EXTERNAL CATHETER IS FULLY REMOVED FROM THE PATIENT TO AVOID URINE BACKFLOW. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. PRIOR TO CONNECTING THE PUREWICKTM FEMALE EXTERNAL CATHETER TO HOSPITAL WALL SUCTION TUBING, VERIFY SUCTION FUNCTION BY COVERING THE OPEN END OF THE SUCTION TUBING WITH ONE HAND AND OBSERVING THE PRESSURE DIAL. IF THE PRESSURE DOES NOT INCREASE WHEN THE LINE IS COVERED, VERIFY THAT THE TUBING IS SECURED, CONNECTED, AND NOT KINKED. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02267 | 2023-04-03 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THERE WAS A PINHOLE AND STERILE WATER LEAKED. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED A 2-WAY CATHETER. VISUAL INSPECTION NOTED A 0.033" PINHOLE ON THE SHAFT OF THE CATHETER LOCATED 0.4740" FROM THE BIFURCATION. ALSO RECEIVED 2 PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWS OVERVIEW OF CATHETER. SECOND PHOTO SAMPLE ZOOMS IN ON POINT OF HOLE PRESENT ON CATHETER. BASED ON PHYSICAL SAMPLE RECEIVED PRODUCT DOES NOT MEET SPECIFICATIONS WHICH STATES "CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE." ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS, THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02266 | 2023-04-03 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT ABOUT 5 HOURS AFTER THE CATHETER WAS INDWELLED, IT CAME OUT SPONTANEOUSLY. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 29MAR2022, IT WAS NOTED THAT WATER LEAK INFORMATION NOT PROVIDED FROM CUSTOMER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00216 | 2023-04-03 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING PREPARATION FOR A RECANALIZATION PROCEDURE, THE TIP OF THE DEVICE ALLEGEDLY SEPARATED. THERE WAS NO PATIENT CONTACT. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 10/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01126 | 2023-04-03 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN CV CATHETER,TRIPLE-LUMEN REPAIR KIT, COMPLETE EXTERNAL SEGMENT, 12.5F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 0601740 | 0601740 |  | IT WAS REPORTED THAT DURING THE PREPARATION OF PORT PLACEMENT, THE MEDICAL ADHESIVE GLUE WAS ALLEGEDLY HARD. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: A VOLUNTARY RECALL HAS BEEN INITIATED FOR CATHETER REPAIR KITS WHICH ARE PRODUCT CATALOG/LOT NUMBER SPECIFIC. REPORTEDLY THE CATHETER REPAIR KITS HAVE HARDENED/COAGULATED ADHESIVE. HARDENED OR COAGULATED ADHESIVE HAS THE POTENTIAL TO CAUSE DELAYS WHILST AN ALTERNATIVE REPAIR KIT, ADHESIVE OR REPLACEMENT CATHETER IS OBTAINED; THEREFORE, POTENTIALLY PROLONGING SURGERY OR NECESSITATING EXCHANGE. TO DATE THERE HAVE BEEN NO ADVERSE EVENTS WORLDWIDE REPORTED RELATED TO THIS ISSUE. AS RESULT OF THE FIELD ACTION, THIS EVENT IS BEING REPORTED AS A MALFUNCTION REPORTABLE EVENT. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE SILICONE ADHESIVE TUBE WAS RETURNED FOR EVALUATION. GROSS VISUAL EVALUATION WAS PERFORMED. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED NON STANDARD DEVICE ISSUE, AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED FROM THE RETURNED PHYSICAL SAMPLE. HOWEVER, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED HARDENED ADHESIVE ISSUE, AS THE SILICONE ADHESIVE TUBE WAS FELT STIFF, NOT SPONGY. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 09/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00217 | 2023-04-03 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE TIP OF THE CATHETER ALLEGEDLY SEPARATED. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 01/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01128 | 2023-04-03 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SMALL SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770353 |  | IT WAS REPORTED THE CORD IS PULLING AWAY AT THE STRAIN RELIEF, EXPOSING WIRING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF CORD IS PULLING AWAY AT THE STRAIN RELIEF, EXPOSING WIRING WAS CONFIRMED. THE REPORTED ISSUE ROOT CAUSE IS DUE TO USE OF THE DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01129 | 2023-04-03 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (5F) DL (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8295118 |  | IT WAS REPORTED THAT DURING THE PROCESS OF CATHETER PLACEMENT, WHITE FLOCCULE WAS FOUND ADHERED TO THE STYLET INSIDE THE POWERPICC SOLO CATHETER (8295118). THE SHITE FLOCCULE COULD BE FLUSHED OFF. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF MATERIAL ON THE STYLET WAS CONFIRMED; HOWEVER, THE EXACT ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE STYLET. IT HAD BEEN FULLY REMOVED FROM THE CATHETER, WHICH WAS NOT RETURNED FOR EVALUATION. BIOLOGICAL RESIDUES WERE IN SEVERAL REGIONS ALONG THE LENGTH OF THE STYLET. SEVERAL BENDS WERE OBSERVED IN THE STYLET. MATERIAL APPEARED TO BE FLAKING OFF OF THE WIRE. MICROSCOPIC INSPECTION OF THE SAMPLE CONFIRMED THE PRESENCE OF WHITE MATERIAL ON THE STYLET. THE MATERIAL APPEARED BUNCHED AND PEELING. THE MATERIAL APPEARED TO PARTIALLY COAT THE STYLET WIRE. IT APPEARED THAT THE HYDROPHILIC COATING APPLIED TO THE STYLET WIRE DURING PRODUCT MANUFACTURE WAS BUNCHING AND PEELING. IT COULD NOT BE DETERMINED WHAT CAUSED THE DELAMINATION OF THE COATING. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE EXPOSURE TO INCOMPATIBLE CHEMICALS AND UNIDENTIFIED ENVIRONMENTAL FACTORS AFFECTING THE PROPERTIES OF THE COATING. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01132 | 2023-04-03 | 2022-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED BY THE CUSTOMER THAT "WHEN INTRODUCING THE CATHETER, NURSE MET RESISTANCE. INTRODUCER INSPECTED AND FOUND TO HAVE SLIGHT JAGGED EDGES ON EXTERNAL CANNULA." | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01131 | 2023-04-03 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC FT CATHETER 5F RADIOLOGY BASIC TRAY DL (70CM NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275355F |  | IT WAS REPORTED BY THE CUSTOMER THAT "PICC PLACED 2/15/23 FOR DAILY INFUSION. NOTED LEAK AT PURPLE PORT ON (B)(6) 2023 PICC LINE REMOVED. MIDLINE PLACED TO COMPLETE TREATMENT. NO PATIENT HARM." ADDITIONAL INFORMATION RECEIVED 03/28/2023: 1) DID THE EVENT OCCUR DURING AN URGENT OR LIFE-THREATENING MEDICAL SITUATION? IF YES, PLEASE PROVIDE DETAILS. NO 2) DID THE EVENT RESULT IN A CLINICALLY SIGNIFICANT DELAY IN THERAPY THAT CAUSED A NEGATIVE OUTCOME FOR THE PATIENT? NO | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01136 | 2023-04-03 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7957505J |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS INSERTED AND THE CONTINUOUS ADMINISTRATION WAS PERFORMED THROUGH BOTH THE LUMENS, WHEN THE CATHETER WAS FLUSHED WITH SALINE SOLUTION DUE TO BLOOD SAMPLING, A CRACK WAS FORMED ON THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 5 FR DUAL LUMEN GROSHONG NXT CATHETER. THE DAMAGE OBSERVED IN THE RETURNED SAMPLE WAS CHARACTERISTIC OF OVER-PRESSURIZATION (BURST) DAMAGE. THIS CAN OCCUR THROUGH THE USE OF SYRINGES SMALLER THAN 10ML, BY FLUSHING AGAINST AN OCCLUSION, OR DUE TO EXCESSIVE FORCE APPLIED DURING INFUSION PROCEDURES. THE RETURNED PRODUCT SAMPLE WAS EVALUATED, AND A SPLIT WAS OBSERVED IN THE EXTENSION LEG TUBING OF THE WHITE LUMEN. THIS CATHETER DAMAGE WAS CHARACTERISTIC OF A BURST FAILURE, AND THE EVIDENCE OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ 'S' SHAPED SPLIT ¿ TENSILE WEAKNESS AT THE FRACTURE SITE (DUE TO MATERIAL BALLOONING PRIOR TO BURST) ¿ SMOOTH AND GRANULAR FRACTURE SURFACE TEXTURE (TYPICAL OF TEARING FAILURE MODES) ¿ LINEAR, LONGITUDINAL CRACKS ON THE INNER WALL OF THE TUBING IN THE AREA OF THE SPLIT (INDICATING OVER-STRETCHING AND TEARING OF THE TUBING MATERIAL) AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01137 | 2023-04-03 | 2023-03-15 | I, n, j, u, r, y | Manufacturer report | BD PROVENA DL PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT, "REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE." THE EVENT INVOLVED A BD PROVENA DOUBLE LUMEN PICC IN UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION AND SPECIFIC PATIENT DETAILS WERE NOT DISCLOSED AS PART OF THE SURVEY. AS SUCH, ADDITIONAL INFORMATION AND THE SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02270 | 2023-04-03 | 2023-03-17 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETERS CAUSED URINARY TRACT INFECTION POSSIBLE. THE CUSTOMER ASKED TO HAD PATIENT TO STOP USING AND HAD THE PATIENT CHECK WITH THE DOCTOR. THE PATIENT HAD BEEN USING THIS PRODUCT FOR MORE THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE SETUP: 1. CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE DRYDOCTM VACUUM STATION, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE DRYDOCTM VACUUM STATION USER GUIDE FOR FURTHER INFORMATION. 2. USING STANDARD SUCTION TUBING, CONNECT THE PUREWICKTM FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER. PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. REMOVAL: 5. TO REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER, FULLY SEPARATE THE LEGS, GLUTEUS, AND LABIA. TO AVOID POTENTIAL SKIN INJURY UPON REMOVAL, GENTLY PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER DIRECTLY OUTWARD. ENSURE SUCTION IS MAINTAINED WHILE REMOVING THE PUREWICKTM FEMALE EXTERNAL CATHETER. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. MAINTENANCE: 6. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. INDICATION FOR USE: THE PUREWICKTM FEMALE EXTERNAL CATHETER IS INTENDED FOR NON-INVASIVE URINE OUTPUT MANAGEMENT IN FEMALE PATIENTS. CONTRAINDICATIONS: ¿ PATIENTS WITH URINARY RETENTION WARNINGS: ¿ DO NOT USE THE PUREWICKTM FEMALE EXTERNAL CATHETER WITH BEDPAN OR ANY MATERIAL THAT DOES NOT ALLOW FOR SUFFICIENT AIRFLOW. ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. ¿ AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: ¿ AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER ¿ HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE ¿ EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ¿ EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON ¿ DO NOT USE BARRIER CREAM ON THE PERINEUM WHEN USING THE PUREWICKTM FEMALE EXTERNAL CATHETER.BARRIER CREAM MAY IMPEDE SUCTION. ¿ PROCEED WITH CAUTION IN PATIENTS WHO HAVE UNDERGONE RECENT SURGERY OF THE EXTERNAL UROGENITAL TRACT. ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. ¿ MAINTAIN SUCTION UNTIL THE PUREWICKTM FEMALE EXTERNAL CATHETER IS FULLY REMOVED FROM THE PATIENT TO AVOID URINE BACKFLOW. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ PRIOR TO CONNECTING THE PUREWICKTM FEMALE EXTERNAL CATHETER TO HOSPITAL WALL SUCTION TUBING, VERIFY SUCTION FUNCTION BY COVERING THE OPEN END OF THE SUCTION TUBING WITH ONE HAND AND OBSERVING THE PRESSURE DIAL. IF THE PRESSURE DOES NOT INCREASE WHEN THE LINE IS COVERED, VERIFY THAT THE TUBING IS SECURED, CONNECTED, AND NOT KINKED. ¿ ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. ¿ PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." | Y | N | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01118 | 2023-04-03 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5616000 | 5616000 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED OOZING AROUND THE PORT, BRUISED SKIN ON HANDS AND FEET, ABDOMINAL SWELLING, BLURRED VISION, AND NOSE BLEEDS. HOWEVER, THERE IS NO FURTHER INFORMATION AVAILABLE REGARDING ADDITIONAL MEDICAL INTERVENTION OR MEDICATION TO TREAT THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 04/2023. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01133 | 2023-04-03 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER (20G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202252 |  | IT WAS REPORTED BY THE CUSTOMER THAT "SAFETY NEEDLE WOULD NOT ENGAGE. NO NEEDLE STICK INJURY OCCURRED, CLINICIAN WAS ABLE TO SAFELY RECAP NEEDLE." IT WAS STATED THERE WAS NO ADVERSE PATIENT OUTCOME. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, TREND ANALYSIS, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF SAFETY MECHANISM FAILURE TO ACTIVATE WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 20GA X2.25¿ ACCUCATH ACE PERIPHERAL IV CATHETER ASSEMBLY. THE SAMPLE WAS RECEIVED ASSEMBLED, WITH THE CATHETER OVERLAYING THE NEEDLE SHAFT. BLOOD RESIDUE WAS OBSERVED IN THE CATHETER, NEEDLE AND HOUSING. THE NEEDLE SHAFT WAS BENT NEAR THE CARRIER. THE GUIDEWIRE WAS FULLY WITHDRAWN INTO THE NEEDLE. THE SAFETY BUTTON WAS PRESSED; HOWEVER, THE NEEDLE REMAINED OUTSIDE OF THE HOUSING. ATTEMPTS TO ADVANCE THE GUIDEWIRE WERE INITIALLY UNSUCCESSFUL. THE WIRE WAS ADHERED WITHIN THE NEEDLE SHAFT BY COAGULATED BLOOD RESIDUES. MICROSCOPIC INSPECTION OF THE SAMPLE CONFIRMED THAT THE NEEDLE SHAFT WAS BENT JUST DISTAL OF THE CARRIER. GUIDEWIRE MOBILITY WAS ESTABLISHED THROUGH GENTLE MANIPULATION. MICROSCOPIC INSPECTION OF THE GUIDEWIRE REVEALED A PERMANENT BENT NEAR THE COIL REGION. THE COILS WERE DEFORMED AND MISALIGNED, BUT APPEARED INTACT. THE BEND IN THE NEEDLE SHAFT WAS CONSISTENT WITH SHEAR STRESS PLACED ON THE NEEDLE SHAFT. THE GUIDEWIRE DEFORMATION WAS CONSISTENT WITH ATTEMPTED ADVANCEMENT AGAINST RESISTANCE, SUCH AS INTO TISSUE. THE ABUNDANT BLOOD RESIDUE SUGGESTED THAT DAMAGE OCCURRED DURING ATTEMPTED DEVICE PLACEMENT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02263 | 2023-04-03 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 119314 |  | IT WAS REPORTED THAT THERE WAS A HOLE IN THE BALLOON. THE EVENT HAPPENED DURING PRETEST AND THERE WAS NO BALLOON RUPTURE/TEAR. THE WATER LEAKED FROM THE BALLOON AND NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01123 | 2023-04-03 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, PIECE THAT EXTENDS FROM THE BODY OF THE PORT TO ATTACH THE CATHETER ALLEGEDLY BROKE. IT WAS FURTHER REPORTED THAT THE PHYSICIAN WAS ABLE TO INSERT HEMOSTATS INTO THE PORT, LOCATE THE BROKEN PIECE AND REMOVED IT. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWER PORT CLEARVUE ISP WAS RECEIVED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC AND FUNCTIONAL EVALUATIONS WERE PERFORMED. THE PORT STEM WAS NOTED TO BE DEFORMED AS IT APPEARED TO BE FLATTENED. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED DEFORMATION ISSUES. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED PORT STEM FRACTURE AND THE MATERIAL SEPARATION ISSUE AS THERE IS NO OBJECTIVE EVIDENCE PROVIDED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 04/2024), G3, H6 (DEVICE) H11: H6 (METHOD, RESULT, CONCLUSION) H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02231 | 2023-04-03 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT ONE OF INTERMITTENT CATHETERS HAD AN INTRUSION ON THE END OF IT AND IT CAUSED THE PATIENT TO BLEEDING. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "MECHANICAL FAILURE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA INSTRUCTIONS FOR USE REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02233 | 2023-04-03 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER (SILICONE) | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WERE NO ABNORMALITIES IN THE EXPANSION OF THE BALLOON BEFORE INSERTION, BUT AFTER INSERTION INTO THE PATIENT, THE BALLOON WAS DAMAGED AND SPONTANEOUS WITHDRAWAL OCCURRED FOUR TIMES. ALL THE EVENT WAS CONFIRMED WITH THE SAME LOT AND THE REMAINING ONE WAS COLLECTED WITHOUT BEING USED. THE USER THEN USED 29030J14 DUE TO THE LACK OF 119314. SIMILARLY, THERE WAS A REPORT DRAFTED SEPARATELY FOR SPONTANEOUS WITHDRAWAL DUE TO BALLOON DAMAGE. PER INVESTIGATOR NOTIFICATION RECEIVED VIA EMAIL ON 12MAR2023 STATED THAT PER SAMPLES RECEIVED. (B)(4) SAMPLES WERE RETURNED. (B)(4) SAMPLES NOT IN PACKAGING APPEARED TO BE USED WITH MATERIAL NUMBER 119314 AND MANUFACTURING LOT NUMBER NGGS2164. (B)(4) SAMPLES IN UNOPENED PACKAGING WITH MATERIAL NUMBER 119314, (B)(4) WITH BATCH NUMBER NGGT3722, AND (B)(4) WITH BATCH NUMBER NGGS3548. PER INVESTIGATOR NOTIFICATION RECEIVED VIA EMAIL ON 15MAR2023 STATED THAT (B)(4) SAMPLES NOT IN PACKAGING APPEARED TO BE USED WITH MATERIAL NUMBER 119314 AND MANUFACTURING LOT NUMBER NGGS2164. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 15MAR2023, STATED THAT THERE WAS NO BALLOON RUPTURE. THE CATHETER HAD A BALLOON PINHOLE. | THE REPORTED EVENT WAS CONFIRMED HOWEVER THE CAUSE WAS UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND SOLUTION IMMEDIATELY LEAKED OUT OF A 0.013" PINHOLE FOUND ON THE BALLOON. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS. PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER. WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02234 | 2023-04-03 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | URINARY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THESE MAGIC 3 GO INTERMITTENT CATHETERS WERE HARD TO INSERT. PER FOLLOW UP VIA PHONE ON 13MAR2023, CUSTOMER STATED THAT THE MAGIC 3 GO WERE TOO DRY FOR THEM TO USE AND WERE TOO FLIMSY WHICH MADE IT DIFFICULT TO INSERT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DEVICE DESCRIPTION: THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS A SILICONE TUBE INSERTED INTO THE URETHRA FOR THE PURPOSE OF DRAINING THE BLADDER OF URINE. NOT MADE WITH NATURAL RUBBER LATEX DOES NOT CONTAIN DEHP. INDICATIONS FOR USE: THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. INSTRUCTIONS FOR USE: REVIEW THE SELFCATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS." | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02236 | 2023-04-03 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN STERILIZED WATER WAS INJECTED DURING THE PRETEST, THE WATER DID NOT RETURN FROM THE BALLOON, AND AFTER LEAVING IT FOR A WHILE, THE WATER RETURNED. ABOUT 10ML OF WATER WAS ASPIRATED BY SYRINGE. THE BALLOON WAS DEFLATED BY SPONTANEOUS/ GENTLE DEFLATION BY SYRINGE. USUALLY BALLOON CAN BE DEFLATED WITHIN AROUND 5 SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16662839 | 2023-04-03 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® LUBRICATH® FOLEY CATHETER | CATHETER, COUDE | C. R. BARD, INC. |  |  |  | DURING CYSTOURETHROSCOPY AND URETHROPLASTY, THE FOLEY BUB WAS DEFLATED AND THERE WAS RESISTANCE MET DURING ATTEMPT TO REMOVE. THE BULB WAS FURTHER DEFLATED AND RESISTANCE PERSISTED. THE BULB WAS INFLATED AND DEFLATED AGAIN WITH SALINE AND DESPITE USE OF LUBRICATE IT WAS DIFFICULT TO REMOVE. THIS RESULTED IN A THIN RIM NOTED TO BE PRESENT AFTER REMOVAL AND SUPERFICIAL LACERATION ALONG POSTERIOR EDGE OF THE URETHRAL MEATUS. THIS AREA REQUIRED REPAIR USING INTERRUPTED STITCHES PARALLEL TO THE LUMEN OF THE URETHRA. |  | N | Y | OTHER |  | 2023-04
1018233-2023-02239 | 2023-04-03 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THE CUSTOMER WENT THROUGH ALMOST A WHOLE BOX THAT DID NOT HAVE THE LUBRICANT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02237 | 2023-04-03 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 14MAR2023, IT WAS STATED THAT THERE WERE NO OBVIOUS VISIBLE DEFECTS TO THE RETURNED SAMPLE. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE COULD BE LUMEN COMPROMISED BY A PUNCTURE OR CUT. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE SOLUTION IMMEDIATELY WENT INTO THE DRAINAGE LUMEN AT BIFURCATION AND CREATED LUMEN TO LUMEN LEAK. THIS IS OUT OF SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D,F,H. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02240 | 2023-04-03 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0125RL20 | 0125RL20 |  | IT WAS REPORTED THAT CUSTOMER WAS STATING THAT THIS WAS NOT A 3 WAY AND NEEDS ITEM RETURN. PER FOLLOW UP VIA PHONE (B)(6) 2023, THERE WAS NO ISSUE WITH THE PRODUCT ITSELF. CUSTOMER STATED THE DISTRIBUTOR HAD THE PRODUCT LISTED INCORRECTLY ON THE SITE FOR ORDERING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02241 | 2023-04-03 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE LOT NUMBERS OF 123422C ON THE PACKAGING CARTON BOTTOM AND SIDE ARE DIFFERENT (MANUFACTURE DATE: (B)(6) 2018, EXPIRY DATE: (B)(6) 2023), AND THE LOT NUMBER ON THE CARTON BOTTOM WAS MYCVRE43 COMPARED TO MYCVRB43 ON THE SIDE. THE LOT NUMBER OF EACH PRODUCT INSIDE WAS MYCVRE43. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2022, IT WAS ALSO STATED THAT THERE WAS NO PATIENT INVOLVED. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. BASED ON THE ATTACHED PHOTO, OBSERVED THAT THE LABELLING OF PACKAGING LOT SEEMS LIKE MYCVRB43. THIS WAS DUE TO ILLEGIBLE PRINTING BY THE OLD PRINTER THAT CAUSE THE LOT# MISINTERPRETED. THE ALPHABET ¿E¿ WAS WRONGLY INTERPRETED AS ALPHABET ¿B¿ BY THE CUSTOMER AS LOT#MYCVRB43. THEREFORE, THIS COMPLAINT IS CONFIRMED MANUFACTURING RELATED. FURTHER INVESTIGATION WAS RECORDED IN CID# (B)(4). THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO ILLEGIBLE PRINTING. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" CORRECTIONS: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02268 | 2023-04-03 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998606 | 998606 |  | IT WAS REPORTED THAT THE BALLOON OF THE BALLOON DILATION CATHETER WAS FOUND TO BE LEAKING WATER DURING INFLATION AND PRE-FLUSHING. | THE REPORTED EVENT IS CONFIRMED, CAUSE UNKNOWN. NO PHYSICAL SAMPLE WAS RETURNED, HOWEVER A 20 SECOND VIDEO SAMPLE WAS PROVIDED. VISUAL EVALUATION OF THE VIDEO SAMPLE NOTED WATER APPEARED TO BE LEAKING FROM THE BASE OF THE BALLOON WHERE IT CONNECTS TO THE SHAFT OF THE CATHETER. UNABLE TO DETERMINE WHETHER THE LEAK WAS COMING FROM THE SHAFT ITSELF OR THE THE BALLOON BASED ON THE VIDEO SAMPLE. THOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, A POTENTIAL ROOT CAUSES FOR THIS EVENT COULD BE, "INADEQUATE MATERIAL SELECTION" OR "POOR SHAFT DESIGN". AS NO PATIENT INVOLVEMENT WAS REPORTED THE DEVICE WAS NOT USED, HOWEVER, IS INTENDED FOR TREATMENT PURPOSES. IT WAS UNKNOWN IF THE DEVICE HAD MET ALL RELEVANT SPECIFICATIONS. NO MANUFACTURING ISSUES OR NON-CONFORMANCES WERE NOTED DURING REVIEW OF THE DHR THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "THE X-FORCE® BALLOON DILATION CATHETER INSTRUCTIONS FOR USE, WAS REVIEWED AND FOUND TO BE ADEQUATE AS IT STATES: DO NOT EXCEED THE RECOMMENDED RATED BURST PRESSURE (RBP) FOR THIS DEVICE. BALLOON RUPTURE MAY OCCUR IF THE RBP RATING IS EXCEEDED. PLEASE REFER TO THE DEVICE LABEL FOR THE RECOMMENDED MAXIMUM RATED BURST PRESSURE. EXTREME CAUTION SHOULD BE USED WHEN CONSIDERING DILATION OF IRREGULAR, NOT-COMPLIANT TISSUE OR IN THE PRESENCE OF CERTAIN CALCULI. INSPECTION PRIOR TO USE: THE X-FORCE® BALLOON DILATION CATHETER IS A STERILE, SINGLE USE DEVICE. CAREFULLY INSPECT THE CATHETER AND THE STERILE PACKAGING FOR SIGNS OF DAMAGE THAT MAY HAVE OCCURRED DURING SHIPMENT. DO NOT USE THE PRODUCT IF DAMAGE IS EVIDENT. STORAGE: STORE THE BALLOON DILATION CATHETER IN A COOL, DRY, DARK PLACE. DO NOT STORE NEAR RADIATION OR ULTRAVIOLET LIGHT SOURCES AS THESE MAY DAMAGE PRODUCT MATERIALS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02242 | 2023-04-03 | 2022-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589, NGDY1685 AND 2120842164. AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON 29MAR2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022 AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON (B)(6) 2022. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED . ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED. ON (B)(6) 2022 THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED AS THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON 19FEB2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02245 | 2023-04-03 | 2023-01-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589, NGDY1685 AND 2120842164. AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON 29MAR2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022 AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON (B)(6) 2022. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED. ON (B)(6) 2022, THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED AS THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE WRONG PRODUCT SIZE SELECTED CAUSE URINE LEAKAGE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02262 | 2023-04-03 | 2023-01-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE CATHETER BALLOON BURST. THE CATHETER SPONTANEOUSLY FELL OUT 15 DAYS AFTER PLACEMENT AND ON CHECKING THE CATHETER, THERE WERE TRACES OF BALLOON BURST OBSERVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02258 | 2023-04-03 | 2023-03-13 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD A INFECTION DUE TO USE OF PUREWICK URINE COLLECTION SYSTEM. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02259 | 2023-04-03 | 2023-03-16 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION FROM USING THE PUREWICK FEMALE EXTERNAL CATHETER. CUSTOMER STATED THAT CUSTOMER WAS NOT SURE WHAT HAPPENED. CUSTOMER STATED PATIENT WAS DISCHARGED FROM HOSPITAL EARLIER TODAY. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION. THE PATIENT HAD BEEN USING THIS PRODUCT FOR LESS THAN 90 DAYS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ." THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE SETUP: 1. CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE DRYDOCTM VACUUM STATION, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE DRYDOCTM VACUUM STATION USER GUIDE FOR FURTHER INFORMATION. 2. USING STANDARD SUCTION TUBING, CONNECT THE PUREWICKTM FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER. PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. REMOVAL: 5. TO REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER, FULLY SEPARATE THE LEGS, GLUTEUS, AND LABIA. TO AVOID POTENTIAL SKIN INJURY UPON REMOVAL, GENTLY PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER DIRECTLY OUTWARD. ENSURE SUCTION IS MAINTAINED WHILE REMOVING THE PUREWICKTM FEMALE EXTERNAL CATHETER. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. MAINTENANCE: 6. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. INDICATION FOR USE: THE PUREWICKTM FEMALE EXTERNAL CATHETER IS INTENDED FOR NON-INVASIVE URINE OUTPUT MANAGEMENT IN FEMALE PATIENTS. CONTRAINDICATIONS: PATIENTS WITH URINARY RETENTION WARNINGS: DO NOT USE THE PUREWICKTM FEMALE EXTERNAL CATHETER WITH BEDPAN OR ANY MATERIAL THAT DOES NOT ALLOW FOR SUFFICIENT AIRFLOW. TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER, HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE , EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE, EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON: DO NOT USE BARRIER CREAM ON THE PERINEUM WHEN USING THE PUREWICKTM FEMALE EXTERNAL CATHETER. BARRIER CREAM MAY IMPEDE SUCTION. PROCEED WITH CAUTION IN PATIENTS WHO HAVE UNDERGONE RECENT SURGERY OF THE EXTERNAL UROGENITAL TRACT. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. MAINTAIN SUCTION UNTIL THE PUREWICKTM FEMALE EXTERNAL CATHETER IS FULLY REMOVED FROM THE PATIENT TO AVOID URINE BACKFLOW. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. PRIOR TO CONNECTING THE PUREWICKTM FEMALE EXTERNAL CATHETER TO HOSPITAL WALL SUCTION TUBING, VERIFY SUCTION FUNCTION BY COVERING THE OPEN END OF THE SUCTION TUBING WITH ONE HAND AND OBSERVING THE PRESSURE DIAL. IF THE PRESSURE DOES NOT INCREASE WHEN THE LINE IS COVERED, VERIFY THAT THE TUBING IS SECURED, CONNECTED, AND NOT KINKED. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02260 | 2023-04-03 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | TEMP SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314M | 119314M |  | IT WAS REPORTED THAT THE FOLEY TRAY 14FR HAD NO FOLEY CATHETER ATTACHED TO DRAINAGE SYSTEM. PER FOLLOW UP VIA EMAIL RESPONSE ON 13MAR2023, IT WAS REPORTED THAT THE FOLEY CATHETER EXPELLED WITH FUDUS EVALUATION IN OPERATING ROOM 3 SPECIAL CASES. STATED THAT 14 FRENCH FOLEY REPLACED IN OPERATING ROOM DRAINING WELL 200CC INTIAL AND URINE OUTPUT IN CASE 150CC, INFLATION OF BALLOON WITNESSED AND DOCUMENTED ON 1ST INSERTION. THIS WAS A THEME THAT HAS BEEN REPORTED BEFORE OF FOLEYS BEING EXPELLED SPONTANEOUSLY. IN THIS CASE THE BALLOON WAS DEFLATED WHEN IT FELL OUT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿WATER LOST VIA PERMEATION¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE PROVIDED LOT NUMBER WAS INVALID THEREFORE DHR REVIEW CAN¿T BE COMPLETED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: STERILE UNLESS PACKAGE IS OPENED OR DAMAGED, EXCEPT FOR ANY INDIVIDUALLY PACKAGED COMPONENTS WITHIN THE TRAY WHICH ARE NOT LABELED AS STERILE. THESE COMPONENTS ARE NOT TERMINALLY STERILIZED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02252 | 2023-04-03 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER SPONTANEOUSLY DRAINED OUT OF THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02246 | 2023-04-03 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589, NGDY1685 AND 2120842164 . AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON 29MAR2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022 AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON (B)(6) 2022. (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED. ON (B)(6) 2022 THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED AS THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02247 | 2023-04-03 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT SOME OF THE INTERMITTENT CATHETERS HAD NO LUBE IN THE PACKET. NO ADDITIONAL INFORMATION PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INCORRECT PROCESS PARAMETERS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED ISSUE. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02251 | 2023-04-03 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT DURING A PRETEST IN THE OPERATION ROOM, WATER BACK FLOWED FROM THE INFLATION VALVE AND LEAKED. THERE WAS NO BALLOON TEAR/RUPTURE. NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER AND THE CATHETER WAS INSERTED BY HAND. AS PER THE ADDITIONAL FOLLOW UP INFORMATION RECEIVED VIA IBC ON 24MAR2023, IT WAS STATED THAT THE EVENT WAS JUST A LEAK ISSUE DURING PRETEST AND THERE WAS NO RESISTANCE REPORTED. | THE REPORTED EVENT IS UNCONFIRMED AS THE PRODUCT MEETS SPECIFICATIONS. VISUAL EVALUATION NOTED RECEIVED A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING AND A SYRINGE. INFLATED BALLOON WITH 10 ML OF SOLUTION AND THE CATHETER RESTED WITH NO LEAKS NOTED AND NO BACKFLOW NOTED. DHR REVIEW IS NOT REQUIRED AS THE REPORTED EVENT IS UNCONFIRMED. LABELLING REVIEW IS NOT REQUIRED AS THE REPORTED EVENT IS UNCONFIRMED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02261 | 2023-04-03 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MALE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53810G | 53810G |  | IT WAS REPORTED THAT THE ORDERS HAVE CONTAINED DIFFERENT SIZE MALE EXTERNAL CATHETERS BUT ARE MARKED AS 10FR. REPRESENTATIVE ADVISED THAT WHEN THIS OCCURS PATIENT NEEDS TO LET US KNOW SO WE CAN RESOLVE THAT ISSUE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR-MANDREL LACK OF IDENTIFICATION/MISIDENTIFICATION". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA INSTRUCTIONS FOR USE REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01156 | 2023-04-04 | 2023-03-20 | I, n, j, u, r, y | Manufacturer report | BARD PROVENA CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED "THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE." NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01148 | 2023-04-04 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, THERE IS A HOLE IN THE CATHETER RIGHT ABOVE THE WINGS, FOUND DURING USE. NO PATIENT HARM. UPDATE 4-APR-2023: THE PATIENT WOULD TAKE SAMPLES FROM THIS ON HIS HEALTHCARE CENTRAL AND THEN THE NURSE COULD NOT FLUSH IN OR GET ANY BACKFLOW IN THE PICC LINE. IT WHEEZED AND SEEMED TO HAVE AIR COMING OUT OF IT. AND THE LIQUID THAT WAS FLUSHED IN RAN OUT UNDER THE DRESSING. THE PATIENT WAS SENT TO THE ONCOLOGY WARD WHERE HE WAS TREATED WHERE THE SAME THING WAS NOTICED AND THE PICC LINE WAS DISCONTINUED. LUCKILY THE PATIENT MANAGED TO BE PUSHED IN FOR A NEW PICC LINE JUST 30 MINUTES AFTER THE WRONG ONE WAS PULLED, SO SHE GOT A NEW PICC LINE THE SAME DAY THE FIRST ONE WAS TAKEN OUT. THE PATIENT WAS THEREFOR ABLE RECEIVE HIS TREATMENT THE NEXT DAY AS PLANNED. THE ONLY SYMPTOM OF THE PICC LINE NOT WORKING AS IT SHOULD WAS WHEN THE NURSE ON THE HEALTH CARE CENTRAL WOULD FLUSH IT AND FLUID LEAKED OUT. OTHERWISE THE PATIENT EXPERIENCED NO PROBLEMS. THE CATHETER WAS SECURED WITH A STATLOCK AND THE PATIENT WAS GIVEN CHEMOTHERAPY, FOLFIRINOX WAS THE COMBINATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A HOLE IN THE CATHETER RIGHT ABOVE THE WINGS WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR S/L POWERPICC SOLO. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED JUST DISTAL OF THE MOLDED SUTURE WINGS. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING) AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01150 | 2023-04-04 | 2023-03-13 | I, n, j, u, r, y | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2174108 |  | IT WAS REPORTED, DURING THE ADMINISTRATION OF THERAPY (SALINE/SOLDESAM) AN EXTRAVASATION OCCURRED. THE PICC WAS REMOVED WITHOUT ANY CONSEQUENCES FOR THE PATIENT. THE PICC WAS DAMAGED. THE DEVICE HAS NEVER BEEN USED WITH CONTRAST MEDIUM. THE PATIENT REPORTED THAT HE MADE A PARTICULARLY ABRUPT MOVEMENT AND EXPERIENCED A SLIGHT TRANSIENT PAIN IN THE ARM WHERE THE DEVICE WAS POSITIONED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAKING CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR S/L POWERPICC CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED JUST DISTAL OF THE 6CM DEPTH MARKING. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED. ¿ FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR. ¿ 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER. ¿ OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING). AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01151 | 2023-04-04 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT,4FR POEWRMIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P4154108D |  | IT WAS REPORTED VIA MEDWATCH, "ASSESSED MIDLINE THAT IS LEAKING AT THE HUB. FLUSHED MIDLINE IN RIGHT UPPER EXTREMITY WITH SALINE AND IT IS LEAKING AT THE DISTAL END OF THE CATHETER." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01155 | 2023-04-04 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER THAT AFTER INSERTION OF THE CATHETER, THE CATHETER WAS TRIMMED. WHEN CONNECTING IT TO THE FITTING, AFTER A ¿CLICK¿ WAS HEARD, THE OPERATOR CHECKED IF THE CONNECTION WAS FIRM WITH A PULL OF THE FITTING, THEN THE FITTING WAS DISENGAGED. THE FITTING WAS REPLACED WITH A NEW ONE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING INVESTIGATION ELEMENTS WERE DETERMINED TO BE APPLICABLE AND WERE COMPLETED AS PART OF THIS INVESTIGATION: ¿ COMPLAINT HISTORY REVIEW, ¿ DHR RECORDS, ¿ SAMPLE ANALYSIS, ¿ REVIEW OF RISK DOCUMENTATION AND ¿ LABELING REVIEW. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A FAULTY GROSHONG CONNECTOR IS INCONCLUSIVE DUE TO THE STATE OF THE RETURNED SAMPLE. ONE 4 FR GROSHONG NXT TWO-PIECE CONNECTOR WAS RETURNED FOR EVALUATION. THE CONNECTOR WAS RETURNED ASSEMBLED AND NO DAMAGE WAS FOUND ON THE RETURNED SAMPLE. THE CONNECTION OF THE TWO-PIECE CONNECTOR WAS FOUND FIRM AND WAS TO BE ABLE TO BE DISASSEMBLED BY HAND WITH MODERATE FORCE, AND AN AUDIBLE CLICK COULD BE HEARD WHEN REASSEMBLING THE CONNECTOR. THE COMPLAINT OF DIFFICULTY CONNECTING A TWO-PIECE CONNECTOR COULD NOT BE CONFIRMED BECAUSE THE GROSHONG CONNECTOR WAS RETURNED ASSEMBLED AND THE CONNECTION WAS FOUND TO BE FIRM. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01154 | 2023-04-04 | 2023-03-20 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT "THE PATIENT HAD A GROSHONG CATHETER IMPLANTED IN 2021. ON (B)(6) 2023, THE CATHETER COULD NOT BE REMOVED WHEN THE GROSHONG CATHETER WAS ATTEMPTED TO BE REMOVED. WHEN THE PATENCY OF THE CATHETER WAS CONFIRMED, THERE WAS NO PROBLEM. CT EXAMINATION WAS PERFORMED AND REVEALED THAT THE CATHETER HAD ADHERED TO THE VESSEL AND HAD INSERTED INTO THE RIGHT ATRIUM. CONSIDERING THE RISK OF REMOVAL, THE DOCTOR DECIDED THAT THE CATHETER REMAINS INSIDE OF THE PATIENT'S BODY AND CONTINUES THE FOLLOW-UP OBSERVATION. THE HEALTH INJURY TO THE PATIENT WAS UNKNOWN." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01157 | 2023-04-04 | 2023-03-20 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT "THE PATIENT HAD A GROSHONG CATHETER IMPLANTED IN 2021. ON (B)(6) 2023, BECAUSE INFECTION DUE TO CATHETER RELATED BLOOD STREAM INFECTION WAS SUSPECTED, THE CATHETER WAS ATTEMPTED TO BE REMOVED. THE HEALTH INJURY TO THE PATIENT WAS UNKNOWN." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01162 | 2023-04-04 | 2023-03-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER FULL KIT (5F)(DUAL-LUMEN) | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P6255208 |  | IT WAS REPORTED BY THE CUSTOMER "WHEN RINSING WITH PHYSIOLOGICAL SERUM, WE NOTICED THAT THE DEVICE WAS DEFECTIVE BECAUSE IT NO LONGER ADMINISTERED CORRECTLY. CALL TO THE ON-CALL MANAGER WHO ADVISED US TO CALL THE ON-CALL ANAESTHETIST, WHO ASKED US TO REMOVE THE DEVICE AND KEEP IT IF NEEDED FOR ANALYSIS AND TO PUT IN AN INFUSION IN THE MEANTIME TO BE ABLE TO ADMINISTER THE ATB." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01159 | 2023-04-04 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER BASIC KIT (4F) (SINGLE-LUMEN) (20CM) | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P6154118 |  | IT WAS REPORTED THAT "POWER MIDLINE INSERTED ON (B)(6) 2023. NOTED TO BE LEAKING ON (B)(6) 2023. PATIENT WAS ON A SLOW INFUSION FROM 08-13TH MARCH, AND THEN INFUSION CHANGED TO A FAST RATE ON (B)(6). THIS WAS WHEN THE LEAKING WAS NOTICED. IT IS UNCLEAR WHEN THE LEAKING BEGAN. ON (B)(6) HOLE NOTED NEAR '0 CM' MARKING ON THE MIDLINE. THE MIDLINE HAD TO BE REMOVED. ALTERNATIVE ACCESS REQUIRED." IT WAS STATED THERE WAS NO HARM TO THE PATIENT, NO ERRONEOUS RESULTS BECAUSE OF THE EVENT, AND NO CHANGE IN COURSE OF TREATMENT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01160 | 2023-04-04 | 2023-03-04 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIG JAPAN HANG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405J |  | IT WAS REPORTED THAT "THE PATIENT HAD A GROSHONG CATHETER IMPLANTED VIA THE LEFT UPPER ARM ON (B)(6) 2023. THE CATHETER WAS CUT TO A LENGTH OF 34 CM, AND THE CATHETER PART WHICH WAS OUTSIDE OF THE PATIENT'S BODY WAS APPROXIMATELY 4 TO 5 CM. THE CATHETER WAS NOT U-SHAPED AND SECURED STRAIGHT. THE CATHETER WAS SECURED ONLY WITH THE STATLOCK WITHOUT USING THE SUTURE WING. A DRESSING WAS APPLIED TO COVER THE INSERTION SITE, THE CATHETER, AND THE STATLOCK. ON (B)(6) 2023, CT EXAMINATION WAS PERFORMED AND REVEALED THAT THE CATHETER WAS DETACHED FROM THE CATHETER CONNECTOR AND THE CATHETER MIGRATED INTO THE PULMONARY ARTERY. THE CATHETER WAS REMOVED THROUGH THE GROIN. THE OVERSLEEVE AND CATHETER CONNECTOR OF THE RETURNED SAMPLE ARE DETACHED, BUT THIS WAS DETACHED BY THE DOCTOR AFTER REMOVING THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF CATHETER DETACHMENT IS CONFIRMED. ONE 4 FR GROSHONG NXT CLEARVUE CATHETER WITH THE TWO-PIECE CONNECTOR AND A STATLOCK STABILIZATION DEVICE WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED HEAVY BLOOD USE RESIDUES THROUGHOUT THE DEVICE, AND THE CATHETER WAS RETURNED DETACHED FROM THE DISASSEMBLED TWO-PIECE CONNECTOR AND STATLOCK STABILIZATION DEVICE. THE RETURNED CATHETER TERMINATED AT THE 34 CM DEPTH MARKER, JUST AS STATED IN THE EVENT DESCRIPTION; THEREFORE, NO ADDITIONAL SECTION OF CATHETER WAS TRIMMED FOLLOWING THE INITIAL INSERTION OF THE DEVICE. A MICROSCOPIC OBSERVATION REVEALED THE PROXIMAL TIP OF THE TRIMMED CATHETER TO BE INTACT WITH NO DAMAGE OBSERVED FROM BEING INSERTED INTO THE DISTAL CONNECTOR PIECE. A TACTILE EVALUATION OF THE RETURNED CATHETER REVEALED NO TENSILE WEAKNESS ALONG THE DEVICE WHERE THE CATHETER MAY HAVE BEEN PULLED FROM THE DISTAL CONNECTOR PIECE. A FUNCTIONAL TEST OF ATTEMPTING TO INSERT THE CATHETER INTO THE DISTAL PIECE OF THE TWO-PIECE CONNECTOR WAS UNSUCCESSFUL AS THE CATHETER WOULD NOT ADVANCE PAST THE CLEAR, DISTAL PORTION OF THE OVERSLEEVE. THE DISTAL CONNECTOR PIECE WAS THEN CUT IN HALF LENGTHWISE TO OBSERVE THE INSIDE OF THE DEVICE. A MICROSCOPIC OBSERVATION OF THE DISTAL CONNECTOR PIECE AFTER IT WAS CUT LENGTHWISE REVEALED THE COMPRESSION SLEEVE HAD BEEN ADVANCED INTO THE LOCKED POSITION WITH NO REMAINS OF THE CATHETER INSIDE THE DISTAL CONNECTOR. THE COMPLAINT OF A CATHETER DETACHING FROM THE TWO-PIECE CONNECTOR IS CONFIRMED. OBSERVATIONS OF THE PROXIMAL SEGMENT OF A RETURNED CATHETER THAT HAD BEEN PROPERLY ASSEMBLED SHOULD EXHIBIT SOME CATHETER SEGMENTS POTENTIALLY LEFT IN THE DEVICE, AND THE CATHETER SHOULD EXHIBIT MARKS ON THE PROXIMAL SEGMENT WHERE IT WAS TRIMMED FROM BEING COMPRESSED CIRCUMFERENTIALLY BY THE COMPRESSION SLEEVE INSIDE THE CONNECTOR AND PULLED LONGITUDINALLY DEVELOPING SOME TENSILE WEAKNESS. THE RETURNED CATHETER, HOWEVER, WAS OBSERVED TO BE INTACT WITH NO TENSILE WEAKNESS OR DAMAGE TO THE PROXIMAL END THAT ATTACHES TO THE TWO-PIECE CONNECTOR, AND THE RETURNED DISTAL PIECE OF THE TWO-PIECE CONNECTOR WAS OBSERVED TO HAVE NO CATHETER SEGMENTS LEFT INSIDE THE ENGAGED COMPRESSION SLEEVE. THESE OBSERVATIONS SUGGESTED THAT THE CATHETER WAS NOT INSERTED THROUGH THE COMPRESSION SLEEVE PRIOR TO CONNECTOR ASSEMBLY. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01149 | 2023-04-04 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000954 |  | IT WAS REPORTED BY THE CUSTOMER THAT WHEN FLUSHING THE LINE THEY NOTICED SALINE COMING OUT OF THE CLOSED SYSTEM THAT HOUSES THE WIRE ON THE RED LUMEN. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01161 | 2023-04-04 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F DUAL-LUMEN POLYURETHANE CATHETER RADIOLOGY BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275335 |  | IT WAS REPORTED BY THE CUSTOMER THAT "POWER PICC CAME OUT LAST NIGHT." NO OTHER INFORMATION WAS PROVIDED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01163 | 2023-04-04 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER (20G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202252 |  | IT WAS REPORTED THAT "DURING ACCUCATH TRAINING TODAY, CLINICIAN WAS STERILE ABOUT TO PICK UP THE CATHETER OUT OF THE INTERMEDIATE KIT AND THE SAFETY WAS ALREADY ENGAGED, LEAVING NO NEEDLE TO ACCESS THE VEIN. WE HAD TO OPEN A NEW KIT." | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01164 | 2023-04-04 | 2022-12-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000375 |  | IT WAS REPORTED BY THE CUSTOMER "PICC LEAKING DURING ADMINISTRATION, WAS FOUND TO BE FRACTURED ON REMOVAL. PATIENT REQUIRED REPLACEMENT PICC TO CONTINUE THERAPY. IT WAS STATED THE EVENT DID NOT INVOLVE AN URGENT/LIVE THREATENING MEDICAL SITUATION. THE LINE WAS REPLACED IN TIME FOR NEXT TREATMENT. PATIENT EXPERIENCED ANXIETY BUT NO HARM CAUSED TO PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01169 | 2023-04-04 | 2023-01-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SIM-CATH DIALYSIS CATHETER KIT 12F X 20CM (STRAIGHT) (SHORT TERM) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5553200 |  | IT WAS REPORTED THAT GUIDEWIRE GOT STUCK WHEN INSERTING CATHETER. A NEW CATHETER WAS USED AND INSERTED. NO OTHER INFORMATION WAS PROVIDED. (B)(6) 2023 - A PHOTO SAMPLE OF THE AFFECTED DEVICE EXHIBITED A KINKED GUIDEWIRE. | THE MODEL DESCRIPTION OF TJF-260V AND JF-260V WAS SAME. ONE OF REPRESENTATIVE PRODUCT WAS CHOSEN FOR PROCESSING PURPOSES. THE SUSPECT DEVICE HAS NOT BEEN RETURNED TO OLYMPUS FOR EVALUATION. THE INVESTIGATION IS IN PROCESS. THE LITERATURE ARTICLE IS ATTACHED FOR ADDITIONAL INFORMATION. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED WITH DEVICE EVALUATION RESULTS. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01170 | 2023-04-04 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 3F SL SOLO MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1193108D5 |  | IT WAS REPORTED BY THE CUSTOMER "WHILE PLACING PICC LINE DILATOR IT WAS DIFFICULT TO ADVANCE AND CAUSING DISCOMFORT BEYOND WHAT IS EXPECTED DURING THE PROCEDURE. PROCEDURE WAS HALTED, DILATOR REMOVED AND INSPECTED. IT WAS NOTED THE DILATOR HAD BENT BACK AT THE TAPER AT THE POINT OF INSERTION. THIS CAUSED THE DILATOR NOT TO ADVANCE AND CAUSE INCREASED PAIN. A NEW KIT WAS BROUGHT TO THE BEDSIDE AND PICC WAS SUCCESSFULLY PLACED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01158 | 2023-04-04 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 20CM (SHORT-TERM) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5593200 |  | IT WAS REPORTED, WIRE COULD BE INSERTED WITHOUT ANY PROBLEMS. A CORRECTION WAS DESIRED/NECESSARY IN THE CASE OF A VASCULAR OBSTRUCTION. WHEN THE WIRE WAS WITHDRAWN, THE TIP OF THE WIRE APPARENTLY JAMMED IN THE CANNULA (NOT AT THE TIP), MOBILIZATION IN ONE DIRECTION WAS NO LONGER POSSIBLE. AFTER PULLING THE CANNULA AND WIRE, THE PROBLEM PERSISTED EX-SITU, EVEN UNDER FORCED PRESSURE. NO IMPACT ON PATIENT, USER OR 3RD PERSON REPORTED, BUT DELAY IN CATHETER INSERTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF THE GUIDEWIRE BECOMING STUCK WITHIN THE INTRODUCER NEEDLE WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 18GA INTRODUCER NEEDLE AND ONE 0.037" J-TIP GUIDEWIRE. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE CAUSED BY RETRACTION OF THE GUIDEWIRE AGAINST THE BEVEL OF THE INTRODUCER NEEDLE. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND THE GUIDEWIRE CORE WIRE WAS CONFIRMED TO BE BROKEN WHICH ALLOWED THE OUTER COIL WIRE TO BECOME LOOSE AND DEFORMED. MICROSCOPIC EXAMINATION OF THE FRACTURE SITES REVEALED THE FOLLOWING: NARROWING OF THE WIRE CROSS-SECTION NEAR THE FRACTURE SITE, WHICH IS A CHARACTERISTIC FEATURE OF A STRONG PULL ON THE WIRE DAMAGE TO THE INSIDE EDGE OF THE INTRODUCER NEEDLE WHICH CAN OCCUR IF THE GUIDEWIRE IS FORCEFULLY RETRACTED AGAINST THE NEEDLE, DAMAGING THE SHAPE OF THE SHARPENED BEVEL BIOLOGICAL MATERIAL WAS ALSO SEEN ON THE WIRE WHICH MAY HAVE CONTRIBUTED TO THE OBSERVED GUIDEWIRE FRACTURE AS RETRACTION OF THE WIRE TOGETHER WITH THE BIOLOGICAL MATERIAL COULD HAVE CAUSED THE PIECES TO BECOME STUCK. NORMAL MOVEMENT OF THE GUIDEWIRE IS AWAY FROM THE SHARPENED BEVEL AND WILL NOT DAMAGE THE WIRE. IF THE GUIDEWIRE DIRECTION IS REVERSED, THE GUIDEWIRE IS THEN PULLED AGAINST THE SHARPENED EDGE OF THE NEEDLE BEVEL AND CAN CAUSE SHEARING DAMAGE OF THE WIRE AND/OR CAUSE THE WIRE TO BECOME STUCK WITHIN THE NEEDLE. AN EXAMINATION OF THE WIRE STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01147 | 2023-04-04 | 2023-03-07 | I, n, j, u, r, y | Manufacturer report | FULL KIT 20G X 10 CM WITH PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F220108PT |  | IT WAS REPORTED VIA MEDWATCH, "ATTEMPTED MIDLINE CATHETER INSERTION USING ULTRASOUND GUIDANCE IN THE LEFT CEPHALIC VEIN. WHEN ATTEMPTING TO DEPLOY GUIDEWIRE THEN CATHETER, I MET RESISTANCE, WHICH WAS ALSO HURTING THE PATIENT. SO, I STOPPED, PULLED CATHETER BACK, PULLED GUIDEWIRE BACK, THEN REMOVED NEEDLE FROM HIS ARM. INSPECTED THE NEEDLE/CATHETER/GUIDEWIRE AND FOUND THAT PART OF THE CATHETER BROKE OFF FROM THE NEEDLE. I USED THE ULTRASOUND AND WAS ABLE TO VISUALIZE THE CATHETER PARTIALLY IN THE CEPHALIC VEIN." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01143 | 2023-04-04 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5304150 | 5304150 |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CLIP OF VENOUS LIMB OF HEMODIALYSIS CATHETER ALLEGEDLY BROKE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE 10F GLIDEPATH D/L CATHETER WAS RECEIVED FOR EVALUATION. GROSS VISUAL, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. UPON FUNCTIONAL TESTING, BOTH THE LUERS WERE PATENT TO THE INFUSION AND ASPIRATION TESTS PERFORMED. NO LEAKS WERE OBSERVED THROUGHOUT TESTS PERFORMED. THEREFORE, THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED FRACTURE ISSUE AS NO FRACTURE WAS NOTED TO THE CLAMP. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 08/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01107 | 2023-04-04 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER, WHEN CATHETER WAS FLUSHED PRIOR TO USE ON PATIENT AND THERE IS A HOLE IN THE EXTERNAL TUBE. BEFORE USE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING EXTERNAL TUBE ON A POWERPICC SOLO WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR S/L POWERPICC SOLO WITH THE STYLET INSERTED INTO THE DEVICE. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH EXTENSION TUBING DAMAGE CAUSED BY CONTACT WITH A SHARP EDGED INSTRUMENT SUCH AS A SCALPEL. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND THREE SPLITS WERE OBSERVED ALONG THE EXTENSION TUBING OF THE STYLET. MICROSCOPIC EXAMINATION OF THE TUBING SPLITS CONFIRMED IT WAS TYPICAL OF CONTACT WITH A SHARP BLADED INSTRUMENT, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: THE FRACTURE SURFACE WAS REFLECTIVE IN NATURE AND CONTAINED A STRIATED PATTERN. THIS CAN OCCUR DUE TO PATTERN TRANSFER OF THE SHARPENED EDGE TYPICALLY FOUND ON A SCALPEL-TYPE INSTRUMENT. SHARPLY FORMED FRACTURE EDGES. PLANAR SHAPE TO THE FRACTURE SURFACE. AN EXAMINATION OF THE EXTENSION TUBING STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01139 | 2023-04-04 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | A710962 | A710962 |  | IT WAS REPORTED THAT DURING PORT PLACEMENT PROCEDURE, THE AIRGUARD VALVED INTRODUCER WAS ALLEGEDLY CRACKED AND SILICONE VALVE WAS ALLEGEDLY BROKE WHILE UNSCREWING THE INNER STYLET TO PLACE THE CATHETER THOUGH THE PEEL AWAY SHEATH. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE FULLY PEELED 10.0F INTRODUCER PEEL-APART SHEATH WAS RECEIVED FOR EVALUATION. VISUAL EVALUATION WAS PERFORMED. A BEND AND DEFORMATION WAS NOTED TO THE SAMPLE. THE VALVE CAP WAS NOTED TO BE MISSING FROM ONE OF THE HALF PEEL-APART SHEATH SHAFTS AND RETURNED. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FRACTURE ISSUE. THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED LOSS OF OR FAILURE TO BOND AND DEFORMATION DUE TO COMPRESSIVE STRESS ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 02/2024). H11: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00219 | 2023-04-04 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CRUOS106 | CRUOS106 |  | IT WAS REPORTED THAT PRIOR TO A RECANALIZATION PROCEDURE, THE CATHETER WAS ALLEGEDLY BROKEN WHEN IT WAS OPENED. IT WAS FURTHER REPORTED WHEN THE CATHETER WAS FLUSHED, THERE WAS SOME FLUID COMING OUT OF THE SIDE OF THE CATHETER JUST ABOUT AN INCH FROM THE PORT. REPORTEDLY UPON CLOSER OBSERVATION, THERE WAS A SMALL CRACK IN THE SIDE OF THE CATHETER THAT WAS LEAKING FLUID. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 06/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01141 | 2023-04-04 | 2023-01-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK APHERESIS PORT ACCESS SYS |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE PATIENT EXPERIENCED PAIN. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING INTERVENTION OR MEDICATION TO TREAT PAIN. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02238 | 2023-04-04 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE LOT NUMBERS OF 123422C ON THE PACKAGING CARTON BOTTOM AND SIDE ARE DIFFERENT (MANUFACTURE DATE: 2018/8, EXPIRY DATE: 2023/6), AND THE LOT NUMBER ON THE CARTON BOTTOM WAS MYCVRE43 COMPARED TO MYCVRB43 ON THE SIDE. THE LOT NUMBER OF EACH PRODUCT INSIDE WAS MYCVRE43. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 24MAR2022, IT WAS NOTED THAT THE NO PATIENT INVOLVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02274 | 2023-04-04 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT AFTER THE SURGERY, DUE TO THE BALLOON RUPTURE, THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT IN THE WARD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02276 | 2023-04-04 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER WAS INSERTED WITHOUT ANY PROBLEMS, AND WHEN THE URINE IN THE UROBAG WAS DISCARDED AFTER THE OPERATION, THE BALLOON WAS RUPTURED AND THE FOLEY CATHETER WAS PULLED OUT. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 29MAR2022, IT WAS STATED THAT THE FOLLEY CATHETER FELL OUT NOT DRAGGED OUT. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. VISUAL INSPECTION NOTED SAC WAS IRREGULAR BURST WITHOUT MISSING PIECES. INTRODUCED WATER INTO THE INFLATION LUMEN AND DID NOT EXPERIENCE ANY OBSTRUCTIONS TO IMPEDE FLOW. MICROSCOPIC EXAMINATION FOUND THERE WAS A ADDITIONAL LAYER OF LATEX THAT ADHERED TO THE CATHETER REINFORCEMENT SHAFT. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN PROBLEM OCCURRED COULD NOT DETERMINED. A POTENTIAL ROOT CAUSE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ THIN SAC. A POTENTIAL ROOT CAUSE FOR THIS POTENTIAL FAILURE MODE COULD BE DUE TO THIN SAC DUE LOW LATEX VISCOSITY. A REVIEW OF THE DEVICE HISTORY RECORD COULD NOT BE PERFORMED BECAUSE LOT NUMBER IS UNKNOWN. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® I.C. SILVER FOLEY TRAY B SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02278 | 2023-04-04 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, THE CATHETER BALLOON ASYMMETRICALLY INFLATED AND ALSO WAS DIFFICULT TO REMOVE WATER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT WAS UNCONFIRMED A LABELING REVIEW WAS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02277 | 2023-04-04 | 2023-01-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING TESTING, THE BALLOON WAS INFLATED, BUT THE BALLOON DID NOT DEFLATE WHEN THE SYRINGE WAS PULLED. | THE REPORTED EVENT UNCONFIRMED AS THE PRODUCT MEETS SPECIFICATIONS. VISUAL EVALUATION NOTED RECEIVED A 3-WAY TEMP SENSING CATHETER AND A STRAIGHT TIPPED SYRINGE. REPLACED RETURNED INFLATION VALVE WITH IN-HOUSE VALVE AND RE-ATTEMPTED DEFLATION WITH BOTH SYRINGES. ATTEMPTED TO DEFLATE BALLOON USING RETURNED SYRINGE AND IN-HOUSE INFLATION VALVE AND ONLY 1 ML COULD BE WITHDRAWN. DEFLATED BALLOON WITH IN-HOUSE LUER LOCK SYRINGE IN 3 MINUTES AND 10 SECONDS WITH NO CUFFING OR LEAKS NOTED. DISSECTED BALLOON AND FOUND THAT THE NOTCH WAS PERFORATED COMPLETELY. DISSECTED INFLATION FUNNEL AND PIN GAUGES 0.028¿ AND BELOW COULD FREELY PASS THROUGH LUMEN. 0.029¿ - 0.037¿ TOOK SOME EFFORT TO GET THROUGH AND 0.038¿ AND ABOVE CANNOT PASS THROUGH. THEREFORE, THE PRODUCT MEETS SPECIFICATIONS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE LABELLING REVIEW IS NOT REQUIRED AS THE REPORTED EVENT IS UNCONFIRMED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02279 | 2023-04-04 | 2023-02-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER (SILICONE) | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WERE NO ABNORMALITIES IN THE EXPANSION OF THE BALLOON BEFORE INSERTION, BUT AFTER INSERTION INTO THE PATIENT, THE BALLOON WAS DAMAGED AND SPONTANEOUS WITHDRAWAL OCCURRED FOUR TIMES. ALL THE EVENT WAS CONFIRMED WITH THE SAME LOT AND THE REMAINING ONE WAS COLLECTED WITHOUT BEING USED. THE USER THEN USED 29030J14 DUE TO THE LACK OF 119314. SIMILARLY, THERE WAS A REPORT DRAFTED SEPARATELY FOR SPONTANEOUS WITHDRAWAL DUE TO BALLOON DAMAGE. PER INVESTIGATOR NOTIFICATION RECEIVED VIA EMAIL ON 12-MAR-2023 STATED THAT PER SAMPLES RECEIVED. 8 SAMPLES WERE RETURNED. 4 SAMPLES NOT IN PACKAGING APPEARED TO BE USED WITH MATERIAL NUMBER 119314 AND MANUFACTURING LOT NUMBER NGGS2164. 4 SAMPLES IN UNOPENED PACKAGING WITH MATERIAL NUMBER 119314, 3 WITH BATCH NUMBER NGGT3722, AND 1 WITH BATCH NUMBER NGGS3548. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 15-MAR-2023, STATED THAT THERE WAS NO BALLOON RUPTURE. THE CATHETER HAD A BALLOON PINHOLE. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02282 | 2023-04-04 | 2023-03-14 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT URINARY TRACT INFECTION WITHIN TWO WEEKS OF USING THE PUREWICK FEMALE EXTERNAL CATHETER. IT WAS STATED THAT THE PATIENT GOTTEN THEM IN THE PAST BUT NOT RECENTLY UNTIL USING THIS. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "WARNINGS: DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS." "PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE". " RECOMMENDATIONS: ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02284 | 2023-04-04 | 2023-02-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT ON FEBRUARY 22, DURING THE OPERATION, A CATHETER WAS USED WITHOUT LUBRICANT AND A 15ML OF NORMAL SALINE WAS INJECTED. THEN THE THREE-CAVITY CATHETER WAS CONNECTED TO THE BLADDER FOR CONTINUOUS IRRIGATION. ON FEB 26, THE CATHETERIZATION WAS REQUESTED TO STOP. WHEN THE CATHETER WAS REMOVED FROM THE PATIENT, IT WAS FOUND THAT THE BALLOON WAS RUPTURED. THEN THE PATIENT WAS ABLE TO URINATE ON SELF AND THE URINE WAS RELIEVED SMOOTHLY. THERE WAS NO HARM CAUSED TO THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02286 | 2023-04-04 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, THE CATHETER BALLOON ASYMMETRICALLY INFLATED AND ALSO WAS DIFFICULT TO REMOVE WATER. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR EVENT IS NOT REPORTABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02287 | 2023-04-04 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE TUBE HAD BLACK DOTS OR AIR BUBBLES AND WAS NOT USED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 22MAR2023, THE TUBE WAS A CATHETER. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. BASED ON THE EVALUATION, IT WAS OBSERVED THAT THERE WAS INK OF CATHETER CAP WAS STICK AT THE FRONT SIDE OF DRAINAGE BAG. THIS COMPLAINT CONFIRMED RELATED TO MANUFACTURING ISSUE SINCE THIS IS A PART OF PACKAGING PROCESS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE OPERATOR HANDLING METHOD (DID NOT DIP PROPERLY), SOLVENT OVERFLOW AT DIP STATION OR CLOTH IS WET OF DIP STATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE(S) THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® SILVER TSC TRAY SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02291 | 2023-04-04 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MALE INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THERE WAS NO HOLE IN THE TIP OF THE INTERMITTENT CATHETER UPON INSERTION. PER FOLLOW-UP VIA PHONE ON 02APR2023, IT WAS REPORTED THAT THE CUSTOMER UNDERSTOOD THAT THEY WERE SENT A SUBSTITUTE CATHETER BECAUSE THE ONES THEY NORMALLY USED WERE ON BACKORDER WITHOUT A RELEASE DATE. ALSO STATED THAT MATERIAL #53814G WITH LOT # JUGV1681 HAD NO HOLES IN THE TIP. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. VISUAL INSPECTION RECEIVED ONE (1) MAGIC3 GO IN OPEN PACKAGING WITH NO BURRS, SHARPS, NICKS, OR FLASHES PRESENT. CATHETER DID NOT HAVE ANY EYELETS ON TIP. THEREFORE, PRODUCT DOES NOT MEET SPECIFICATION. ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED, A POTENTIAL ROOT CAUSE COULD BE MECHANICAL FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02295 | 2023-04-04 | 2023-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK MALE EXTERNAL CATHETER KIT | PUREWICK MALE EXTERNAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWM030 |  | IT WAS REPORTED THAT THE MANUFACTURER INSTRUCTIONS DID NOT ADDRESS THE NEED TO DISCONNECT THE PUREWICK MALE EXTERNAL CATHETER FROM SUCTION. STATED THAT THERE WERE INSTANCES WHERE THIS DEVICE WOULD NEED TO BE DISCONNECTED FROM SUCTION IN A REAL-LIFE SITUATION SUCH AS THERAPY, AMBULATION, TRANSPORTING TO AND FROM IMAGING. IN THESE CASES, THEY COULD NOT LOOK AT THE MANUFACTURER INSTRUCTIONS FOR ANY GUIDANCE FOR WHAT TO DO. IF THEY JUST DISCONNECT THE TUBING, IT COULD LEAD TO THE URINE COMING OUT OF THE DEVICE IF THE PATIENT VOIDS WHICH WAS A POSSIBILITY SINCE THESE COULD BE USED FOR INCONTINENCE MANAGEMENT. THIS LEADS TO THEIR NURSES TRYING TO GET CREATIVE TO SOLVE THE PROBLEM OF DRIBBLING OR FLOW OF URINE. THE MOST LOGICAL IDEA BEING, PLUG THE PORT SO URINE DID NOT COME OUT AND HOOK BACK UP TO SUCTION UPON RETURN. STATED THAT THEY WOULD LIKE TO BE ABLE TO UNHOOK PUREWICK FROM SUCTION TO MOVE PATIENT WITH THERAPY. CURRENTLY THERE WERE NO GUIDELINES IN THE MANUFACTURER INSTRUCTIONS ON HOW TO DO THAT, NO WAY TO CAP OR PLUG THE PUREWICK SO EITHER THE URINE LEAKS OUT THE BOTTOM OF THE PUREWICK WHILE THE PATIENT WAS UP, OR THEY NEED TO TAKE OFF AND PUT A NEW PUREWICK ON AFTER WHICH COULD CAUSE SKIN ISSUES AND EXCESSIVE USE OF SUPPLIES LEADING TO INCREASED COST AND WASTE. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT/ MISSING TRANSLATION; MISSING INSTRUCTIONS; VENDOR/PRINTER ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: " CONTRAINDICATIONS: PATIENTS WITH URINARY RETENTION. WARNINGS: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. DO NOT COVER FRESH SURGICAL WOUNDS WITH THE DEVICE. DO NOT USE THE DEVICE WITH ANY MATERIAL OR IN ANY POSITION THAT DOES NOT ALLOW FOR SUFFICIENT AIRFLOW. TO AVOID POTENTIAL SKIN INJURY, NEVER PULL THE DEVICE DIRECTLY AWAY FROM THE PATIENT. ALWAYS PEEL IN THE DIRECTION FROM HEAD TO FOOT. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE, AND FEDERAL LAWS AND REGULATIONS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE DEVICE. HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE. DO NOT USE BARRIER CREAM ON THE PENIS OR PUBIC SKIN AROUND THE BASE OF THE PENIS WHEN USING THE DEVICE. CREAMS MAY IMPEDE SUCTION OR WEAKEN ADHESIVE. PROCEED WITH CAUTION IN PATIENTS WHO HAVE UNDERGONE RECENT SURGERY OF THE EXTERNAL UROGENITAL TRACT, PENIS, OR PUBIC AREA. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW DEVICE. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. PRIOR TO CONNECTING THE DEVICE TO HOSPITAL WALL SUCTION TUBING, VERIFY SUCTION FUNCTION BY COVERING THE OPEN END OF THE SUCTION TUBING WITH ONE HAND AND OBSERVING THE PRESSURE DIAL. IF THE PRESSURE DOES NOT INCREASE WHEN THE LINE IS COVERED, VERIFY THAT THE TUBING IS SECURED, CONNECTED, AND NOT KINKED. ENSURE THE DEVICE REMAINS CONNECTED AFTER TURNING THE PATIENT, MONITORING FOR PULLING AND TENSION ON THE DEVICE. REMOVE THE DEVICE PRIOR TO AMBULATION. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS. SETUP: CONNECT A CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. USING STANDARD SUCTION TUBING, CONNECT THE DEVICE TO A COLLECTION CANISTER. NOTE: COLLECTION CANISTER AND SUCTION TUBING ARE NOT INCLUDED. IF HOSPITAL POLICY ALLOWS AND IF USING A GRADUATED CANISTER, CAPTURED URINE MAY BE USED FOR APPROXIMATE URINE OUTPUT MEASUREMENT. PERI-CARE AND PLACEMENT: IF THERE IS SIGNIFICANT PUBIC HAIR, TRIM OR CLIP THE PUBIC HAIR. PERFORM PERINEAL CARE USING THE INCLUDED WIPES AND ASSESS SKIN INTEGRITY. FOLLOW AND DOCUMENT PER HOSPITAL PROTOCOL. NOTE: USE THE INCLUDED DRYING WIPE TO ENSURE SKIN IS DRY PRIOR TO PLACEMENT OF THE DEVICE. MOISTURE ON SKIN WILL WEAKEN THE ADHESIVE. ORIENT THE DEVICE, ALIGNING DRAINAGE FITTING TOWARDS THE FEET OF THE PATIENT. SLIDE THE OPENING OF THE DEVICE OVER THE PENIS UNTIL CLOSE TO THE PELVIC SKIN. CENTER PENIS WITHIN THE OPENING. REMOVE ADHESIVE LINER AND PRESS THE DEVICE AGAINST THE PELVIC SKIN TO ADHERE. ENSURE BASE HAS FULLY ADHERED AROUND THE BASE OF THE PENIS. NOTE: THE SCROTUM SHOULD NOT BE PLACED INSIDE THE DEVICE. REMOVAL: TO REMOVE THE DEVICE, GENTLY LIFT THE TOP EDGE OF THE DEVICE OFF THE SKIN. IN A SLOW, DOWNWARD PEELING MOTION, REMOVE THE DEVICE. IF NECESSARY, USE A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) TO HELP LOOSEN ADHESIVE. WARNING: TO AVOID POTENTIAL SKIN INJURY, NEVER PULL THE DEVICE DIRECTLY AWAY FROM THE PATIENT. ALWAYS PEEL IN THE DIRECTION FROM HEAD TO FOOT. MAINTENANCE: REPLACE THE DEVICE AT LEAST EVERY 24 HOURS OR IF SOILED WITH FECES, BLOOD, OR SEMEN." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02296 | 2023-04-04 | 2023-03-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT 6 DAYS AFTER PLACEMENT, THE CATHETER BROKE AND FELL OUT OF THE PATIENT. THE CUSTOMER REQUESTED TO INVESTIGATE THE BROKEN SURFACE AND TO CONFIRM THAT THE BREAKAGE WAS CAUSED BY THE MANUFACTURING OR NOT. THE EVENT OCCURRED AROUND 16:45 ON MARCH 12, 2023. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 20MAR2023, THE CATHETER GOT BROKEN, HOWEVER, NO BROKEN PIECE LEFT IN THE PATIENT BODY AT THE BREAKAGE. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. SAMPLE WAS EVALUATED AND OBSERVED THE CATHETER WAS CUT OFF AT SHAFT. THE SURFACE WAS NORMAL IN APPEARANCE WITH THE PRESENCE OF SMOOTH AND CLEAR CUT. THE CATHETER SHAFT LOOKS LIKE HAS BEEN CUT WITH APPLIED SOME FORCE THAT COULD CAUSE THE CATHETER BREAK. NONE OF THE CONDITIONS ABOVE WERE EVIDENT ON THE SAMPLE. THE EXACT CAUSE ON HOW AND WHEN PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED.( EXAMPLE: ROUGH HANDLING OR USE SHARP OBJECT). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02297 | 2023-04-04 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® FLIP-FLO¿ CATHETER VALVE | CATHETER VALVE | C.R. BARD, INC. (COVINGTON) -1018233 | BFF5 | BFF5 |  | IT WAS REPORTED THAT THE CUSTOMER HAD BROUGHT SOME FLIP FLOW CATHETER VALVES AND STATED THAT THEY WERE NOT DRAINING PATIENT'S BLADDER UNLESS ATTACHED TO A BAG. STATED THAT THEY HAD PROBLEMS WITH THEIR CURRENT CATHETER VALVES AS IT WAS NOT EMPTYING UNLESS ATTACHED TO A LEG OR NIGHT BAG. WHEN THE TAP WAS OPENED, THERE WAS NO FLOW BUT, AS SOON AS THEY ATTACHED IT TO A BAG, THEIR BLADDER EMPTIED. STATED THAT THEY PREVIOUSLY HAD VALVES WHICH WOULD EMPTY THEIR BLADDER VIA THE CATHETER DIRECTLY JUST BY TURNING THE LEVER. THE CUSTOMER ASKED WHETHER THERE WAS SOMETHING WRONG WITH THE VALVE OR WAS THERE ANOTHER POSSIBLE REASON FOR THIS. THE CUSTOMER HAD TRIED SEVERAL WITH THE SAME RESULT. THE REPRESENTATIVE GAVE THEM THE NURSE ADVICE LINE AND ADVISED THAT THEY MIGHT BE ABLE TO HELP. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INSPECTION RESULTS (MEASUREMENT, TESTING DATA) NOT VERIFIED BY IQA ASSIGNEE ERROR OF INSPECTOR". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "1. WASH HANDS. 2. ATTACH FLIP-FLO¿ VALVE TO CATHETER ENSURING THAT YOU DO NOT TOUCH EITHER THE CATHETER END OR THE VALVE CONNECTOR. 3. WASH HANDS THOROUGHLY BEFORE AND AFTER OPERATING FLIP-FLO¿ VALVE. 4. EMPTY BLADDER AS FREQUENTLY AS ADVISED BY YOUR DOCTOR OR NURSE. 5. TO OPEN THE FLIP-FLO¿ VALVE, PUSH LEVER DOWN. 6. ALLOW URINE TO DRAIN INTO TOILET OR AN APPROPRIATE RECEPTACLE. 7. TO CLOSE THE FLIP-FLO¿ VALVE, PULL LEVER UP. 8. A NIGHT DRAINAGE BAG MAY BE ATTACHED TO THE FLEXIBLE CONNECTOR TO ALLOW CONTINUOUS URINE DRAINAGE OVERNIGHT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01145 | 2023-04-04 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D2 |  | IT WAS REPORTED BY THE CUSTOMER, "WE EXPERIENCED LEAKING FROM AROUND THE WIRE AS FLUSH WAS DONE THROUGH THE SIDE PORT EVEN AFTER THE HUB CONNECTION WAS CHECKED. THE FLUID IS COMING DIRECTLY FROM THE WIRE PUNCTURE THROUGH THE RUBBER INJECTION SITE. THERE WAS NO REPORTED SAFETY CONCERN FOR CLINICIAN." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED AND THE CAUSE OF THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE SAMPLE RETURN. EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02283 | 2023-04-04 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | MALE INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THERE WAS A CUT FROM THE SELF-INSERTING CATHETER THAT CAUSED PROFUSE BLEEDING AND NOTED BURRED/SHARP TIP UPON INSERTION OF THE INTERMITTENT CATHETER. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW-UP VIA PHONE ON 02APR2023, IT WAS REPORTED THAT THE CUSTOMER UNDERSTOOD THAT THEY WERE SENT A SUBSTITUTE CATHETER BECAUSE THE ONES THEY NORMALLY USED WERE ON BACKORDER WITHOUT A RELEASE DATE. ALSO STATED THAT MATERIAL #53814G WITH (BATCH # JUGV1681) HAD NO HOLES IN THE TIP. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "MANDRELS TOUCHING EACH OTHER WITHIN THE PALLET." IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION; BLEEDING FROM THE URETHRA; IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02289 | 2023-04-04 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS INDWELLED AFTER THE OPERATION, IT FELL OFF SHORTLY AFTER MOVING TO THE WARD. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. BASED ON THE EVALUATION, A TEAR WAS FOUND ON THE RUBBERIZED LAYER CAUSING INTERLUMEN LEAKAGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON UNIFORM RUBBERIZED THICKNESS ETC CAUSING TORN RUBBERIZE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® BIOCATH® FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D,H. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01142 | 2023-04-04 | 2023-03-03 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808561 | 1808561 |  | IT WAS REPORTED THAT EIGHTEEN DAYS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED BODY SPASMS, HEART PALPITATIONS AND PAIN ON LEFT SIDE. IT WAS FURTHER REPORTED THAT WHEN THE PORT WAS ATTEMPTED TO FLUSH, THE SALINE WAS NOTED TO BE LEAKING AROUND THE PORT PLACEMENT SITE. IT WAS ALSO REPORTED THAT UPON IMAGING, IT WAS CONFIRMED THAT THE TIP OF THE PORT ALLEGEDLY BROKE AND THE PORT TIP ALLEGEDLY PIERCED RIGHT VENTRICLE. REPORTEDLY, THE PATIENT UNDERWENT EMERGENCY SURGERY WHERE THE PORT AND BROKEN PIECE WAS REMOVED, AND RIGHT VENTRICLE WAS TREATED. THE DETAILS ON THIS SURGERY WERE NOT AVAILABLE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 09/2025. DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01144 | 2023-04-04 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT MRI ISP, IMPLANTABLE PORT, GROSHONG SINGLE LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE DEVICE ALLEGEDLY HAD A BLOCKAGE. THE PATIENT ALLEGEDLY EXPERIENCED SWELLING UPON FLUSHING. THERE WAS NO REPORTED PATIENT INJURY. | H10: THIS SUPPLEMENTAL MDR IS BEING SUBMITTED TO REPORT THAT MFR RPT# 3006260740-2023-01144 WAS A DUPLICATE RECORD AND WAS OPENED IN ERROR. THE EVENT DETAILS ARE BEING CAPTURED UNDER COMPLAINT FILE #(B)(4) AND WAS REPORTED TO THE FDA UNDER MFR RPT# 3006260740-2023-01121. H10: D4 (EXPIRY DATE: 03/2023). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3012307300-2023-03459 | 2023-04-04 | 2023-03-14 | I, n, j, u, r, y | Manufacturer report | DELTEC POWER PORT-A-CATH II VENOUS ACCESS | PORT-CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCU | ST PAUL | 21-4485-24 | 21-4485-24 |  | IT WAS REPORTED THAT THE CATHETER ATTACHED TO PORT CHAMBER LEAKED. THE LEAK WAS DISCOVERED DURING EVALUATION OF IMPLANTED DEVICE IN-SITU VIA CONTRAST INJECTION. SEDATION WAS PROVIDED, THE DEFECTIVE ASSEMBLY WAS REMOVED AND A NEW DEVICE WAS INSTALLED. NO PATIENT INJURY. | DEVICE EVALUATION: ONE DEVICE WAS RETURNED FOR INVESTIGATION. VISUAL INSPECTION FOUND A CUT WAS IDENTIFIED ON THE CATHETER, CONFIRMING THE COMPLAINT. A LEAK WAS DETECTED DURING FUNCTIONAL TESTING. A ROOT CAUSE CANNOT BE ASSOCIATED WITH MANUFACTURING PROCESS SINCE THE FAILURE COULD NOT BE REPRODUCED FOLLOWING ASSEMBLY PROCESS. THE MOST PROBABLE ROOT CAUSE WAS THAT THE DAMAGE OCCURRED AFTER THE PRODUCT LEFT THE MANUFACTURING FACILITIES. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) SHOWS THERE WERE NO OBSERVATIONS OR NONCONFORMITIES RECORDED DURING MANUFACTURE TO SUGGEST AN ISSUE OF THIS NATURE WOULD OCCUR WITH THIS LOT OF PRODUCTS. NO CORRECTIVE ACTIONS ARE REQUIRED SINCE THE ROOT CAUSE COULD NOT BE ATTRIBUTED TO THE MANUFACTURING PROCESS. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01175 | 2023-04-05 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT TWO YEARS TEN MONTHS AND TWENTY DAYS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY HAD SWELLING ABOVE THE PORT DURING FLUSHING. IT WAS FURTHER REPORTED THAT THERE WAS NO BLOOD RETURN WITH SLIGHT TENDERNESS UPON FLUSHING. FURTHERMORE, A CHEST X-RAY WAS TAKEN, AND BLURRY CATHETER WALL WAS FOUND. REPORTEDLY, THE CATHETER WAS ALLEGEDLY BROKEN AT APPROXIMATELY 4 CM ABOVE THE PORT CONNECTION AND CATHETER WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. ALSO, TWO PHOTOS AND TWO IMAGES WERE PROVIDED FOR REVIEW. THE IMAGES SHOW A BLURRY VIEW OF THE CATHETER PASSING THE CLAVICLE REGION. THE PHOTOS SHOW THE FRACTURE IN THE CATHETER. VISUAL, GROSS VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPOUND BREAK WAS NOTED ON THE CATHETER AND A SMALL SPLIT WAS NOTED IN ONE OF THE BREAK REGIONS. THE EDGES OF THE COMPOUND BREAK WERE NOTED TO BE UNEVEN AND THE SURFACE WAS NOTED TO BE ROUND AND SMOOTH IN BOTH REGIONS. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND THE IDENTIFIED WEAR AND DEFORMATION ISSUES. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED SUCTION ISSUE AS NO OBJECTIVE EVIDENCE WAS PROVIDED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 01/2022). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01174 | 2023-04-05 | 2021-07-14 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 4808560 | 4808560 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE DEVICE ALLEGEDLY BROKE AND MIGRATED INTO INFERIOR VENA CAVA. IT WAS FURTHER REPORTED THAT THE PHYSICIAN HAD CUT THE STYLET. FURTHERMORE, PATIENT EXPERIENCED HEART PALPITATIONS. REPORTEDLY, THE DEVICE WAS REMOVED. THE CURRENT PATIENT STATUS WAS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: H6 (METHOD). H11: H6 (DEVICE). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01172 | 2023-04-05 | 2021-07-23 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 4808560 | 4808560 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE DEVICE ALLEGEDLY BROKE AND MIGRATED TO LUNG AND EVENTUALLY TO THE PERICARDIAL SPACE. REPORTEDLY, THE DEVICE WAS REMOVED. THE CURRENT PATIENT STATUS WAS UNKNOWN. | H10: THE INITIAL MDR WAS INADVERTENTLY SUBMITTED WITH A G3 DATE OF 03/08/2023. THE CORRECT G3 DATE IS 03/07/2023. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED ISSUES, AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01188 | 2023-04-05 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118J |  | IT WAS REPORTED THAT THE PATIENT HAD A POWERPICC IMPLANTED VIA THE RIGHT UPPER ARM ON (B)(6) 2023. ON (B)(6) 2023, WHEN THE DRESSING WAS REMOVED BECAUSE THE DRESSING WAS PEELING OFF AND ATTEMPTED TO BE CHANGED, IT WAS FOUND THAT THE CATHETER WAS DAMAGED. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED. ONE 5 FR D/L POWERPICC WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUES THROUGHOUT THE RETURNED DEVICE, AND NEEDLELESS INJECTION CAPS RETURNED ATTACHED TO EACH LUER OF THE DEVICE. RESIDUES FROM SECUREMENT TECHNIQUES WERE DRIED ALONG THE PROXIMAL END OF THE CATHETER SHAFT. THE POWERPICC WAS FUNCTIONALLY TESTED BY FLUSHING EACH LUMEN WITH WATER IN A 12 ML SYRINGE AFTER PRESSURIZING THE PICC, AND THE DEVICE LEAKED AT THE 4 CM DEPTH MARKER DURING THIS TEST ALONG THE CATHETER SECUREMENT RESIDUES. A MICROSCOPIC OBSERVATION REVEALED A CIRCUMFERENTIALLY ALIGNED SPLIT IN THE CATHETER AT THE 4 CM DEPTH MARKER. THE SPLIT WAS OBSERVED AT THE SECUREMENT AREA OF THE CATHETER WHERE THE GLUE WAS STUCK ONTO THE DEVICE. THE CATHETER AROUND THE SPLIT HAD WHITENING AROUND THE EDGES OF THE SPLIT AND A GRANULAR FRACTURE SURFACE. THE SPLIT APPEARED TO HAVE ROUGH SURFACE EDGES AS WELL. POLISHED MARKS LIKELY FROM SECUREMENT TECHNIQUES WERE OBSERVED CIRCUMFERENTIALLY ALONG THE EDGES OF THE SPLIT. THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED AND WAS LIKELY DUE TO MATERIAL FATIGUE, POTENTIALLY RELATED TO SECUREMENT TECHNIQUES. OBSERVATIONS OF THE SURROUNDING AREA OF THE SPLIT AND THE LOCATION OF THE SPLIT SUGGEST SECUREMENT COULD CONTRIBUTE TO THE FAILURE; HOWEVER OTHER MODES OF FAILURE MAY HAVE CONTRIBUTED SUCH AS KINKING OR SOME MATERIAL FATIGUE. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01187 | 2023-04-05 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, PRIOR TO USE ON PATIENT HOLE WAS FOUND IN PICC LINE. ADDITIONAL INFORMATION RECEIVED 24-MAR-23: IT WAS REPORTED BY THE CUSTOMER "INSPECTION OF PICC LINE (B)(6) 2023 DUE TO NO VENOUS BACKFLOW DURING SAMPLING. DRESSING THAT MOISTENS WHEN INFUSING NACL, NO VENOUS BACKFLOW, TAKING OUT PICC LINE AND CATHETER DEFECT AT 6CM MARK ON PICC LINE. DAMAGE TO CATHETER WHERE NACL HEALS OUT, DAMAGE LOCATED 4CM FROM SKIN LEVEL. DEVICE CONTAMINATION: NO CHEMO, ONLY ANTIBIOTIC. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED, BUT THE EXACT CAUSE OF FAILURE COULD NOT BE DETERMINED. ONE 4 FR S/L POWERPICC SOLO CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUES THROUGHOUT THE RETURNED DEVICE, AND THE GROSS VISUAL INSPECTION WAS UNREMARKABLE. A TACTILE EVALUATION OF THE DEVICE REVEALED THE COMPLAINANT CATHETER TO SHOW NO SIGNS OF TENSILE WEAKNESS. A MICROSCOPIC OBSERVATION REVEALED A CIRCUMFERENTIALLY ALIGNED SPLIT JUST DISTAL OF THE MOLDED JOINT. THE FRACTURE SURFACE WAS GRANULAR AND THE FRACTURE EDGES WERE SHARP. WHITENING OCCURRED AT THE CORNERS OF THE SPLIT, BUT NO ROUNDING OR POLISHED EDGES WERE FOUND ALONG THE DEVICE. NOTHING REMARKABLE WAS OBSERVED OPPOSITE THE SPLIT. THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED; HOWEVER, THE EXACT CAUSE OF FAILURE COULD NOT BE DETERMINED. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS 'CAUSE UNKNOWN' AT THIS TIME. THE GRANULAR SURFACE TEXTURE WAS CONSISTENT WITH MATERIAL TEARING, WHICH SUGGESTED THAT MECHANICAL FORCES CONTRIBUTED TO THE SPLIT; HOWEVER, THE UNUSUAL SHAPE OF THE SPLIT SUGGESTED THAT AN ADDITIONAL UNIDENTIFIED FACTOR(S) ALSO CONTRIBUTED. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01186 | 2023-04-05 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE STRAIN RELIEF WIRES ARE EXPOSED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01180 | 2023-04-05 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC HF CATHETER 5F RADIOLOGY BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3385355Q |  | IT WAS REPORTED BY THE CUSTOMER "AFTER INSERTION OF TRIPLE LUMEN PICC, ASPIRATED AND FLUSHED LINE AND AIR WAS PULLED BACK." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02315 | 2023-04-05 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT SAME AS LAST TIME 23020607-IT WAS REPORTED THAT THE URINE LEAKED FROM THE INFLATION LUMEN IN A LARGER AMOUNT ON FEBRUARY 7TH AND IN A SMALLER AMOUNTS ON FEBRUARY 12TH , 19TH AND 21ST. ON FEBRUARY 13TH, URINE LEAKAGE HAS OCCURRED. ON 21ST FEBRUARY, WHEN THE CATHETER WAS REMOVED, AIR CAME INTO THE SYRINGE VIGOROUSLY. ATTEMPTING SEVERAL TIMES TO REMOVE WATER IN THE BALLOON, IT WAS DIFFICULT TO REMOVE. AFTER ABOUT THE FOURTH ATTEMPT OF REMOVAL, PULLING THE SYRINGE PLUNGER SLOWLY, 6 ML OF FIXATION WATER WAS REMOVED. HOWEVER, THE BALLOON REMAINED SLIGHTLY INFLATED AFTER THE CATHETER REMOVAL, IT WAS ALSO REPORTED THAT EVEN THOUGH THE SYRINGE WAS CONNECTED TO THE INFLATION VALVE OF THE FOLEY CATHETER, THE WATER IN THE BALLOON COULD NOT BE DRAINED EASILY. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING:"[WARNINGS]1. METHOD FOR USE(1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.](2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.]2. APPLICABLE PATIENTSPATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]1. METHOD FOR USE:(1) DO NOT REUSE.(2) DO NOT RESTERILIZE.(3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS.(4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.](5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.]2. APPLICABLE PATIENTS(1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX(2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER[SHAPE, CONFIGURATION AND PRINCIPLES]BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY.THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT.<MATERIAL>BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATINGTHIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING.<SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR1. FOLEY CATHETER 2. ACCESSORIESCLOSED DRAINAGE BAG(THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATERWATER SOLUBLE LUBRICANTANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTIONTWEEZERSGAUZE PADSWATERPROOF SHEETCOTTON BALLS GLOVESSTATLOCK® FOLEY[INTENDED USE & EFFECT- EFFICACY]THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG.[DIRECTIONS FOR USE]1. METHOD OF USETHE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE.(1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER.(2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1)(3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2(4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3)(5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4(6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5(7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT.(8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6)(9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7(10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED.<DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT>1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT.2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE.3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 84) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION.5) UNKINK TUBING.<HOW TO DISCONNECT CATHETER FROM TUBING>CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 92. PRECAUTIONS FOR USE(1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER.(2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.](3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ](4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL.(5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL.(6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY.(7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN.(8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL.(9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.](10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING;1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG.2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER.3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN.2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL.3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE.4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL.5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS.6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT.7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS]1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS)(1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR.2. IMPORTANT PRECAUTIONS(1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER.(2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE.(3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿.<TROUBLESHOOTING>WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES.THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES.A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.)B. BALLOON-RUPTURE METHODWITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST.A. NON-RUPTURE METHOD1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER.2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 103) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 114) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 125) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13B. BALLOON-RUPTURE METHOD1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 142) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES.A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 173. MALFUNCTION AND ADVERSE EVENTS1) MALFUNCTION- CATHETER KINKING, DAMAGE, RUPTURE- DIFFICULTY OR FAILURE TO REMOVE THE DEVICE- OCCLUSION OF CATHETER INNER LUMENS- ENCRUSTATION- ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE2) ADVERSE EVENTS- URINARY-TRACT INFECTION- HEMORRHAGE, HEMATURIA- ALLERGY REACTION TO THE DEVICE- CALCULUS FORMATION - EDEMA- PAIN- DISCOMFORT- INJURY OF BLADDER OR URETHRAL- URETHRITIS, URINARY INCONTINENCE- RETAINED BALLOON FRAGMENTS[STORAGE METHOD AND EXPIRATION DATE]1. STORAGESTORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT.2. EXPIRATION DATEINDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX."CORRECTION: D.H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02318 | 2023-04-05 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DIFFICULT TO DRAIN THE WATER FROM THE FOLEY CATHETER BALLOON DURING THE PRETEST. IT WAS STATED THAT THE WHEN USER INFLATED THE BALLOON, IT WAS POINTED OUT THAT THE INFLATION WAS ONE-SIDED, CLOSE TO 10:00. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON CONCENTRICITY WAS OBSERVED TO BE 100:00. THE BALLOON PASSIVELY DEFLATED IN 3 MINUTES AND 21 SECONDS WITH NO CUFFING OR LEAKS NOTED. THIS IS OUT OF SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE FOREIGN MATTER DURING PROCESSES (FLAP OR FLASH DURING PUNCHING). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS: ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING. THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR. 1. BALLOON CATHETER, FOLEY CATHETER, TEMPERATURE-SENSING CATHETER. 2. ACCESSORIES: CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE: (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE: (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2: (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS. 1) MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE ; DIFFICULTY OR FAILURE TO REMOVE THE DEVICE ; OCCLUSION OF CATHETER INNER LUMENS ; ENCRUSTATION ; ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE ; DEVICE DAMAGE DUE TO INAPPROPRIATE USE ; FAILURE TO MEASURE TEMPERATURE ; IMPROPER TEMPERATURE INDICATION. 2) ADVERSE EVENTS: URINARY-TRACT INFECTION, HEMORRHAGE, HEMATURIA, ALLERGY REACTION TO THE DEVICE, CALCULUS FORMATION, EDEMA, PAIN, DISCOMFORT, INJURY OF BLADDER OR URETHRAL, URETHRITIS, URINARY INCONTINENCE, RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16681522 | 2023-04-05 | 2023-03-17 | I, n, j, u, r, y | User Facility report | BARDPORT, SLIMPORT, X-PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | BARD ACCESS SYSTEMS, INC. |  |  |  | LITERATURE REVIEW AND MIFUS INCLUDED IN THIS LINE CORPORATE SUPPLY CHAIN TO ESCALATE AN ISSUE BROUGHT TO US BY OUR CLINICAL LEADERS, RISK MANAGEMENT, AND REGULATORY. SPECIFICALLY, THEY ARE LOOKING TO ESCALATE THE ISSUE THAT YOUR MIFU INDICATES THAT HEPARIN FLUSHES BE USED TO MAINTAIN THE PORT PATENCY. THE CLINICAL LEADERSHIP HAVE REVIEWED THE LITERATURE ON BEST PRACTICE COMING OUT OF ONCOLOGY NURSING SOCIETY AND INS NURSING SOCIETIES AND WOULD LIKE TO FOLLOW THE RECOMMENDATION OF REMOVING HEPARIN FROM THE PORT LINE FLUSHES. PROPOSED CHANGE: CURRENT PRACTICE - TO FLUSH WITH SALINE THEN HEPARIN FOR ALL NON-VALVED PORTS. PROPOSED PRACTICE CHANGE - TO FLUSH WITH ONLY NORMAL SALINE IN A PULSATILE METHOD. THE ISSUE REQUIRING YOUR ASSISTANCE: THE ISSUE IS THAT THE INS AND ONS GUIDELINES TO NO LONGER USE HEPARIN WHEN FLUSHING IMPLANTED PORTS DOES NOT MATCH THE MIFUS. A DIFFERENT MANUFACTURER, ANGIODYNAMICS, OFFERS OPTIONALITY BETWEEN HEPARIN OR SALINE FOR MAINTAINING PATENCY. NEXT STEPS: WE WOULD NEED THE MIFU UPDATED IN ORDER TO EXECUTE THIS CHANGE AND STANDARDIZE PRACTICE ACROSS OUR HEALTH SYSTEM. THIS HAS BEEN CLEARED WITH NURSING SHARED GOVERNANCE, SMILOW ONCOLOGY NURSING LEADERSHIP, NON-ONCOLOGY NURSING LEADERSHIP (NEUROSCIENCES, HOIT, INTERVENTIONAL IMMUNOLOGY), PHARMACY, AND THE EDUCATION CONSORTIUM. THERE ARE CONCERNS FOR PATIENT SAFETY IF WE STAY WITH THE CURRENT PRACTICE, AS WELL AS A SIGNIFICANT COST SAVINGS ASSOCIATED WITH THE PRACTICE CHANGE. I HAVE INCLUDED THE LITERATURE CITED BY THE TEAMS PROPOSING THIS CHANGE FOR YOUR REFERENCE BELOW AS WELL AS ALL MIFUS INCLUDING THE ANGIODYNAMICS ONE THAT OFFERS OPTIONALITY. MANUFACTURER RESPONSE FOR PORT ACCESS DEVICES, CONVERSATION STARTED WITH BARD, THEY ARE DISCUSSING INTERNALLY. THE REP SAID THIS IS NEW INFORMATION THEY HAVE BEEN WORKING ON, BUT MAY OFFER PRODUCT ALTERNATIVE RATHER THAN CHANGING THE MIFU TO REFLECT THE LATEST RESEARCH. REP IS WORKING ON PUTTING TOGETHER A CALL WITH BARD MEDICAL AND OUR YNHHS STAKEHOLDERS. |  | Y | N |  |  | 2023-04
1018233-2023-02311 | 2023-04-05 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT THE FOLEY WAS INSERTED IN EMERGENCY DEPARTMENT (ED) ON 07MAR AND REMOVED FOLEY ON (B)(6) 2023. STATED THAT THE FOLEY CATHETER HAD AN ISSUE AS IT WAS BEING REMOVED FROM THE PATIENT, SO THEY HAD TO CUT THE INFLATION LUMEN TO DEFLATE THE FOLEY BALLOON. | THE REPORTED EVENT WAS CONFIRMED HOWEVER THE CAUSE WAS UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. THE POTENTIAL FAILURE MODES ARE COLLAPSED LUMEN AND BLOCKED LUMEN. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER A SYRINGE. VISUAL INSPECTION NOTED THAT INFLATION CAP AND VALVE WERE CUT OFF UPON RECEIVAL OF SAMPLE. DISSECTED BALLOON AND FOUND THAT THE NOTCH WAS PERFORATED COMPLETELY. DISSECTED INFLATION FUNNEL AND PIN GAUGES 0.031¿ AND BELOW COULD FREELY PASS THROUGH LUMEN. 0.032¿ - 0.042¿ TOOK SOME EFFORT TO GET THROUGH AND 0.043¿ AND ABOVE CANNOT PASS THROUGH. AN IN-HOUSE CATHETER WAS USED TO TEST FOR VALVE FAILURE. THE RETURNED VALVE WAS PLACED INTO THE IN-HOUSE CATHETER AND WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. DEFLATED IN-HOUSE BALLOON PASSIVELY WITH THE RETURNED VALVE SUCCESSFULLY WITH NO CUFFING OR LEAKS NOTED. BASED ON PHYSICAL SAMPLE RECEIVED THE PRODUCT DOES NOT MEET SPECIFICATIONS ACCORDING TO INSPECTION PROCEDURE WHICH STATES "SHAFT MUST NOT BE DAMAGED OR PINCHED." A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE: ¿ INCORRECT DRAIN FORMER SELECTION ¿ GLUE FROM TEMP SENSING ASSEMBLY ¿ PHYSIOLOGICAL OBSTRUCTION ¿ PATIENT POSITION ¿ KINKED CATHETER ¿ BLOCKED BY CLOTS, ETC. ¿ TUBING DESIGN (LUMEN ID UNDERSIZED) ¿ LUMEN WALL THICKNESS UNDERSIZED ¿ INADEQUATE MATERIAL SELECTION ¿ LUMEN COLLAPSE ¿ FOREIGN MATTER IN LUMEN ¿ KINKED OR PINCHED SHAFT THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02314 | 2023-04-05 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE IRRIGATION AND FLUSHING FAILED WHILE USING FOLEY CATHETER. THE USER CHANGED THE DEVICE TO A NEW ONE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. HOWEVER, THE POTENTIAL ROOT CAUSE COULD BE MANUFACTURING RELATED DUE TO NO IRRIGATION EYE OR POOR IRRIGATION EYE CREATED DURING EYE BURNING PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02306 | 2023-04-05 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON DID NOT DEFLATE. THE FOLEY CATHETER WAS CUT OFF. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE THIN RUBBERIZE WALL THAT CAUSING COLLAPSE/PINCH INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT¿ AND MIGHT BE CONTRIBUTED BY USER RELATED, EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02304 | 2023-04-05 | 2023-03-14 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT WHEN 10CC OR MORE OF WATER WAS INJECTED, THE BALLOON BURST. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 23MAR2023, THE BALLOON OVERFILLED TO INFLATE. AS PER THE NOTIFICATION RECEIVED ON 04APR2023, THE BALLOON HAS RUPTURED DURING USE ON 14TH MARCH AND THE PATIENT DEVELOPED CYSTITIS, AND IT WAS FOUND THAT THERE WAS A BROKEN PIECE OF THE DAMAGED PIECE LEFT INSIDE THE PATIENT'S BODY. REMOVAL OF THE REMAINING FRAGMENTS HAS BEEN COMPLETED, AND THERE WERE NO FRAGMENTS LEFT IN THE BODY AT PRESENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01177 | 2023-04-05 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY THE CUSTOMER, "ACCUCATH ACE INTRAVASCULAR CATHETER 20G RETRACTING NEEDLED DEVICE DID NOT RETRACT. STAFF HAD TO MANUALLY PULLED THE NEED OUT. STAFF WAS SUCCESSFUL WITH NO INJURY TO THAT PATIENT OR STAFF." 03/17/2023 ADDITIONAL INFORMATION: IT WAS REPORTED BY THE CUSTOMER "INSERTION WAS COMPLETED. THE SAFETY ISSUE WAS MOSTLY WITH RISK TO THE STAFF BECAUSE THE STAFF HAD TO MANUALLY RETRACT THE NEEDLE." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01176 | 2023-04-05 | 2022-02-09 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR AND SIX MONTHS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED MILD ERYTHEMA AND SWELLING ALONG THE CATHETER AFTER LAST CYCLE OF CHEMO. IT WAS FURTHER REPORTED THAT HYPERPIGMENTATION AND SCARRING WAS NOTED ALONG THE CATHETER. FURTHERMORE, AFTER INFUSION, SLIGHT PALPABLE PUFFINESS WAS NOTED ADJACENT TO LINE AROUND RIGHT CLAVICLE. REPORTEDLY, THE CATHETER WAS BROKEN AT POSTERIOR SIDE OF THE CATHETER AND THE CATHETER WAS REMOVED AND REPLACED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION AND ONE IMAGE AND TWO PHOTOS WERE PROVIDED FOR REVIEW. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPOUND BREAK WAS NOTED ON THE ATTACHED CATHETER. THE EDGES OF THE COMPOUND BREAK ON THE ATTACHED CATHETER WERE NOTED TO BE UNEVEN. THE SURFACE WAS NOTED TO BE ROUND AND SMOOTH IN BOTH REGIONS. THE PHOTOS SHOWS THE FRACTURED CATHETER ATTACHED WITH THE POWERPORT. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. FURTHERMORE, THE CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 12/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02301 | 2023-04-05 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE CATHETER WAS SPONTANEOUSLY PULLED OUT DURING WIPING, SO WHEN THE BALLOON WAS EXAMINED, IT WAS FOUND TO BE DAMAGED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02299 | 2023-04-05 | 2022-12-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE SURVEY RESPONDENT STATED NO, THEY DID NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCTS PROVIDES SUFFICIENT INFORMATION ABOUT THE PRODUCTS AND ITS MEDICAL APPLICATION BECAUSE THE INSTRUCTIONS FOR USE (IFU) HAD INADEQUATE INSTRUCTIONS FOR HEALTHCARE PROCESSIONAL AND INADEQUATE OR INSUFFICIENT TRAINING IN RELATION TO BIOCATH® FOLEY CATHETER PRECONNECTED TO 500ML BARDIA® LEG BAG, FEMALE LENGTH, FRENCH SIZE 12. | PER ADDITIONAL INFORMATION RECEIVED, BD HAS DETERMINED THAT THIS MDR EVENT IS NOT REPORTABLE. H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02300 | 2023-04-05 | 2022-12-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER TUBING WAS NOT HOLDING. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01173 | 2023-04-05 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN DIALYSIS | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | UNK DIALYSIS |  | IT WAS REPORTED THAT THREE MONTHS TWENTY ONE DAYS POST DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CLAMP OF THE CATHETER ALLEGEDLY BROKE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, TWO PHOTOS WERE PROVIDED FOR REVIEW. THE PHOTOS SHOWS THE UNKNOWN DIALYSIS CATHETER IMPLANTED IN THE PATIENT IN WHICH THE CLAMP WAS NOTED TO BE BROKEN. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01179 | 2023-04-05 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS FOUR MONTHS AND TWO DAYS POST PORT PLACEMENT PROCEDURE, DURING PORT REMOVAL PROCEDURE, THE CATHETER ALLEGEDLY BROKE AND TOP OF IT REMAINED IN THE VEIN. IT WAS FURTHER REPORTED THAT THE PORT WAS REMOVED BECAUSE IT HAD BEEN INACCESSIBLE FOR SEVERAL WEEKS. REPORTEDLY THE CATHETER ALLEGEDLY HAD SEVERAL TEARS DOWN THE SHAFT. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02305 | 2023-04-05 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V24S |  | IT WAS REPORTED THAT THE PATIENT USED A THREE-LUMEN FOLEY CATHETER FOR BLADDER IRRIGATION AND COMPRESSION TO STOP BLEEDING ON (B)(6) 2023. ON THE SIXTH DAY AFTER THE PATIENT RETURNED TO THE WARD, THE BALLOON RUPTURED AND THE THREE-LUMEN CATHETER FELL OUT, RESULTING IN THE INABILITY OF THE PATIENT TO PERFORM BLADDER IRRIGATION AND COMPRESSION TO STOP BLEEDING. IMMEDIATELY, A NEW CATHETER WAS REPLACED. AS PER DESCRIPTION OF EVENT CAUSE ANALYSIS, STATED ABOUT THE BALLOON QUALITY PROBLEM. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01178 | 2023-04-05 | 2022-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED ERYTHEMA, PAIN AND ITCHING ALONG THE PORT TUNNEL. IT WAS FURTHER REPORTED THAT MICRO LEAKAGE WAS ALLEGEDLY NOTED. REPORTEDLY, IT WAS DECIDED TO REMOVE AND REPLACE THE OLD CATHETER WITH ANOTHER CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION AND ONE PHOTO WAS PROVIDED FOR REVIEW. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED AND THE EDGES WERE NOTED TO BE UNEVEN. THE SURFACE OF THE BREAK WAS NOTED TO BE ROUND AND SMOOTH. KINKS WERE NOTED THROUGHOUT THE CATHETER. UPON INFUSION, A LEAK WAS OBSERVED FROM THE PARTIAL CIRCUMFERENTIAL BREAK. A LEAK WAS NOTED UPON PHOTO REVIEW. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED LEAK AND THE IDENTIFIED FRACTURE, DEFORMATION AND NATURALLY WORN ISSUES. THE IDENTIFIED BREAK IS TYPICAL OF FLEXURAL FATIGUE, WHICH IS DUE TO THE REPETITIVE, CYCLIC KINKING OF THE CATHETER. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 02/2023), G3, H6 (DEVICE, RESULT). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02325 | 2023-04-06 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRETEST, WATER COULD BE INJECTED BUT COULD NOT BE DRAINED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE POOR SAMPLE CONDITION. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO GENERATED AT SUPPLIER SITE OR DURING TRANSIT TO BARD, (EXAMPLE: DEFECTIVE / TORN / BROKEN COMPONENTS). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR. 1. BALLOON CATHETER: FOLEY CATHETER. TEMPERATURE-SENSING CATHETER. 2. ACCESSORIES: CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.), SYRINGE PREFILLED WITH STERILE WATER, WATER SOLUBLE LUBRICANT, ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION, TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS, GLOVES, AND STATLOCK FOLEY. [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE: (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE: (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD: 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS: 1) MALFUNCTION: - CATHETER KINKING, DAMAGE, RUPTURE, - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE, - OCCLUSION OF CATHETER INNER LUMENS, - ENCRUSTATION, - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE, - DEVICE DAMAGE DUE TO INAPPROPRIATE USE, - FAILURE TO MEASURE TEMPERATURE, AND - IMPROPER TEMPERATURE INDICATION. 2) ADVERSE EVENTS: - URINARY-TRACT INFECTION, - HEMORRHAGE, HEMATURIA, - ALLERGY REACTION TO THE DEVICE, - CALCULUS FORMATION, - EDEMA, - PAIN, - DISCOMFORT, - INJURY OF BLADDER OR URETHRAL, - URETHRITIS, URINARY INCONTINENCE, AND - RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01195 | 2023-04-06 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN, WHITE, 4.2F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 0601610 | 0601610 |  | IT WAS REPORTED THAT SOMETIME POST CHRONIC CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY STARTED OUT AS AN ANUERYSM AND EVENTUALLY TORN. IT WAS FURTHER REPORTED THAT THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00228 | 2023-04-06 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U4150530RX | U4150530RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A CHRONIC TOTAL OCCLUSION IN THE SUPERFICIAL FEMORAL ARTERY VIA AN IPSILATERAL ANTEGRADE APPROACH, THE PTA BALLOON ALLEGEDLY RUPTURED UNDER NOMINAL PRESSURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 06/2023), G3. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16688828 | 2023-04-06 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERHICKMAN POLYURETHANE CATHETER WITH SURECUFF INGROWTH | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 0805810 | 0805810 |  | DURING REMOVAL OF HICKMAN CATHETER, CUFF REMAINED IN PLACE REQUIRING SURGICAL REMOVAL. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16688812 | 2023-04-06 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | MALFUNCTION IN PICC LINE WIRE, STOPCOCK PERMITTED AIR INTRUSION WHEN ASPIRATED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02323 | 2023-04-06 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT WHITE DEPRESSED DAMAGES WITH THE BALLOON PART FOR A TOTAL OF 5350 FOLEY CATHETERS FROM TWO BATCHES ( BATCH# MYGN5842 AND BATCH# MYGR5328) WERE FOUND DURING USE AND INSPECTION. THE QUALITY DEPARTMENT OF COMPANY HAD DETERMINED THAT THERE WERE RISKS OF BALLOON BURSTING AND THAT THESE PRODUCTS COULD NOT BE USED IN THE PRODUCTION. THE ISSUE WAS DETECTED DURING REUSE OF A SINGLE USE MEDICAL DEVICE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01192 | 2023-04-06 | 2023-03-16 | I, n, j, u, r, y | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D5 |  | THE FOLLOWING WAS REPORTED BY RN "I WAS ATTEMPTING TO FLUSH PICC LINE WHEN SUDDENLY A SPRAY OF SOMETHING WET WAS FELT ON MY FACE, FOREHEAD AND EYE AREA. AFTER LOOKING DOWN, THIS RN REALIZED THE "SPRAY" WAS A MIX OF BLOOD PRODUCT AND SALINE FROM THE PICC LINE. UPON FURTHER INSPECTION, IT WAS DISCOVERED THAT THE RUBBER STOPPER AT THE END OF THE PICC WHICH HOLDS THE GUIDE WIRE DISLODGED. PATIENT WAS KNOWN HIV AND NURSE WENT TO ED FOR TREATMENT." BLOOD AND SALINE SPRAYING OUT OF PICC LINE INTO NURSE'S EYE. ADDITIONAL INFORMATION RECEIVED 04/06/2023: IT WAS STATED "RN WAS TREATED WITH ANTIVIRALS FOR 3 DAYS. ADVISED TO STOP ANTIVIRAL TREATMENT BY CX IN EMPLOYEE HEALTH. DISCONTINUATION OF MEDICATION (B)(6) 2023. CLINICIAN CLARIFIED THAT THE RUBBER STOPPER CAME OUT PARTIALLY. RN NOTICED SOMETHING WAS WRONG WHEN THEY FLUSHED PICC LINE VIA THE PORT CONNECTED TO THE T LOCK AND BLOOD MIXED WITH SALINE SPRAYED ONTO THEIR FACE. IT WAS UPON FURTHER REVIEW OF THE DEVICE THAT THEY NOTICED A POSSIBLE DEFECTIVE. THERE WAS NO RESISTANCE NOTED WHEN ATTEMPTING TO FLUSH THROUGH THE T-LOCK." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02322 | 2023-04-06 | 2023-03-14 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD TO TAPE THE ENDS OF THE HOSES AND THE PATIENT WAS STILL WAKING UP SOAKED BECAUSE THE PUREWICK URINE COLLECTION SYSTEM HAD NO SUCTION, ADVISED THEY NEEDED TO TROUBLE SHOOT THE MACHINE SO THEY COULD PINPOINT THE ISSUE AND IF NEEDED SEND A REPLACEMENT. THE REPRESENTATIVE DID NOT DO A WATER TEST. THE PATIENT HAD BEEN USING THIS PRODUCT FOR LESS THAN 90 DAYS. PER ADDITIONAL INFORMATION RECEIVED VIA LIBERATOR ON 01MAR2023, STATED THAT THE PUREWICK URINE COLLECTION SYSTEM WAS NOT SUCTIONING ENOUGH. PATIENT DID TROUBLESHOOT WATER TEST AND THE MACHINE FAILED. PATIENT SAID THAT THEY WASTED 1/2 BOX OF WICKS. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. PER SAMPLE FORM RECEIVED ON 14MAR2023, IT WAS STATED THAT THE DEVICE WAS REPLACED AND THE PATIENT EXPERIENCED URINARY TRACT INFECTION (UTI) AND WAS TREATED WITH ANTIBIOTICS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. WARNINGS: TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02329 | 2023-04-06 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON DID NOT DEFLATE. THE FOLEY CATHETER WAS CUT OFF. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE THIN RUBBERIZE WALL THAT CAUSING COLLAPSE/PINCH INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT¿ AND MIGHT BE CONTRIBUTED BY USER RELATED, EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. STERILE: UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE LATEX AND MAY BURST BALLOON. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT DESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02328 | 2023-04-06 | 2023-03-16 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER WAS TORN AND THE FRAGMENT REMAINED IN THE PATIENT'S BODY AND WAS SURGICALLY REMOVED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, CATHETER WAS BROKEN. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. SAMPLE WAS EVALUATED AND OBSERVED THE CATHETER WAS CUT OFF AT SHAFT. THE SURFACE WAS NORMAL IN APPEARANCE WITH THE PRESENCE OF SMOOTH AND CLEAR CUT. THE CATHETER SHAFT LOOKS LIKE HAS BEEN CUT WITH APPLIED SOME FORCE THAT COULD CAUSE THE CATHETER BREAK. NONE OF THE CONDITIONS ABOVE WERE EVIDENT ON THE SAMPLE. THE EXACT CAUSE ON HOW AND WHEN PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED.( EXAMPLE: ROUGH HANDLING OR USE SHARP OBJECT). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D, H. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02366 | 2023-04-06 | 2023-03-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT HAD AN INDWELLING FOLEY CATHETER AFTER THE OPERATION, THE DRAINAGE WAS SMOOTH, AND THE COLOR OF URINE WAS LIGHT RED. ON (B)(6), THE PATIENT'S FOLEY CATHETER FELL OUT. AFTER CHECKING, IT WAS FOUND THAT THE BALLOON OF THE FOLEY CATHETER BURST. AFTER COMFORTING THE PATIENT, A NEW CATHETER WAS USED FOR CONTINUED CATHETERIZATION. THEN THE PATIENT WAS KEPT ON OBSERVATION AND THE PATIENT'S URINE COLOR WAS PAID ATTENTION TO. THE PATIENT HAD NO COMPLAINTS OF DISCOMFORT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATER ALLERGY". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01203 | 2023-04-06 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 5F 55CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | IT WAS REPORTED BY THE CUSTOMER THEY WERE NOT ABLE TO ACCESS THE VEIN WITH INTRODUCER, THE SHEATH STARTED TO BUCKLE AS IT WAS ENTERING THE PATIENT'S SKIN. HAD TO OPEN A NEW KIT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, COMPLAINT AND BATCH HISTORY, APPLICABLE PREVIOUS INVESTIGATION(S), APPLICABLE MANUFACTURING RECORDS, PHOTO ANALYSIS AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED (SPLIT) SHEATH TIP WAS CONFIRMED AND WAS LIKELY RELATED TO THE MANUFACTURING PROCESS. TWO PHOTOGRAPHS OF A 5FR MICROINTRODUCER/SHEATH/DILATOR WAS RETURNED FOR EVALUATION. THE SHAFT OF THE DILATOR APPEARED SLIGHTLY BOWED AND REDDISH RESIDUAL MATERIAL WAS PRESENT WITHIN THE LUMEN OF THE DILATOR, INDICATING THAT THE DEVICE HAD BEEN USED. IN THE FIRST PHOTOGRAPH, THE DISTAL END OF THE SHEATH APPEARED UNREMARKABLE; NO BUNCHING OF THE SHEATH MATERIAL WAS NOTED AND THE SHEATH APPEARED TO BE APPROPRIATELY TAPERED. THE SECOND PHOTOGRAPH SHOWED A CLOSE-UP OF THE DISTAL END OF THE SHEATH ON THE OTHER SIDE. THE DISTAL END OF THE SHEATH CONTAINED A LONGITUDINAL SPLIT WITH STRAIGHT SPLIT EDGES. A SMALL AMOUNT OF BALLOONING WAS SEEN IN THE SHEATH MATERIAL ON ONE SIDE OF THE SPLIT. THE MANUFACTURING PROCESS LIKELY CAUSED OR CONTRIBUTED TO THE OBSERVED SPLIT. CHARACTERISTICS OF THE RAW MATERIAL AND/OR STORAGE CONDITIONS MAY HAVE ALSO CONTRIBUTED TO THE DAMAGE. BD IS WORKING CLOSELY WITH MANUFACTURING TO HELP PREVENT RECURRENCE OF THE REPORTED EVENT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02365 | 2023-04-06 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS FOUND BLOCKED SO THAT A NEW URINARY CATHETER WAS REPLACED ON (B)(6) 2023. ON (B)(6), THE CATHETER FELL OUT FROM THE PATIENT. UPON CHECKING, IT WAS FOUND THAT THE CATHETER BALLOON WAS EMPTY. WHEN THE USER INJECTED IT AGAIN FOR TESTING THE BALLOON, SOME LIQUID LEAKED FROM THE INJECTION VALVE. THEN A NEW CATHETER WAS REPLACED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01198 | 2023-04-06 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT "AFTER THE GROSHONG CATHETER WAS IMPLANTED, A PINHOLE WAS FOUND ON THE CATHETER, AND SALINE SOLUTION LEAKED. THERE WAS NO REPORTED PATIENT INJURY." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK IN THE GROSHONG CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR GROSHONG NXT CLEARVUE CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH CATHETER DAMAGE CAUSED BY CONTACT WITH A SHARP EDGED INSTRUMENT SUCH AS A SCALPEL. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED AT THE INTERSECTION OF THE CATHETER TO THE DISTAL CONNECTOR PIECE JUST PROXIMAL OF THE 53 CM DEPTH MARKER. MICROSCOPIC EXAMINATION OF THE CATHETER SPLIT CONFIRMED IT WAS TYPICAL OF CONTACT WITH A SHARP BLADED INSTRUMENT, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: THE FRACTURE SURFACE WAS REFLECTIVE IN NATURE AND CONTAINED A STRIATED PATTERN. THIS CAN OCCUR DUE TO PATTERN TRANSFER OF THE SHARPENED EDGE TYPICALLY FOUND ON A SCALPEL-TYPE INSTRUMENT. SHARPLY FORMED FRACTURE EDGES. PLANAR SHAPE TO THE FRACTURE SURFACE. THE SHORT TIME FRAME BETWEEN THE INSERTION DATE AND THE DATE OF DISCOVERY SUGGESTS THIS DAMAGE LIKELY OCCURRED DURING INSERTION. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02362 | 2023-04-06 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT DRAIN BAG SEPARATES FROM THE CATHETER WHILE BEING USED. IT HAD HAPPENED ON MANY OCCASIONS DUE TO FORCE OF URINE AND IT WAS NOTED IN TRAUMA ICU, 3ACU, AND 6T ALSO IT WAS REPORTED THAT IF LET GO OF TUBING, FOLEY CATHETER CAME OUT OF PATIENT AND IT DID NOT STAY IN PLACE, NEEDED TO HOLD IT IN PLACE WITH ONE HAND. NURSE ALSO STATED THAT THE GREEN SPIGOT HARD TO TURN WHICH MAKES EMPTYING THE DRAIN BAG TO DIFFICULT ALSO THE WHITE CLAMP UNCLAMPED AND LEAKED ONCE OUT OF THE PATIENT, IT WAS HARD TO GET A SAMPLE FROM DRAIN BAG. IT WAS REPORTED THAT FOLEY CATHETER HAS SLOWER FLOW WHEN OBTAINING URINE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: SALT ACCUMULATION)/BLOCK DRAINAGE LUMEN/NO DRAINAGE EYE). THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY WAS UNKNOWN, THE FOLEY CATHETER IFUS WERE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01201 | 2023-04-06 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3175335 |  | IT WAS REPORTED BY THE CUSTOMER THAT A DEFECTIVE SINGLE PICC LINE WAS RECOVERED. THE PATIENT HAD TO HAVE THE PICC LINE REPLACED WITH DOUBLE LUMEN THAT DID NOT LEAK. ADDITIONAL INFORMATION RECEIVED ON 3/20/2023: A SL 5FR PICC LINE WAS INSERTED AND FOUND TO BE BACKING UP WITH BLOOD AND LEAKING AT THE HUB/NEEDLELESS CONNECTOR. THE CATHETER WAS SECURED WITH STATLOCK. THERE WAS NO BREAK. PATIENT RECEIVED ROCEPHIN INFUSIONS THROUGH THE PICC AS AN OUTPATIENT. CLINICIAN STATES THE PICC WAS NOTED TO BARELY LEAK WITH FLUSHING SALINE, BUT NOT WITH INFUSION OF ROCEPHIN FOR SOME TIME. THE IR DOCTOR WAS NOTIFIED AND CLEARED THE LINE FOR USE. THE CLINICIAN STATES THAT FROM THEIR UNDERSTANDING, SEVERAL WEEKS LATER AFTER RETURNING HOME FROM HIS DAILY ROCEPHIN INFUSION, THE PATIENT CALLED, STATING THAT THE PICC WAS LEAKING AND BACKED UP WITH BLOOD. PATIENT WAS INSTRUCTED TO RETURN TO THE CATH LAB AND THE SL PICC WAS REMOVED AND A NEW DL PICC WAS PLACED. THERE WERE NO SIGNS OR SYMPTOMS OR OTHER COMPLICATIONS. THE EVENT DID NOT INVOLVE AN URGENT OR LIFE-THREATENING MEDICAL SITUATION. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01200 | 2023-04-06 | 2022-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F SINGLE-LUMEN POLYURETHANE CATHETER RADIOLOGY 70 CM GUIDEWIRE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3175335 |  | IT WAS REPORTED BY THE CUSTOMER THAT A DEFECTIVE SINGLE PICC LINE WAS RECOVERED. THE PATIENT HAD TO HAVE THE PICC LINE REPLACED WITH DOUBLE LUMEN THAT DID NOT LEAK. ADDITIONAL INFORMATION RECEIVED 3/20/2023: SINGLE LUMEN 5F PICC LINE INSERTED, FOUND TO BE LEAKING AT HUB/NEEDLELESS CONNECTOR. THE CATHETER WAS SECURED WITH STATLOCK. THERE WAS NO BREAK IN THE CATHETER; THE NURSE NOTICED IT LEAKED AT THE HUB WHEN FLUSHED. THERE WAS NO CLINICAL DELAY IN TREATMENT BUT THE PATIENT DID HAVE ANOTHER PROCEDURE TO EXCHANGE THE PICC FOR A DL PICC. NO SIGNS OR SYMPTOMS OR OTHER COMPLICATIONS OCCURRED. ANCEF WAS INFUSED THROUGH THE PICC BUT THE CLINICIAN REPORTS THE PICC WAS NOTED TO LEAK JUST WITH FLUSHING WITH SALINE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01199 | 2023-04-06 | 2023-03-20 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7957505J |  | IT WAS REPORTED THAT "DURING THE INSERTION PROCEDURE OF THE GROSHONG CATHETER VIA THE RIGHT UPPER ARM FOR THE TPN ADMINISTRATION, THE INTRODUCER NEEDLE WAS ACCESSED THE VESSEL AND THE GUIDEWIRE WAS INSERTED AFTER BLOOD RETURN COULD BE CONFIRMED. BECAUSE THE GUIDEWIRE STOPPED ADVANCING NEAR THE AXILLA AND COULD NOT BE WITHDRAWN, THE GUIDEWIRE WAS WITHDRAWN SLOWLY OVER TIME, AND THE STUCK GUIDEWIRE WAS REMOVED. A LARGE AMOUNT OF THROMBUS ADHERED TO THE INSIDE AND OUTSIDE OF THE REMOVED INTRODUCER NEEDLE. THE PROCEDURE WAS COMPLETED BY REINSERTING VIA THE LEFT SIDE. THERE WAS NO REPORTED PATIENT INJURY." | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02355 | 2023-04-06 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS USED ON A LABORING PATIENT AND WHEN IT WAS DC¿D THE REPRESENTATIVE ONLY GOT OUT COUPLE OF ML¿S. PER ADDITIONAL INFORMATION RECEIVED VIA SAMPLE FORM ON 29MAR2023, IT WAS REPORTED THAT PER REPRESENTATIVE THE BALLOON INSUFFLATION ON INSERTION WAS VERY STIFF AND DIFFICULT TO INSTILL THE NORMAL SALINE. REPRESENTATIVE ENSURED PROPER PLACEMENT AND HAD COPIOUS URINE RETURN. ALSO VERIFIED BALLOON WAS NOT IN URETHRA. UPON REMOVAL ONLY A COUPLE OF CCS WERE ASPIRATED FROM BALLOON ON REMOVAL. CATHETER EASILY REMOVED AND BALLOON APPEARS TO BE FULLY DEFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02352 | 2023-04-06 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS DIFFICULT TO INFLATE AND WHEN IT WAS DC IT ONLY HAD COUPLE OF ML'S OF STERILE WATER. PER ADDITIONAL INFORMATION RECEIVED VIA SAMPLE FORM ON 29MAR2023, IT WAS REPORTED THAT PER REPRESENTATIVE THE BALLOON INSUFFLATION ON INSERTION WAS VERY STIFF AND DIFFICULT TO INSTILL THE NORMAL SALINE. REPRESENTATIVE ENSURED PROPER PLACEMENT AND HAD COPIOUS URINE RETURN. ALSO VERIFIED BALLOON WAS NOT IN URETHRA. UPON REMOVAL ONLY A COUPLE OF CCS WERE ASPIRATED FROM BALLOON ON REMOVAL. CATHETER EASILY REMOVED AND BALLOON APPEARS TO BE FULLY DEFLATED. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED A 2-WAY CATHETER ATTACHED TO A DRAINAGE BAG. INFLATED BALLOON WITH 10 ML OF SOLUTION SUCCESSFULLY WITH SOME DIFFICULTY AND THE CATHETER RESTED WITH NO LEAKS NOTED. ATTEMPTED TO DEFLATE BALLOON PASSIVELY AND SOLUTION COULD NOT BE WITHDRAWN. SOLUTION WAS WITHDRAWN BY ASPIRATION SUCCESSFULLY BUT WITH EXTREME DIFFICULTY. THEREFORE, PRODUCT DOES NOT MEET SPECIFICATIONS, WHICH STATES "CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE." ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED, A POTENTIAL ROOT CAUSE COULD BE VALVE DAMAGED, POOR MOLDING, INCORRECT DIMENSIONS. A REVIEW OF THE DHR DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿DIRECTIONS FOR USE ¿ PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD. 1. WASH HANDS AND DON CLEAN GLOVES. 2. EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CAREKIT. 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT'S PERI-URETHRAL AREA 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL 5. USE PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP. 6. DON STERILE GLOVES. 7. PLACE UNDERPAD BENEATH PATIENT PLASTIC/"SHINY¿ SIDE DOWN. 8. POSITION FENESTRATED DRAPE ON PATIENT. 9. STRUCTURES 3 FOAM SWABSTICKS IN POVIDONE IODINE. 10. REMOVE FOLEY CATHETER FROM WRAP AND LUBRICATE CATHETER. 11. ATTACH THE WATER FILLED SYRINGE 12. PREPARE PATIENT WITH 3 FOAM SWABSTICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS. NOTE: USE EACH SWABSTICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE FAR LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE NEAR LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA MINORA. MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD. 13. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND. (A) WHEN THE CATHETER TIP HAS ENTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE. (B) INSERT CATHETER TWO MORE INCHES AND INFLATE CATHETER BALLOON. 14. INFLATE CATHETER BALLOON USING ENTIRE 10CC OF STERILE WATER PROVIDED IN THE PREFILLED SYRINGE. NOTE: USE OF LESS THAN 10CC CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON 15. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK 16. SECURE THE FOLEY CATHETER TO THE PATIENT. USE THE SMLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED (SEE SMLOCK®FOLEY STABILIZATION DEVICE IFU). NOTE: PLEASE MAKE SURE PATIENT IS APPROPRIATE FOR USE OF STATLOCK® STABILIZATION DEVICE 17. POSITION HANGER ON BED RAIL AT THE FOOT OF THE BED. NOTE: EXERCISE CARE TO KEEP BAG OFF THE FLOOR 18. USE GREEN SHEETING CLIP TO SECURE DRAINAGE TUBE TO THE SHEET. MAKE SURE TUBE IS NOT KINKED. 19. INDICATE TIME AND DATE OF CATHETER INSERTION ON PROVIDED LABELS. PLACE DESIGNATED LABELS ON PATIENT CHART AND DRAINAGE SYSTEM. 20. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL. FOLEY CATHETER REMOVAL: 1. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. 2. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE. THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. 3. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. 4. IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. 5. SHOULD BALLOON RUPTURE OCCUR, CARESHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE ¿ SECURE THE FOLEY CATHETER, USE THE SMLOCK®FOLEY STABILIZATION DEVICE IF PROVIDED ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT. ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE. ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED. STERILIZED USING ETHYLENE OXIDE STERILE UNLESS PACKAGE IS OPENED OR DAMAGED, EXCEPT FOR ANY INDIVIDUALLY PACKAGED COMPONENTS WITHIN THE TRAY WHICH ARE NOT LABELED AS STERILE. THESE COMPONENTS ARE NOT TERMINALLY STERILIZED. BARD®EZ·LOK'" SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK OR SLIP TIP COLLECTION DEVICES. BO VACUTAINER" URINE PRODUCTS CAN BE USED FOR COLLECTION, STORAGE, AND TRANSPORT OF URINE SPECIMENS FROM A FOLEY CATHETER. ¿ OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. ¿ SWAB SURFACE OF BAA¿®EZ-LO,·· SAMPLING PORT WITH ANTISEPTIC (FIG. 2) E.G., SITE-SCRUB'" IPA DEVICE WHICH IS AN EFFECTIVE DISINFECTING LUER HUBS. ¿ USING ASEPTIC TECHNIQUE, POSITION DEVICE IN THE CENTER OF THE SAMPLING DEVICE FIRMLY AND TWIST GENTLY TO SAMPLING PORT. ¿ IF USING BO VACUTAINER®, ACCESS SHEATH COLLECT URINE INTO THE BD PER HOSPITAL PROTOCOL. IF USING SYRINGE, SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE COLLECTION CUP OR FOLLOW HOSPITAL. ¿ AFTER SAMPLE IS OBTAINED CONFIRM UNKINKED AND URINE IS FLOWING FREELY OBSTRUCTED. DISCARD COLLECTION DEV HOSPITAL PROTOCOL. ¿ FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETER USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02354 | 2023-04-06 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | A948314 | 175814 |  | IT WAS REPORTED THAT THE INSERTED FOLEY AND BAG KEPT LEAKING ON FLOOR DESPITE BEING CLOSED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02353 | 2023-04-06 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123526A | 123526A |  | IT WAS REPORTED THAT THE FOLEY CATHETERS WERE LEAKING FROM THE BULB. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON BURST. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE PATIENT USE ONLY. DO NOT REUSE AND RESTERILIZE. FOR UROLOGICAL USE ONLY. USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PROTOCOL, THE MAY BE CUT OFF. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02351 | 2023-04-06 | 2023-03-17 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTIONS AND THEY HAD PROBLEMS WITH DISCOMFORT AND HAD FELT THEY HAD INJURED WHEN TAKING THE PUREWICK FEMALE EXTERNAL CATHETER OFF. THE PATIENT HAD BEEN USING THIS PRODUCT FOR LESS THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION, DISCOMFORT AND INJURY AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. PER FOLLOW UP VIA PHONE ON 07APR2023, IT WAS REPORTED THAT PATIENT DID SEEK MEDICAL ATTENTION FOR URINARY TRACT INFECTIONS, JUST FINISHED A ROUND OF ANTIBIOTICS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE COMPLETED DUE TO NO LOT NUMBER PROVIDED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02348 | 2023-04-06 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON SPONTANEOUSLY FELL OUT OF THE PATIENT. THE CATHETER WAS PLACED IN A BEDRIDDEN PATIENT AND SPONTANEOUSLY FELL OFF AFTER 7 DAYS. THIS EVENT OCCURRED IN THE SAME WARD AS THAT OF PREVIOUS REPORT (B)(4). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02347 | 2023-04-06 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE FOLEY CATHETER BALLOON WHEN STERILIZED WATER WAS INJECTED DURING THE CUFF TEST . THE EVENT OCCURRED DURING PRETEST. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE VALVE MENTIONED AS BACKFLOW . IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. AS PER MAIL RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT THE THEY CONFIRMED IT TO THE SALES REPRESENTATIVE ABOUT THE WATER LEAK LOCATION, AND THIS OPTION FOR BACKFLOW VALVE WAS CHECKED BY MISTAKE IN INQUIRY SHEET. THE CORRECT LEAK POINT WAS FROM BALLOON, NOT FROM VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02346 | 2023-04-06 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE WATER WAS DIFFICULT TO BE INJECTED INTO THE CATHETER AND THE BALLOON DID NOT INFLATE. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. THE AFFECTED SYRINGE WAS MANUFACTURED. FURTHER INVESTIGATION IS DOCUMENTED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. PACKAGING LOTS NUMBER MANUFACTURED WITH AFFECTED SYRINGE BATCHES IDENTIFIED AND DOCUMENTED. THE DHR REVIEW IS REQUIRED AND THE RISK REVIEW AND LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01197 | 2023-04-06 | 2023-01-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 20CM (SHORT-TERM) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5593200 |  | BD VASCULAR ACCESS FIELD ASSURANCE RECEIVED A RETURNED GUIDEWIRE THAT WAS STUCK IN AN INTRODUCER NEEDLE AND THE WIRE WAS BROKEN. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE ANALYSIS, PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A BROKEN GUIDEWIRE IS CONFIRMED. ONE NIAGARA CATHETER KIT WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED THE KIT WAS RETURNED OPEN AND USE RESIDUE WAS OBSERVED ON THE COMPONENTS WITHIN THE KIT. THE CORE WIRE OF THE GUIDEWIRE WITHIN THE KIT WAS OBSERVED TO BE BROKEN AND THE COILED WIRE WAS FOUND TO BE UNRAVELED. THE GUIDEWIRE WAS RETURNED THREADED THROUGH THE INTRODUCER NEEDLE BUT WAS NOT FOUND TO BE STUCK WITHIN THE NEEDLE. A MICROSCOPIC OBSERVATION REVEALED SOME DAMAGE ON THE BEVEL OF THE INTRODUCER NEEDLE. BOTH WELD TIPS OF THE GUIDEWIRE WERE OBSERVED TO BE PRESENT, AND NO BREAKS WERE FOUND IN THE COILED WIRE. THE BREAK IN THE CORE WIRE OF THE GUIDEWIRE WAS OBSERVED TO BE TAPERED WITH A MOSTLY SMOOTH AND GRANULAR FRACTURE SURFACE, WHICH IS INDICATIVE OF TENSILE FAILURE CAUSED BY EXCESSIVE PULLING FORCES ON THE GUIDEWIRE. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01196 | 2023-04-06 | 2022-11-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REPLACEMENT CONNECTOR FOR 4F GROSHONG NXT CLEARVUE PICC WITH EXTENSION LEG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7812400 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE CENTRAL VENOUS CATHETER KIT WITH PERIPHERAL CANNULA CAN BE RELEASED WITH SLIGHT FORCE AFTER REPLACEMENT OF JOINT AND CANNOT BE FASTENED TIGHTLY. REPLACE WITH A NEW PRODUCT. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02345 | 2023-04-06 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT THERE WAS A CATEGORICAL LEAKAGE. ON OPENING THE STERILIZATION PACK, THERE WAS A PART OF THE LUBRICANT PACKAGE THAT WAS OPEN AND THERE WAS NO CONTENT OF CATEJELLY INSIDE. SOME BAGS WERE EMPTY. | THE REPORTED EVENT WAS CONFIRMED ¿ UNKNOWN CAUSE. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. BASED ON THE EVALUATION, IT WAS OBSERVED THE TRAY WAS WET WITH TRACES OF JELLY LEAKED IN THE TRAY. OBSERVED THE JELLY SACHET WAS NOT PROPERLY SEALED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE, ¿DEFECT GENERATED AT SUPPLIER'S SITE OR DURING TRANSFER TO BARD". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS · PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRI-SIL® FOLEY CATHETER SYSTEM CONSISTS OF A SYRINGE PREFILLED WITH STERILE WATER AND WATER-SOLUBLE LUBRICANT. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. THE SURFACE OF CATHETER IS COATED FIRST WITH A TRACE AMOUNT OF METALLIC SILVER, AND THEN WITH POLYURETHANE ONTO THAT, TO INHIBIT PROLIFERATION OF BACTERIA THAT MAY ADHERE TO THE CATHETER. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. 3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE 1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. 4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTIONS: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02344 | 2023-04-06 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER FELL OUT OF THE PATIENT DURING A SURGERY. ACCORDING TO THE INQUIRY SHEET ATTACHED, STATED THAT THE BALLOON RUPTURE OR TEAR WAS OBSERVED ON THE FOLEY CATHETER. THE EVENT OCCURRED HALF AN HOUR AFTER PLACEMENT. STATED THAT WATER LEAKED FROM THE BALLOON. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. THE CATHETER WAS NOT PULLED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 22MAR2023, STATED THAT NO MISSING PIECES OF THE BALLOON WERE OBSERVED. | THE REPORTED EVENT IS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED A 3-WAY TEMP SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED BALLOON BURST MEASURING 0.3170" WITH NO MISSING PIECES NOTED. RECEIVED 2 PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWS OVERVIEW OF CATHETER. SECOND PHOTO SAMPLE ZOOMS IN ON END OF CATHETER WITH SLIGHT WRINKLE PRESENT ON END OF CATHETER. THE BALLOON O.D. MUST CONFORM TO DRAWING." ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE BURST BALLOONS. THE PRODUCT WAS USED FOR PATIENT DIAGNOSTIC OR TREATMENT. IT IS UNKNOWN IF THE PRODUCT WAS INFLUENCED BY THE REPORTED EVENT. DHR REVIEW IS NOT REQUIRED AS NO LOT NUMBER WAS PROVIDED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02332 | 2023-04-06 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE PATIENT USED A 2MOON16A DISPOSABLE FOLEY CATHETER FOR URINARY CATHETERIZATION BEFORE THE OPERATION ON (B)(6), AND THE CATHETER WAS RETAINED AFTER THE OPERATION. WHEN THE CATHETER WAS PLANNED TO BE REMOVED ON (B)(6), IT WAS FOUND THAT THE CATHETER HAD SLIPPED OUT OF THE PATIENT BODY. THE BALLOON HAD RUPTURED, POLLUTING THE QUILT, AND THE QUILT WAS REPLACED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 22MAR2023, THERE WERE NO MISSING PIECES OF THE BALLOON LEFT IN PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02333 | 2023-04-06 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE PATIENT USED A DISPOSABLE STERILE CATHETER FOR INTUBATION AND CATHETERIZATION ON (B)(6). AT 16:22 PM ON (B)(6), THE DISPOSABLE STERILE CATHETER OF THE PATIENT SLIPPED. THE INSPECTION FOUND THAT THE DISPOSABLE STERILE CATHETER WAS DRY AND FLAT, AND THE FLUID LEAKED OUT. THE USER IMMEDIATELY REPORTED TO THE DOCTOR ON DUTY FOR INSPECTION. IF THERE WAS NO ABNORMALITY, ASKED TO RE-INTUBATE THE CATHETER AND TAKE IT IMMEDIATELY. THE PATIENT WAS RE-CATHETERIZED WITH A STERILE CATHETER. THEN REPLACED A NEW CATHETER FOR THE PATIENT. THE PATIENT WAS RE-CATHETERIZED WITH A STERILE CATHETER. THE SECONDARY INTUBATION INCREASED PATIENT PAIN. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILIZATION METHOD: STERILE, STERILIZED BY G RADIATION. DO NOT USE IF PACKAGE IS OPENED OR DAMAGED. DISPOSABLE. DESTROY AFTER USE. DO NOT USE THIS PRODUCT IF YOU HAVE A LATEX ALLERGY. STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN SERIES IS COMPOSED OF TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. MATERIAL OF COMPONENTS: CATHETER BODY ¿ NATURAL LATEX. INFLATION CONICAL INTERFACE ¿ NATURAL LATEX. DEFLATION CONICAL INTERFACE ¿ NATURAL LATEX. FLUSHING CONICAL INTERFACE ¿ NATURAL LATEX. BALLOON ¿ NATURAL LATEX. VALVE ¿ POLYPROPYLENE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTIONS: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX-BASED URETHRAL CATHETERS, SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. DO NOT ASPIRATE URINE THROUGH THE DRAINAGE FUNNEL WALL. SINGLE USE ONLY. DO NOT RE-STERILIZE. FOR UROLOGICAL USE ONLY. PLEASE INSPECT VISUALLY FOR COMPLETENESS OR SURFACE WEAR OF THE PRODUCT BEFORE IT IS USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM DIRECT LIGHT, PREFERABLY STORED IN THE ORIGINAL PACKING BOX. PRECAUTIONS: FEDERAL (USA) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02335 | 2023-04-06 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123620A |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON BURST AND FELL OUT OF THE PATIENT. IT WAS NOTED THAT THERE WAS NO IMPACT ON A PATIENT, HCP, USER, OR OTHERS. PER THE FOLLOWUP INFORMATION RECEIVED ON 20MAR2023, IT WAS CONFIRMED THAT WATER WAS USED AS LUBRICANT. IT WAS STATED THAT THE CATHETER FELL OUT ONE DAY AFTER THE PLACEMENT. IT WAS ALSO STATED THAT THE THERE NO FRAGMENTS OF THE CATHETER TUBE BALLOON REMAINING IN THE PATIENT'S BODY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02327 | 2023-04-06 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT DRAIN BAG SEPARATES FROM THE CATHETER WHILE BEING USED. IT HAD HAPPENED ON MANY OCCASIONS DUE TO FORCE OF URINE AND IT WAS NOTED IN TRAUMA ICU, 3ACU, AND 6T .ALSO ,IT WAS REPORTED THAT IF LET GO OF TUBING, FOLEY CATHETER CAME OUT OF PATIENT AND IT DID NOT STAY IN PLACE, NEEDED TO HOLD IT IN PLACE WITH ONE HAND. NURSE ALSO STATED THAT THE GREEN SPIGOT HARD TO TURN WHICH MAKES EMPTYING THE DRAIN BAG DIFFICULT ALSO THE WHITE CLAMP UNCLAMPED AND LEAKED ONCE OUT OF THE PATIENT, IT WAS HARD TO GET A SAMPLE FROM DRAIN BAG. IT WAS REPORTED THAT FOLEY CATHETER HAS SLOWER FLOW WHEN OBTAINING URINE . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02307 | 2023-04-06 | 2023-02-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V24S |  | IT WAS REPORTED THAT THE PATIENT USED A THREE-LUMEN FOLEY CATHETER FOR BLADDER IRRIGATION AND COMPRESSION TO STOP BLEEDING ON (B)(6) 2023. AFTER A WHILE, THE PATIENT RETURNED TO THE WARD, THE BALLOON WAS FOUND RUPTURED AND THE THREE-LUMEN TUBE FELL OFF, WHICH MADE IT IMPOSSIBLE FOR THE PATIENT TO PERFORM BLADDER IRRIGATION AND COMPRESSION HEMOSTASIS, WHICH COULD POTENTIALLY CAUSE LIFE-THREATENING HEAVY BLEEDING. IMMEDIATELY A NEW THREE-LUMEN CATHETER SHOULD BE REPLACED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02349 | 2023-04-06 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. THE CATHETER WAS INSERTED INTO THE PATIENT IN THE OPERATING ROOM. AFTER RETURNING TO ROOM 6B AFTER SURGERY, URINE DID NOT FLOW OUT, SO BY CHECKING THE CATHETER, THE CATHETER HAD FALLEN OUT OF THE PATIENT DUE TO FIXATION WATER LEAKAGE. WHEN THE CATHETER WAS REMOVED, IT WAS FOUND TO BE A DEFECTIVE PRODUCT THAT DID NOT INFLATE THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01193 | 2023-04-06 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT WAS UNABLE TO BE ASPIRATED DURING IMPLANTED CENTRAL VENOUS ACCESS PORT ACCESS WITH THE NON CORING NEEDLE. IT WAS FURTHER REPORTED THAT THE CATHETER FRACTURE WAS SUSPECTED AS PER THE CHEST X-RAY EXAMINATION. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 12/2023. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01194 | 2023-04-06 | 2022-09-30 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY TWO YEARS POST PORT PLACEMENT PROCEDURE IN THE RIGHT ANTERIOR CHEST VIA RIGHT INTERNAL JUGULAR VEIN, AN X-RAY EXAMINATION WAS PERFORMED AND REVEALED THAT THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE DISTAL CATHETER SEGMENT MIGRATED INTO THE PULMONARY ARTERY. REPORTEDLY, THE PORT BODY AND THE DISTAL CATHETER SEGMENT WERE REMOVED. THE PATIENT IS STABLE NOW. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED AT THE DISTAL END OF THE CATHETER. THE DISTAL CATHETER SEGMENT WAS NOT RETURNED. THE EDGES OF THE BREAK WAS NOTED TO BE JAGGED AND THE SURFACE WAS NOTED TO BE ROUGH IN ONE REGION AND SMOOTH ON THE OTHER REGION. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND MATERIAL SEPARATION ISSUE. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MIGRATION ISSUE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01242 | 2023-04-07 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 3275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT AFTER THE POWERPICC CATHETER WAS INSERTED IN THIS PATIENT, LIQUIDS WERE FOUND LEAKING FROM THE DUAL-LUMEN PURPLE CONNECTOR WHEN USING THE SYRINGE AND CONNECTING WITH THE TRANSFUSION DEVICE. THE NURSE REPLACED IT WITH THE Q-SYTE AND LEAKAGE STILL OCCURRED. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01245 | 2023-04-07 | 2023-03-16 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE CATHETER IS LEAKING FROM THE WINGS WHEN FLUSHING. DISCOVERED WHEN THEY CHANGE THE DRESSING OF THE CATHETER. THE CATHETER WAS PULLED OUT WHEN THEY DISCOVERED THE HOLE. PATIENT MISSED ONE ANTIBIOTIC DOSE. NO PATIENT HARM. 6-APR-2023 UPDATE: SAMPLES CONTAMINATION: "PICC-LINEN HASN'T BEEN IN CONTACT WITH CHEMO, ONLY ANTIBIOTIC."- CLOXACILLIN NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE PRODUCT RETURNED FOR EVALUATION WAS A 4FR SINGLE LUMEN POWERPICC SOLO CATHETER WITH A NEEDLELESS INJECTION CAP. USE RESIDUE WAS OBSERVED ON THE CATHETER. THE CATHETER SHAFT WAS TERMINATED AT THE 41CM DEPTH MARKING. AN ATTEMPT TO FLUSH THE CATHETER WITH WATER USING A 12ML SYRINGE REVEALED A LEAK BETWEEN THE MOLDED JOINT AND THE CATHETER SHAFT. MICROSCOPIC INSPECTION OF THE LEAK SITE CONFIRMED A SPLIT. THE SURFACE OF THE SPLIT WAS GRANULAR AND UNEVEN. THE FEATURES OF THE SPLIT WERE CONSISTENT WITH MATERIAL FATIGUE DUE TO CATHETER MANIPULATION SUCH AS CYCLIC KINKING. THE LOCATION OF THE DAMAGE SUGGESTED THAT MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01238 | 2023-04-07 | 2023-03-16 | I, n, j, u, r, y | Manufacturer report | PRISTINE LONG-TERM HEMODIALYSIS CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 5403190 |  | IT WAS REPORTED THAT SOMETIME POST CHRONIC CATHETER PLACEMENT, CENTRAL LINE-ASSOCIATED BLOODSTREAM INFECTION WAS ALLEGEDLY DEVELOPED IN THE PATIENT¿S CENTRAL LINE. IT WAS FURTHER REPORTED THAT CULTURE WAS TAKEN FROM THE LINE AND LABORATORY INFECTION CONFIRMED THE GROWTH OF SPHINGOMONAS PAUCIMOBILIS BACTERIA. FURTHERMORE, PATIENT ALLEGEDLY EXPERIENCED SYMPTOMS OF TACHYCARDIA, HYPOTENSION, LETHARGY, AND FEVER. REPORTEDLY, ANTIBIOTICS WERE ADMINISTERED TO TREAT CENTRAL LINE-ASSOCIATED BLOODSTREAM INFECTION. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02392 | 2023-04-07 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS FOUND BLOCKED SO THAT A NEW URINARY CATHETER WAS REPLACED ON (B)(6) 2023. ON (B)(6), THE CATHETER FELL OUT FROM THE PATIENT. UPON CHECKING, IT WAS FOUND THAT THE CATHETER BALLOON WAS EMPTY. WHEN THE USER INJECTED IT AGAIN FOR TESTING THE BALLOON, SOME LIQUID LEAKED FROM THE INJECTION VALVE. THEN A NEW CATHETER WAS REPLACED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO DRAINAGE EYE OCCLUSION BLOCKED DRAINAGE LUMEN CAUSED BY PLC FAILURE AND ALSO MIGHT BE CONTRIBUTE BY NO DRAINAGE EYE OR USER RELATED (SALT ACCUMULATION). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01236 | 2023-04-07 | 2022-11-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT IN THE PROCESS OF CATHETERIZING THE PATIENT, THE NEEDLE REAMER IN THE CENTRAL VENOUS CATHETER PACK AFTER PERIPHERAL INTUBATION WAS PUSHED FORWARD WITH NORMAL FORCE AND FELT STUCK. LATER, IT WAS FOUND THAT THERE WAS A TEAR IN THE ANTERIOR SEGMENT, WHICH WAS EASY TO DAMAGE THE BLOOD VESSEL. TIMELY REPLACEMENT OF A NEW PUNCTURE KIT. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16701528 | 2023-04-07 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 1174108D5 | 1174108D5 |  | RN WAS ATTEMPTING TO FLUSH PICC LINE WHEN SUDDENLY A SPRAY OF SOMETHING WET WAS FELT ON HER FACE, FOREHEAD AND EYE AREA. AFTER LOOKING DOWN, THE RN REALIZED THE "SPRAY" WAS A MIX OF BLOOD PRODUCT AND SALINE FROM THE PICC LINE. UPON FURTHER INSPECTION, IT WAS DISCOVERED THAT THE RUBBER STOPPER AT THE END OF THE PICC WHICH HOLDS THE GUIDE WIRE DISLODGED. BARD REP ON SITE AND TOOK ITEM. |  | N | Y |  |  | 2023-04
3006260740-2023-01222 | 2023-04-07 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE CUSTOMER'S SENSOR IS HAVING ISSUES WHERE THE YELLOW WAVE FORM IS COMING UP BEFORE THE PICC IS INSERTED INTO THE PATIENT. THEY VERIFIED THAT IT'S ONLY HAPPENING WITH THIS ONE SENSOR AS THEIR OTHER SENSORS ARE WORKING FINE ON THE SAME TABLET. ALSO VERIFIED THEY ARE FLUSHING THE PICCS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF SENSOR IS HAVING ISSUES WHERE THE YELLOW WAVE FORM IS COMING UP BEFORE THE PICC IS INSERTED INTO THE PATIENT WAS UNCONFIRMED. THE ECG FUNCTIONALITY FUNCTIONS PROPERLY. THE ROOT CAUSE OF THE REPORTED FAILURE IS INCONCLUSIVE AS THE REPORTED ISSUE COULD NOT BE REPRODUCED DURING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02377 | 2023-04-07 | 2023-03-17 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTIONS AND THEY HAD PROBLEMS WITH DISCOMFORT AND HAD FELT THEY HAD INJURED WHEN TAKING THE PUREWICK FEMALE EXTERNAL CATHETER OFF . THE PATIENT HAD BEEN USING THIS PRODUCT FOR LESS THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION, DISCOMFORT AND INJURY AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT PATIENT DID SEEK MEDICAL ATTENTION FOR URINARY TRACT INFECTIONS, JUST FINISHED A ROUND OF ANTIBIOTICS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "USER HAS LIMITED HAND DEXTERITY AND STRENGTH." IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "1. CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE DRYDOCTM VACUUM STATION, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE DRYDOCTM VACUUM STATION USER GUIDE FOR FURTHER INFORMATION. 2. USING THE STANDARD SUCTION TUBING, CONNECT THE PUREWICKTM FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER. PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTES: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. REMOVAL: 5. TO REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER, FULLY SEPARATE THE LEGS, GLUTEUS, AND LABIA. TO AVOID POTENTIAL SKIN INJURY UPON REMOVAL, GENTLY PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER DIRECTLY OUTWARD. ENSURE SUCTION IS MAINTAINED WHILE REMOVING THE PUREWICKTM FEMALE EXTERNAL CATHETER. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. MAINTENANCE: 6. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02374 | 2023-04-07 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT WHITE DEPRESSED DAMAGES WITH THE BALLOON PART FOR A TOTAL OF 5350 FOLEY CATHETERS FROM TWO BATCHES WERE FOUND DURING USE AND INSPECTION. THE QUALITY DEPARTMENT OF THE COMPANY HAD DETERMINED THAT THERE WERE RISKS OF BALLOON BURSTING AND THAT THESE PRODUCTS COULD NOT BE USED IN THE PRODUCTION. THE ISSUE WAS DETECTED DURING REUSE OF A SINGLE USE MEDICAL DEVICE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02371 | 2023-04-07 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51620 | 51620 |  | IT WAS REPORTED THAT THE CUSTOMER HAVING CONSISTENT ISSUES WITH THE INTERMITTENT CATHETER INDIVIDUAL WRAPPERS. THEY STATED EITHER THE PULL TABS WERE COMPLETELY MISSING AND DID NOT WORK OR THE WRAPPER HOLES WERE NOT CUT IN THE RIGHT AREA. THEY WOULD EXPERIENCE THIS EVERY NOW AND THEN A FEW YEARS AGO, BUT NOW IT WAS ALMOST EVERY CATHETER THEY USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-04
2020394-2023-00235 | 2023-04-07 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CQ7564J |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A HIGHLY STENOTIC LESION, THE PTA BALLOON ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CONQUEST PTA DILATATION CATHETER PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CONQUEST PTA DILATATION CATHETER PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 08/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02370 | 2023-04-07 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0102L16 | 0102L16 |  | IT WAS REPORTED THAT THE BALLOON OF THE FOLEY CATHETER RUPTURED AFTER 4 DAYS OF USE. THE CUSTOMER MENTIONED THAT THEY EXPERIENCED PAIN WITH THIS ISSUE. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01208 | 2023-04-07 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT "DURING USE OF THE SELDINGER, AFTER SUCCESSFUL PUNCTURE, INSERTION OF THE GUIDE WIRE AND REMOVAL OF THE INTRODUCER NEEDLE, WHEN THE GUIDE WIRE WAS USED TO GUIDE THE INTRODUCER NEEDLE, IT WAS FOUND THAT THE INTRODUCER SHEATH WAS UNABLE TO PASS OVER THE GUIDE WIRE. THE DISTAL END OF THE GUIDE WIRE WAS THICKENED, LEADING TO TRIMMING WITH SCISSORS DURING USE. BUT, IT WAS DIFFICULT TO CUT OFF AND THERE WERE POWDERS ON THE CUT GUIDE WIRE." IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01209 | 2023-04-07 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT "DURING USE OF THE SELDINGER, AFTER SUCCESSFUL PUNCTURE, INSERTION OF THE GUIDE WIRE AND REMOVAL OF THE INTRODUCER NEEDLE, WHEN THE GUIDE WIRE WAS USED TO GUIDE THE INTRODUCER NEEDLE, IT WAS FOUND THAT THE INTRODUCER SHEATH WAS UNABLE TO PASS OVER THE GUIDE WIRE. THE DISTAL END OF THE GUIDE WIRE WAS THICKENED, LEADING TO TRIMMING WITH SCISSORS DURING USE. BUT, IT WAS DIFFICULT TO CUT OFF AND THERE WERE POWDERS ON THE CUT GUIDE WIRE." IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE THIRD DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01210 | 2023-04-07 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT "DURING USE OF THE SELDINGER, AFTER SUCCESSFUL PUNCTURE, INSERTION OF THE GUIDE WIRE AND REMOVAL OF THE INTRODUCER NEEDLE, WHEN THE GUIDE WIRE WAS USED TO GUIDE THE INTRODUCER NEEDLE, IT WAS FOUND THAT THE INTRODUCER SHEATH WAS UNABLE TO PASS OVER THE GUIDE WIRE. THE DISTAL END OF THE GUIDE WIRE WAS THICKENED, LEADING TO TRIMMING WITH SCISSORS DURING USE. BUT, IT WAS DIFFICULT TO CUT OFF AND THERE WERE POWDERS ON THE CUT GUIDE WIRE." IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE FOURTH DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01207 | 2023-04-07 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT "DURING USE OF THE SELDINGER, AFTER SUCCESSFUL PUNCTURE, INSERTION OF THE GUIDE WIRE AND REMOVAL OF THE INTRODUCER NEEDLE, WHEN THE GUIDE WIRE WAS USED TO GUIDE THE INTRODUCER NEEDLE, IT WAS FOUND THAT THE INTRODUCER SHEATH WAS UNABLE TO PASS OVER THE GUIDE WIRE. THE DISTAL END OF THE GUIDE WIRE WAS THICKENED, LEADING TO TRIMMING WITH SCISSORS DURING USE. BUT, IT WAS DIFFICULT TO CUT OFF AND THERE WERE POWDERS ON THE CUT GUIDE WIRE." IT WAS REPORTED THIS OCCURRED WITH FOUR DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01214 | 2023-04-07 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1809660 | 1809660 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE THE CLEAR PLASTIC VALVED INTRODUCER ALLEGEDLY CRACKED AND BROKE OFF WHEN DILATOR WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RECEIVED FOR EVALUATION. HOWEVER, ONE PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS THE INTRODUCER SHEATH IN WHICH THE VALVE CAP WAS NOTED TO BE SEPARATED FROM THE SHEATH. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED MATERIAL SEPARATION ISSUE. HOWEVER THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED FRACTURE ISSUE AS NO EVIDENCE OF THE REPORTED ALLEGATION WAS NOTED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 12/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01215 | 2023-04-07 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1809660 | 1809660 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE CLEAR PLASTIC PIECE ALLEGEDLY CRACKED AND BROKE OFF FROM THE VALVED INTRODUCER, WHEN DILATOR WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01216 | 2023-04-07 | 2023-01-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE GUIDEWIRE ALLEGEDLY COULD NOT BE INSERTED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: A J-TIP GUIDEWIRE LOADED INTO A GUIDEWIRE HOOP WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC VISUAL AND DIMENSIONAL EVALUATION WERE PERFORMED. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO INSERT ISSUE AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT ARE UNKNOWN. HOWEVER, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED DEFORMATION, MATERIAL PROTRUSION AND STRETCHED ISSUES AS A BEND WITH UNCOILING WAS NOTED THROUGHOUT CENTER PORTION OF THE GUIDEWIRE. THE ROUND CORE WIRE WAS NOTED TO BE PROTRUDING FROM THE DISTAL END OF THE UNCOILED PORTION OF THE GUIDEWIRE. BASED ON THE MEASUREMENTS RECORDED DURING SAMPLE EVALUATION AND APPLICABLE DRAWINGS, NO MEASUREMENTS RECORDED COULD BE CONFIRMED TO BE OUT OF TOLERANCE. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: (EXPIRY DATE: 09/2024). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01241 | 2023-04-07 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1174108D |  | IT WAS REPORTED BY THE CUSTOMER THAT "LONG BLACK HAIR WAS FOUND WHEN FIRST FLAP OF PICC KIT WAS OPENED. NOT TOUCHING ANY PICC SUPPLIES SO COVERED WITH STERILE GAUZE AND OPENED REMAINDER OF KIT." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01243 | 2023-04-07 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER AN UNIDENTIFIED SUBSTANCE WAS NOTED ON THE INTRODUCER SHEATH. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF AN UNIDENTIFIED SUBSTANCE ON THE MICROINTRODUCER WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. TWO PHOTOGRAPHS OF A 4.5FR ORANGE HANDLED MICROINTRODUCER WAS RETURNED FOR EVALUATION. BOTH PHOTOGRAPHS WERE OF THE HANDLE THAT CONTAINED THE 4.5FR SIZE INDICATOR. IT APPEARED THAT THE DEVICE HAD BEEN USED AS THE DILATOR HAD BEEN REMOVED FROM THE SHEATH AND THE SHEATH HAD BEEN SPLIT. A DARKISH RED RESIDUAL MATERIAL WAS SEEN ON THE HANDLE OF THE MICROINTRODUCER. HOWEVER, THE SUBSTANCE AND ITS ORIGIN COULD NOT BE DETERMINED FROM THE RETURNED PHOTOGRAPHS. IF THE PHYSICAL SAMPLE IS RECEIVED AT A LATER DATE, THIS INVESTIGATION WILL BE UPDATED WITH THE RELEVANT INFORMATION. POSSIBLE CONTRIBUTING FACTORS COULD INCLUDE IMPURITIES OR RESIDUAL MATERIAL INTRODUCED DURING THE MANUFACTURING OR PACKAGING PROCESS OR RESIDUAL MATERIAL FROM THE CLINICAL PROCEDURE. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. MANUFACTURING CONTROLS ADDRESSING CLEANLINESS ARE IN PLACE TO MITIGATE THE OCCURRENCE OF THIS TYPE OF FAILURE. THE REPORT OF A FOREIGN MATERIAL HAS BEEN DOCUMENTED AND INCLUDED IN COMPLAINT TRENDING. THIS TYPE OF COMPLAINT WILL CONTINUE TO BE MONITORED AS PART OF ONGOING EFFORTS TO IDENTIFY POTENTIAL ISSUES. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02379 | 2023-04-07 | 2023-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 319416AM | 129416M |  | IT WAS REPORTED THAT THERE WAS POSSIBLY A MANUFACTURING ISSUE IN PRODUCTION, GOT OFF LINE AND THE HOLES WERE PUNCTURED IN THE WRONG PLACE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02400 | 2023-04-10 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT BEFORE USING THE DEVICE ON THE PATIENT, THERE WAS A HOLE ON THE CATHETER TIP FOUND. THE LOCATION OF THE HOLE WAS AT THE TEMPERATURE SENSOR AREA AND THE USER DID NOT USE THE DEVICE BECAUSE OF BEING SUPPOSED THAT IT WOULD CAUSE TEMPERATURE DEFECTS. THE SAME EVENT OCCURRED THREE TIMES IN A ROW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01256 | 2023-04-10 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1676300 | 1676300 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE ACCESS NEEDLE ALLEGEDLY CRACKED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRY DATE: 03/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01254 | 2023-04-10 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT THREE YEARS THREE MONTHS AND EIGHT DAYS POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY FOUND TO BE CRACKED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE, AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. A PARTIAL COMPOUND BREAK WAS NOTED APPROXIMATELY 7.3CM FROM THE DISTAL END OF THE CATH-LOCK AND THE EDGES OF THE PARTIAL COMPOUND BREAK ON THE ATTACHED CATHETER WERE NOTED TO BE UNEVEN. UPON INFUSION, A LEAK WITH THROMBUS WAS OBSERVED FROM THE PARTIAL COMPOUND BREAK ON THE ATTACHED CATHETER. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND IDENTIFIED WEAR AND DEFORMATION ISSUE. BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01253 | 2023-04-10 | 2023-03-11 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT ONE MONTH AND NINE DAYS POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED NECK PAIN FROM THE CHEST TO THE SIDE OF HER HEAD AND FEVER. IT WAS FURTHER REPORTED THAT PATIENT WAS TAKEN TO EMERGENCY ROOM AND DIAGNOSED WITH DUAL STAGE THROMBOSIS WITH BLOOD CLOTS. REPORTEDLY, ANTICOAGULANT MEDICATIONS WERE PRESCRIBED TO TREAT THE CLOT. FURTHERMORE, PATIENT ALSO EXPERIENCED INFECTION THAT MIGHT BE CAUSED BY MISHANDLING AT THE HEALTH CARE FACILITY AND PATIENT GOT RELIEVED, AND IMPROVED WITH THE MEDICATION. THE STATUS OF THE PATIENT IS REPORTED TO BE STABLE AND WAS CONTINUING WITH THE CHEMOTHERAPY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 01/2024) | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02416 | 2023-04-10 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS DAMAGED. A FEW MINUTES AFTER INSERTING THE CATHETER, THE USER FOUND THAT THE CATHETER CAME OFF. THE USER CHECKED AND IT WAS EMPTY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01258 | 2023-04-10 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 5F DUAL-LUMEN CATHETER BASIC TRAY W/135CM NITINOL GUIDEWIRE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6295335 |  | IT WAS REPORTED "LEAK AT THE LEVEL OF THE SCREW THREAD OF ONE OF THE 2 WAYS OF THE PICC LINE. THE CONSEQUENCES FOR THE PATIENTS WHO HAVE BEEN AFFECTED BY THE DEFECT OF THE DEVICES ARE THE WITHDRAWAL OF THE PICC LINE RESULTING IN A DELAY IN THE TREATMENT." A SPECIFIC NUMBER OF AFFECTED DEVICES OR PATIENTS WAS NOT PROVIDED BY THE COMPLAINANT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02429 | 2023-04-10 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK MALE EXTERNAL CATHETER | PUREWICK MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWM030 |  | IT WAS REPORTED THAT THEY HAVE HAD SEVERAL MALE PATIENTS WITH THE NEWER PUREWICK MALE EXTERNAL CATHETER ON AND THE ADHESIVE WAS SO STICKY IT TEARS THE SKIN OFF THE SCROTUM WHEN REMOVED. CUSTOMER STATED THAT THE OLD SYSTEM THEY USED HAD LARGE ENOUGH OPENING BOTH SCROTUM AND SHAFT COULD BE PLACED IN THE BAG AND ADHESIVE WAS ON LESS SENSITIVE AREA. THEY WONDER IF THE NEWER PUREWICK SYSTEM HAD LARGER OPENING TO ELIMINATE THE ADHESIVE ON THE SCROTUM. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. IF THERE IS SIGNIFICANT PUBIC HAIR, TRIM OR CLIP THE PUBIC HAIR. PERFORM PERINEAL CARE USING THE INCLUDED WIPES AND ASSESS SKIN INTEGRITY. FOLLOW AND DOCUMENT PER HOSPITAL PROTOCOL. NOTE: USE THE INCLUDED DRYING WIPE TO ENSURE SKIN IS DRY PRIOR TO PLACEMENT OF THE DEVICE. MOISTURE ON SKIN WILL WEAKEN THE ADHESIVE. 4. ORIENT THE DEVICE, ALIGNING DRAINAGE FITTING TOWARDS THE FEET OF THE PATIENT. SLIDE THE OPENING OF THE DEVICE OVER THE PENIS UNTIL CLOSE TO THE PELVIC SKIN. CENTER PENIS WITHIN THE OPENING. REMOVE ADHESIVE LINER AND PRESS THE DEVICE AGAINST THE PELVIC SKIN TO ADHERE. ENSURE BASE HAS FULLY ADHERED AROUND THE BASE OF THE PENIS. NOTE: THE SCROTUM SHOULD NOT BE PLACED INSIDE THE DEVICE. REMOVAL: 5. TO REMOVE THE DEVICE, GENTLY LIFT THE TOP EDGE OF THE DEVICE OFF THE SKIN. IN A SLOW, DOWNWARD PEELING MOTION, REMOVE THE DEVICE. IF NECESSARY, USE A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) TO HELP LOOSEN ADHESIVE. WARNING: TO AVOID POTENTIAL SKIN INJURY, NEVER PULL THE DEVICE DIRECTLY AWAY FROM THE PATIENT. ALWAYS PEEL IN THE DIRECTION FROM HEAD TO FOOT. MAINTENANCE: 6. REPLACE THE DEVICE AT LEAST EVERY 24 HOURS OR IF SOILED WITH FECES, BLOOD, OR SEMEN." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01257 | 2023-04-10 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602680 | 0602680 |  | IT WAS REPORTED THAT DURING A CHRONIC CATHETER PLACEMENT PROCEDURE, LEAKAGE WAS ALLEGEDLY FOUND FROM THE DRUM WHEN THE PORT WAS UNITED WITH THE CATHETER. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2026). DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02426 | 2023-04-10 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI16 |  | IT WAS REPORTED THAT UPON ATTEMPTING TO INJECT FIXATION WATER INTO THE CATHETER DURING A PRETEST, THE WATER LEAKED FROM A PINHOLE ON THE BALLOON AREA AND THE BALLOON DID NOT INFLATE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02422 | 2023-04-10 | 2023-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT NEEDED TO INDWELL A FOLEY CATHETER AFTER THE OPERATION, AND THE DRAINAGE WAS NOT SMOOTH AFTER THE CATHETERIZATION ON (B)(6) 2023. THE CATHETER WAS FOUND TO BE BLOCKED, AND A NEW CATHETER WAS REPLACED IMMEDIATELY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INCORRECT PARAMETER SETTING - BLOCKED DRAINAGE EYE/LUMEN OR NO DRAINAGE EYE OR USER RELATED (SALT ACCUMULATION). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT IF HAVE LATEX ALLERGY" | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01251 | 2023-04-10 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, KIT, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602833 | 0602833 |  | IT WAS REPORTED THAT ONE YEAR THREE MONTHS AND TWENTY THREE DAYS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN FOR CHEMOTHERAPY, THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE DISTAL CATHETER SEGMENT WAS REMOVED. REPORTEDLY, THE PORT BODY WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE BARD PORT MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED APPROXIMATELY 12.9CM FROM THE DISTAL END OF THE CATH-LOCK. THE EDGES OF THE PARTIAL CIRCUMFERENTIAL BREAK WERE NOTED TO BE ROUND. A SMALL SPLIT WAS NOTED UNDER THE BREAK REGIONS. UPON INFUSION, A LEAK FROM THE PARTIAL CIRCUMFERENTIAL BREAK ON THE ATTACHED CATHETER WAS OBSERVED. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE. HOWEVER THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED MATERIAL SEPARATION ISSUE AS THERE WAS NO OBJECTIVE EVIDENCE OBTAINED FROM THE EVALUATION. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: (EXPIRATION DATE: 07/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01250 | 2023-04-10 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT, THE PORT WAS ALLEGEDLY OCCLUDED. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN DEVICE NAME AND PMA/510K. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16706752 | 2023-04-10 | 2023-02-06 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | PUREWICK FEMALE EXTERNAL CATHETER, NOT MADE NATURAL RUBBER LATEX | COLLECTOR, URINE, POWERED, NON INDWELLING CATHETER | C. R. BARD, INC. | PWF030K | PWF030K |  | PUREWICK WAS PLACED WHILE THE INCONTINENT PATIENT WAS IN THE EMERGENCY DEPARTMENT ON [DATE REDACTED]. ON [DATE REDACTED], THE NURSE COMPLETED THE MORNING SKIN CHECK AND NOTICED MACERATED REDNESS TO GROIN. THE PUREWICK WAS IN PLACE. AT SHIFT CHANGE, NURSE NOTICED REDDENED AREA AROUND DEVICE. IT WAS UNCLEAR IF IT THE AREA IS OPEN. IT WAS DIFFICULT TO ASSESS DUE TO INCONTINENT URINE. URINE VS WOUND DRAINAGE. SUCTION SET UP FOR ONE CANISTER IN ROOM. PARACENTESIS SET UP TO SUCTION. UNABLE TO CONNECT PUREWICK TO SUCTION. PUREWICK REPOSITIONED FREQUENTLY THROUGHOUT SHIFT DUE TO PATIENT GOING TO ULTRASOUND FOR PARACENTESIS AND RETURNING. WOUND AND OSTOMY CARE (WOC) NURSE WAS CONSULTED. PUREWICK REMOVED. PATIENT ALREADY ON SPECIALTY BED. SKIN CLEANSED AND BARRIER CREAM APPLIED. WOC CONSULT COMPLETED [DATE REDACTED] WITH WOC NURSE'S FINDINGS AS FOLLOWS: PARTIAL THICKNESS TISSUE LOSS NOTED TO THE LABIA MAJORA THAT IS LINEAR IN PRESENTATION CONSISTENT WITH THE PUREWICK EXTERNAL CATHETER. AREA MEASURES 3CM X 0.2CM X 0.1CM AND 2.6CM X 0.2CM X 0.1CM TO THE LEFT AND RIGHT OF THE LABIA MAJORA. RED WOUND BASE. NO DRAINAGE NOTED. PUREWICK HAS BEEN REMOVED AT THIS TIME. INCONTINENCE BRIEF IN PLACE AT TIME OF ASSESSMENT. NURSING HAS SPECIALTY BED IN PLACE AS WELL AS TAPS AND HEEL BOOTS. WOUND CLASSIFIED AS A STAGE 2 PRESSURE INJURY BY WOC. THE PRESSURE INJURY HAD IMPROVED DURING HOSPITALIZATION. |  | N | Y |  |  | 2023-04
1018233-2023-02407 | 2023-04-10 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED THAT THE PATIENT CALLED TO SAY THAT THEY HAS RECEIVED ORDER 23 BUT THEY WERE ALL DRY AND NOT LUBRICATED. AND THE PATIENT HAS OPENED ALL THE 5 BOXES AND THE ISSUE HAS HAPPENED WITH THEM ALL. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTERMITTENT CATHETER HYDROPHILIC COATED SILICONE HOW TO PREPARE AND USE YOUR HYDROSILTM CATHETER FOR UROLOGICAL USE ONLY. INTENDED FOR USE BY PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION AND MEASUREMENT. THE DEVICES ARE PASSED TO THE URINARY BLADDER VIA THE URETHRA. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. 1. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. 2. PRIOR TO OPENING THE SEALED CATHETER POUCH, APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. 3. WET THE CATHETER BY HOLDING THE PACKAGE WITH THE PRINTED SIDE UP AND TIP THE PACKAGE END-TO-END THREE TO SIX TIMES TO WET THE CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. 4. PEEL OPEN THE PACK AT THE FUNNEL/HANDLE END JUST ENOUGH TO EXPOSE 5 CM (2¿) OF THE CATHETER, OR FOR PRODUCTS WITH AN INSERTION SLEEVE, EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 5. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. 6. WASH YOUR HANDS AGAIN. 7. USING THE CATHETER WITH INSERTION SLEEVE (GRIPPER): A. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. B. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. C. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. WITHOUT INSERTION SLEEVE (GRIPPER): A. HOLD THE HANDLE/FUNNEL END AND REMOVE THE CATHETER FROM THE PACKAGING. B. GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA AND ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. 8. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. 9. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02411 | 2023-04-10 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX®LUBRICATH® FOLEY CATHETER ALCOCK MODEL | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0134L18 | 0134L18 |  | IT WAS REPORTED THAT THE BALLOON BURST. THE CATHETER FELL OFF AFTER INSERTION TO THE PATIENT. UPON CHECKING THE CATHETER, THE BALLOON WAS FOUND BURST. THE EVENT OCCURRED AFTER 5 HOURS OF PLACEMENT TO THE PATIENT. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. THERE WAS NO PATIENT URINE STONE, AND THE CATHETER WAS NOT PULLED BY HAND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02398 | 2023-04-10 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT BEFORE USING THE DEVICE ON THE PATIENT, THERE WAS A HOLE FOUND ON THE CATHETER TIP. THE LOCATION OF THE HOLE WAS AT THE TEMPERATURE SENSOR AREA AND THE USER DID NOT USE THE DEVICE BECAUSE OF BEING SUPPOSED THAT IT WOULD CAUSE TEMPERATURE DEFECTS. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE TIP GLUING-INSUFFICIENT GLUE. GLUE IS TOO HARD TO DISPENSE CAUSING THE " V SHAPE" IS TEARING THAT CONTRIBUTE TO THE REPORTED EVENT. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS : 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS : 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER SHAPE, CONFIGURATION AND PRINCIPLES: BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL : BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. SIZES OF CATHETER : AVAILABLE IN SIZES 14FR, 16FR 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY INTENDED USE & EFFECT- EFFICACY : THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. DIRECTIONS FOR USE : 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT : 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING : CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS : 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING : WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) MRI SAFETY : 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3.MALFUNCTION AND ADVERSE EVENTS : 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE : 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01252 | 2023-04-10 | 2016-11-11 | I, n, j, u, r, y | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR AND NINE MONTHS POST PORT PLACEMENT IN THE RIGHT ANTERIOR CHEST, THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE PORT BODY AND THE CATHETER WERE REMOVED. REPORTEDLY, THE DISTAL CATHETER SEGMENT WAS FOUND TO BE MIGRATED INTO THE RIGHT ATRIUM. BECAUSE, THE DISTAL CATHETER SEGMENT WAS ADHERED TO THE RIGHT ATRIUM, CONSIDERING THE RISK OF REMOVAL, THE DISTAL PART OF CATHETER REMAINED INSIDE OF THE BODY. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02405 | 2023-04-10 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILE WATER COULD BE INJECTED DURING PRETEST IN THE OPERATING ROOM, BUT THE WATER COULD NOT BE DRAINED FROM THE FOLEY CATHETER. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 30MAR2023, IT WAS STATED THAT THE DETAILS REGARDING THE PHOTO SAMPLE WEREN'T PROVIDED FROM CUSTOMER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02395 | 2023-04-10 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS REQUESTED TO BE INSERTED FOR THE PATIENT ON (B)(6) 2023. ON (B)(6) 2023, THE FOLEY CATHETER SLIPPED OUT. UPON CHECKING, THE BALLOON OF THE FOLEY CATHETER WAS EMPTY. THE WATER WAS INJECTED AGAIN TO CHECK THE BALLOON, AND IT WAS FOUND THAT THERE WAS SLIGHT LEAKAGE AT THE WATER INJECTION PORT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02402 | 2023-04-10 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT THE FOLEY CATHETER PLACED INTRAOPERATIVE WITHOUT DIFFICULTY AROUND 12.30 PM. IT WAS STATED THAT THE PATIENT WAS FLIPPED PRONE, AND THE FOLEY CATHETER WAS INTACT, SECURED AND DRAINING CLEAR YELLOW URINE. CERTIFIED REGISTERED NURSE ANESTHETISTS (CRNA) EMPTIED APPROXIMATELY 200-300CC AT ABOUT 1500, IT WAS NOTED THAT THE FOLEY CATHETER HAD FALLEN OUT OF THE PATIENT'S PENIS AND WAS ON THE FLOOR UNDER THE TABLE. WHEN STAFF EXAMINED THE BALLOON OF THE FOLEY, IT WAS DEFLATED AND APPEARED TO HAVE A HOLE IN IT. STAFF USED A SYRINGE TO TRY TO ADD AIR TO THE BALLOON, AND IT WAS IN FACT BROKEN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿IMPERFECTION IN BALLOON DUE TO PROCESS ¿. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED DIRECTIONS FOR USE WASH HANDS AND DON CLEAN GLOVES EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERIURETHRAL AREA 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL 5. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP 6. DON STERILE GLOVES 7. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE 8. POSITION FENESTRATED DRAPE ON PATIENT 9. SATURATE 3 FOAM SWAB STICKS IN POVIDONE IODINE 10.ATTACH THE WATER FILLED SYRINGE TO THE INFLATION PORT NOTE: IT IS NOT NECESSARY TO PRE-TEST THE FOLEY CATHETER BALLOON 11. REMOVE FOLEY CATHETER FROM WRAP AND LUBRICATE CATHETER 12. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA INORA MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD 13. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND A. WHEN CATHETER TIP HAS NTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE B. INSERT CATHETER TWO MORE INCHES AND INFLATE CATHETER BALLOON 14. INFLATE CATHETER BALLOON USING ENTIRE 10CC OF STERILE WATER PROVIDED IN THE PREFILLED SYRINGE NOTE: USE OF LESS THAN 10CC CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON 15. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK 16. SECURE THE FOLEY CATHETER TO THE PATIENT USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED (SEE STATLOCK® FOLEY STABILIZATION DEVICE IFU) NOTE: PLEASE MAKE SURE PATIENT IS APPROPRIATE FOR USE OF STATLOCK® STABILIZATION DEVICE 17. POSITION HANGER ON BED RAIL AT THE FOOT OF THE BED NOTE: EXERCISE CARE TO KEEP BAG OFF THE FLOOR 18. USE GREEN SHEETING CLIP TO SECURE DRAINAGE TUBE TO THE SHEET. MAKE SURE TUBE IS NOT KINKED 19. INDICATE TIME AND DATE OF CATHETER INSERTION ON PROVIDED LABELS. PLACE DESIGNATED LABELS ON PATIENT CHART AND DRAINAGE SYSTEM 20. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL FOLEY CATHETER REMOVAL 1. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE 2. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY 3. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION 4. IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL 5. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT¿. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02403 | 2023-04-10 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT 8 DAYS AFTER PLACEMENT, THE BALLOON BURST. THE CATHETER SPONTANEOUSLY FELL OFF SEVERAL DAYS AFTER PLACEMENT TO THE PATIENT. UPON CHECKING THE CATHETER, THE BALLOON APPEARED TO HAVE RUPTURED. TWO CASES OCCURRED IN A ROW, AND IT IS UNKNOWN IF THE EVENT OCCURRED TO THE SAME PATIENT AS THE OTHER COMPLAINT PREVIOUSLY REPORTED. BUT LATER SINCE THE TWO SIZES VARIED THE USER STATED THAT THE PATIENT WAS NOT SAME, AND THE ISSUE COULD HAVE BEEN WITH THE LOT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02399 | 2023-04-10 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE FOLEY CATHETER HAD COMPLETELY FALLEN OUT WHEN THE USER CHECKED THE CONDITION OF THE PATIENT'S CATHETER ON THE MORNING OF THE (B)(6) 2023. THE FIXATION WAS STILL IN PLACE AND NO EXTERNAL FORCE WAS USED TO PULL. THE BALLOON WAS BURST AND THEN A NEW CATHETER WAS USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01275 | 2023-04-11 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THE HEPARIN CAP OF THE GROSHONGPICC CATHETER KIT LEAKED LIQUID. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE VIDEO OF AN INJECTION CAP WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE VIDEO SHOWED A LENGTH OF EXTENSION TUBING CONNECTED TO THE DISTAL END OF AN INJECTION CAP. WHAT APPEARS TO BE THE NEEDLE OF AN INFUSION SET WAS OBSERVED TO BE INSERTED INTO THE PROXIMAL END OF THE INJECTION CAP. IN THE VIDEO, A DRIPPING LEAK WAS OBSERVED WHEN A CLEAR LIQUID WAS INFUSED THROUGH THE SYSTEM. THE LEAK APPEARED TO EMANATE FROM THE CONNECTION BETWEEN THE INJECTION CAP AND THE HUB OF THE EXTENSION TUBING. NO FURTHER DETAILS OF THE LEAK COULD BE SEEN DUE TO THE QUALITY OF THE RETURNED VIDEO. WHILE THE EXACT CAUSE OF THE OBSERVED LEAK COULD NOT BE DETERMINED FROM THE RETURNED VIDEO, POSSIBLE CAUSES INCLUDE INCORRECT/INCOMPLETE CONNECTION, INCOMPATIBLE COMPONENTS, MECHANICAL INTERFERENCE, AND DAMAGED COMPONENTS. HOWEVER, THE ROOT CAUSE COULD NOT BE DETERMINED WITHOUT THE RETURN OF THE PHYSICAL SAMPLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02463 | 2023-04-11 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123518A |  | IT WAS REPORTED THAT THE BALLOON WAS FOUND TO BE BURST WHEN THE CATHETER WAS REMOVED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 25MAR2023, NO MISSING PIECES OF BALLOON WERE FOUND. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01273 | 2023-04-11 | 2023-03-22 | I, n, j, u, r, y | Manufacturer report | POWERPICC PROVENA CATHETER W/ SOLO VALVE TECHNOLOGY RADIOLOGY BASIC KIT (5F) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S3395355 |  | IT WAS REPORTED BY THE CUSTOMER "PATIENTS CODING OR SHORTNESS OF BREATHE AFTER PLACING PICC LINE IN IR SUITE". NO OTHER INFORMATION WAS PROVIDED. A SPECIFIC NUMBER OF AFFECTED DEVICES OR PATIENTS WAS NOT PROVIDED BY THE COMPLAINANT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01274 | 2023-04-11 | 2023-03-22 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO FT CATHETER 5F RADIOLOGY BASIC TRAY (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3295355F |  | IT WAS REPORTED BY THE CUSTOMER "PATIENTS CODING OR SHORTNESS OF BREATHE AFTER PLACING PICC LINE IN IR SUITE". NO OTHER INFORMATION WAS PROVIDED. A SPECIFIC NUMBER OF AFFECTED DEVICES OR PATIENTS WAS NOT PROVIDED BY THE COMPLAINANT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01276 | 2023-04-11 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER WITH SHERLOCK 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 92655517 |  | CUSTOMER REPORTED THAT THE DUAL GROSHONG IS BREAKING/LEAKING AT THE CONNECTION OF THE WHITE LUMEN BELOW THE INJECTION CAP. IT WAS REPORTED THAT THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF LEAKING AT THE CONNECTION OF THE WHITE LUMEN BELOW THE INJECTION CAP WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 5 FR D/L GROSHONG NXT CATHETER. THE DAMAGE OBSERVED IN THE RETURNED SAMPLE WAS CHARACTERISTIC OF OVER-PRESSURIZATION (BURST) DAMAGE. THIS CAN OCCUR THROUGH THE USE OF SYRINGES SMALLER THAN 10ML, BY FLUSHING AGAINST AN OCCLUSION, OR DUE TO EXCESSIVE FORCE APPLIED DURING INFUSION PROCEDURES. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED ON THE PROXIMAL END OF THE DISTAL EXTENSION LEG. THIS CATHETER DAMAGE WAS TYPICAL OF A BURST, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: 'C' SHAPED SPLIT TENSILE WEAKNESS AT THE FRACTURE SITE (DUE TO MATERIAL BALLOONING PRIOR TO BURST) GRANULAR FRACTURE SURFACE TEXTURE (TYPICAL OF TEARING FAILURE MODES) THE CATHETER MATERIAL WAS VISIBLY LIGHTER IN COLOR AT THE FRACTURE SITE AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01277 | 2023-04-11 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER WITH SHERLOCK 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 92655517 |  | CUSTOMER REPORTED THAT THE DUAL GROSHONG IS BREAKING/LEAKING AT THE CONNECTION OF THE WHITE LUMEN BELOW THE INJECTION CAP. IT WAS REPORTED THAT THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02464 | 2023-04-11 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT THE TIP OF THE FOLEY CATHETER WHERE THE BALLOON WAS, ROLLED WHEN PULLED OUT. THIS CAUSED A ROUGH EDGE TO CATCH WHEN REMOVING AND THE PATIENT EXPERIENCED BLEEDING. CUSTOMER TESTED SOME OTHER TRAYS AND HAD THE SAME ISSUE. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01279 | 2023-04-11 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT DURING DRESSING CHANGE IT WAS SEEN THAT POWER PICC 5FR DL HUB IS BROKEN. DEVICE REMAINS IN PATIENT'S BODY. PATIENT STATUS/ INTERVENTION: OK. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. . THE MANUFACTURER HAS RECEIVED THE SAMPLE AND, IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01272 | 2023-04-11 | 2022-10-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F DUAL-LUMEN POLYURETHANE CATHETER RADIOLOGY 135 CM GUIDEWIRE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275135 |  | IT WAS REPORTED THAT "THE PATIENT WAS HAVING A DOUBLE LUMEN PICC LINE INSERTED. THE RADIOLOGIST ACCESSED THE PATIENTS VEIN WITH THE INTRODUCER NEEDLE AND THE NURSE TOOK THE WIRE OUT OF ITS PROTECTIVE PLASTIC COVER AND HANDED THE SOFT END OF THE WIRE TO THE RADIOLOGIST . THE TIP OF THE WIRE WAS CURLED/BENT BEFORE WE EVEN USED IT AND AFTER MANY ATTEMPTS OF TRYING TO INSERT THE GUIDEWIRE INTO THE INTRODUCER NEEDLE UNSUCCESSFULLY, WE ABANDONED USING THIS WIRE AND OPENED UP A DIFFERENT SUBSTITUTE WIRE. WE THEN PROCEEDED WITH INSERTING THE PICC LINE INTO THE PATIENT SUCCESSFULLY. IT WAS THE GUIDEWIRE THAT WAS THE ISSUE IN THIS KIT. IMPACT OF INCIDENT: CAUSED AN INCONVENIENCE BUT NO HARM TO THE PATIENT AND WE WERE ABLE TO SUCCESSFULLY INSERT THE PICC LINE." (B)(6) 2023 - THE GUIDEWIRE OF THE RETURNED SAMPLE EXHIBITED A KINK. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE ANALYSIS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A KINKED GUIDEWIRE WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 135CM IR GUIDEWIRE. THE SAMPLE WAS RECEIVED LOADED BACKWARD IN ITS PLASTIC HOOP. THE COIL/CORE TIP OF THE GUIDEWIRE WAS VISIBLE. A SHARP KINK WAS OBSERVED NEAR THE TIP. USAGE RESIDUES WERE OBSERVED ON THE WIRE. A BEND WAS OBSERVED IN THE CORE WIRE NEAR THE COIL/CORE TRANSITION. MICROSCOPIC INSPECTION OF THE WIRE CONFIRMED A KINK JUST PROXIMAL OF THE WELD TIP. THE COILS WERE MISALIGNED AT THE KINK. BLOOD RESIDUE WAS OBSERVED IN THE VICINITY OF THE KINK. GUIDEWIRE DAMAGE WAS OBSERVED; HOWEVER, EVIDENCE WAS NOT DISCOVERED THAT INFORMED THE EXACT ROOT CAUSE OF THE DAMAGE. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE GUIDEWIRE MANIPULATION PRIOR TO OR DURING ATTEMPTED USE. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01282 | 2023-04-11 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER, "THE RUBBER PIECE THAT HOLDS THE TECHNOLOGY WIRE, HOLE IS TOO BIG. IT SUCKS AIR IN AND FLUID COMES OUT. THEY STATED IT HAS BEEN HAPPENING SINCE JANUARY AND ALL THE PLACERS ON THE TEAM HAVE HAD IT HAPPEN MULTIPLE TIMES." THE CUSTOMER REPORTED MULTIPLE EVENTS OCCURRED HOWEVER THE EXACT QUANTITY OF EVENTS WAS NOT PROVIDED TO BD VASCULAR ACCESS FIELD ASSURANCE. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01281 | 2023-04-11 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8275118 |  | IT WAS REPORTED THAT DURING DRESSING CHANGE IT WAS SEEN THAT POWER PICC 5FR DL HUB IS BROKEN. THE PATIENT IS A CHILD. THE FLOW IS INTACT. THE DEVICE REMAINS IN THE PATIENT'S BODY. PATIENT STATUS/ INTERVENTION: OK. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED CATHETER IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE 5 FR POWERPICC DL CATHETER WAS PROVIDED FOR EVALUATION. THE CATHETER WAS TRIMMED AT THE 42 CM DEPTH MARKER. THE HUB WAS OBSERVED TO CONTAIN A HORIZONTAL FRACTURE ON THE MAIN STEM AND APPEARED TO PROPAGATE TOWARDS A LONGITUDINAL FRACTURE TOWARDS THE WINGS. ONE PHOTO SAMPLE OF A 5 FR POWERPICC CATHETER WAS ALSO PROVIDED AND THE CONDITION OF THE DEVICE CORRESPONDED WITH THE RETURNED PHYSICAL SAMPLE. MICROSCOPIC INSPECTION OF THE FRACTURE SURFACES REVEALED THEM TO BE ROUGH AND UNEVEN. THE HUB WAS OBSERVED TO CONTAIN DISCOLORATION AND EVIDENCE OF MATERIAL DEGRADATION. WHILE A BREAK IN THE CATHETER HUB WAS CONFIRMED, EXACT CAUSE AND CONDITIONS OF USE RESULTING IN THE REPORTED EVENT COULD NOT BE DETERMINED. THEREFORE, THIS COMPLAINT WILL BE CONFIRMED AS CAUSE UNKNOWN AT THIS TIME. POSSIBLE CAUSES OF THE OBSERVED DAMAGE INCLUDE EXPOSURE TO INCOMPATIBLE CHEMICALS, AND MATERIAL FATIGUE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01286 | 2023-04-11 | 2023-03-22 | I, n, j, u, r, y | Manufacturer report | BARD POWERGLIDE MIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED DURING A CLINICAL ASSESSMENT: "REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE MIDLINE INSERTION SITE. 18G MIDLINE IN UPPER ARM." NO OTHER INFORMATION WAS PROVIDED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION AND SPECIFIC PATIENT DETAILS WERE NOT DISCLOSED AS PART OF THE SURVEY. AS SUCH, ADDITIONAL INFORMATION AND THE SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01283 | 2023-04-11 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8295118 |  | IT WAS REPORTED THAT WHEN PRE-FLUSHING THE CATHETER, THE OPERATOR FOUND WHITE EXTRANEOUS SUBSTANCE ADHERED TO THE SURFACE OF THE STYLET OF THE CATHETER. THE ¿EXTRANEOUS SUBSTANCE¿ CAN BE REMOVED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01284 | 2023-04-11 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 5F DUAL-LUMEN CATHETER BASIC TRAY W/ 70CM NITINOL GUIDEWIRE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6295355 |  | IT WAS REPORTED "LEAK AT THE LEVEL OF THE SCREW THREAD OF ONE OF THE 2 WAYS OF THE PICC LINE. THE CONSEQUENCES FOR THE PATIENTS WHO HAVE BEEN AFFECTED BY THE DEFECT OF THE DEVICES ARE THE WITHDRAWAL OF THE PICC LINE RESULTING IN A DELAY IN THE TREATMENT." A SPECIFIC NUMBER OF AFFECTED DEVICES OR PATIENTS WAS NOT PROVIDED BY THE COMPLAITANT. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01290 | 2023-04-11 | 2023-03-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH 15 CM PRE-CURVE KIT | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5554150 |  | IT WAS REPORTED THE PATIENT REQUIRED HEMODIALYSIS CATHETERIZATION DUE TO ACUTE EXACERBATION OF CHRONIC RENAL FAILURE. WHEN THE CATHETER DILATOR WAS USED, IT WAS FOUND THAT THE CATHETER DILATED IRREGULARLY, WHICH MAY CAUSE DAMAGE TO THE PATIENT. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01289 | 2023-04-11 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG 4F SINGLE-LUMEN PICC (60CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7717405 |  | IT WAS REPORTED IMPROVED SEIDINGER PICC CATHETERIZATION WAS PERFORMED ON THE RIGHT UPPER LIMB UNDER ULTRASONIC GUIDANCE ON (B)(6) 2022, AND 7 TIMES OF INTRAVENOUS CHEMOTHERAPY VIA PICC CATHETER HAVE BEEN COMPLETED. DURING OUTPATIENT MAINTENANCE OF THE PICC CATHETER TODAY, IT WAS FOUND THAT THERE WERE TWO LEAKAGE OF THE CATHETER, NOT IN THE IN VITRO JOINT, BUT IN THE VENOUS BLOOD DUCT OF THE HEAD, WITH SMALL HOLES AT 34CM AND 35CM RESPECTIVELY ON THE SCALE OF THE CATHETER. REPLACE THE CENTRAL VENOUS CATHETER WITH THE SAME BRAND AND BATCH NUMBER IMMEDIATELY AND USE THE NORMAL CANNULA. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01291 | 2023-04-11 | 2022-12-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THE PICC LINE LEAKING FLUID AT ENTRY SITE. PICC LINE REMOVED AND FRACTURE IDENTIFIED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01293 | 2023-04-11 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD GROSHONG PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | CUSTOMER REPORTED DUAL GROSHONG BREAKING/LEAKING AT THE CONNECTION OF THE WHITE LUMEN BELOW THE INJECTION CAP. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
MW5116550 | 2023-04-11 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARD SURESTEP INTERMITTENT CATHETER TRAY PRECONNECTED DRAINAGE BAG | CATHETER, UROLOGICAL | BECTON DICKINSON/C.R. BARD, INC. | INTS14 | INTS14 |  | NO URINE OUTPUT FOLLOWING CATHETER INSERTION. CATHETER REMOVED AND REVEALED NO PERFORATIONS IN THE END OF THE TUBE- DEFECTIVE PRODUCT NOTED BY NURSING STAFF. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01278 | 2023-04-11 | 2023-03-23 | I, n, j, u, r, y | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED EARLIER IN DAY PICC LINE DRESSING NOTED TO BE NON ADHERENT AROUND HUB OF LINE. I REQUESTED PICC DRESSING CHANGE WHICH WAS ACCOMPLISHED. IV FLUIDS WERE RUNNING AT 1 ML/HR ON 1 FR PICC LINE UNTIL 1300 WHEN RATE INCREASED TO 3.5 ML /HR DUE TO NPO STATUS. AT 1810 PICC LINE PUMP ALARMED OCCLUSION AND RN SUMMONED NNP TO BEDSIDE. UPON ARRIVAL I NOTED THE PICC LINE TO BE KINKED UNDER THE DRESSING. RN STATED IT HAD BEEN SLIGHTLY KINKED WHEN DRESSING WAS CHANGED BUT FLUSHED AT THAT TIME. I LOOSENED DRESSING AND PULLED OUT KINK BUT LINE WAS CLOTTED OFF AND HAD TO BE REMOVED. I DISCUSSED WITH RN ANY KINKING BEFORE DRESSING IS PLACED WILL LIKELY ONLY WORSEN WHEN DRESSING PLACED AND BECOMES ADHERENT TO LINE. ALSO DUE TO THE SMALL SIZE OF BABY AND LOW IV RATE IT CAN TAKE SEVERAL HOURS TO READ OCCLUSION AND LIKELY THE LINE WILL HAVE CLOTTED OFF BY THEN DUE TO SMALL DIAMETER. FOLLOW UP BLOOD SUGAR WAS 27 SO LINE LIKELY WASN'T INFUSING WELL FOR QUITE SOME TIME. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02462 | 2023-04-11 | 2023-03-21 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION BUT THEY DID KNOW HOW. THE PATIENT ALSO STATED THAT THE PUREWICK FEMALE EXTERNAL CATHETER DID NOT WORK WELL FOR THE PATIENT. THE PATIENT HAD BEEN USING THIS PRODUCT FOR LESS THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01260 | 2023-04-11 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT NEEDLE WAS ALLEGEDLY DISLODGED FROM PORT WHEN RECEIVING CHEMOTHERAPY THROUGH THE PORT. IT WAS FURTHER REPORTED THAT EXTRAVASATION OF CHEMO SOLUTION OCCURRED. REPORTEDLY, THE NEEDLE LOCKED INTO THE PORT AND DID NOT COME OUT. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRY DATE: 01/2023). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01270 | 2023-04-11 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROVPPSOLO 4FDL MBP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1294108D2 |  | IT WAS REPORTED BY THE CUSTOMER "THE RUBBER PIECE THAT HOLDS THE TECHNOLOGY WIRE, THE HOLE IS TOO BIG. IT SUCKS AIR IN AND FLUID COMES OUT. THIS ACCOUNT HAS SEEN THIS SAME PROBLEM OCCUR ON ALL CONFIGURATIONS OF THEIR PROVENA SOLO PICCS." 03/24/2023 ADDITIONAL INFORMATION RECEIVED: THE PROCEDURE WAS A PICC PLACEMENT. THE PROCEDURE WAS ABLE TO BE COMPLETED. IF YOU PLACE YOUR FINGER OVER THE LARGE HOLE IT STOPS THE AIR FROM SUCKING IN/FLUID FROM COMING OUT. THERE WAS NO PATIENT HARM. THE EVENT DID NOT INTERRUPT THE ADMINISTRATION OF ANY MEDICATION, OR CAUSE A CLINICALLY SIGNIFICANT DELAY IN MEDICATION, THAT NEGATIVELY IMPACTED THE PATIENT. THERE WERE NO ADDITIONAL, UNPLANNED MEDICAL OR SURGICAL INTERVENTION (E.G., ADDITIONAL/ALTERNATIVE MEDICATIONS, IMAGING, ETC.) REQUIRED TO AVOID PATIENT HARM. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01266 | 2023-04-11 | 2023-03-22 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO CATHETER 4F RADIOLOGY BASIC TRAY SL 70 CM NITINOL GUIDEWIRE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3194355 |  | IT WAS REPORTED BY THE CUSTOMER "PATIENTS CODING OR SHORTNESS OF BREATHE AFTER PLACING PICC LINE IN IR SUITE". NO OTHER INFORMATION WAS PROVIDED. A SPECIFIC NUMBER OF AFFECTED DEVICES OR PATIENTS WAS NOT PROVIDED BY THE COMPLAINANT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00238 | 2023-04-11 | 2023-03-13 | I, n, j, u, r, y | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80144 | ATG80144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE MID BASILIC VEIN ARTERIOVENOUS FISTULA STENOSIS VIA DISTAL BASILIC VEIN ARTERIOVENOUS FISTULA, THE PTA BALLOON ALLEGEDLY BURST AT 14 ATM. IT WAS FURTHER REPORTED THAT THE BALLOON WAS ALLEGEDLY SEPARATED FROM THE SHAFT. FURTHERMORE, THERE WAS DIFFICULTY IN RETRACTING THE BALLOON THRU THE INTRODUCER SHEATH AS THE BALLOON SHEARED OFF THE CATHETER AND THE BALLOON DID NOT COME OUT. REPORTEDLY, THE PATIENT WAS INJURED DUE TO RUPTURE AND THE BALLOON COULD NOT BE RETRIEVED WHICH REQUIRED DEPLOYMENT OF TWO ADDITIONAL STENTS TO EXCLUDE THE FOREIGN BODY. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 02/2026) . | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02430 | 2023-04-11 | 2023-03-23 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE NEXT DAY AFTER PLACEMENT, INTRAURETHRAL BALLOON PLACEMENT AND BALLOON INVERSION WAS FOUND. SINCE NO URINE FLOW WAS OBSERVED DURING AND AFTER CATHETER PLACEMENT, A CT SCAN REVEALED THAT A BALLOON HAD BEEN PLACED IN THE URETHRA. IN ADDITION, THE BALLOON HAD BEEN INVERTED IN THE URETHRA AND BLEEDING WAS OBSERVED WHEN THE CATHETER WAS REMOVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02437 | 2023-04-11 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER INSERTION TRAY | FOLEY CATHETER INSERTION TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | 802010 | 802010 |  | IT WAS REPORTED THAT THE FOLEY CATHETER INSERTION TRAY WAS MISSING THE BETADINE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿MACHINE SPEED OUT OF PARAMETERS¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. BASED ON PATIENT CODE 2645 THE PRODUCT WAS NOT USED ON THE PATIENT. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02458 | 2023-04-11 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 52606 |  | IT WAS REPORTED THAT A PEDIATRIC UROLOGY NURSE SENT AN EMAIL REPORTING THAT THE NURSE HAD BEEN CONTACTED BY ONE OF THE PATIENT'S PARENT STATING THAT THEY HAD BEEN FACING ISSUES WITH THE SIZE DIFFERENCE IN DIFFERENT BOXES OF 6CH HYDROSIL PEDIATRIC INTERMITTENT CATHETER-62606P. SOME PHOTOS ARE ATTACHED TO ILLUSTRATE THE ISSUE. PER ADDITIONAL INFORMATION RECEIVED FROM CUSTOMER ON 21MAR2023, STATED THAT THE LOT NUMBER WAS JUDY0063, BUT STATED THAT THIS ISSUE WAS PRESENT IN MOST BOXES AND THERE WAS A SIZE DISCREPANCY. ON DISCUSSING WITH CUSTOMER, THEY STATED THAT ANOTHER PATIENT ALSO HAD MENTIONED THIS ISSUE. THEY WOULD CLARIFY THIS AS THIS PATIENT WAS CURRENTLY AN INPATIENT. THERE HAD BEEN NO IMMEDIATE HARM BUT POTENTIALLY THERE WOULD BE IF THE PARENT INSERTED A CATHETER WHICH WAS NOT THE APPROPRIATE SIZE. THIS CATHETER WAS PRIMARILY THE MAIN CATHETER THE CUSTOMER USED WHEN STARTING TO TEACH PARENTS ABOUT CATHETERIZATION ON THEIR NEW BABIES. THIS OBVIOUSLY WAS VERY WORRYING, AND THE CUSTOMER WOULD HAVE TO LOOK AT ALTERNATIVE CATHETERS. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 28MAR2023, STATED THAT FOLLOWING DISCUSSION WITH THE OTHER PARENT, THEY HAD NOT NOTICED ANY DIFFERENCE IN CATHETER SIZES. THE PARENT STATED THAT THEIR HUSBAND HAS NOTICED. THE REPRESENTATIVE HAVE JUST ADVISED TO LOOK FROM NOW ON AND SEE IF THERE WAS ANY DIFFERENCE. THE FIRST COMPLAINT WAS FOR A MALE CHILD, THEIR WEIGHT WAS 3.8KG AND THEY WERE 16 WEEKS OLD. | THE REPORTED EVENT WAS INCONCLUSIVE AS THE PHYSICAL SAMPLE WAS NOT RETURNED TO BE MEASURED. THE POTENTIAL ROOT CAUSE COULD BE OPERATOR ERROR. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "FOR UROLOGICAL USE ONLY. INTENDED FOR USE BY PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION AND MEASUREMENT. THE DEVICES ARE PASSED TO THE URINARY BLADDER VIA THE URETHRA. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. PRIOR TO OPENING THE SEALED CATHETER POUCH, APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. WET THE CATHETER BY HOLDING THE PACKAGE WITH THE PRINTED SIDE UP AND TIP THE PACKAGE END-TO-END THREE TO SIX TIMES TO WET THE CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. PEEL OPEN THE PACK AT THE FUNNEL/HANDLE END JUST ENOUGH TO EXPOSE 5 CM (2¿) OF THE CATHETER, OR FOR PRODUCTS WITH AN INSERTION SLEEVE, EXPOSE THE INSERTION SLEEVE. DON'T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. USING THE CATHETER. WITH INSERTION SLEEVE (GRIPPER): HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. WITHOUT INSERTION SLEEVE (GRIPPER): HOLD THE HANDLE/FUNNEL END AND REMOVE THE CATHETER FROM THE PACKAGING. GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA AND ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02435 | 2023-04-11 | 2023-03-13 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER WAS PLACED ON (B)(6) AND ON (B)(6), THE CATHETER WAS REMOVED OUT FROM THE PATIENT AND WAS FOUND TO BE TORN. THERE WAS A VISUAL CUT OBSERVED. THE CATHETER TIP WAS REMOVED WITH A CYSTOSCOPE BY A UROLOGIST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02397 | 2023-04-11 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 52606 |  | IT WAS REPORTED THAT A PEDIATRIC UROLOGY NURSE SENT AN EMAIL REPORTING THAT THE NURSE HAD BEEN CONTACTED BY ONE OF THE PATIENT'S PARENT STATING THAT THEY HAD BEEN FACING ISSUES WITH THE SIZE DIFFERENCE IN DIFFERENT BOXES OF 6CH HYDROSIL PEDIATRIC -62606P. SOME PHOTOS ARE ATTACHED TO ILLUSTRATE THE ISSUE (BATCH# JUDY0063). PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON(B)(6) 2023, STATED THAT THE LOT NUMBER WAS JUDY0063, BUT STATED THAT THIS ISSUE WAS PRESENT IN MOST BOXES AND THERE WAS A SIZE DISCREPANCY. ON DISCUSSING WITH CUSTOMER, THEY STATED THAT ANOTHER PATIENT ALSO HAD MENTIONED THIS ISSUE (BATCH# UNK). THEY WOULD CLARIFY THIS AS THIS PATIENT WAS CURRENTLY AN INPATIENT. THERE HAD BEEN NO IMMEDIATE HARM BUT POTENTIALLY THERE WOULD BE IF THE PARENT INSERTED A CATHETER WHICH WAS NOT THE APPROPRIATE SIZE. THIS CATHETER WAS PRIMARILY THE MAIN CATHETER THE CUSTOMER USED WHEN STARTING TO TEACH PARENTS ABOUT CATHETERISATION ON THEIR NEW BABIES. THIS OBVIOUSLY WAS VERY WORRYING AND THE CUSTOMER WOULD HAVE TO LOOK AT ALTERNATIVE CATHETERS. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON (B)(6) 2023, STATED THAT FOLLOWING DISCUSSION WITH THE OTHER PARENT, THEY HAD NOT NOTICED ANY DIFFERENCE IN CATHETER SIZES. THE PARENT STATED THAT THEIR HUSBAND HAS NOTICED. THE REPRESENTATIVE HAVE JUST ADVISED TO LOOK FROM NOW ON AND SEE IF THERE WAS ANY DIFFERENCE. THE FIRST COMPLAINT WAS FOR A MALE CHILD, THEIR WEIGHT WAS 3.8KG AND THEY WERE 16 WEEKS OLD. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02442 | 2023-04-11 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE CATHETER WAS INSERTED INTO THE PATIENT, AND THE BALLOON OF THE CATHETER WAS FOUND TO BE LEAKING, AND A NEW ONE WAS REPLACED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02438 | 2023-04-11 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT THE PATIENT HAD CATHETERIZATION ON (B)(6) 2023, AND RETURNED TO THE ROOM SAFELY AFTER THE OPERATION. ON MARCH 15, THE FOLEY CATHETER BALLOON RUPTURED BY ITSELF, AND THE CATHETER FELL OUT AUTOMATICALLY. IMMEDIATELY THE USER REPORTED TO THE DOCTOR AND OBSERVED THE PATIENT'S URINATION. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 04APR2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01271 | 2023-04-11 | 2023-03-22 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO FT CATHETER 5F RADIOLOGY BASIC TRAY DL (70 CM GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3295355F |  | IT WAS REPORTED BY THE CUSTOMER "PATIENT CODED OR HAD SHORTNESS OF BREATHE AFTER PLACING PICC LINE IN IR SUITE". NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02443 | 2023-04-11 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123620A |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON RUPTURED AND THE CATHETER FELL OUT FROM THE PATIENT BY ITSELF. THEREFORE, A NEW CATHETER WAS REPLACED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 25MAR2023, NO MISSING PIECES OF BALLOON WERE FOUND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02444 | 2023-04-11 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK MALE EXTERNAL CATHETER | PUREWICK MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | PWM030 |  | IT WAS REPORTED THAT CUSTOMER HAD SEVERAL MALE PATIENTS WITH THE NEWER PUREWICK MALE EXTERNAL CATHETER ON AND THE ADHESIVE WAS SO STICKY IT TORE THE SKIN OFF THE SCROTUM WHEN REMOVED. CUSTOMER STATED THAT THE OLD SYSTEM THEY USED HAD A LARGE ENOUGH OPENING BOTH THE SCROTUM AND SHAFT COULD BE PLACED IN THE BAG, AND ADHESIVE WAS ON LESS SENSITIVE AREA. THEY WONDER IF THE NEWER PUREWICK SYSTEM HAD LARGER OPENING TO ELIMINATE THE ADHESIVE ON THE SCROTUM. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02445 | 2023-04-11 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123620A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INDWELLED AFTER THE OPERATION AND THE BALLOON WAS FOUND TO BE RUPTURED. THEREFORE, A NEW CATHETER WAS REPLACED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 25MAR2023, NO MISSING PIECES OF BALLOON WERE FOUND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02451 | 2023-04-11 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166V14S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED FOR THE PATIENT BEFORE THE OPERATION ON (B)(6) 2023. WHEN THE FOLEY CATHETER WAS CONNECTED TO THE DRAINAGE BAG, IT WAS FOUND THAT THE VALVE OF CATHETER DETACHED. THUS THE CATHETER COULD NOT BE USED AND REMOVED. STATED THAT THE RESETTING OF THE URINARY CATHETER INCREASED THE PATIENT'S PHYSICAL DISCOMFORT. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02449 | 2023-04-11 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123614A |  | IT WAS REPORTED THAT ONE HALF OF THE BALLOON WAS INFLATED FOLLOWING INJECTION OF LIQUID AND THE OTHER HALF COULD NOT BE INFLATED. IT CONTRIBUTED TO AN INCOMPLETE BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE ¿ POOR SAMPLE CONDITION. BASED ON THE PHOTOS, IT WAS OBSERVED THAT THE CATHETER BALLOON WAS NOT FULLY INFLATED. UNABLE TO CONFIRMED BALLOON ASYMMETRICAL AS REPORTED. BESIDES, NO FUNCTIONAL TESTING COULD BE PERFORMED AS ONLY PHOTO SAMPLE IS PROVIDED. THEREFORE, THE REPORTED EVENT IS INCONCLUSIVE DUE TO POOR SAMPLE CONDITION. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: UNDER-INFLATION DURING USE)/ ASSEMBLY OF THE BALLOON TO THE SHAFT/ UNEVEN THICKNESS OF BALLOON. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATER ALLERGY" | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02456 | 2023-04-11 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX¿ LUBRICATH¿ FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0166L26 | 0166L26 |  | IT WAS REPORTED THAT THE CUSTOMER BROUGHT A FOLEY CATHETER 26 FR, 30 CC (REF 0166L26) THAT DID NOT EXPIRE UNTIL 31OCT2026. WHEN THEY OPENED IT, THE CATHETER WAS CUT IN 2 DIFFERENT PLACES. THEY ASKED IF THE PRODUCT WAS RECALLED AND HAD QUITE A FEW ON HAND AND WAS NOT SURE WHAT THEY NEED TO DO. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01262 | 2023-04-11 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO FT CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1295108FD |  | IT WAS REPORTED THE CLINICIAN REPORTED THAT THE MAGNETIC PORTION OF THE SHERLOCK TIP LOCATION SYSTEM WAS FAILING, RESULTING IN A MISREPRESENTATION OF THE CATHETER TIP LOCATION. THE CLINICIAN REPORTED THE 3CG PORTION OF THE SHERLOCK, IN HER OPINION, SEEMS TO BE FUNCTIONING APPROPRIATELY, RESULTING IN A MAXIMALLY HEIGHTEN P WAVE, AS WELL AS A DIAMOND RESPONSE FROM THE SYSTEM. THE CLINICIAN SENT OVER A SCREENSHOT FROM THE SR8. SEE ATTACHMENT. UPON EXAMINATION OF THE EKG STRIP, THE INTRINSIC EKG COMING FROM THE DISTAL PORTION OF THE STYLET APPEARS TO BE HIGHLIGHTED GREEN OVER THE QRS COMPLEX AS OPPOSED TO THE P WAVE. WHEN COMPARED WITH THE EXTERNAL EKG, THE ¿UN-HEIGHTENED¿ P WAVES ON THE INTRINSIC WIRE CORRELATED WITH THE P WAVES ON THE EXTERNAL LEADS. OTHERWISE, MYSELF, AND THE CLINICIAN CAN SEE AND UN-HEIGHTENED P WAVE AND A QRS COMPLEX THAT APPEAR TO BE HIGHLIGHTED GREEN, INDICATING DIAMOND PLACEMENT, FOLLOWED BY A T WAVE. THE DEPTH FINDER WAS GIVEN A RESPONSE OF MID DEPTH. DUE TO THE REPORT, CONCERNS, AND QUESTIONABLE ECG READING, THE CLINICIAN WAS ADVISED TO PERFORM A CHEST X-RAY. THE CHEST X-RAY SHOWED A LEFT-SIDED APPROACHED PICC LINE WITH THE DISTAL TIP AT OR NEAR THE SVC/R ATRIAL JUNCTION, INDICATING A PERFECT PLACEMENT. THE CLINICIAN DID REPORT THAT SHE FELT AS THOUGH THIS WAS HAPPENING MORE ON THE DOUBLE LUMEN PICC LINES AS OPPOSED TO THE OTHER PICC LINES BEING PLACED. SHE ALSO THINKS THAT THE LOT NUMBER FOR THOSE PARTICULAR PICC LINES HAD CHANGED RECENTLY, AND IT IS HER OPINION THAT, THESE ISSUES STARTED AROUND THE SAME TIME SHE NOTICED THE LOT NUMBERS CHANGE. THE CLINICIAN ALSO REPORTED THAT THIS HAPPENS ON BOTH SR8 MACHINES AS WELL AS BOTH SHERLOCK CHEST PLATE DEVICES | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01264 | 2023-04-11 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 20GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202250 |  | IT WAS REPORTED BY THE CUSTOMER THAT "IV WAS BEING PLACED, GUIDEWIRE ADVANCED, CATHETER ADVANCED OVER GUIDEWIRE, NEEDLE RETRACTED, WHEN GUIDEWIRE WAS BEING WITHDRAWN, CATHETER CAUGHT AND BECAME DEFORMED. THE GUIDEWIRE WAS UNABLE TO WITHDRAW AND WHOLE PRODUCT HAD TO BE REMOVED. A SECOND STICK WAS REQUIRED, NO PATIENT HARM REPORTED." | "BECTON DICKINSON AND COMPANY'S (BD) MEDICATION DELIVERY SOLUTIONS (MDS) BUSINESS UNIT WILL DISCONTINUE MALFUNCTION MDR REPORTING FOR CERTAIN DEVICE FAILURES THAT HAVE NOT CAUSED OR CONTRIBUTED TO DEATHS OR SERIOUS INJURIES IN THE PAST TWO YEARS, AND WHERE THE LIKELIHOOD OF A DEATH OR SERIOUS INJURY AS A RESULT OF THESE MALFUNCTIONS IS REMOTE. THIS DECISION FOLLOWS FDA GUIDELINES (MEDICAL DEVICE REPORTING FOR MANUFACTURERS GUIDANCE FOR INDUSTRY AND FOOD AND DRUG ADMINISTRATION STAFF ISSUED NOV 8, 2016, REFERENCE SECTION 2.15). BD NOTIFIED FDA OF THIS DECISION ON MAR 3, 2025. FDA HAS REVIEWED AND RESPONDED TO BD¿S NOTIFICATION (REFERENCE FDA DOCUMENT # (B)(4)) ON MARCH 7, 2025. THIS DOCUMENTATION IS AVAILABLE IN BD DOCUMENT MANAGEMENT SYSTEM (SAP). THIS SUPPLEMENTAL MDR IS BEING FILED TO DOCUMENT THAT THE BELOW DEVICE FAILURE WILL NO LONGER BE CONSIDERED A REPORTABLE MALFUNCTION MDR FOR THE PRODUCT FAMILY BELOW: PRODUCT FAMILY: ACCCATH ACE. DEVICE FAILURE: INSERTION DIFFICULTY. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01267 | 2023-04-11 | 2023-03-22 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO CATHETER 4F RADIOLOGY BASIC TRAY SL 70 CM NITINOL GUIDEWIRE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3194355 |  | IT WAS REPORTED BY THE CUSTOMER "PATIENTS CODING OR SHORTNESS OF BREATHE AFTER PLACING PICC LINE IN IR SUITE". NO OTHER INFORMATION WAS PROVIDED. A SPECIFIC NUMBER OF AFFECTED DEVICES OR PATIENTS WAS NOT PROVIDED BY THE COMPLAINANT. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01269 | 2023-04-11 | 2023-03-23 | I, n, j, u, r, y | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED DURING CLINICAL ASSESSMENT: REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02450 | 2023-04-11 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE TIP OF THE FOLEY CATHETER WHERE THE BALLOON WAS, ROLLED WHEN PULLED OUT. THIS CAUSED A ROUGH EDGE TO CATCH WHEN REMOVING AND THE PATIENT EXPERIENCED BLEEDING. CUSTOMER TESTED SOME OTHER TRAYS AND HAD THE SAME ISSUE. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE COULD BE TO THIN SAC COULD CAUSE POOR SHAPE/ BAGGY SAC & NON- CONCENTRIC BALLOON. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE THE PRODUCT CATALOG NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00242 | 2023-04-11 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG120144 | ATG120144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN PULMONARY ARTERY, THE BALLOON WAS ALLEGEDLY FAILED TO DEFLATE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01265 | 2023-04-11 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC KIT 5F DL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 92655517 |  | IT WAS REPORTED, DUAL GROSHONG BREAKING/LEAKING AT THE CONNECTION OF THE WHITE LUMEN BELOW THE INJECTION CAP. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02505 | 2023-04-12 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE PATIENT HAD INDWELLING FOLEY CATHETER AFTER BLADDER EXAMINATION ON (B)(6) 2023. THEN THE FOLEY CATHETER BALLOON RUPTURED, WHICH AFFECTED THE DRAINAGE OF URINE. THEN A NEW CATHETER WAS REPLACED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 04APR2023, STATED THAT THERE WERE NO MISSING PIECES OF BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
MW5116603 | 2023-04-12 | 2023-03-31 | I, n, j, u, r, y | Voluntary report | SOLO PICC LINE | CATHETER, INTRAVASCULAR, THERAPEUTIC, LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 4 FR SINGLE LUMEN |  |  | PICC WAS MALFUNCTIONING, ONCE REMOVED, THERE WAS BREAK BETWEEN THE 5-6CM MARK ON THE CATHETER. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01301 | 2023-04-12 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN, WHITE, 4.2F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 0601600 | 0601610 |  | IT WAS REPORTED THAT SOMETIME POST CATHETER PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY FRACTURED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01302 | 2023-04-12 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BROVIAC CV CATHETER REPAIR KIT, SINGLE-LUMEN, WHITE, 6.6F | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 0601620 | 0601620 |  | IT WAS REPORTED THAT SOMETIME POST CATHETER PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY FRACTURED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02509 | 2023-04-12 | 2023-01-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT A PATIENT'S RECENT CATHETER HAD BURST INSIDE THEIR BLADDER (BATCH# NGGT2360). ALSO STATED THAT SEVERAL CATHETER BALLOON WERE DEFLATED INSIDE THEIR BLADDER PREVIOUSLY (BATCH# NGFX3156 AND BATCH# NGFV3848). THE CUSTOMER WANTED TO MAKE THE REPRESENTATIVE AWARE IN CASE OF A PRODUCTION OR BATCH ERROR. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 23MAR2023, STATED THAT THE FOLEY CATHETER BALLOON OF BATCH # NGGT2360 WAS SPLIT, CATHETER BALLOON OF BATCH # NGFX3156 WAS DEFLATED AND CATHETER BALLOON OF BATCH # NGFV3848 WAS ALSO DEFLATED. THE BALLOON APPEARED TO BE CLEAN SPLIT. THE CUSTOMER WAS UNSURE IF ANY MATERIAL WAS LEFT IN BLADDER, SO THEY ESCALATED TO GENERAL PRACTITIONER WHO DIDN¿T FEEL APPROPRIATE FOR A ULTRASOUND SCAN TO CONFIRM THIS UNLESS THE PATIENT BECOMES SYMPTOMATIC. ON (B)(6) 2023, THE PREVIOUS FOLEY CATHETER BALLOON INSERTED WITH BATCH # NGGT2360 WAS SPLIT, ON (B)(6) 2023 THE BALLOON DEFLATED FOR WHICH THEY WERE UNSURE OF LOT NUMBER AS THIS WAS NOT DOCUMENTED (BATCH# UNK) WHEN IT WAS INSERTED AND ON (B)(6) 2022, THE CATHETER BALLOON WITH BATCH # NGFV3848 DEFLATED. THEY WERE UNSURE IF BALLOON WITH BATCH NGGT2360 WAS EITHER DEFLATED OR SPLIT FIRST AND IT WAS UNABLE TO DETERMINE. UPON ARRIVAL THIS CATHETER WAS ALREADY OUT DUE TO FALLING OUT AND ALREADY HAD SPLIT AS PER THE PICTURE PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02508 | 2023-04-12 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT THE CUSTOMER HAD ENCOUNTERED ISSUE WITH THE BALLOONS OF SILICONE INDWELLING CATHETERS LOSING FLUID AND DEFLATING EARLIER THIS WEEK. THE FOLEY CATHETER WAS PLACED ON (B)(6) 2023 IN SURGERY, AND IT WAS DOCUMENTED THAT THE BALLOON WAS INFLATED. HOWEVER, AT THE TIME OF CATHETER REMOVAL, THE BALLOON WAS FOUND TO BE COMPLETELY DEFLATED (LOT# NGGR3688). THIS WAS THE FIRST TIME THEY HAD HEARD OF THIS ISSUE, BUT THE REGISTERED NURSE ON THIS PATIENT CARE UNIT NOTED THIS MULTIPLE TIMES RECENTLY (LOT# UNK). ALTHOUGH THEY DID NOT HAVE ANY SPECIFICS OR THE CATHETER. THEY DID RE-INFLATE THE BALLOON YESTERDAY AND PLACED IT IN A BAG TO SAVE IT. TODAY, THERE WAS SOME FLUID IN THE BAG, WHICH WOULD INDICATE THE BALLOON WAS SLOWLY LEAKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02506 | 2023-04-12 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PLACED IN PATIENT AS ORDERED FOR THE C-SECTION. DURING FUNDUS MASSAGE AFTER C-SECTION SURGERY, THE FOLEY CATHETER PLACED IN PATIENT CAME OUT AND BALLOON WAS FOUND DEFLATED AND IT APPEARED LIKE THE EXTERIOR LAYER OF THE TUBING WAS SPLIT AND SEPARATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02503 | 2023-04-12 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. COMPLETE CARE® FOLEY TRAY | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 900014A |  | IT WAS REPORTED THAT THE POUCH OF THE POVIDONE IODINE APPEARED TO BE RUPTURED INSIDE THE UNOPENED STERILE PACKAGE OF THE FOLEY TRAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02502 | 2023-04-12 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON SHAFT WAS DAMAGED DURING THE PRETEST. THE CATHETER WAS NOT USED. STATED THAT BALLOON RUPTURE OR TEAR WAS OBSERVED ON THE CATHETER. NONE OF THE MATERIALS POSSIBLY CAME IN CONTACT TO THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02501 | 2023-04-12 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE CATHETER HAD A DAMAGE. THE FOLEY CATHETER WAS INDWELLED A DAY BEFORE, AND AFTER ABOUT 16 HOURS, WHEN THE USER LOOKED INSIDE THE DIAPER, IT NATURALLY CAME OUT. UPON CLOSER INSPECTION, A CRACK WAS OBSERVED ON THE LEFT SIDE OF THE CATHETER. SOME PARTS OF THE CATHETER WERE DISCOLORED. PER MAIL NOTIFICATION RECEIVED FROM IBC ON 28MAR2023, STATED THAT THE BALLOON RUPTURE WAS OBSERVED TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02497 | 2023-04-12 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | MAGIC 3 INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614 | 53614 |  | IT WAS REPORTED THAT THE CUSTOMER ORDERED THROUGH LIBERATOR AND RECEIVED A TOTAL OF 12 BOXES OF INTERMITTENT CATHETER DID NOT HAVE SHIPPING SHEET WITH THE ORDER NUMBER. ALSO STATED THAT FOR 7 OF THE 12 BOXES RECEIVED, NONE OF THE CATHETERS HAD SLEEVES ON THEM AND THEY WOULD NEED REPLACEMENTS. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT CUSTOMER UNDERSTANDS THAT THEY WERE SENT A SUBSTITUTE CATHETER BECAUSE THE ONES THEY NORMALLY USED WERE ON BACKORDER WITHOUT A RELEASE DATE. THEY ALSO STATED THAT ITEM #53814G LOT # JUGV1681 HAD NO HOLES IN THE TIP, AND THEY HAD SAMPLES OF THAT ONE AVAILABLE . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02494 | 2023-04-12 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE BALLOON DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02492 | 2023-04-12 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER HAD BEEN FALLEN OUT OF THE PATIENT ON THE NEXT DAY AFTER PLACEMENT, WHILE THE PATIENT WAS SLEEPING. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02489 | 2023-04-12 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON GOT DAMAGED DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02486 | 2023-04-12 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER COULD NOT BE WITHDRAWN AFTER BALLOON INFLATION IN THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
MW5116595 | 2023-04-12 | 2023-04-06 | I, n, j, u, r, y | Voluntary report | FOLEY CATHETER | CATHETER, UROLOGICAL (ANTIMICROBIAL) AND ACCESSORIES | C.R. BARD, INC. | A319514AM | A319514AM |  | PATIENT HAD A TEMP SENSING FOLEY CATHETER IN AND IT WAS NOTED THAT THERE WAS NO URINE OUTPUT FOR THE CURRENT HOUR COMPARED TO THE PREVIOUS HOUR. WHEN LOOKING AT THE 14FR BALLOON END, THE BALLOON WAS RUPTURED AND NOT INTACT. THIS PARTICULAR FOLEY WAS ONLY PLACED SEVERAL HOURS PRIOR IN THE SAME DAY. NEW 14FR BALLOON TEMPERATURE PROBE FOLEY INSERTED WITHOUT DIFFICULTY AND PATIENT DID NOT EXPERIENCE ANY FURTHER ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02485 | 2023-04-12 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | FOL0102 | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER PATIENT WAS USING FOR HIS STOMA WAS LEAKING WHEN PATIENT USES IT. PATIENT CALLED THE DO AND CONFIRMED THAT IT WAS THE SAME FR SIZE THAT THE DR HAD USED AND DID NOT HAVE ANY ISSUES THEN. THE DOCTOR'S OFFICE ADVISED PATIENT TO CALL US, BUT REPRESENTATIVE EXPLAINED TO PATIENT THAT HE SHOULD CALL THEM AND EXPLAIN WHAT WAS HAPPENING TO SEE IF THEY HAVE ANY MEDICAL ADVICE OR A SOLUTION TO WHY IT WAS LEAKING. REPRESENTATIVE EXPLAINED TO HIM IT COULD BE THAT HE NEEDS LARGER FR SIZE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01300 | 2023-04-12 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT PRIOR TO PORT PLACEMENT PROCEDURE, PACKAGING ALLEGEDLY HAD A STERILITY PROBLEM AROUND THE TRAY. IT WAS FURTHER REPORTED THAT BOTH THE TOP LABEL AND THE INNER TRAY COVER ALLEGEDLY SEEM TO BE SLIT BY A MACHINE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 10/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01299 | 2023-04-12 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0602680 | 0602680 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE NEEDLE TIP WAS ALLEGEDLY DEFECTIVE. IT WAS FURTHER REPORTED THAT GUIDEWIRE WAS ALLEGEDLY FRAYED. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRATION DATE: 03/2026. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01298 | 2023-04-12 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT WAS ALLEGEDLY DAMAGED DUE TO THE PHYSICAL ATTACK THAT HAPPENED YEARS AGO. IT WAS FURTHER REPORTED THAT THE PATIENT HAD NUMEROUS VASCULAR ACCESS DEVICES IN THE PAST. REPORTEDLY, THE PATIENT ALLEGEDLY EXPERIENCED PAIN WHEN THE PORT WAS ACCESSED AND THE BLOOD RETURN WAS ALLEGEDLY NOT POSSIBLE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01294 | 2023-04-12 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNK BROVIAC REPAIR KIT | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 |  | UNK BROVIAC REPAIR KIT |  | IT WAS REPORTED THAT SOMETIME POST CHRONIC CATHETER PLACEMENT, THE DEVICE LINE ALLEGEDLY BROKE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01296 | 2023-04-12 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 2.7F BROVIAC CV CATHETERS 0.5 MM LUMEN REPAIR KIT | PORT AND CATHETER ACCESSORIES | C.R. BARD, INC. (BASD) -3006260740 | 0601600 | 0601600 |  | IT WAS REPORTED THAT SOMETIME POST CATHETER PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY FRACTURED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01295 | 2023-04-12 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT APPROXIMATELY SIX YEARS AND SEVEN MONTHS POST PORT PLACEMENT VIA THE INTERNAL JUGULAR VEIN, THE CATHETER WAS ALLEGEDLY BROKEN. REPORTEDLY, THE PORT BODY AND THE DISTAL CATHETER SEGMENT WERE REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02473 | 2023-04-12 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PATIENT WAS INDWELLED WITH A THREE-LUMEN FOLEY CATHETER IN ORDER TO STOP THE BLEEDING AND DRAIN URINE ON (B)(6) 2023. ON (B)(6) 2023, THE PATIENT COMPLAINED THAT THE CATHETER FELL OUT. UPON CHECKING, IT WAS FOUND THAT THE BALLOON OF THE FOLEY CATHETER WAS DAMAGED, AND A SMALL AMOUNT OF BLOOD WAS FLOWING OUT OF THE URETHRA. IMMEDIATELY INFORMED THE DOCTOR, AND RE-INDWELLED A THREE-LUMEN CATHETER FOR THE PATIENT. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 25MAR2023, NO MISSING PIECES OF BALLOON WERE FOUND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02476 | 2023-04-12 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ALL SILICONE CATHETER | UNKNOWN SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT THE CUSTOMER HAD ENCOUNTERED ISSUE WITH THE BALLOONS OF SILICONE INDWELLING CATHETERS LOSING FLUID AND DEFLATING EARLIER THIS WEEK. THE FOLEY CATHETER WAS PLACED ON (B)(6) 2023 IN SURGERY, AND IT WAS DOCUMENTED THAT THE BALLOON WAS INFLATED. HOWEVER, AT THE TIME OF CATHETER REMOVAL, THE BALLOON WAS FOUND TO BE COMPLETELY DEFLATED. THIS WAS THE FIRST TIME THEY HAD HEARD OF THIS ISSUE, BUT THE REGISTERED NURSE ON THIS PATIENT CARE UNIT NOTED THIS MULTIPLE TIMES RECENTLY. ALTHOUGH THEY DID NOT HAVE ANY SPECIFICS OR THE CATHETER. THEY DID RE-INFLATE THE BALLOON YESTERDAY AND PLACED IT IN A BAG TO SAVE IT. TODAY, THERE WAS SOME FLUID IN THE BAG, WHICH WOULD INDICATE THE BALLOON WAS SLOWLY LEAKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02477 | 2023-04-12 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314M | 119314M |  | IT WAS REPORTED THAT FOLEY CATHETER NOT DRAINING. WHEN ASSESSED FOUND THAT CATHETER HAD COME OUT BALLOON WAS NOT INFLATED, AND BALLOON NOTED TO BE RUPTURED. FOLEY CATHETER HAD DRAINED MINIMAL OVERNIGHT BUT WAS UNABLE TO BE ASSESSED AS PATIENT WAS PRONE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02481 | 2023-04-12 | 2023-03-19 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 123512A | 123512A |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED ON (B)(6) 2023. ON (B)(6) 2023, THE NURSE DISCOVERED THAT THE CATHETER WAS DISLODGED DURING ROUTINE CARE. THE CATHETER WAS INSPECTED, THE BALLOON WAS COMPLETE BUT BURST. AS PER THE PHOTO SAMPLE RECEIVED, IT WAS FOUND THAT THE CATHETER BALLOON WAS BURST. STATED THAT THE USER WAS EXPOSED TO BLOOD, OR BODILY FLUIDS AND THE DETAILED EXPOSURE WAS URINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02484 | 2023-04-12 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119416M | 119416M |  | IT WAS REPORTED THAT THE BLADDER TEMPERATURE WOULD NOT CAPTURE OR REGISTER, CHARGED OUT TEMPERATURE SENSING CORDS 3 TIMES AND MONITOR MODULE. ALSO STATED THAT FOLEY CATHETER WAS MAINTAINED AND NOT REPLACED IN PATIENT AND USED ESOPHAGEAL PROBE FOR TEMPERATURE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿SENSOR WIRE TOO WEAK ¿. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿PROPER TECHNIQUES FOR URINARY CATHETER INSERTION. ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION. ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE. ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT. ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE. ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED. DIRECTIONS FOR USE: WASH HANDS AND DON CLEAN GLOVES. EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT. 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERIURETHRAL AREA. 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL. 5. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP. 6. DON STERILE GLOVES. 7. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN. NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE. 8. POSITION FENESTRATED DRAPE ON PATIENT. 9. SATURATE 3 FOAM SWAB STICKS IN POVIDONE IODINE. 10.ATTACH THE WATER FILLED SYRINGE TO THE INFLATION PORT. NOTE: IT IS NOT NECESSARY TO PRE-TEST THE FOLEY CATHETER BALLOON. 11. REMOVE FOLEY CATHETER FROM WRAP AND LUBRICATE CATHETER. 12. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS. NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA INORA. MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD. 13. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND A. WHEN CATHETER TIP HAS NTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE. B. INSERT CATHETER TWO MORE INCHES AND INFLATE CATHETER BALLOON. 14. INFLATE CATHETER BALLOON USING ENTIRE 10CC OF STERILE WATER PROVIDED IN THE PREFILLED SYRINGE. NOTE: USE OF LESS THAN 10CC CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON 15. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK. 16. SECURE THE FOLEY CATHETER TO THE PATIENT USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED (SEE STATLOCK® FOLEY STABILIZATION DEVICE IFU). NOTE: PLEASE MAKE SURE PATIENT IS APPROPRIATE FOR USE OF STATLOCK® STABILIZATION DEVICE 17. POSITION HANGER ON BED RAIL AT THE FOOT OF THE BED. NOTE: EXERCISE CARE TO KEEP BAG OFF THE FLOOR. 18. USE GREEN SHEETING CLIP TO SECURE DRAINAGE TUBE TO THE SHEET. MAKE SURE TUBE IS NOT KINKED. 19. INDICATE TIME AND DATE OF CATHETER INSERTION ON PROVIDED LABELS. PLACE DESIGNATED LABELS ON PATIENT CHART AND DRAINAGE SYSTEM. 20. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL FOLEY CATHETER REMOVAL. 1. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. 2. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. 3. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. 4. IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. 5. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT¿. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02466 | 2023-04-12 | 2023-02-18 | I, n, j, u, r, y | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT PLACED AN INDWELLING CATHETER IN THE BLADDER FOR IRRIGATION AFTER SURGERY IN THE MORNING OF (B)(6) 2023. NO LUBRICANT WAS USED, IT WAS STERILIZED WITH POVIDONE IODINE AND THEN INSERTED. THE BALLOON WAS INJECTED WITH 30ML OF NORMAL SALINE. THEN, THE NURSE FOUND THAT THE INDWELLING CATHETER HAD PROTRUDED, AND FOUND THAT THE BALLOON WAS LEAKING. THE INCISION COULD NOT BE PRESSED, SO A NEW CATHETER WAS REINSERTED. STATED THAT THE URINARY CATHETER PROLAPSE COULD NOT BE CUT BY COMPRESSION, RE-INSERTION OF CATHETER WAS DONE WHICH INCREASED TREATMENT TIME, AND RISK OF INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02545 | 2023-04-13 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT THE EYELETS ON SEVEN OF THE INTERMITTENT CATHETERS WERE NOT CUT OR PUNCHED PROPERLY. PATIENT DID NOT HAVE THE LOT NUMBER OR ANY SAMPLES TO SEND. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02547 | 2023-04-13 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE MAGIC 3 HYDROPHILIC INTERMITTENT CATHETER WOULD NOT INSERT ALL THE WAY. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OUTSIDE DIAMETER IS TOO LARGE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE HYDROPHILIC COATED SILICONE INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02544 | 2023-04-13 | 2023-03-24 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER LF HC | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT A URINARY TRACT INFECTION SO DOCTOR TOLD THE PATIENT TO DISCONTINUE USE. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE SETUP: 1. CONNECT THE CANISTER TO WALL SUCTION AND SET TO A MINIMUM OF 40MMHG CONTINUOUS SUCTION. ALWAYS USE THE MINIMUM AMOUNT OF SUCTION NECESSARY. IF USING THE DRYDOCTM VACUUM STATION, CONNECT THE CANISTER TO THE UNIT AND TURN THE UNIT ON. PLEASE CONSULT THE DRYDOCTM VACUUM STATION USER GUIDE FOR FURTHER INFORMATION. 2. USING STANDARD SUCTION TUBING, CONNECT THE PUREWICKTM FEMALE EXTERNAL CATHETER TO THE COLLECTION CANISTER. PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. REMOVAL: 5. TO REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER, FULLY SEPARATE THE LEGS, GLUTEUS, AND LABIA. TO AVOID POTENTIAL SKIN INJURY UPON REMOVAL, GENTLY PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER DIRECTLY OUTWARD. ENSURE SUCTION IS MAINTAINED WHILE REMOVING THE PUREWICKTM FEMALE EXTERNAL CATHETER. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. MAINTENANCE: 6. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. INDICATION FOR USE: THE PUREWICKTM FEMALE EXTERNAL CATHETER IS INTENDED FOR NON-INVASIVE URINE OUTPUT MANAGEMENT IN FEMALE PATIENTS. CONTRAINDICATIONS: ¿ PATIENTS WITH URINARY RETENTION WARNINGS: ¿ DO NOT USE THE PUREWICKTM FEMALE EXTERNAL CATHETER WITH BEDPAN OR ANY MATERIAL THAT DOES NOT ALLOW FOR SUFFICIENT AIRFLOW. ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. ¿ AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER; HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE; EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE; EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON. ¿ DO NOT USE BARRIER CREAM ON THE PERINEUM WHEN USING THE PUREWICKTM FEMALE EXTERNAL CATHETER.BARRIER CREAM MAY IMPEDE SUCTION. ¿ PROCEED WITH CAUTION IN PATIENTS WHO HAVE UNDERGONE RECENT SURGERY OF THE EXTERNAL UROGENITAL TRACT. ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. ¿ MAINTAIN SUCTION UNTIL THE PUREWICKTM FEMALE EXTERNAL CATHETER IS FULLY REMOVED FROM THE PATIENT TO AVOID URINE BACKFLOW. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ PRIOR TO CONNECTING THE PUREWICKTM FEMALE EXTERNAL CATHETER TO HOSPITAL WALL SUCTION TUBING, VERIFY SUCTION FUNCTION BY COVERING THE OPEN END OF THE SUCTION TUBING WITH ONE HAND AND OBSERVING THE PRESSURE DIAL. IF THE PRESSURE DOES NOT INCREASE WHEN THE LINE IS COVERED, VERIFY THAT THE TUBING IS SECURED, CONNECTED, AND NOT KINKED. ¿ ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. ¿ PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS.". THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02539 | 2023-04-13 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123518A | 123518A |  | IT WAS REPORTED THAT FOLEY CATHETER WAS DEFECTIVE. THEY WENT TO CHANGE THE CATHETER, USUALLY THE BALLOON OPENED UP BUT INSTEAD IT WERE ALL BUNCHED UP WHERE THE CATHETER GOES IN ON THAT END. | THE REPORTED EVENT WAS NOT CONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS NOT CONFIRMED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE ''UNDER CURED SAC'' (EXAMPLE: POOR SHAPE/ BAGGY SAC-NONCONCENTRIC BALLOON). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "CDC GUIDELINES FOR APPROPRIATE INDICATIONS FOR INDWELLING URETHRAL CATHETER USE: ¿ PATIENT HAS ACUTE URINARY RETENTION OR BLADDER OUTLET OBSTRUCTION ¿ NEED FOR ACCURATE URINE OUTPUT MEASUREMENTS ¿ USE FOR SELECTED SURGICAL PROCEDURES ¿ TO ASSIST IN HEALING OF OPEN SACRAL OR PERINEAL WOUNDS ¿ PATIENT REQUIRES PROLONGED IMMOBILIZATION ¿ TO IMPROVE COMFORT FOR END OF LIFE CARE PROPER TECHNIQUES FOR URINARY CATHETER INSERTION: - PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION - INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT - USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE - DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN THE PATIENT RECORD PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: - SECURE THE FOLEY CATHETER. USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED - MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS - MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING - KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES - EMPTY THE COLLECTION BAG REGULARLY USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT * GENERALLY, DRAINAGE IS ACCOMPLISHED BY INSERTING THE CATHETER THROUGH THE URETHRA AND INTO THE BLADDER. HOWEVER, DRAINAGE IS SOMETIMES ACCOMPLISHED BY SUPRAPUBIC OR OTHER PLACEMENT OF THE CATHETER, SUCH AS NEPHROSTOMY TRACT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02541 | 2023-04-13 | 2023-03-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD ISSUES WITH THE PUREWICK FEMALE EXTERNAL CATHETER IN THE HOSPITAL BECAUSE OF IMPROPERLY TRAINED STAFF. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02551 | 2023-04-13 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THERE WAS A CRACK IN THE TIP OF THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00248 | 2023-04-13 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U4150215RX | U4150215RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE SUPERFICIAL FEMORAL ARTERY VIA A CONTRALATERAL APPROACH, THE PTA BALLOON ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 03/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01320 | 2023-04-13 | 2023-02-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT THEY HAD A PROBLEM WITH THE END CAP ON THE SINGLE LUMEN PICC. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01317 | 2023-04-13 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK CV PLUS SLIDING POST, TRICOT HYDROCOLLOID CRESCENT PAD, PIGTAIL PAD | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | CV0220CE |  | IT WAS REPORTED THAT "ON (B)(6) 2023, DURING THE FIRST CYCLE OF CHEMOTHERAPY FOR THE PATIENT, THE CATHETER RETAINER WAS OPENED DURING THE MAINTENANCE OF THE PICC CATHETER FOR THE PATIENT. HOWEVER, IT WAS FOUND THAT THE FASTENER WAS NOT PROPERLY CONNECTED AND WAS NOT USED, SO THE CATHETER RETAINER WAS REPLACED." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01313 | 2023-04-13 | 2023-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT "ON (B)(6) 2023: IN THE PROCESS OF PICC CATHETER PLACEMENT FOR THE PATIENT, WHEN THE GUIDE WIRE WAS NEEDED, THE TOP WAS BLUNT AND COULD NOT ENTER THE NEEDLE GUIDE. AFTER CUTTING, THE GUIDE WIRE WAS USED." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01318 | 2023-04-13 | 2022-12-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 15CM (PRE-CURVED) (SHORT-TERM) DL | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5594150 |  | IT WAS REPORTED BY THE ATTENDING PHYSICIAN THAT "THE METALLIC GUIDE OF THE CATHETER PRESENTS RESISTANCE, THEREFORE IT IS NOT POSSIBLE FOR IT TO PASS, FOR WHICH REASON ANOTHER CATHETER IS REQUESTED." ADDITIONAL INFORMATION RECEIVED ON 06.JAN.2023: RESISTANCE WAS NOTED BEFORE USE ON THE PATIENT. THERE WAS NO PATIENT HARM REPORTED, NO NEED FOR MEDICAL OR SURGICAL INTERVENTION. (B)(6) 2023 - THE RETURNED SAMPLE EXHIBITED A BREAK IN THE COIL WIRE OF THE GUIDEWIRE. | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE ANALYSIS, PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF DIFFICULTY PASSING A GUIDEWIRE IS CONFIRMED. ONE NIAGARA KIT TRAY WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED THE KIT TRAY WAS RETURNED WITHOUT A LID AND ALL KIT COMPONENTS EXCEPT ONE END CAP AND TWO TRANSPARENT DRESSINGS WERE RETURNED WITH THE TRAY. USE RESIDUE WAS OBSERVED ON THE KIT COMPONENTS, AND THE GUIDEWIRE WAS OBSERVED TO BE DAMAGED. THE CORE WIRE OF THE GUIDEWIRE WAS OBSERVED TO BE BROKEN DURING TACTILE TESTING. A MICROSCOPIC OBSERVATION REVEALED SOME DAMAGE ON THE BEVEL OF THE INTRODUCER NEEDLE. BOTH WELD TIPS OF THE GUIDEWIRE WERE OBSERVED TO BE PRESENT, AND NO BREAKS WERE FOUND IN THE COILED WIRE. THE BREAK IN THE CORE WIRE OF THE GUIDEWIRE WAS OBSERVED TO BE TAPERED WITH A MOSTLY SMOOTH AND GRANULAR FRACTURE SURFACE, WHICH IS INDICATIVE OF TENSILE FAILURE CAUSED BY EXCESSIVE PULLING FORCES ON THE GUIDEWIRE. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02555 | 2023-04-13 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER DIFFICULT TO BE REMOVED FROM THE FOLEY CATHETER BALLOON. THE USER ALSO FELT RESISTANCE WHEN INJECTING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01319 | 2023-04-13 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT DURING USE OF THE SELDINGER, AFTER SUCCESSFUL PUNCTURE, THE GUIDE WIRE WAS INSERTED. FOLLOWING REMOVAL OF THE INTRODUCER NEEDLE, IT WAS FOUND THAT THE INTRODUCER SHEATH COULD NOT PASS THROUGH THE GUIDE WIRE SINCE THE DISTAL END OF THE GUIDE WIRE WAS THICKENED. THIS LED TO TRIMMING WITH SCISSORS DURING USE. BUT THE TRIMMING WAS DIFFICULT AND THERE WERE POWDERS RISING FROM THE TRIMMING ON THE GUIDE WIRE. NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01323 | 2023-04-13 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC SV 4F DL BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 3274118J |  | IT WAS REPORTED THAT THE PATIENT HAD A POWERPICC IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, LEAKAGE OCCURRED FROM THE CATHETER NEAR THE LUER HUB. THEREFORE, THE DEVICE WAS REMOVED FROM THE PATIENT¿S BODY. THE CATHETER WAS SECURED BY THE SECURING METHOD USED IN PEDIATRICS, NOT BY THE STATLOCK SECURING. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01324 | 2023-04-13 | 2023-02-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PER-Q-CATH PLUS PICC WITH STIFFENING STYLET WIRE 5F (DUAL-LUMEN)(65CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 4235105 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE CATHETER PERFORATION DURING GUIDEWIRE REMOVAL. NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. . | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02549 | 2023-04-13 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT EVEN THOUGH THE SYRINGE WAS INSERTED INTO THE INFLATION VALVE AND INJECTED, THE FIXED WATER DOES NOT ENTER INTO CATHETER AND LEAKS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02554 | 2023-04-13 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THERE WAS A HOLE ON THE CATHETER TIP BEFORE USING IT. THE LOCATION OF THE HOLE WAS AT THE TEMPERATURE SENSOR AREA AND THE USER ALLEGEDLY THOUGHT TO REFRAIN FROM USING IT BECAUSE OF BEING SUPPOSED THAT IT WOULD CAUSE TEMPERATURE DEFECTS, BUT USED IT. THE PATIENT WAS INFECTED, SO THE SAMPLE WAS UNABLE TO BE RECOVERED. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE INFECTION AT THIS TIME AND MEDICAL INTERVENTION WAS UNKNOWN. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 27MAR2023, IT WAS STATED THAT THE A CRACK ON THE CATHETER TIP WAS REPORTED. THERE IS NO ASSOCIATION BETWEEN A PATIENT HAVING AN INFECTION AND CATHETER ISSUE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02532 | 2023-04-13 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN WATER WAS INJECTED DURING THE BALLOON TEST AND AN ATTEMPT WAS MADE TO DRAIN THE STERILIZED WATER INTO THE SYRINGE, THE WATER DID NOT RETURN EVEN WHEN NEGATIVE PRESSURE WAS APPLIED BY PULLING THE SYRINGE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN 2 MINUTES AND 34 SECONDS WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01309 | 2023-04-13 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED ON (B)(6) 2023, AT 10:00 THE CATHETER WAS FOUND TO LEAK LIQUID DURING INFUSION, AND THE DRUG LIQUID WAS FOUND TO LEAK OUTSIDE. AT 10:30 THE CENTRAL VENOUS CATHETER ATTACHMENT WAS REPLACED. REPLACEMENT WAS COMPLETED AT 11:30. PICC WAS USED FOR INTRAVENOUS CHEMOTHERAPY. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01307 | 2023-04-13 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED BY THE CUSTOMER" T-LOCK CONNECTOR WHERE WIRE PASSES THROUGH SEPTUM. SOME LEAKING IS OCCURRING THROUGH SEPTUM." NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01311 | 2023-04-13 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT SOMETIME POST PLACEMENT PROCEDURE, THE PORT WAS ALLEGEDLY OCCLUDED AND EVEN AFTER HAVING ANTI-THROMBOLYTIC AGENTS AND DYE STUDIES PERFORMED, THE PORT WAS UNABLE TO BE INJECTED DUE TO THE OCCLUSION. IT WAS FURTHER REPORTED THAT X-RAY THAT SHOWED THE BEND AT THE END OF THE CLEAR TUBE LOCATED IN THE HEART BOTTOM CHAMBER. THE PORT WAS REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 01/2024. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02542 | 2023-04-13 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616 | 53616 |  | IT WAS REPORTED THAT THE MAGIC3 GO MALE INTERMITTENT CATHETERS DID NOT HAVE MUCH LUBE. CUSTOMER ASKED IF THE CATHETER WOULD DRY OUT AND HOW LONG IT WOULD BE GOOD. REPRESENTATIVE ADVISED CUSTOMER ON USAGE. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTERMITTENT CATHETER HYDROPHILIC COATED SILICONE" INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. THE CATHETER BECOMES SLIPPERY WHEN WETTED WITH WATER, ELIMINATING THE NEED FOR A SEPARATE LUBRICANT. FOR YOUR ADDED CONVENIENCE, THIS CATHETER IS PACKAGED WITH ITS OWN STERILE WATER. SIMPLY RELEASE THE WATER FROM ITS FOIL PACKET AND THEN TIP THE UN-OPENED CATHETER PACKAGE END-TO-END. THE CATHETER ACTS LIKE A MAGNET TO ATTRACT THE WATER AND ACTIVATE ITS SLIPPERY COATING. FOLLOW THESE STEPS FOR BEST RESULTS. RELEASE THE WATER PRIOR TO OPENING THE SEALED CATHETER POUCH: APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. WET THE CATHETER. HOLD PACKAGE WITH PRINTED SIDE UP. TIP PACKAGE END-TO-END THREE TO SIX TIMES TO WET CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. USE THE CATHETER. PEEL BACK PACKAGE TO EXPOSE FUNNEL END OF CATHETER. IF DESIRED, USE THE SELF-ADHESIVE TAPE ON THE PACKAGE TO TEMPORARILY ATTACH THE POUCH TO ANY DRY VERTICAL SURFACE. REMOVE CATHETER AND USE ACCORDING TO PHYSICIAN¿S INSTRUCTIONS. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE, DO NOT RESTERILIZE. MADE OF SILICONE ELASTOMER.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
2020394-2023-00246 | 2023-04-13 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U35130630 | U35130630 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN RIGHT SUPERFICIAL FEMORAL ARTERY THROUGH CONTRALATERAL APPROACH, THE PTA BALLOON WAS SEEMED TO BE ALLEGEDLY TWISTED WHILE IN THE PATIENT. IT WAS FURTHER REPORTED THAT WHILE INFLATING TO NOMINAL PRESSURE, THERE WAS ONE CENTIMETER SEGMENT WHICH WAS NOT DILATED. REPORTEDLY, THE BALLOON WAS RE POSITIONED TO POST DILATE THE SEGMENT THAT WAS NOTICED AS NOT DILATED AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 06/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01303 | 2023-04-13 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | A710962 | A710962 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PORT ALLEGEDLY BLED THROUGH THE CATHETER. HOWEVER, THERE IS NO INFORMATION REGARDING TREATMENT PROVIDED FOR BLEEDING. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 11/2022). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02514 | 2023-04-13 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE CATHETER WAS SPONTANEOUSLY WITHDRAWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02510 | 2023-04-13 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT STERILIZED WATER LEAKED FROM NEAR THE VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02519 | 2023-04-13 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50000000 | 119416M |  | IT WAS REPORTED THAT THE BLADDER TEMPERATURE WOULD NOT CAPTURE OR REGISTER, CHARGED OUT TEMPERATURE SENSING CORDS 3 TIMES AND MONITOR MODULE. ALSO STATED THAT FOLEY CATHETER WAS MAINTAINED AND NOT REPLACED IN PATIENT AND USED ESOPHAGEAL PROBE FOR TEMPERATURE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02523 | 2023-04-13 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON DEFLATED AND THE FOLEY CATHETER FELL OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02512 | 2023-04-13 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0119L20 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE FOLEY CATHETER. THE CATHETER WAS PLACED TO THE PATIENT IN THE OPERATING ROOM, AND THE STERILE WATER LEAK LOCATION WAS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01306 | 2023-04-13 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D |  | IT WAS REPORTED BY THE CUSTOMER THAT ¿WHILE INSERTING PICC-AFTER ONE PASS AND MEETING RESISTANCE, 3CG WIRE WAS REMOVED FROM PATIENT AND TIP WAS NOTED TO BE BENT AT A 90-DEGREE ANGLE. NEW 3CG WIRE WAS ACQUIRED AND USED FOR COMPLETION OF PROCEDURE¿ NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02525 | 2023-04-13 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | MAGIC 3 COUDE INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE HYDROPHILIC WAS TOO GRITTY AND PATIENT DID NOT UNDERSTAND HOW TO USE THE INTERMITTENT CATHETER. REPRESENTATIVE WENT OVER USAGE, BUT PATIENT PREFERS PRELUBED CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02524 | 2023-04-13 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS PLACED IN THE PATIENT ON THE 3RD FLOOR WARD. AFTER FOUR DAYS OF PLACEMENT, A NURSE FOUND THAT THE CATHETER HAD FALLEN OUT OF THE PATIENT, AND UPON CHECKING IT, THE BALLOON WAS FOUND TO BE BURST. PER MAIL RECEIVED FROM IBC ON 10APR2023, STATED THAT THE BALLOON RUPTURE WAS NOT OBSERVED TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02530 | 2023-04-13 | 2023-03-24 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE BALLOON HAS BURST. THE CYSTOSCOPY CONFIRMED NO REMNANTS IN THE BODY. AS PER THE ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON 10APR2023, THE EVENT WAS CATHETER FRACTURE. THE EVENT WAS CONFIRMED AFTER DETENTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02531 | 2023-04-13 | 2023-02-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE PATIENT GOT CATHETERIZATION AFTER THE OPERATION. ON (B)(6), IT WAS FOUND THAT THE CATHETER FELL OUT FROM THE PATIENT AND THEN A NEW CATHETER WAS USED. IT WAS MENTIONED THAT THE CATHETER PROLAPSED AND THE BALLOON WAS RUPTURED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6)2023, IT WAS CONFIRMED THAT THERE WAS NO MISSING PIECES OF THE BALLOON INSIDE THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02535 | 2023-04-13 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THERE WAS A CRACK AT THE TIP OF THE CATHETER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THIS IS A PART OF TSC ASSEMBLY PROCESS AND THE REPORTED EVENT WAS UNCONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER SHAPE, CONFIGURATION AND PRINCIPLES: BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY INTENDED USE & EFFECT- EFFICACY: THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. DIRECTIONS FOR USE: 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDO-ELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. CORRECTION- D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01305 | 2023-04-13 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D2 |  | IT WAS REPORTED BY THE CUSTOMER "PICC OCCLUDED WITHIN SHORT TIME AFTER PLACEMENT. USED ADDITIONAL KIT BECAUSE OF LEAKING LUMENS." NO OTHER INFORMATION WAS PROVIDED. | H11: THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02537 | 2023-04-13 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® NEPHROSTOMY BALLOON DILATION CATHETER | DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 996101 | 996101 |  | IT WAS REPORTED THAT AFTER OPENING THE PRODUCT, WHEN THE CUSTOMER STARTED TO INFLATE THE BALLOON, THE DOCTOR NOTICED THE BALLOON WAS LEAKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01343 | 2023-04-14 | 2023-03-28 | I, n, j, u, r, y | Manufacturer report | BARD DL POWERPICC SOLO | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER DURING CLINICAL ASSESSMENT THAT THERE IS "REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE, DOUBLE LUMEN POWERPICC SOLO IN UPPER ARM." | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01345 | 2023-04-14 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET BASIC TRAY (5F) (DL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4275118 |  | IT WAS REPORTED THE PICC TRAY SCALPEL SHOWING PHYSICAL ABNORMALITIES. NO PATIENT INVOLVEMENT IDENTIFIED BEFORE USING IN PATIENT. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01347 | 2023-04-14 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET BASIC TRAY (4F) (SL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4174118 |  | IT WAS REPORTED THE CATHETER CLAMP CAME ASSEMBLED THE WRONG WAY, NOT PLAYING THE ROLE OF CLAMPING. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01348 | 2023-04-14 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV 3F SL MAX 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1173108D |  | IT WAS REPORTED, 3FR PROVENA PICC, THE MICROINTRODUCER BUCKLED IN ON ITSELF. "I MADE A SMALL NICK, THEN TOOK THE WHITE INTERNAL PIECE OF THE DIALATOR TO PRE-DIALATE BECAUSE I ALWAYS HAVE ISSUES WITH THE 3 FR ONES. I PUT THE TWO PIECES BACK TOGETHER TO DIALATE AGAIN AND IT BUCKLED BACK ON ITSELF." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01351 | 2023-04-14 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD DL POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED "FLUID LEAKING FROM CENTRAL LINE INSERTION SITE." THE EVENT INVOLVED A DOUBLE LUMEN POWERPICC PLACED IN THE UPPER ARM. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01349 | 2023-04-14 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | FULL KIT 18G X 10 CM WITH BIOPATCH AND PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F318108PT |  | IT WAS REPORTED BY THE CUSTOMER "SAFETY NEEDLE WOULD NOT ENGAGE/SAFETY WHEN DEPLOYING CATHETER." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01358 | 2023-04-14 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 6174118 |  | IT WAS REPORTED BY THE CUSTOMER THAT INSERTING PICC IT IS NOT SAFE AS THE GUIDEWIRE CAP IS NOT TIGHT. IN DURING INSERTION PLACER FOUND OUT THAT LEAKAGE WAS COMING FROM THE GUIDEWIRE CAP. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01357 | 2023-04-14 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 6174118 |  | IT WAS REPORTED BY THE CUSTOMER THAT INSERTING PICC IT IS NOT SAFE AS THE GUIDEWIRE CAP IS NOT TIGHT. IN DURING INSERTION PLACER FOUND OUT THAT LEAKAGE WAS COMING FROM THE GUIDEWIRE CAP. IT WAS REPORTED THIS OCCURRED WITH TWO DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01340 | 2023-04-14 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F NURSE BASIC TRAY WITH MICROINTRODUCER (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275118 |  | IT WAS REPORTED THAT THIS PATIENT HAD A CATHETER PLACED FROM THE RIGHT BASILIC VEIN ON (B)(6) 2023. THE CATHETER HAD NO PROBLEM DURING INSERTION AND WAS MAINTAINED PROPERLY (DISINFECTED WITH CHLORHEXIDINE GLUCONATE DISINFECTANT). DURING MAINTENANCE ON (B)(6) 2023, THE PICC CATHETER WING FRACTURED, BUT NO LIQUID LEAKED. TO PREVENT THE RISK OF LIQUID LEAKAGE FROM CATHETER BREAKAGE, THE DEPARTMENT IMMEDIATELY REMOVED THE CATHETER AND INSERTED ANOTHER NEW KIT OF CATHETER. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01356 | 2023-04-14 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1295108D |  | IT WAS REPORTED BY THE CUSTOMER "I HAD AN ISSUE WITH A DOUBLE LUMEN PICC THE OTHER DAY. THE SILICONE PIECE THE STOPPER WITH THE WIRE AT THE END OF THE WIRE PORT WOULD START TO "BALLOON" WHEN FLUSHED. IF FLUSHED HARD ENOUGH, I IMAGINE IT WOULD HAVE RUPTURED. THEY DID LEAVE THE PICC IN AND HAD NO ISSUES. SAID ONLY FELT A LITTLE RESISTANCE ON INSERTION." ADDITIONAL INFORMATION RECEIVED 4/10/2023: IT WAS REPORTED THAT THERE WAS NO PRESENT ISSUES OR ANY HARM THAT WAS EVIDENT WITH THE PATIENT. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01337 | 2023-04-14 | 2023-03-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA CATHETER W/ SHERLOCK 3CG (TPS) STYLET (5F) (TRIPLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1385108D |  | IT WAS REPORTED, LEAKING AT T PORT CONNECTOR UPON FLUSHING. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01326 | 2023-04-14 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT APPROXIMATELY THREE MONTHS POST DIALYSIS CATHETER PLACEMENT, THE EXTENSION PART ALLEGEDLY BROKE. IT WAS FURTHER REPORTED THAT THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01330 | 2023-04-14 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "PATIENT RANG TO SAY PICC LEAKING. PATIENT ATTENDED FOR ASSESSMENT, HE STATED HE HAD FELT SOMETHING AT HIS SHOULDER WHEN THE LINE WAS FLUSHED. LINE REMOVED, HOLE PRESENT AT 22CM (SO INTERNAL)" ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: THERE WAS NO LIFE-THREATENING SITUATION. LEAK WAS DISCOVERED WHEN THE LINE WAS BEING FLUSHED, PATIENT HAD NO BLOOD RETURN FROM PICC AND HAD TAURALOK INSTILLED ON VARIOUS OCCASIONS. THE DAY OF THE PICC REMOVAL HE COMPLAINED OF A STRANGE SENSATION AT HIS CHEST AFTER HIS PICC WAS FLUSHED. PICC FLUSHED POST REMOVAL HOLE PRESENT AT 22CM. PATIENT HAD TO HAVE A REPLACEMENT PICC PLACED FOR HIS LAST CHEMO. PATIENT WAS RECEIVING TWO WEEKLY IMDG WITH CETUXIMAB ¿ IRINOTECAN ,CALCIUM FOLINATE AND 5FU WITH PRE MEDS OF PIRITON. PICC WAS SECURED WITH SECURACATH. NO FURTHER INJURY OR HARM TO PATIENT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02567 | 2023-04-14 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE PATIENTS WAS UNSOUND AFTER A FEW HOURS AND THE BALLOON DID NOT REMAIN INFLATED, EVEN AFTER CHECKING THE BALLOON BEFORE PROBING ON (B)(6) 2023. STATED THAT THERE WAS AN INFECTIOUS RISK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02565 | 2023-04-14 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER CAME OUT IMMEDIATELY ABOUT AN HOUR AFTER BEING INDWELLED. THE FOLEY CATHETER WAS INSERTED BY HAND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02564 | 2023-04-14 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT FOLEY CATHETER WAS INSERTED BEFORE SURGERY AND SPONTANEOUSLY REMOVED AFTER SURGERY AFTER ABOUT 2 HOURS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02561 | 2023-04-14 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE PATIENT GOT CATHETERIZATION FOR CONTINUED DRAINAGE OF URINE AND BLADDER IRRIGATION, WHILE THE FLOW WAS FLUENT AND CLEAR. AT 15:00PM, IT WAS FOUND THAT THE BALLOON WAS BURST AND THE CATHETER FELL OUT. WITH CHECK BY DOCTOR, NO DAMAGE ON URETHRA AND THE CATHETER WAS REMOVED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 30MAR2023, IT WAS NOTED THAT THERE WAS NO MISSING PIECES OF THE BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02560 | 2023-04-14 | 2023-03-24 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE BALLOON HAD BURST. THROUGH CYSTOSCOPY IT WAS CONFIRMED THAT THERE WAS NO REMNANTS IN THE BODY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00252 | 2023-04-14 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE OF A CALCIFIED TARGET LESION IN THE ANTERIOR TIBIAL ARTERY, THE TIP OF THE CATHETER ALLEGEDLY SEPARATED. IT WAS FURTHER REPORTED THAT THE DISTAL TIP OF THE CATHETER STILL ATTACHED TO THE CORE WIRE. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRY DATE: 02/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01325 | 2023-04-14 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HICKMAN CV CATHETER, TRIPLE-LUMEN, 12.5F | CHRONIC CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 0600650 | 0600650 |  | IT WAS REPORTED THAT SOMETIME POST CHRONIC CATHETER PLACEMENT, THERE WAS ALLEGEDLY A TINY HOLE FOUND IN THE ONE OF THE LUMENS. IT WAS FURTHER REPORTED THAT THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 01/2027). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01327 | 2023-04-14 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808560 | 1808560 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY PARTIALLY FRACTURED. REPORTEDLY, THE DEVICE WAS COMPLETELY REMOVED FROM PATIENT . THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. D4 (EXPIRATION DATE: 03/2023). SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01354 | 2023-04-14 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1274108D2 |  | IT WAS REPORTED BY THE CUSTOMER THAT MULTIPLE SAME LOT NUMBERS OCCLUDING WITHIN SHORT TIME AFTER PLACEMENT. USING ADDITIONAL KITS BECAUSE OF LEAKING LUMENS. NO OTHER INFORMATION WAS PROVIDED. THE CUSTOMER REPORTED MULTIPLE EVENTS HOWEVER THE EXACT NUMBER OF EVENTS WAS NOT PROVIDED TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01334 | 2023-04-14 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/ MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER THAT "AT 11:00 ON (B)(6), THE NURSE FOUND THAT THE PICC EXTENSION TUBE WAS LEAKING WHEN THE NURSE WAS CHECKING THE EQUIPMENT FOR PICC CATHETERIZATION AT THE BEDSIDE AND REPLACED IT IMMEDIATELY." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01361 | 2023-04-14 | 2023-03-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1174108D |  | IT WAS REPORTED BY THE CUSTOMER THERE WAS LEAKING AT T PORT CONNECTOR UPON FLUSHING. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01360 | 2023-04-14 | 2023-03-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 8275118 |  | IT WAS REPORTED BY THE CUSTOMER THAT DURING DRESSING CHANGE IT WAS SEEN THAT POWER PICC 5FR DL LUMEN HUB IS BROKEN. PICC STILL IN PATIENT. NO OTHER INFORMATION WAS PROVIDED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00257 | 2023-04-15 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | AT75164 | AT75164 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE VIA UPPER LEFT VEIN, THE PTA BALLOON ALLEGEDLY DID NOT INFLATE COMPLETELY. IT WAS FURTHER REPORTED THAT THE BALLOON WAS ALLEGEDLY CONSTRICTED IN THE MIDDLE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ATLAS PTA DILATATION CATHETER WAS RETURNED FOR EVALUATION. THE RETURNED BALLOON WAS NOTED TO BE A CONSTRICTOR AND CONTAINS UNRAVELED, PEELED PEEBAX AND FRAYED FIBERS, WHICH WERE NOTED ON THE BALLOON UNDER MICROSCOPIC OBSERVATION; NO OTHER SPECIFIC ANOMALIES WERE NOTED DURING THE VISUAL EVALUATION. DURING THE FUNCTIONAL TESTING, THE BALLOON WAS INFLATED WITH THE IN-HOUSE PRESTO INFLATION DEVICE UP TO 6 ATM. UPON INFLATING THE BALLOON, A CONSTRICTION WAS NOTED ON THE BALLOON NEAR THE DISTAL END, THEN THE BALLOON WAS DEFLATED WITHOUT ANY ISSUE. NO OTHER FUNCTIONAL TESTING WAS PERFORMED. ALTHOUGH DURING THE FUNCTIONAL TESTING THE BALLOON WAS ABLE TO INFLATE AND DEFLATE, A CONSTRICTION WAS NOTED AT THE DISTAL END UPON INFLATING THE BALLOON, AND IT WAS ALSO NOTED DURING THE VISUAL EVALUATION; AND THUS THE INVESTIGATION IS CONFIRMED FOR THE REPORTED ABNORMAL INFLATION. PEELED PEEBAX, UNRAVELED AND FRAYED FIBERS WERE ABLE TO BE OBSERVED, HENCE THE INVESTIGATION IS ALSO CONFIRMED FOR THESE IDENTIFIED ANOMALIES. A DEFINITIVE ROOT CAUSE FOR THE REPORTED ABNORMAL INFLATION ISSUE AND IDENTIFIED PEELED PEEBAX, UNRAVELED AND FRAYED FIBERS COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 09/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00254 | 2023-04-15 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 014 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | US14150510 | US14150510 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE SUPERFICIAL FEMORAL ARTERY, THE PTA BALLOON ALLEGEDLY RUPTURED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 06/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00253 | 2023-04-15 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRASCORE 014 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | US141502H30 | US141502H30 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A CALCIFIED TARGET LESION BELOW THE KNEE, THE PTA BALLOON ALLEGEDLY RUPTURED AT 6 ATM. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 08/2024), G3. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01363 | 2023-04-15 | 2023-02-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LOW ARTIFACT POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK LOW ARTIFACT POWER PORT |  | IT WAS REPORTED THAT APPROXIMATELY FIVE YEARS POST PORT PLACEMENT PROCEDURE, THE PORT ALLEGEDLY HAD SPLINT. IT WAS FURTHER REPORTED THAT THE PORT ALLEGEDLY STOPPED WORKING AND A COMPUTED TOMOGRAPHIC EXAMINATION SHOWED THAT THE PORT APPEARED TO HAVE A BLOCKAGE. FURTHERMORE, THE PORT WAS REMOVED AND UPON REMOVAL IT WAS IDENTIFIED THAT THE PORT SEPTUM WAS ALLEGEDLY PARTIALLY DISLODGED FROM THE PORT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01365 | 2023-04-15 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE BROKEN CATHETER WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01366 | 2023-04-15 | 2023-03-08 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY SIX YEARS AND THREE MONTHS POST PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN, A FOLLOW UP COMPUTED TOMOGRAPHIC EXAMINATION WAS PERFORMED AND REVEALED THAT THE CATHETER WAS ALLEGEDLY BROKEN AND THE DISTAL CATHETER SEGMENT WAS MIGRATED INTO THE PULMONARY VEIN. REPORTEDLY, THE PORT BODY AND THE DISTAL CATHETER SEGMENT WERE REMOVED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01364 | 2023-04-15 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT ISP M.R.I., 8FR GROSHONG | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 4808570J |  | IT WAS REPORTED THAT THE APPROXIMATELY FOUR YEARS AND SIX MONTHS POST PORT PLACEMENT VIA THE LEFT SUBCLAVIAN VEIN, AN X-RAY EXAMINATION WAS PERFORMED AND REVEALED THAT THE TOTAL PARENTERAL NUTRITION ALLEGEDLY LEAKED. IT WAS FURTHER REPORTED THAT WHEN THE PORT SYSTEM WAS REMOVED, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN SUBCUTANEOUSLY AT APPROXIMATELY EIGHTEEN CENTIMETERS FROM THE TIP OF THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT ISP M.R.I., 8FR GROSHONG THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWER PORT ISP M.R.I., 8FR GROSHONG ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01368 | 2023-04-17 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK APHERESIS PORT ACCESS SYS |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY EXPERIENCED SKIN SCABBING ON THE PORT SITE. IT WAS FURTHER REPORTED THAT THE PORT SITE ALLEGEDLY WAS NOT FULLY HEALED. REPORTEDLY, THERE WAS NO MEDICAL INTERVENTION OR MEDICATION PRESCRIBED TO TREAT THE SCABBING. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02602 | 2023-04-17 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT AFTER A CATHETER WAS PLACED IN THE EMERGENCY DEPARTMENT AND TRANSFERRED TO THE NEUROSURGERY WARD, THE CATHETER SPONTANEOUSLY FELL OUT IN THE HOSPITAL ROOM. WHEN THE REMOVED BALLOON WAS REFILLED WITH WATER, THE WATER FLOWED INTO THE DRAINAGE LUMEN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01373 | 2023-04-17 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT "THE GUIDE WIRE WAS NOT SMOOTH AND HAD DEFECTS SUCH THAT THE PROXIMAL END OF THE MICRO-INTRODUCER SHEATH WAS TORN WHEN REMOVING THE GUIDE WIRE. BREACHES OF THE INNER VASCULAR WALL WAS CAUSED DURING REMOVAL OF THE SHEATH." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED MICROINTRODUCER IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. THREE PHOTO SAMPLES OF A 4.5 FR MICROEZ MICROINTRDOCUER WERE PROVIDED FOR EVALUATION. THE FIRST AND SECOND IMAGES SHOWED THE MICROINTRODUCER SHEATH WHICH HAS NOT BEEN PEELED. THE DISTAL END OF THE SHEATH REVEALED DAMAGE AND DEFORMATION. THE DEFORMATION APPEARED TO CONSIST OF INWARD BUNCHING OF THE SHEATH TIP. THE THIRD IMAGE SHOWED AN IMPLANTED 4 FR GROSHONG NXT CATHETER SECURED BY A STATLOCK DEVICE. THE CONDITION OF THE GUIDEWIRE WAS NOT SHOWN IN THE IMAGES. BASED ON THE PHOTO SAMPLES PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE INADEQUATE SKIN INCISION, STEEP INSERTION ANGLE, AND DAMAGE PRIOR TO USE. A REVIEW OF THE MANUFACTURING RECORDS DID NOT REVEAL EVIDENCE TO SUGGEST A MANUFACTURING RELATED ROOT CAUSE. PHOTO SAMPLES OF THE GUIDEWIRE WERE NOT PROVIDED AND THE CONDITION REMAINS UNKNOWN. THE PRESENCE OF DEFORMATION ON THE SHEATH TIP IS INDICATIVE OF A DIFFICULT INSERTION; THEREFORE, THE COMPLAINT IS CONFIRMED BUT THE EXACT FACTORS LEADING TO THE ISSUE REMAIN UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01375 | 2023-04-17 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG 4F SINGLE-LUMEN PICC (60CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7717405 |  | IT WAS REPORTED THAT "DOING VASCULAR ACCESS TO THE PATIENT, WHEN PUMPING TO TEST THE PATENCY OF THE CATHETER, THE ANTERIOR SECTION OF THE CATHETER LEAKED AND NO FLUID CAME OUT AT THE OUTLET, CAUSING NO HARM TO THE PATIENT, AND THE CATHETER WAS REPLACED." | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02607 | 2023-04-17 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT NURSE STATED THAT SILICONE FOLEY CATHETER WAS DIFFICULT TO DEFLATE AND HARD TO REMOVE AFTER A PATIENT PROCEDURE. THEY DID NOT SAVE THE SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02606 | 2023-04-17 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT FOLEY CATHETER WAS PLACED IN PATIENT AS ORDERED FOR THE C SECTION. DURING FUNDUS MASSAGE AFTER C SECTION SURGERY, THE FOLEY CATHETER PLACED IN PATIENT CAME OUT AND BALLOON WAS FOUND DEFLATED AND IT APPEARED LIKE THE EXTERIOR LAYER OF THE TUBING WAS SPLIT AND SEPARATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02608 | 2023-04-17 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THESE INTERMITTENT CATHETERS DID NOT HAVE ENOUGH LUBRICANT. THE REPRESENTATIVE OFFERED TO CHANGE THE CATHETER TO SOMETHING ELSE AND THE PATIENT DECLINED. THE PATIENT STATED THAT THEY WANTED TO KEEP ORDERING THE SAME CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR, VISCOMETER FAILURE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01385 | 2023-04-17 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED "FOR THE PATIENT TO DO VASCULAR CHANNEL DREDGING, IN THE PUMPING TEST WHETHER THE CATHETER IS PATENCY, THE ANTERIOR SEGMENT OF THE CATHETER LEAKAGE. NO MORE OUTLET OUTLET. PROLONG PATIENT TREATMENT TIME. REPLACE THE CATHETER." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01391 | 2023-04-17 | 2023-03-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROV 4F DL MAX 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D |  | IT WAS REPORTED BY THE CUSTOMER "LEAKING AT T PORT CONNECTOR UPON FLUSHING." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01386 | 2023-04-17 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE CORD HAS COMPLETELY PULLED OUT OF THE SENSOR EXPOSING WIRING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF THE CORD HAS COMPLETELY PULLED OUT OF THE SENSOR EXPOSING WIRING WAS CONFIRMED. THE USB CABLE IS PULLED FROM THE SHERLOCK SENSOR HOUSINGS EXPOSING BARE WIRE ENDS. THE ROOT CAUSE OF THE REPORTED FAILURE WAS IDENTIFIED AS DEVICE USE. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02597 | 2023-04-17 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWKIT03 | PWKIT03 |  | IT WAS REPORTED THAT THE PATIENT WOKE UP WET BY USING THE PUREWICK FEMALE EXTERNAL CATHETER AND HAD BEEN WEARING A DIAPER. PATIENT ALSO WANTED TO KNOW HOW TO TAKE OFF THE LID AND THE CANNISTER IN PUREWICK ACCESSORIES REPLACEMENT KIT. THE REPRESENTATIVE ADVISED PATIENT TO BEND THE WICK INTO A U-SHAPE AND TO NOT PUSH THE TUBING TOO FAR DOWN INTO THE WICK AND TO TRY TWISTING THE TOP AND PULL IT OFF. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02609 | 2023-04-17 | 2023-03-27 | I, n, j, u, r, y | Manufacturer report | UNKNOWN INTERMITTENT CATHETER | UNKNOWN INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER SAID THE INTERMITTENT CATHETER WAS INFERIOR TO THE PREVIOUSLY USED CATHETER 53814G. THEY HAD DEVELOPED TWO INFECTIONS WITH THE CURRENT PRODUCT. THEY HAD BEEN TOLD THAT THE PREVIOUS PRODUCT WAS NO LONGER AVAILABLE OR BEING MADE. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT PATIENT CURRENTLY HAD A URINARY TRACT INFECTION AND THEY WERE GOING TO THE DOCTOR'S OFFICE TODAY FOR PRESCRIPTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS/ MECHANICAL AND /OR CHEMICAL RESPONSES FROM THE PATIENT". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01372 | 2023-04-17 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT TITANIUM ISP 6CF INT WSP ATT SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8706060 |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR AND SEVEN MONTHS POST PORT PLACEMENT IN THE RIGHT CHEST WALL VIA RIGHT SUBCLAVIAN VEIN, BLOOD RETURN COULD NOT BE ALLEGEDLY CONFIRMED. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY FOUND TO BE RUPTURED UPON X-RAY EXAMINATION. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO AND A VIDEO WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. EXPIRY DATE: 02/2023. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01370 | 2023-04-17 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616070 | 1616070 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE DILATOR WAS REMOVED AND THE CATHETER WAS INSERTED. BLOOD BACKFLOW WAS ALLEGEDLY OCCURRED. IT WAS FURTHER REPORTED THAT BLOOD ALLEGEDLY LEAKED FROM THE VALVE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. DEVICE IDENTIFICATION (EXPIRATION DATE: 01/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01369 | 2023-04-17 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616070 | 1616070 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE DILATOR WAS REMOVED AND THE CATHETER WAS INSERTED. BLOOD BACKFLOW WAS ALLEGEDLY OCCURRED. IT WAS FURTHER REPORTED THAT BLOOD ALLEGEDLY LEAKED FROM THE VALVE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 EXPIRATION DATE: 01/2024. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01367 | 2023-04-17 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616070 | 1616070 |  | IT WAS REPORTED THAT THREE YEARS AND ELEVEN DAYS POST PORT PLACEMENT, UPON CONTRAST COMPUTED TOMOGRAPHIC EXAMINATION, SALINE SOLUTION ALLEGEDLY COULD NOT BE INFUSED. IT WAS FURTHER REPORTED THAT THE BLOOD RETURN WAS NOT ALLEGEDLY CONFIRMED AND THE CATHETER WAS ALLEGEDLY OCCLUDED. REPORTEDLY, THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 08/2022). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
MW5116746 | 2023-04-17 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | SURESTEP FOLEY TRAY SYSTEM | CATHETER, RETENTION TYPE, BALLOON | C.R. BARD, INC. |  |  |  | ATTEMPT TO INFLATE BALLOON ON FOLEY. NOTED BALLOON NOT WORKING. UPON INSPECTION, END PIECE MISSING. NO HARM TO PATIENT. |  | N | Y |  |  | 2023-04
3006260740-2023-01384 | 2023-04-17 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2174108 |  | DURING EVALUATION OF A RETURNED STYLET AND T-LOCK ASSEMBLY SHOWS LEAKING THROUGH RUBBER STOPPER. NO KNOWN IMPACT TO THE CLINICIAN OR PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FAILURE WAS FOUND DURING EVALUATION OF A RETURNED SAMPLE. ADDITIONAL INVESTIGATION IS REQUIRED TO DETERMINE THE CAUSE OF THE FAILURE. RESULTS OF THE INVESTIGATION ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02599 | 2023-04-17 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT WHEN THE PACKAGE WAS OPENED, THE TIP OF THE CATHETER HAD BLACK STAINS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01374 | 2023-04-17 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT "ON (B)(6), PICC PUNCTURE AND CATHETER PLACEMENT WERE PERFORMED ON THE PATIENT'S RIGHT UPPER ARM. AFTER SUCCESSFUL PUNCTURE, THE CATHETER WAS SENT TO THE VASCULAR SHEATH. THE VASCULAR SHEATH WAS SENT SMOOTHLY, BUT THE CATHETER COULD NOT BE SENT NORMALLY. AFTER THE CATHETER SHEATH WAS REMOVED, IT WAS FOUND THAT THERE WERE CREASES AT THE JUNCTION BETWEEN THE NECK OF THE SHEATH OPENING AND THE VASCULAR SHEATH, RESULTING IN THE CATHETER COULD NOT BE SENT NORMALLY." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02593 | 2023-04-17 | 2023-01-25 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE NURSE DISCOVERED THE FOLEY CATHETER SLIPPED OUT OF THE PATIENT AND THE BALLOON WAS DEFLATED. POST INCIDENT ON (B)(6) 2023, THE WARD NURSING OFFICER INFLATED 10 MLS OF AIR TO TEST THE BALLOON OF THE URINARY CATHETER AND THERE WAS NO LEAKAGE FOUND. THE WARD NURSING OFFICER RETESTED AGAIN ON (B)(6) 2023 AND OBSERVED THAT THE BALLOON LOOKED SMALLER COMPARED TO THE PREVIOUS DAY, DESPITE INFLATING 10MLS OF AIR. THE WARD NURSING OFFICER SUSPECTED POSSIBLE INTERNAL LEAKAGE OF THE URINARY CATHETER. THE USER WAS EXPOSED TO HAZARDOUS BLOOD OR BODY FLUIDS AND THE DETAILED EXPOSURE WAS URINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
16750977 | 2023-04-17 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | DURING PICC PLACEMENT RUBBER-LIKE STOPPER DISLODGED FROM EXTENSION. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16751001 | 2023-04-17 | 2023-02-19 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | FOLEY BALLOON | CATHETER, UROLOGICAL (ANTIMICROBIAL) AND ACCESSORIES | C.R. BARD INC. |  |  |  | BALLOON DEFLATED WITHIN ABOUT 30SECONDS. |  | N | Y |  |  | 2023-04
1018233-2023-02570 | 2023-04-17 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI18 | 0165SI18 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS NOT STAYING IN PATIENT BLADDER BUT WAS SLIPPING OUT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED.THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02596 | 2023-04-17 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165PL10 |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PLANNED TO BE PERFORMED FOR THE PATIENT ON (B)(6) 2023. BEFORE INSERTION, THE FOLEY CATHETER BALLOON WAS INJECTED WITH LIQUID FOR TESTING THE LEAKAGE, BUT THE LIQUID COULD NOT BE WITHDRAWN AFTER THAT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02571 | 2023-04-17 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRETEST IN THE OPERATING ROOM, WATER WAS DIFFICULT TO BE REMOVED FROM THE CATHETER BALLOON. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE INCORRECT SETUP. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED NO OBVIOUS DEFECTS. THIS IS OUT OF SPECIFICATION WHICH STATES, INSPECT FOR MISSING OR INCOMPLETE FUNNEL FILL, POOR GLUING, DAMAGE OR SCRATCHES. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ CORRECTION- D, F, H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02568 | 2023-04-17 | 2023-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0166V14S |  | IT WAS REPORTED THAT ON (B)(6) 2023, THE CATHETERIZATION WAS PLANNED TO BE PERFORMED FOR THE PATIENT. BEFORE INSERTION, THE BALLOON WAS PREFILLED WITH LIQUID FOR TESTING FUNCTIONS, BUT BIG RESISTANCE WAS FELT AND THE LIQUID CANNOT BE INJECTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02569 | 2023-04-17 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI24 |  | IT WAS REPORTED THAT THE BALLOON WAS ASYMMETRICAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02572 | 2023-04-17 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRETEST IN THE OPERATING ROOM, INJECTED STERILE WATER WAS DIFFICULT TO BE REMOVED FROM THE CATHETER BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02575 | 2023-04-17 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT SINCE WATER COULD NOT BE INJECTED THROUGH THE INFLATION VALVE, WHEN THE SYRINGE WAS INSERTED DEEPLY INTO THE VALVE, THE SYRINGE DID NOT ENTER AND BROKE OFF . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02574 | 2023-04-17 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT SINCE THE SYRINGE WAS DEFECTIVE, UPON INJECTING WATER USING THE SYRINGE TO INFLATE THE BALLOON, THE WATER LEAKED FROM THE CONNECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02578 | 2023-04-17 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE WATER DIFFICULT TO BE REMOVED FROM THE FOLEY CATHETER BALLOON. THE USER ALSO FELT RESISTANCE WHEN INJECTING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02583 | 2023-04-17 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THERE WAS A CRACK AT THE TIP OF THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02582 | 2023-04-17 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BROKE DURING SELF-EXTRACTION BY PATIENT. APPROXIMATELY 16 CM FROM THE TIP OF THE CATHETER WAS TORN OFF DURING SELF-REMOVAL BY THE PATIENT. THE REST OF THE CATHETER CAME OUT NATURALLY AFTER THAT, AND THERE WAS NO RESIDUAL BODY. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 05APR2023, STATED THAT THERE WAS NO MEDICAL INTERVENTION REQUIRED. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 28APR2023, STATED THAT THE BROKEN PIECES WERE RETAINED IN PATIENT AND CAME OUT AFTER SOMETIME. | THE REPORTED EVENT WAS CONFIRMED AS USE RELATED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE MISHANDLING OF DEVICE LEADING TO THE FAILURE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: F, H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02586 | 2023-04-17 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THERE WAS A CRACK AT THE TIP OF THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02590 | 2023-04-17 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE BULGE OF THE CUFF REMAINED SMALL IN A DONUT SHAPE SO THE WATER IN THE BALLOON COULD NOT BE DRAINED DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02591 | 2023-04-17 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123518A |  | IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL AND WAS UNABLE TO URINATE ON THEIR OWN, SO THE CATHETERIZATION WAS PERFORMED. ON MARCH 23 IN THE MORNING, WHEN THE BLADDER IRRIGATION WAS PERFORMED, IT WAS FOUND THAT THE FOLEY CATHETER FELL OUT AND UPON CHECKING, THE BALLOON WAS BURST. THEN, A NEW CATHETER WAS USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01371 | 2023-04-17 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PRISTINE LONG-TERM HEMODIALYSIS CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5403190 | 5403190 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE DEVICE SHEATH ALLEGEDLY CRACKED. IT WAS FURTHER REPORTED THAT THE DEVICE PEELED AWAY AND THE HANDLE OF THE SHEATH BROKE OFF. REPORTEDLY, THE DEVICE WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 04/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02573 | 2023-04-17 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING PRETEST IN THE OPERATING ROOM, INJECTED STERILE WATER WAS DIFFICULT TO BE REMOVED FROM THE CATHETER BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01399 | 2023-04-18 | 2022-12-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS FOUND THAT THE BLOOD COULD NOT BE RECOVERED DURING THE INFUSION OF LIQUID. AFTER THE INJECTION OF LIQUID, WATER WAS FOUND AT THE CATHETER JOINT. CHECK THE LOOSENING AND SLIP OF THE CATHETER JOINT. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01393 | 2023-04-18 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT DURING PORT DEVICE IMPLANT PROCEDURE, THE PEEL AWAY SHEATH WAS ALLEGEDLY CRACKED PREMATURELY CAUSING HEMORRHAGE. THE PROCEDURE WAS COMPLETED USING THE SAME DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE FULLY PEELED 8.0FR INTRODUCER PEEL-APART SHEATH AND ONE VESSEL DILATOR WERE RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL EVALUATIONS WERE PERFORMED. BOTH SHEATH SHAFTS WERE NOTED TO BE TWISTED WITH BUNCHING THROUGHOUT AND THE VESSEL DILATOR WAS NOTED TO BE BENT. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED PREMATURE SEPARATION ISSUE, AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRY DATE: 04/2024). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02648 | 2023-04-18 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998006 | 998006 |  | IT WAS REPORTED THAT WHEN PRESSURE WAS APPLIED WITH AN INFLATION DEVICE, AN EXPLOSION-LIKE SOUND WAS HEARD, THE METER PART FLEW TO THE ASSISTANT NURSE'S EYES. PATIENT STATED THAT WHEN HE CHECKED THE FLYING OBJECT, IT WAS THE METER PART, AND THE MANOMETER PART WAS DAMAGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01400 | 2023-04-18 | 2023-03-24 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE PATIENT HAD A SENSATION OF SWELLING AT THE INSERTION SITE DURING MAINTENANCE OF THE CATHETER. THE DEPARTMENT HAS REMOVED THE CATHETER FROM THE PATIENT. THIS PATIENT UNDERWENT DIAGNOSIS BY ULTRASOUND UPON AWARENESS OF SWELLING AT THE INSERTION SITE, WHICH INDICATED THROMBOSIS SURROUNDING THE BASILIC VEIN AND THE AXILLARY VEIN. ADDITIONAL INFORMATION RECEIVED 04/24/2024: AFTER THROMBOSIS WAS FOUND, THE HEAD NURSE IMMEDIATELY EXTUBED THE TUBE AND GAVE THE PATIENT ANTI-INFLAMMATORY AND ANALGESIC TREATMENT. THE PATIENT WAS DISCHARGED FROM HOSPITAL WITHIN A WEEK. PATIENT EXPERIENCED PAIN AND STIFFNESS AT THE PUNCTURE POINT ON THE RIGHT ARM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER CAUSING SWELLING IN THE ARM NEAR THE INSERTION SITE WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR S/L GROSHONG CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED BETWEEN THE 37 CM AND 38 CM DEPTH MARKINGS. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED. FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR. 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER. OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING). AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01401 | 2023-04-18 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO FT CATHETER WITH SHERLOCK 3CG TPS STYLET 5F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1295108FD |  | IT WAS REPORTED "NOT PICKING UP ECG WAVE FORMS, P WAVE SPECIFICALLY. CANNOT TELL IF IT'S THE SHERLOCK SENSOR OR THE PICC. AN ECHO WAS TAKEN DUE TO PATIENT SYMPTOMS AND SHOWED THE PICC TIP WAS IN THE RIGHT VENTRICLE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF THE INTERNAL WIRE LEADS WOULD NOT GIVE READINGS/HOOK UP PROPERLY TO THE SHERLOCK DEVICE IS INCONCLUSIVE DUE TO THE NATURE OF THE RETURNED SAMPLE. ONE 3CG TLS STYLET WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED THAT THE RETURNED 3CG TLS STYLET APPEARED INTACT, AND THE LEADS ASSEMBLY WAS NOT RETURNED FOR EVALUATION. NO SHARP BENDS WERE OBSERVED ALONG THE STYLET. THE PRODUCT WAS FUNCTIONALLY TESTED BY ATTACHING THE STYLET TETHER TO A NON-COMPLAINANT LEADS ASSEMBLY AND SHERLOCK SENSOR TOP PLATE. THE ATTACHMENT WAS UNREMARKABLE. A DIGITAL MULTIMETER TESTER DID NOT REVEAL ANY DISCONTINUITIES WITH THE COMPLAINANT STYLET. A STEADY RESISTANCE READ AT ABOUT 3.3 OHMS. NO DEFICIENCIES WERE DISCOVERED DURING EVALUATION OF THE RETURNED SAMPLE; HOWEVER, THE POTENTIALLY IMPLICATED LEADS ASSEMBLY WAS NOT RETURNED FOR EVALUATION. CONSEQUENTLY, THIS COMPLAINT IS INCONCLUSIVE AT THIS TIME. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01402 | 2023-04-18 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT W/BIOPATCH AND PROBE COVER | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108BDP |  | IT WAS REPORTED BY THE CUSTOMER THAT "SHE PLACED A MIDLINE ON (B)(6) AND IT STARTED LEAKING AND NEEDED REPLACEMENT ON (B)(6). NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01406 | 2023-04-18 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY W/MICROINTRODUCER 55CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE GUIDEWIRE CAP IS LEAKING. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: PATIENT GOT ANOTHER PICC. THIS INCIDENT HAPPENED DURING INSERTION AND WAS NOT SUCCEEDING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01405 | 2023-04-18 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108D |  | IT WAS REPORTED BY THE CUSTOMER THAT "RN'S HAVE TROUBLE WITH IT KINKING IN SOME CIRCUMSTANCES [WITH THE 3.5FR SHEATH IN THE PICC KIT]. " A SPECIFIC NUMBER OF AFFECTED DEVICES WAS NOT PROVIDED BY THE COMPLAINANT. | H11: THE INITIAL COMPLAINT APPEARED TO BE A REPORTABLE OCCURRENCE. UPON RECEIVING ADDITIONAL INFORMATION ABOUT THE EVENT, IT WAS DETERMINED THAT A REPORTABLE EVENT HAD NOT OCCURRED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01404 | 2023-04-18 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PROVENA MIDLINE 3F SL | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000879 |  | IT WAS REPORTED BY THE CUSTOMER THAT "A COUPLE OF WEEKS AGO I DID A 3 FR AND COULDN'T ADVANCE THE INTRODUCER BECAUSE A LIP FORMED ON THE END AND WOULDN'T PENETRATE THE VEIN." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01407 | 2023-04-18 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE GUIDE WIRE COULD NOT ENTER THE INTRODUCER NEEDLE, LEADING TO REBOUNDING OF THE GUIDE WIRE. THE PATIENT¿S BLOOD WAS SPLASHED ONTO THE FACE OF THE OPERATOR. ADDITIONAL INFORMATION RECEIVED 4/10/2023: IT WAS REPORTED BY THE CUSTOMER "BLOOD MADE CONTACT WITH MUCOUS MEMBRANES WHEN IT SPLASHED IN THE CLINICIAN'S FACE. IT WAS A LITTLE BLOOD SCATTERED ACROSS THE FACE AND EYES OF THE NURSE. THE PATIENT DOES NOT HAVE ANY BLOOD-BORNE DISEASES AND THE NURSE IS NOT RECEIVING PROPHYLACTIC OR THERAPEUTIC TREATMENT." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DIFFICULTY PASSING A GUIDEWIRE THROUGH AN INTRODUCER NEEDLE WAS CONFIRMED AND APPEARS TO BE SUPPLIER RELATED. ONE 21 GA INTRODUCER NEEDLE AND GUIDEWIRE W/HOOP WERE RETURNED FOR EVALUATION. BLOOD RESIDUE AND EVIDENCE OF USE WAS NOTED ON THE DEVICE. AN ATTEMPT TO ADVANCE THE GUIDEWIRE THROUGH THE NEEDLE WAS UNSUCCESSFUL DUE TO AN OBSTRUCTION EXPERIENCED WITHIN THE LUER HUB/NEEDLE ORIFICE TRANSITION. MICROSCOPIC INSPECTION OF THE LUER HUB REVEALED AN OFF-CENTER TRANSITION WITH THE NEEDLE ORIFICE. A CLEAR RESIDUE WAS OBSERVED TO BE OUTLINING THE NEEDLE TRANSITION. THE MATERIAL APPEARED TO BE ADHESIVE USED IN THE MANUFACTURING PROCESS; THEREFORE, THE COMPLAINT IS CONFIRMED. BD IS WORKING CLOSELY WITH THE SUPPLIER FACILITY TO PREVENT REOCCURRENCE OF THE REPORTED EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01408 | 2023-04-18 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 5FR 55 CM DL FT 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | BD VASCULAR ACCESS DEVICES FIELD ASSURANCE RECEIVED A SHEATH FOR EVALUATION THAT HAD A SPLIT AT THE TIP. NO OTHER INFORMATION WAS PROVIDED BY THE CUSTOMER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF AN INTRODUCER NOT PENETRATING THE SKIN WITHOUT BREAKING IS CONFIRMED AND WAS DETERMINED TO BE MANUFACTURING RELATED. THREE 5.0 FR MICROINTRODUCERS WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED BLOOD USE RESIDUES ON ALL THREE OF THE RETURNED DEVICES. SAMPLE ONE HAD A KINK TOWARDS THE DISTAL END OF THE GRAY SHEATH, SAMPLE TWO HAD TWO KINKS ALONG THE GRAY MICROINTRODUCER SHEATH, AND SAMPLE THREE HAD A SLIGHT BEND THROUGHOUT THE ENTIRE DEVICE. A MICROSCOPIC OBSERVATION REVEALED SAMPLE ONE TO HAVE A LONGITUDINALLY ALIGNED SPLIT ON THE DISTAL END OF THE GRAY SHEATH THAT APPEARS TO BE MANUFACTURING RELATED. SAMPLES TWO AND THREE HAD NO SPLITS IN THE GRAY SHEATHS AND NO OBVIOUS DEFORMATION TO THE DISTAL ENDS OF THE DEVICES. SAMPLES TWO AND THREE HAD SMOOTH TRANSITIONS FROM THE DISTAL END OF THE GRAY SHEATH TO THE INTRODUCER. BASED ON THE OBSERVATIONS MADE FOR EACH MICROINTRODUCER, ONLY SAMPLE ONE WAS RELATED TO A MANUFACTURING DEFECT. THE COMPLAINT OF AN INTRODUCER NOT PENETRATING THE SKIN WITHOUT BREAKING IS CONFIRMED AND WAS DETERMINED TO BE MANUFACTURING RELATED. THE INVESTIGATION WAS FORWARDED TO THE MANUFACTURING FACILITY FOR FURTHER EVALUATION, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
MW5116825 | 2023-04-18 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | FOLEY URINARY CATHETER | COLLECTOR, URINE, (AND ACCESSORIES) FOR INDWELLING CATHETER | C. R. BARD, INC. |  |  |  | URINARY CATHETER BALLOON WITH LEAK/COLLAPSE OF BALLOON WHILE ATTEMPTING TO INSERT FOLEY URINARY CATHETER. WHILE PLACING NEW FOLEY CATHETER (18 FR NON LATEX 30ML BALLOON), SN USED 10ML OF STERILE SALINE TO START INFLATING THE BALLOON. THERE WAS A MILD REPORT OF DISCOMFORT; 5ML MORE STERILE SALINE INSTILLED WHEN PULLING BACK NO SALINE RETURN. FOLEY REMOVED AND NOTED THE BALLOON WAS NOT INFLATED AND THERE WERE S SHAPED FLAPS WHERE THE BALLOON WAS SUPPOSED TO BE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01398 | 2023-04-18 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE INTRODUCER WILL NOT SLIDE IN AS EXPECTED, CAUSING TRAUMA TO THE SKIN AND VESSEL BECAUSE OF FORCE NEEDED TO INSERT THE DEVICE. NO ADDITIONAL PATIENT INFORMATION IS AVAILABLE, THERE WERE NO SERIOUS ADVERSE EVENTS, BUT IT DID CAUSE PATIENT DISCOMFORT AND DELAY IN TREATMENT TIME. ADDITIONAL INFORMATION RECEIVED: THERE WAS NO TREATMENT NEEDED FOR THE SKIN OR VESSEL TRAUMA. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DIFFICULT INSERTION IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE 5 FR X 7 CM MICROEZ MICROINTRODUCER WAS RETURNED FOR EVALUATION. A PRODUCT LABEL INDICATING LOT: REGZ0186. DRIED BLOOD RESIDUE AND EVIDENCE OF USE WAS NOTED ON THE DEVICE. THE DILATOR WAS PRESENT WITHIN THE SHEATH WHICH HAD NOT BEEN PEELED. THE DISTAL END OF THE DILATOR WAS OBSERVED TO BE BENT. ADDITIONAL DEFORMATION WAS NOTED AT THE SHEATH TIP. MICROSCOPIC INSPECTION OF THE SHEATH DEFORMATION REVEALED A SECTION OF THE TIP TO BE BUNCHED INWARDS. EVIDENCE OF THE SHEATH TIP TAPER BEING PRESENT WAS OBSERVED. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE ADVANCEMENT AGAINST RESISTANCE, INADEQUATE SKIN NICK, AND SHEATH TIP TRANSITION. A REVIEW OF THE MANUFACTURING RECORDS DID NOT REVEAL EVIDENCE TO SUGGEST A MANUFACTURING RELATED ROOT CAUSE. THE DAMAGE PRESENT ON THE GREY SHEATH IS INDICATIVE OF A DIFFICULT INSERTION; THEREFORE, THE COMPLAINT IS CONFIRMED BUT THE EXACT CONTRIBUTING FACTORS REMAIN UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01403 | 2023-04-18 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY (4F) SL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4174118 |  | IT WAS REPORTED THAT "DURING THE PICC CATHETER INSERTION PROCEDURE, A NURSE FROM THE VASCULAR ACCESS PROGRAM SHOWS RESISTANCE IN THE PASSAGE OF THE MICRO INTRODUCER (PEEL AWAY), HAVING TO CARRY OUT A GREATER DERMOTOMY IN THE PATIENT WITHOUT OBTAINING ADMISSION. WHEN SHE CHECKS, SHE OBSERVES THE GRAY PART A ROUGH DAMAGE, WHICH DOES NOT ALLOW SLIDING THROUGH THE SKIN TO CONTINUE THE PROCEDURE, WHICH REQUIRED OPENING ANOTHER PICC TRAY TO USE THE MICRO INTRODUCER OF THAT TRAY." | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02647 | 2023-04-18 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER RECEIVED A STERILE URINE CATHETER KIT WITH A LARGE BLACK HAIR IN IT. PER FOLLOW UP VIA EMAIL ON 19APR2023, IT WAS REPORTED THAT THE STERILE PACKAGE HAD A HAIR IN IT. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND THAT IT WAS NOT AN UCC PRODUCT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02634 | 2023-04-18 | 2023-03-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE CATHETER BALLOON WAS UNABLE TO DEFLATE. WATER INJECTED INTO THE BALLOON WAS UNABLE TO BE WITHDRAWN EVEN IF NEGATIVE PRESSURE WAS APPLIED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "VALVE DAMAGED, POOR MOLDING,". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "[WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02645 | 2023-04-18 | 2023-03-28 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETER CAUSED THE PATIENT TO GET A URINARY TRACT INFECTION. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 30 DAYS. NO MEDICAL INTERVENTION WERE REPORTED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS: DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS." "PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE." " RECOMMENDATIONS: ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. HANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02615 | 2023-04-18 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT BEFORE USING THE FOLEY CATHETER, INJECTING FIXATION WATER INTO THE CATHETER TO CHECK FOR DAMAGE IN THE FOLEY CATHETER, THERE WAS A PINHOLE IN THE CENTER OF THE FOLEY CATHETER AND DISTILLED WATER LEAKED OUT FROM THE PINHOLE. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿WARNINGS 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] CONTRAINDICATIONS 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02613 | 2023-04-18 | 2023-03-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT SILICONE FOLEY CATHETER BALLOON WAS INFLATED WITH FLUID, AND THE FLUID IMMEDIATELY STARTED TO LEAK FROM THE BALLOON. IT WAS FOUND THAT THE BALLOON HAD A HOLE IN IT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "IMPERFECTION IN BALLOON DUE TO PROCESS". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARD® EZ-LOK® SAMPLING PORT (INDICATED BY THE BLUE STEM IN THE PORT) ACCEPTS LUER LOCK (FIG. 1A) OR SLIP TIP SYRINGES (FIG. 1B). 1. OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. 2. SWAB SURFACE OF BARD® EZ-LOK® SAMPLING PORT WITH ANTISEPTIC WIPE. (FIG. 2) 3. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT (AT APPROXIMATELY 80 - 100 DEGREE ANGLE). PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. (FIG. 3) 4. SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND REMOVE SYRINGE FROM SAMPLE PORT. (FIG. 4) 5. UNKINK TUBING AND TRANSFER URINE SPECIMEN INTO SPECIMEN CUP OR FOLLOW HOSPITAL PROTOCOL. DISCARD SYRINGE ACCORDING TO HOSPITAL PROTOCOL. 6. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02617 | 2023-04-18 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT ON THE NEXT DAY AFTER PLACEMENT, INTRAURETHRAL BALLOON PLACEMENT AND BALLOON INVERSION WAS FOUND. SINCE NO URINE FLOW WAS OBSERVED DURING AND AFTER CATHETER PLACEMENT, A CT SCAN REVEALED THAT A BALLOON HAD BEEN PLACED IN THE URETHRA. IN ADDITION, THE BALLOON HAD BEEN INVERTED IN THE URETHRA AND BLEEDING WAS OBSERVED WHEN THE CATHETER WAS REMOVED. | THE REPORTED EVENT WAS INCONCLUSIVE .BASED ON THE ATTACHED CT SCAN PHOTO, IT WAS OBSERVED THAT FOLEY CATHETER BALLOON DID NOT ENTERS THE BLADDER. HOWEVER, WE UNABLE TO DETERMINE THE CATHETER CONDITION BASED ON ATTACHED PHOTO. THEREFORE, THE REPORTED EVENT WAS INCONCLUSIVE DUE TO POOR SAMPLE CONDITION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE DUE TO USER ISSUE SINCE THE CATHETER/BALLOON WAS PLACED IN THE URETHRA AND NOT THE BLADDER. THE CATHETER NOT BEING PLACED IN THE INTENDED LOCATION FOR PROPER URINE DRAINAGE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: H. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02618 | 2023-04-18 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WATER LEAKED FROM AROUND THE VALVE WHEN INJECTING WATER WITH A SYRINGE TO EXPAND THE BALLOON DURING THE TEST. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE ¿INSERTS WITH EDGES (OPERATOR HITS THE TUBE AGAINST THE BOTTOM INSERT WHEN REMOVING THE PIECE FROM THE MOLD)". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. ALLOW THE BALLOON TO DEFLATE SLOWLY ON ITS OWN. DO NOT ASPIRATE OR MANUALLY ACCELERATE THE DEFLATION OF THE BALLOON. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01397 | 2023-04-18 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK APHERESIS PORT ACCESS SYS |  | IT WAS REPORTED THAT SOMETIMES POST PORT PLACEMENT PROCEDURE, PORT ALLEGEDLY HAD A SMALL HOLE. IT WAS FURTHER REPORTED THAT, THE PATIENT EXPERIENCED SWELLING. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02614 | 2023-04-18 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT INJECTING DISTILLED WATER INTO THE BALLOON, WATER LEAKED FROM A SIDE HOLE. CONFIRMED WITH SALES REP, A WATER LEAK FROM A SIDE HOLE MEANS A WATER LEAK FROM THE FUNNEL BIFURCATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02619 | 2023-04-18 | 2023-03-28 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT PURCHASED A PUREWICK FEMALE EXTERNAL CATHETER IN 2021, THE PATIENT HAD GOTTEN 2 URINARY TRACT INFECTIONS BECAUSE OF USING THIS . THE PATIENT ALSO STATED THAT THEIR TUBING WAS CLOUDY WITH A WHITE FILM INSIDE . REPRESENTATIVE ADVISED THAT THE PRODUCT WAS MEANT TO PREVENT INFECTION. THE PATIENT STATED THAT THEY DID NOT PHYSICALLY SEE THEIR DOCTOR BUT DID SEND IN A URINE SAMPLE DETERMINED THEY DID HAVE THE 2 URINARY TRACT INFECTIONS. REPRESENTATIVE ADVISED PATIENT FOR A 24 HOUR USER THEY SHOULD REPLACE THE KIT EVERY 30 DAYS AS SUGGESTED BY THE MANUFACTURER. THE PATIENT HAD BEEN USING THIS PRODUCT FOR MORE THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE URINARY TRACT INFECTIONS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. WARNINGS: ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: ¿ AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER ¿ HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. ¿ EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. ¿ PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02629 | 2023-04-18 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 52606 |  | IT WAS REPORTED THAT THE PATIENT HAD ANOTHER INCIDENT OF HYDROSIL 6CH PAED 62606 BEING DIFFERENT SIZES WITHIN THE SAME BOX. NEONATE WAS UNABLE TO BE CATHETERISED AND HAS HAD TO GO TO THEATRE THIS MORNING. SHE COULD NOT SAY IF THIS WAS DUE TO THE CATHETER ISSUE. DR. ON DUTY HAD NOTICED SIZE DISCREPANCY. | THE REPORTED EVENT IS UNCONFIRMED AS THE PRODUCT MEETS SPECIFICATIONS. THE ROOT CAUSE OF THE REPORTED ISSUE WAS UNKNOWN. VISUAL INSPECTION NOTED TWO (2) PHOTO SAMPLES. FIRST PHOTO SAMPLE SHOWCASES TWO (2) ISCS PACKAGING WITH THE SAME PCN (62606P), LOT NUMBER (JUDY0063), AND EXPIRATION DATE (2024-10-28). SECOND PHOTO SAMPLE SHOWCASES TWO (2) CATHETERS SIDE BY SIDE. OUTER DIAMETER SPECIFICATION FOR 6PEDIATRIC FR CATHETERS IS 2.01 MM ±0.33 MM. MEASUREMENTS FOR BOTH SAMPLES ARE: 1.85 MM AND 2.16 MM. THEREFORE, PRODUCT MEETS SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿FOR UROLOGICAL USE ONLY. INTENDED FOR USE BY PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICES ARE PASSED TO THE URINARY BLADDER VIA THE URETHRA. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. THIS IS A SINGLE USE DEVICE. DO NOT RE-STERILIZE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. PRIOR TO OPENING THE SEALED CATHETER POUCH, APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. WET THE CATHETER BY HOLDING THE PACKAGE WITH THE PRINTED SIDE UP AND TIP THE PACKAGE END-TO-END THREE TO SIX TIMES TO WET THE CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. PEEL OPEN THE PACK AT THE FUNNEL/HANDLE END JUST ENOUGH TO EXPOSE 5 CM (2¿) OF THE CATHETER, OR FOR PRODUCTS WITH AN INSERTION SLEEVE, EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. USING THE CATHETER WITH INSERTION SLEEVE (GRIPPER): HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. WITHOUT INSERTION SLEEVE (GRIPPER): HOLD THE HANDLE/FUNNEL END AND REMOVE THE CATHETER FROM THE PACKAGING. GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA AND ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02628 | 2023-04-18 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI24 |  | IT WAS REPORTED THAT THE BALLOON WAS ASYMMETRICAL. | THE REPORTED EVENT WAS NOT CONFIRMED BECAUSE THE DEVICE MEETS SPECIFICATIONS. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS NOT-CONFIRMED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE ''UNDER CURED SAC'' (EXAMPLE: POOR SHAPE/ BAGGY SAC-NONCONCENTRIC BALLOON). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "BARDEX® SILVER LUBRICATH® FOLEY CATHETER SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02612 | 2023-04-18 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165L16 | 0165L16 |  | IT WAS REPORTED THAT UPON PLACING A FOLEY CATHETER PRIOR TO A SCHEDULE C-SECTION, CUSTOMER NOTICED A BLUE FLECK WHICH APPEARED TO BE PLASTIC IN THE COLLECTION TUBE WHILE CHECKING FOR URINE FOR CORRECT PLACEMENT. A BLUE PIECE OF SOMETHING SEMI-TRANSPARENT WAS IMMEDIATELY FLOATING IN THE URINE. IT HAD A DARKER BLUE LINE DOWN THE MIDDLE. IT WAS ABOUT THE SIZE OF AN ERASER TIP. THERE WERE NO DIFFICULTIES INSERTING IT, SO IT WAS NOT AN OUTSIDE CONTAMINANT. CUSTOMER TOOK PHOTOS OF THE PACKAGING AND TOOK A PHOTO OF THE PLASTIC LATER THAT DAY WHEN THEY EMPTIED THE CATHETER DURING RECOVERY. THEY GUESSED THAT IT WAS PROBABLY PART OF THE PLASTIC SHEATH THAT THE CATHETER TIP IS COVERED WITH PRIOR TO PLACEMENT. THIS MADE SENSE BUT THEY ARE NOT 100% SURE. EITHER WAY THIS COULD POTENTIALLY HARM A PATIENT OR CLOG A CATHETER OR POTENTIALLY AFFECT TEST RESULTS. IT NEEDS TO BE INVESTIGATED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO NO PLASTIC LAYER INSIDE TOTE BIN, AND BOTTOM OF THE BIN DIRTY. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A REVIEW OF THE DEVICE LABELING HAVE BEEN CONDUCTED FOR INVESTIGATION PURPOSE. 1. OCCLUDE DRAINAGE TUBING A MINIMUM OF 3 INCHES BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE IS VISIBLE UNDER THE ACCESS SITE. 2. SWAB SURFACE OF BARD® EZ-LOK® SAMPLING PORT WITH ANTISEPTIC WIPE. (FIG. 2) 3. USING ASEPTIC TECHNIQUE, POSITION THE SYRINGE IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT (AT APPROXIMATELY 80 - 100 DEGREE ANGLE). PRESS THE SYRINGE FIRMLY AND TWIST GENTLY TO ACCESS THE SAMPLING PORT. (FIG. 3) 4. SLOWLY ASPIRATE URINE SAMPLE INTO SYRINGE AND REMOVE SYRINGE FROM SAMPLE PORT. (FIG. 4) 5. UNKINK TUBING AND TRANSFER URINE SPECIMEN INTO SPECIMEN CUP OR FOLLOW HOSPITAL PROTOCOL. DISCARD SYRINGE ACCORDING TO HOSPITAL PROTOCOL. 6. FOLLOW ESTABLISHED HOSPITAL PROTOCOL FOR SPECIMEN LABELING AND TRANSPORT TO LAB. STERILE: CONTENTS OF INNER WRAP ARE STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RE-STERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01396 | 2023-04-18 | 2023-03-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | APHERESIS PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | A710962 | A710962 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED REDNESS, SWELLING AND A SMALL SKIN TEAR. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING ADDITIONAL INTERVENTION OR MEDICATION TO TREAT REDNESS, SWELLING AND A SMALL SKIN TEAR. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL CONDITIONS CANNOT BE REPLICATED TO CONFIRM THESE FAILURES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02637 | 2023-04-18 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS UNABLE TO DEFLATE. THE BALLOON WAS ABLE TO INFLATE BUT NOT DEFLATE. EVEN IF THE BALLOON CAN BE FILLED WITH WATER, IT CANNOT BE DEFLATED. THE SAME EVENT HAS OCCURRED BEFORE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02639 | 2023-04-18 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | INTS14 | INTS14 |  | IT WAS REPORTED THAT NURSE ATTEMPTED TO STRAIGHT CATHETER A PATIENT ON (B)(6) 2023, THE CATHETER DID NOT HAVE ANY OPENINGS AT THE END OF THE TUBE FOR FLUID TO FLOW. | THE REPORTED EVENT IS CONFIRMED MANUFACTURING RELATED. VISUAL EVALUATION NOTED RECEIVED 1 SURESTEP INTERMITTENT CATHETER WITH NO EYELETS PRESENT. THEREFORE, PRODUCT DOES NOT MEET SPECIFICATIONS WHICH STATES "MISSING COMPONENTS ARE NOT ALLOWED." ALTHOUGH AN EXACT ROOT CAUSE COULD NOT BE DETERMINED A POTENTIAL ROOT CAUSE COULD BE: THERMOCOUPLE OOC: - THERMOCOUPLE FAILURE . ELECTRICAL FAILURE. - INCONSISTENT TEMP/HUM PROFILE. DHR REVIEW DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. LABELLING REVIEW IS NOT REQUIRED AS THE REPORTED EVENT IS CONFIRMED MANUFACTURING RELATED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01395 | 2023-04-18 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | UNK HEMOSTAR |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER WAS ALLEGEDLY LEAKED. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS ONE HEMOSTAR DIALYSIS CATHETER IMPLANTED IN A PATIENT BODY AND BLOOD WAS NOTED TO BE LEAKING FROM THE AREA CONNECTING THE CATHETER TO THE BIFURCATION. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FLUID LEAK ISSUE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02643 | 2023-04-18 | 2023-03-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123618A |  | IT WAS REPORTED THAT THE BALLOON OF THE FOLEY CATHETER CANNOT BE DEFLATED DURING INJECTING. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE)/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATER ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02633 | 2023-04-18 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414 |  | IT WAS REPORTED THAT TEMPERATURES COULD NOT BE MEASURED DURING SURGERY. AFTER THE SURGERY, THE USER CUT THE LEAD WIRE PORTION OF THE FOLEY CATHETER AND MANAGED IT IN THE HOSPITAL WARD. HOWEVER, THEY WERE STILL CONCERNED ABOUT WHY SUCH AN EVENT HAD OCCURRED. THE LEAD WIRE PORTION THAT IS DISCONNECTED IS TUCKED BETWEEN THE COVER PAPER. EXPLAINED THAT IT IS DIFFICULT TO REPRODUCE BECAUSE OF THE SEVERED LEAD WIRE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE POOR SAMPLE CONDITION. INSPECTION ON THE EXTERIOR OF THE SAMPLE AND FOUND THERMISTOR WIRE WAS CUT. UNABLE TO MEASURE THE RESISTANCE ACROSS THE THERMISTOR PLUG DUE TO POOR SAMPLE CONDITION. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE, DUE TO SENSOR WIRE TOO WEAK, THERMISTOR CHIP TOO WEAK, HANDLING DAMAGED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS: · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR. 1. BALLOON CATHETER WITH TEMPERATURE-SENSING. 2. ACCESSORIES: CLOSED DRAINAGE BAG. SYRINGE PREFILLED WITH STERILE WATER; WATER SOLUBLE LUBRICANT, ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION, TWEEZERS, GAUZE PADS, WATERPROOF SHEET, COTTON BALLS , GLOVES, AND STATLOCK® FOLEY. [INTENDED USE & EFFECT- EFFICACY]: THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE]: 1. METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE: 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS]: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS). 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING>: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD. WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD: 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS: 1) MALFUNCTION: - CATHETER KINKING, DAMAGE, RUPTURE, - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE, - OCCLUSION OF CATHETER INNER LUMENS, - ENCRUSTATION, - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE, - DEVICE DAMAGE DUE TO INAPPROPRIATE USE, - FAILURE TO MEASURE TEMPERATURE, AND - IMPROPER TEMPERATURE INDICATION. 2) ADVERSE EVENTS: - URINARY-TRACT INFECTION, - HEMORRHAGE, HEMATURIA, - ALLERGY REACTION TO THE DEVICE, - CALCULUS FORMATION, - EDEMA, - PAIN, - DISCOMFORT, - INJURY OF BLADDER OR URETHRAL, - URETHRITIS, URINARY INCONTINENCE, AND - RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02632 | 2023-04-18 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE USER FELT DIFFICULTY IN INFLATING THE BALLOON OF THE FOLEY CATHETER. USER WAS NOT PRE-TESTING THE BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO MANUFACTURING RELATED DUE TO COLLAPSED / PINCHED INFLATION LUMEN UNDER THE BALLOON OR ALONG THE SHAFT. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY WAS UNKNOWN, THE FOLEY CATHETER IFUS WERE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. CORRECTION: A H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01437 | 2023-04-19 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT W/BIOPATCH AND PROBE COVER | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108BDP |  | IT WAS REPORTED BY CUSTOMER THAT A MIDLINE WAS PLACED ON 2/27 AND IT WAS LEAKING ON 2/29 SO IT WAS REPLACED. AFTER PULLING THE LINE, IT WAS FLUSHED AND SEEMED TO BE WORKING WITH NO LEAKING THROUGHOUT THE LINE. THE CATHETER HAD BEEN SECURED WITH STATLOCK. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01450 | 2023-04-19 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE NURSING STAFF INTENDS TO USE A CENTRAL VENOUS CATHETER INSERTED THROUGH THE PERIPHERY TO INFUSE CHEMOTHERAPY DRUGS INTO THE PATIENT, AND INSPECTION REVEALS THAT THE CATHETER IS LEAKING. ANOTHER CATHETER WAS REPLACED FOR USE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01409 | 2023-04-19 | 2023-03-22 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, UPON ACCESSING, THE PATIENT ALLEGEDLY EXPERIENCED BLEEDING AT THE INSERTION SITE. IT WAS FURTHER REPORTED THAT ICE PACKS, PRESSURE DRESSINGS, AND HEMOSTATIC AGENTS WERE USED, BUT THE SITE STILL SEEMED TO HAVE BLEEDING. REPORTEDLY, THE PATIENT HAS BEEN PROVIDED WITH ANTIPLATELET AND BLOOD THINNING MEDICATION. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 06/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01415 | 2023-04-19 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616000 | 1616000 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PORT ALLEGEDLY FAILED TO ASPIRATE BLOOD. IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED PAIN AT THE PORT SITE AND UP THE RIGHT SIDE OF HER NECK. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: (B)(6) 2022). SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01451 | 2023-04-19 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405J |  | IT WAS REPORTED BY THE CUSTOMER THAT AFTER THE CATHETER WAS IMPLANTED, THE CATHETER WAS FLUSHED USING A 10CC SYRINGE AND INFUSION WAS PERFORMED. SIX HOURS AFTER THE CATHETER WAS FLUSHED AND LOCKED WITH HEPARIN SALINE, THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL, AND WHEN THE CATHETER WAS FLUSHED AT THE NEW HOSPITAL, SALINE SOLUTION LEAKED NEAR THE CATHETER CONNECTOR. THEREFORE, WHEN THE PATIENT WAS RETURNED TO THE FACILITY AND THE CATHETER WAS CHECKED AGAIN, LEAKAGE WAS CONFIRMED. THE DEVICE WAS REMOVED FROM THE PATIENT¿S BODY. WHEN FLUSHING THE CATHETER OUTSIDE THE PATIENT'S BODY, NO LEAKAGE WAS OBSERVED. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01440 | 2023-04-19 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D |  | IT WAS REPORTED THAT THE INTRODUCER TIP BENT AND CRACKED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED MICROINTRODUCER SHEATH IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE 4.5 FR X 7CM MICROEZ MICROINTRODUCER, 4 FR SL POWERPICC W/3CG STYLET, AND ACCESSORY COMPONENTS WERE RETURNED FOR EVALUATION. THE DILATOR WAS RETURNED SEPARATE FROM THE SHEATH. BLOOD RESIDUE AND EVIDENCE OF USE WAS NOTED. MICROSCOPIC INSPECTION OF THE DISTAL END OF THE SHEATH REVEALED INWARD BUNCHING AND DEFORMATION. THE SHEATH ORIFICE ALSO APPEARED TO BE COMPRESSED AND CONTAINED AN OVAL SHAPE. THE DISTAL END OF THE DILATOR ALSO APPEARED TO HAVE BEEN CUT AS THE TAPERED REGION WAS NOT PRESENT AND THE DISTAL END OF THE SHAFT WAS COMPRESSED. BASED ON THE INFORMATION PROVIDED AND CONDITION OF THE RETURNED SAMPLE, POSSIBLE CONTRIBUTING FACTORS INCLUDE ADVANCEMENT AGAINST RESISTANCE, STEEP INSERTION ANGLE, AND INADEQUATE SKIN NICK. SINCE DEFORMATION AND EVIDENCE OF A DIFFICULT INSERTION WAS NOTED, THE COMPLAINT IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01453 | 2023-04-19 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE CATHETER WAS PLACED UNDER ULTRASONIC GUIDANCE, AND LEAKAGE OF THE CATHETER WAS FOUND WHEN THE SALINE WAS DELIVERED TO THE CATHETER (NOT FOUND WHEN THE CATHETER WAS PRE-FILLED DUE TO RELATIVELY SMALL TRACHOMA). NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02681 | 2023-04-19 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 1758SI14 | 1758SI14 |  | IT WAS REPORTED THAT THE BALLOONS WERE NOT INFLATING EVENLY SO THAT THE CATHETER WAS LOPSIDED AND PRESSES INTO THE BLADDER WALL THUS NOT DRAINING PROPERLY. PATIENT HAD THEM INFLATE IT FIRST TO TEST THE BALLOONS BEFORE USAGE. PATIENT HAD KEPT 6-8 OCCURRENCE IN THE LAST YEAR WITH THE PACKAGES SO THERE MAY BE MULTIPLE LOTS BUT AS PATIENT WAS IN A WHEELCHAIR, PATIENT DOES NOT WANT TO OPEN THEM UNTIL HE HAS TO SHIP THEM. AS PER INFORMATION RECEIVED VIA EMAIL ON 05APR2023, IT WAS STATED THAT 3 FOLEY CATHETER WAS RECEIVED UNDER LOT# NGGT3581, 2 FOLEY CATHETER RECEIVED UNDER THE LOT#NGCZ051, ONE FOLEY CATHETER RECEIVED UNDER THE LOT#NGCW0564. (B)(4). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01433 | 2023-04-19 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER WITH SHERLOCK 3CG (TPS) (5F) (DL) (55CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 92655517 |  | CUSTOMER REPORTED: ON THE MORNING OF (B)(6), UNIT STAFF ASKED US TO ASSESS WHAT APPEARED TO BE A LEAK IN A DUAL LUMEN GROSHONG PICC. THEY HAD CLAMPED THE LINE AS SOON AS THEY DISCOVERED IT. ON INVESTIGATION/FLUSHING, IT WAS IN FACT LEAKING AT THE CONNECTION OF THE WHITE LUMEN BELOW THE INJECTION CAP. THE PICC LINE WAS REMOVED AND A NEW PICC LINE INSERTED. THERE WAS NO HARM. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01434 | 2023-04-19 | 2023-03-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER WITH SHERLOCK 3CG (TPS) (5F) (DL) (55CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 92655517 |  | THE FOLLOWING WAS REPORTED: PATIENT HAS RIGHT ARM PICC DUAL LUMEN 5 FRENCH BARD GROSHONG WAS ADMINISTERING IV MEDICATION VIA WHITE PORT, POST MEDICATION WAS FLUSHING WITH STERILE 10 CC PREFILLED NORMAL SALINE FLUSH, HEARD A POPPING NOISE, AND LINE SNAPPED CONNECTED TO MY FLUSH, WHITE PORT WAS CLAMPED AND WRAPPED IN GAUZE AND TAPE NOT FOR USE, LINE REQUIRED REPLACEMENT. NO APPARENT HARM - JUST PATIENT INCONVENIENCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A SNAPPED CONNECTION TO THE FLUSH WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 5 FR D/L GROSHONG CATHETER. THE DAMAGE OBSERVED IN THE RETURNED SAMPLE WAS CHARACTERISTIC OF OVER-PRESSURIZATION (BURST) DAMAGE. THIS CAN OCCUR THROUGH THE USE OF SYRINGES SMALLER THAN 10ML, BY FLUSHING AGAINST AN OCCLUSION, OR DUE TO EXCESSIVE FORCE APPLIED DURING INFUSION PROCEDURES. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED ON THE WHITE LUMEN SIDE OF THE EXTENSION TUBING. THIS CATHETER DAMAGE WAS TYPICAL OF A BURST, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: 'S' SHAPED SPLIT. GRANULAR FRACTURE SURFACE TEXTURE (TYPICAL OF TEARING FAILURE MODES). AN OCCLUSION WAS OBSERVED DISTAL TO THE BURST SITE, AND THIS MAY HAVE CONTRIBUTED TO THE FAILURE. FORCEFUL FLUSHING AGAINST AN OCCLUSION CAN CAUSE THIS TYPE OF FAILURE. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01416 | 2023-04-19 | 2023-02-08 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP TI 8F CHRONWSH | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8708060 |  | IT WAS REPORTED THAT THREE WEEKS POST PORT PLACEMENT, EXTRAVASATION WAS CONFIRMED IN THE PORTOGRAM. IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED SHOOTING PAIN INTO HIS SHOULDER WHILE FLUSHING. REPORTEDLY, THE DEVICE HAD EXPIRED BEFORE THE PROCEDURE, AND THE DEVICE HAS BEEN REMOVED SURGICALLY. THE CURRENT PATIENT STATUS IS UNKNOWN. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED IMPROPER OR INCORRECT PROCEDURE OR METHOD USED ISSUE SINCE THE LOT EXPIRATION DATE WAS BEFORE THE SAMPLE IMPLANTATION DATE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED EXTRAVASATION AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRY DATE: 08/2022). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02677 | 2023-04-19 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE PIECE TO HOLD THE INTERMITTENT CATHETER UP ON THE OUTSIDE OF PACKAGING WAS NOT CUT OUT. THE PATIENT STATED THAT IT COULD BE IRRITATING. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿INADEQUATE PRESSURE¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. A DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02657 | 2023-04-19 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER CAUSED BLADDER INFECTION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR BLADDER INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: " THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02652 | 2023-04-19 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER TRAY | C.R. BARD, INC. (COVINGTON) -1018233 | INTS14 | INTS14 |  | IT WAS REPORTED THAT CUSTOMER ATTEMPTED TO STRAIGHT CATHETER A PATIENT ON (B)(6), THE INTERMITTENT CATHETER DID NOT HAVE ANY OPENINGS AT THE END OF THE TUBE FOR FLUID TO FLOW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02659 | 2023-04-19 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT ABNORMAL TEMPERATURE WAS DISPLAYED. DURING A SURGICAL OPERATION, AN ABNORMAL TEMPERATURE OF 39°C OR HIGHER WAS DISPLAYED WHEN AN ELECTROCAUTERY SCALPEL WAS USED WITH THE TSC LEAD WIRE AND CABLE CONNECTION CLOSE TO THE COUNTER ELECTRODE PLATE ATTACHED TO THE THIGH. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR DIAGNOSTIC AND TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "SENSOR WIRE TOO WEAK". THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. CAUTION: AS WITH ALL TEMPERATURE PROBES, IN THE PRESENCE OF RF ENERGY SOURCES, LOCAL HEATING, TEMPERATURE ERRORS, AND PROBE DAMAGE MAY OCCUR. DO NOT STRETCH CATHETER. THIS WILL CAUSE REPOSITIONING OF PROBE. DO NOT USE STYLET. THIS WILL CAUSE STRETCHING OF CATHETER." SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01441 | 2023-04-19 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN EVALUATING THE FUNCTION OF THE PICC CATHETER PRIOR TO PLACEMENT, FOREIGN MATTERS APPEARING HAIR WERE FOUND IN THE PACKAGE OF THE PRODUCT (0668945). | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01443 | 2023-04-19 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY (70CM NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174355 |  | IT WAS REPORTED A LEAK WAS OBSERVED AT THE STATLOCK LEVEL BUT THE PAD UNDER THE ONE-WAY VALVE WAS NOT WET. PICC LINE REMOVED BECAUSE THERE WAS PROBABLY A LEAK ON THE PICC LINE AND THE PAD (UNDER THE PICC LINE) WAS NOT WET. THE PATIENT DOES NOT RECEIVE THE FULL PRESCRIBED DOSE OF TREATMENT BECAUSE OF THE LEAKAGE. THE PATIENT DID NOT SUFFER ANY CLINICAL CONSEQUENCES FROM THIS EVENT BECAUSE THE SERVICE NOTICED IT QUICKLY AND THE TREATMENT GIVEN AT THAT TIME WAS A NUTRIENT TREATMENT TO HYDRATE THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01446 | 2023-04-19 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD SL POWERPICC SOLO | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT: THERE WAS FLUID LEAKING FROM THE CENTRAL LINE INSERTION SITE. EVENT INVOLVED A SL POWERPICC SOLO PLACED IN THE UPPER ARM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01445 | 2023-04-19 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118J |  | IT WAS REPORTED THAT WHEN THE STYLET WAS PULLED BACK WHEN USING IT, THE STYLET WAS EXTENDED. THERE WAS NO REPORTED PATIENT INVOLVEMENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED STYLET IS CONFIRMED AND WAS DETERMINED TO BE USE-RELATED. ONE STIFFENING STYLET THREADED THROUGH A T-LOCK WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED THE CORE WIRE OF THE STYLET WAS BROKEN, AND THE COILED WIRE WAS UNRAVELED PROXIMAL TO THE T-LOCK. THE PORTION OF THE COILED WIRE DISTAL TO THE T-LOCK WAS OBSERVED TO REMAIN TIGHTLY COILED. A MICROSCOPIC OBSERVATION SHOWED THE COILED WIRE ENDED ABRUPTLY AND THE WELD TIP OF THE STYLET WAS MISSING. THE FRACTURE SURFACES OF BOTH THE CORE WIRE AND THE COILED WIRE WERE OBSERVED TO BE ANGLED AND LUSTROUS WITH LINEAR STRIATIONS LEADING UP TO THE PEAK OF THE BREAKS, INDICATING THE STYLET WAS CUT WITH A SHARP INSTRUMENT SUCH AS SCISSORS OR A SCALPEL. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01447 | 2023-04-19 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT W/BIOPATCH AND PROBE COVER | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108BDP |  | IT WAS REPORTED BY CUSTOMER THAT WHEN PLACED A MIDLINE ON 2/9 AND IT STARTED LEAKING AND NEEDED REPLACEMENT ON 2/17. THE CATHETER WAS REMOVED AND REPLACED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01438 | 2023-04-19 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PER-Q-CATH PLUS PICC 3F WITH STIFFENING STYLET WIRE SL (65 CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 4133105 |  | IT WAS REPORTED "THE MATERIAL PRESENTED AN ISSUE DURING ITS PLACEMENT. ACCORDING TO THE REPORTER, TO INSERT THE CATHETER INTO PATIENT IT WAS PERFORMED THE MANDREL REMOVAL AND TESTED WITH SALINE SOLUTION, WHEN THE SOLUTION WAS INFUSED A LEAKAGE WAS NOTED IN THE CATHETER. WHEN THE DEVICE WAS INSPECTED, A CRACK WAS OBSERVED. IT WAS NEEDED ANOTHER CATHETER TO COMPLETE THE PROCEDURE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF AN EXTERNAL LEAK IN THE CATHETER WAS CONFIRMED BUT THE EXACT CAUSE IS UNKNOWN. THREE PHOTOGRAPHS OF THE IMPLICATED 3FR S/L PER-Q-CATH PICC WERE RETURNED FOR EVALUATION. A LONGITUDINAL SPLIT WAS OBSERVED IN THE CATHETER SHAFT, JUST DISTAL TO THE MOLDED SUTURE WINGS BETWEEN THE 0 CM AND 1 CM DEPTH MARKERS. IT WAS REPORTED THAT THE DAMAGE WAS NOTED JUST AFTER THE STYLET WAS REMOVED. THE DAMAGE SEEN IN THE PHOTOGRAPHS COULD BE CONSISTENT WITH DAMAGE FROM THE INTERNAL STIFFENING STYLET. HOWEVER, THIS COULD NOT BE CONCLUSIVELY CONFIRMED WITHOUT MICROSCOPIC EXAMINATION AND DIMENSIONAL ANALYSIS OF THE PHYSICAL SAMPLE. POSSIBLE CONTRIBUTING FACTORS FOR STYLET DAMAGE TO THE CATHETER INCLUDE ADVANCEMENT OR RETRACTION OF THE STYLET AGAINST RESISTANCE, THE STYLET NOT BEING FLUSHED PRIOR TO REMOVAL, BENDS OR KINKS IN THE STYLET, OR COMPRESSION OF THE CATHETER TUBING AGAINST THE INLAID STYLET. THE IFU CONTAINS DETAILED INSTRUCTIONS ON HYDRATING THE STYLET AND REMOVING THE T-LOCK ASSEMBLY AND STYLET AS A UNIT. THE IFU CAUTIONS, ¿NEVER USE FORCE TO REMOVE THE STYLET. RESISTANCE CAN DAMAGE THE CATHETER. IF RESISTANCE OR BUNCHING OF THE CATHETER IS OBSERVED, STOP STYLET WITHDRAWAL AND ALLOW THE CATHETER TO RETURN TO NORMAL SHAPE. WITHDRAW BOTH THE CATHETER AND STYLET TOGETHER APPROXIMATELY 2 CM AND REATTEMPT STYLET REMOVAL. REPEAT THIS PROCEDURE UNTIL THE STYLET IS EASILY REMOVED. ONCE THE STYLET IS OUT, ADVANCE THE CATHETER INTO THE DESIRED POSITION.¿ OTHER POSSIBLE CAUSES FOR AN EXTERNAL LEAK COULD INCLUDE OVER-PRESSURIZATION OF THE CATHETER TUBING, SHARP INSTRUMENT DAMAGE, OR TENSILE (PULLING) DAMAGE. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01449 | 2023-04-19 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE CATHETER IS INDWELLING, AND A 10ML EMPTY NEEDLE IS USED TO FLUSH THE CATHETER DURING WEEKLY ROUTINE CATHETER MAINTENANCE. WHEN THE CATHETER IS FLUSHED, THE CATHETER IS RUPTURED, THE CATHETER IS REMOVED, AND THE CATHETER IS REINSERTED. PATIENT EXPERIENCED PAIN. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01436 | 2023-04-19 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0607540CE |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE DEVICE ALLEGEDLY HAD CONTRAST LEAKAGE AT THE CERVICAL LEVEL DURING RADIOGRAPHIC CONTROL OF A VENOUS ACCESS IMPLANTABLE CHAMBER. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE X-PORT MRI IMPLANTABLE PORT ATTACHED TO A CATHETER WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC VISUAL, TACTILE EVALUATION AND FUNCTIONAL TESTING WERE PERFORMED. AN IN-HOUSE SYRINGE WAS ATTACHED TO A SAFETY INFUSION SET AND THE NON-CORING NEEDLE WAS INSERTED INTO THE PORT SEPTUM. THE PORT BODY WAS PATENT TO BOTH INFUSION AND ASPIRATION WITHOUT ISSUE. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FLUID LEAK ISSUE AS THERE WERE NO VISIBLE SPLIT OR BREAKAGE OF CATHETER WAS IDENTIFIED DURING EVALUATION AND NO LEAKS WERE OBSERVED THROUGHOUT THE TESTS. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 04/2021). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01444 | 2023-04-19 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 4FR DL PROVENA MIDLINE MAX BARRIER KIT | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4254108D |  | IT WAS REPORTED BY THE CUSTOMER THAT PATIENT IN THE ICU HAD A PROVENA 4F DL MIDLINE STARTED LEAKING WHERE THE CATHETER MEETS THE BIFURCATION. CONCERNS WERE THE RISK OF INFECTION AND INCOMPLETE DELIVERY OF MEDICATIONS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01442 | 2023-04-19 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 8295118 |  | IT WAS REPORTED THAT WHEN EVALUATING THE FUNCTION OF THE PICC CATHETER PRIOR TO PLACEMENT, FOREIGN MATTERS LOOKING LIKE HAIR WERE FOUND INSIDE THE CATHETER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01432 | 2023-04-19 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED AT 12:10 ON (B)(6) 2023, THE PICC SPECIALIST NURSE OF BREAST SURGERY FOUND THAT THE SUPPORTING GUIDE WIRE COULD NOT PASS THROUGH THE CATHETER SHEATH DURING THE PICC CATHETER IMPLANTATION (BARDER 4F) FOR THE PATIENT, AND THE OBSERVATION FOUND THAT THE INTERNAL APERTURE RATIO OF THE SHEATH WAS SMALLER THAN THE OUTER DIAMETER OF THE GUIDE WIRE. THE CATHETER SHEATH WAS SUCCESSFULLY INSERTED INTO THE BLOOD VESSEL, BUT WHEN THE GUIDE WIRE WAS WITHDRAWN, IT WAS FOUND THAT THE GUIDE WIRE COULD NOT BE REMOVED. THE PATIENT EXPERIENCED PAIN AND DISCOMFORT WHEN REPEATED ATTEMPTS WERE MADE TO REMOVE THE GUIDE WIRE, AND THE EXTERNAL PART OF THE GUIDE WIRE WAS BROKEN. IT WAS CONSIDERED THAT FORCED REMOVAL OF THE GUIDE WIRE MIGHT LEAD TO THE RISK OF BROKEN WIRE IN THE BLOOD VESSEL, WITH SERIOUS CONSEQUENCES. THE PATIENT WAS SENT TO THE OPERATING ROOM OF THE INTERVENTIONAL DEPARTMENT. IMAGING OBSERVATION SHOWED THAT PART OF THE GUIDE WIRE (ABOUT 7CM LONG) WAS STILL IN THE BLOOD VESSEL. THE BROKEN GUIDE WIRE WAS REMOVED UNDER THE GUIDANCE OF DSA. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02672 | 2023-04-19 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT WHEN STERILE WATER WAS INJECTED TO INFLATE THE BALLOON DURING TESTING, THE BALLOON RUPTURED. | THE REPORTED EVENT WAS CONFIRMED HOWEVER THE CAUSE WAS UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. VISUAL INSPECTION NOTED SAC BURST ALONG LUMEN WITHOUT MISSING PIECES. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/EXPOSURE TO PETROLATUM BASED PRODUCTS MECHANICAL FAILURE/OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01431 | 2023-04-19 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F TL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000886 |  | IT WAS REPORTED BY THE CUSTOMER "THE LINE BROKE UNDER DRESSING, IT APPEARS ONE OF THE EXTERNAL LUMENS WAS CUT OR BROKE WHILE LINE WAS DWELLING." ADDITIONAL INFORMATION RECEIVED 4/11/2023: "CATHETER BREAK DISCOVERED AS CATHETER LEAKING WHEN FLUSHED. BREAK WAS EXTERNAL TO THE PATIENT, NO PIECES RETAINED IN PATIENT. EVENT WAS NOT AN URGENT/LIFE THREATING MEDICAL SITUATION. PICC HAD TO BE REPLACED, NO OTHER NEGATIVE IMPACT TO PATIENT." | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00262 | 2023-04-19 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U35130512 | U35130512 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE LEFT SUPERFICIAL FEMORAL ARTERY, THE PTA BALLOON ALLEGEDLY RUPTURED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRY DATE: 05/2025), G3, H6 (DEVICE). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02650 | 2023-04-19 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT TEMPERATURE WAS NOT CORRECT. AS PER MAIL RECEIVED ON 30-MAR-2023 FROM IBC INITIAL REPORTED EVENT WAS WRONG AND THE CORRECT EVENT WAS THE THERMISTOR WIRE HAS BECOME LOOSE A LITTLE. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. IT IS UNKNOWN IF THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED 1 DRAINAGE BAG WITH INLET TUBING ATTACHED TO A 3-WAY TEMPERATURE SENSING CATHETER. VISUAL INSPECTION NOTED THAT THE TEMPERATURE WIRE WAS PROTRUDING OUT FROM THE CATHETER SHAFT. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE CATHETER IS STRETCHED DURING INSERTION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿BARD® SILVER LUBRI-SIL® FOLEY TRAY SINGLE USE ONLY [WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY.] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ CORRECTIONS: D, F, H. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02649 | 2023-04-19 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123618C |  | IT WAS REPORTED THAT THE FOLEY CATHETER CAUSED ADHESION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01414 | 2023-04-19 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NAVARRE OPTI DRAIN | DRAINAGE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NOD8LPT | NOD8LPT |  | IT WAS REPORTED THAT DURING THE PREPARATION OF AN ULTRASOUND-GUIDED PERCUTANEOUS TRANSHEPATIC CHOLANGIOGRAPHY PROCEDURE, THE LENGTH OF THE TROCAR NEEDLE WAS SHORTER THAN THE CATHETER, RESULTING IN THE TIP OF THE NEEDLE NOT BEING EXPOSED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 02/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01413 | 2023-04-19 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NAVARRE OPTI DRAIN | DRAINAGE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NOD8LPT | NOD8LPT |  | IT WAS REPORTED THAT DURING THE PREPARATION OF AN ULTRASOUND-GUIDED PERCUTANEOUS TRANSHEPATIC CHOLANGIOGRAPHY PROCEDURE, THE LENGTH OF THE TROCAR NEEDLE WAS SHORTER THAN THE CATHETER, RESULTING IN THE TIP OF THE NEEDLE NOT BEING EXPOSED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 02/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01412 | 2023-04-19 | 2022-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSPLIT CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5735190 | 5735190 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY MALFUNCTIONED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE INTRODUCER NEEDLE WITH A LOADED J-TIP GUIDEWIRE WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, DIMENSIONAL AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED GUIDEWIRE STRETCHED, FRACTURE AND MATERIAL SEPARATION ISSUES AS A COMPLETE BREAK WAS NOTED TO THE FLAT CORE WIRE GUIDEWIRE AND IT APPEARED TO BE UNCOILED AND STRETCHED AT THE DISTAL PORTION. BASED ON THE MEASUREMENTS RECORDED DURING SAMPLE EVALUATION AND APPLICABLE DRAWINGS, NO MEASUREMENTS RECORDED COULD BE CONFIRMED TO BE OUT OF TOLERANCE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16768324 | 2023-04-19 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDEX LUBRICATH FOLEY CATHETER COUDE TIP | CATHETER, COUDE | C. R. BARD, INC. | 0168L18 | 0168L18 |  | DEFECTIVE INDWELLING FOLEY CATHETERS- PIN-HOLE IN PORT MAKING INFLATION BALLOON NOT POSSIBLE. CATHETERS UNABLE TO BE USED. NOT ALL IN THE BOX ARE DEFECTIVE BUT WE HAVE HAD 4-5 IN THE SAME BRAND AND SIZE (18F AND 20F) DO THE SAME THING. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01411 | 2023-04-19 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NAVARRE OPTI DRAIN | DRAINAGE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NOD8LPT | NOD8LPT |  | IT WAS REPORTED THAT DURING THE PREPARATION OF AN ULTRASOUND-GUIDED PERCUTANEOUS TRANSHEPATIC CHOLANGIOGRAPHY PROCEDURE, THE LENGTH OF THE TROCAR NEEDLE WAS SHORTER THAN THE CATHETER, RESULTING IN THE TIP OF THE NEEDLE NOT BEING EXPOSED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 02/2025). H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01419 | 2023-04-19 | 2023-01-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668935 |  | IT WAS REPORTED THAT UPON UNPACKING PRIOR TO INSERTION IN THIS PATIENT, THE OPERATOR FOUND THAT THE GREY PORTION OF THE SHEATH WAS TOO SOFT TO BE INSERTED INTO THE INSERTION SITE. 03/30/2023 - THE RETURNED DEVICE EXHIBITED A "HORN OF MATERIAL" AT THE DISTAL END PROTRUDING FROM THE TIP. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF THE GRAY SHEATH BEING TOO SOFT IS CONFIRMED AND WAS DUE TO AN UNKNOWN CAUSE. ONE 3.5 FR MICROINTRODUCER GRAY SHEATH WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUES THROUGHOUT THE RETURNED MICROINTRODUCER SHEATH, AND IT WAS OBSERVED TO BE RETURNED WITHOUT THE WHITE INSERTER. DEFORMATION OF THE DISTAL END OF THE GRAY SHEATH WAS VISIBLE UPON AN INITIAL VISUAL OBSERVATION. A MICROSCOPIC OBSERVATION REVEALED THE DISTAL END OF THE GRAY SHEATH HAD A HORN OF MATERIAL PROTRUDING FROM TIP. THE DISTAL END OF THE GRAY SHEATH WAS DEFORMED AROUND THE ENTIRETY OF THE TIP. THE COMPLAINT OF THE GRAY SHEATH BEING TOO SOFT ON THE MICROINTRODUCER IS CONFIRMED; HOWEVER, DUE TO THE DAMAGE OF THE DEVICE A ROOT CAUSE COULD NOT BE ESTABLISHED. POSSIBLE CAUSES MAY BE DUE TO THE MANUFACTURING PROCESS OF THE GRAY SHEATH, INSERTION PROCEDURE OF THE DEVICE, OR OTHER UNKNOWN CIRCUMSTANCES. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02656 | 2023-04-19 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THERE WAS A PINHOLE IN THE FOLEY CATHETER SHAFT. THE STERILE WATER GUSHED FROM THE SHAFT DURING THE PRETEST. UPON CLOSER INSPECTION, THERE WAS A PINHOLE IN THE SHAFT. | THE REPORTED EVENT WAS CONFIRMED HOWEVER THE CAUSE WAS UNKNOWN. SAMPLE WAS EVALUATED AND FOUND TEAR ON SHAFT OF CATHETER THAT ALLOW WATER TO LEAK OUT. THE TEAR LOOKS LIKE IT HAD BEEN EXPOSED TO SHARP OBJECT. THERE WAS SCRATCH WAS OBSERVED ALONG SHAFT NEAR TORN AREA. THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE (EXAMPLE: USER RELATED/OPERATOR ERROR/ROUGH HANDLING OR EXPOSE TO SHARP OBJECT/MECHANICAL FORCE). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02662 | 2023-04-19 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI22 |  | IT WAS REPORTED THAT THE BALLOON WAS ASYMMETRICAL DURING PRETEST. | THE REPORTED EVENT WAS CONFIRMED AS MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE HAS BEEN EVALUATED AND FOUND THAT THERE WAS ABNORMALITY ON THE BALLOON SHAPE OF CATHETER. CATHETER HAS BEEN INFLATED WITH WATER AND AIR AND FOUND THE BALLOON SHAPE ARE OUT OF MANUFACTURING SPEC. DISSECT THE CATHETER AND THERE ARE UNEVEN SAC THICKNESS FROM SAC AREA THAT CAN CAUSED AS PER REPORTED EVENT. THIS COMPLAINT IS CONFIRMED MANUFACTURING RELATED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE NON-UNIFORM FINISH DIP COVERAGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26EVERY 2FR [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTIONS: D. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01439 | 2023-04-19 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LOW ARTIFACT POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK LOW ARTIFACT POWER PORT |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY EXPERIENCED EXCESSIVE BLEEDING WITH PORT INSERTION. HOWEVER, THERE IS NO FURTHER INFORMATION PROVIDED REGARDING ADDITIONAL MEDICAL INTERVENTION OR MEDICATION TO TREAT EXCESSIVE BLEEDING. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED AS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, CLINICAL EVENT REPORTED IN THE COMPLAINT CANNOT BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INITIAL MDR WAS INADVERTENTLY SUBMITTED WITH A G3 DATE OF 03/23/2023. THE CORRECT G3 DATE IS 03/22/2023. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02663 | 2023-04-19 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD TWO URINARY TRACT INFECTION SINCE USING PUREWICK FEMALE EXTERNAL CATHETER AND WAS NOT USING THE PRODUCT AT THE MOMENT. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE". " RECOMMENDATIONS: ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01435 | 2023-04-19 | 2023-01-30 | I, n, j, u, r, y | Manufacturer report | STATLOCK SLIDING POST TRICOT CRESCENT PAD INCLUDES SKIN PREP | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | PIC0220CE |  | IT WAS REPORTED THE PATIENT DEVELOPED SKIN REACTIONS AND SORES WITH STATLOCK USE. PATIENT IS RECEIVING DIALYSIS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE COMPLAINT OF A SKIN REACTION IS INCONCLUSIVE DUE TO UNKNOWN CLINICAL CONDITIONS AND/OR PATIENT SENSITIVES. ONE HUNDRED AND FIFTY PICC PLUS STATLOCK DEVICES WITH TRICOT PADS AND ADJUSTABLE POSTS WERE RETURNED FOR EVALUATION. THREE BOXES OF FIFTY SAMPLES EACH WERE RETURNED; ONE OF WHICH WAS RETURNED OPEN AND THE OTHER TWO WERE RETURNED SEALED. FIFTY SEALED PACKAGES WERE RETURNED IN EACH BOX. A STATISTICAL SAMPLING OF THE ONE HUNDRED AND FIFTY RETURNED LOT SAMPLES WERE OPENED AND EXAMINED, AND NO DAMAGE OR DEFECTS WHICH COULD CONTRIBUTE TO THE REPORTED EVENT WERE FOUND ON ANY OF THE RANDOMLY SELECTED LOT SAMPLES. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE COMPLAINT OF A SKIN REACTION COULD NOT BE CONFIRMED FROM THE RETURNED LOT SAMPLES. POSSIBLE CONTRIBUTING FACTORS OF A SKIN REACTION COULD INCLUDE PATIENT SENSITIVITY TO THE DEVICE MATERIALS OR TO MEDICATIONS OR OTHER CHEMICALS SUCH AS OINTMENTS OR SKIN PREPS USED AROUND THE SITE. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01410 | 2023-04-19 | 2023-03-18 | I, n, j, u, r, y | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS AND FOUR MONTHS POST A PORT PLACEMENT VIA THE RIGHT INTERNAL JUGULAR VEIN FOR CHEMOTHERAPY, WHEN THE PORT SYSTEM WAS REMOVED, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN. IT WAS FURTHER REPORTED THAT THE DISTAL CATHETER SEGMENT REMAINED FROM THE VENA CAVA TO THE RIGHT ATRIUM. REPORTEDLY, SINCE THE CATHETER WAS BROKEN, AND THERE WAS A HIGH POSSIBILITY THAT THE DISTAL CATHETER SEGMENT HAS ADHERED TO THE TISSUE, CONSIDERING THE RISK OF REMOVAL, THE DOCTOR DECIDED THAT THE DISTAL PART OF CATHETER REMAINS INSIDE OF THE PATIENT'S BODY. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE X-PORT IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WERE RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE, AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED AT THE DISTAL END OF THE CATHETER AND THE EDGES OF THE COMPLETE CIRCUMFERENTIAL BREAK CATHETER WERE NOTED TO BE UNEVEN. THE SURFACE WAS NOTED TO BE ROUGH IN ONE REGION AND SMOOTH IN THE OTHER REGION. THE DISTAL CATHETER SEGMENT WAS NOT RETURNED. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND MATERIAL SEPARATION ISSUES. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MIGRATION ISSUE, AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT CANNOT BE VERIFIED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3, H2, H6 (METHOD) H11: H6 (RESULT, CONCLUSION) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
16768326 | 2023-04-19 | 2023-03-15 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® IC COMPLETE CARE® | CATHETER, UROLOGICAL (ANTIMICROBIAL) AND ACCESSORIES | C. R. BARD, INC. | A303316A | A303316A |  | HAIR FOUND INSIDE UNOPENED STERILE PACKAGE. |  | N | Y |  |  | 2023-04
1018233-2023-02658 | 2023-04-19 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165L18 | 0165L18 |  | IT WAS REPORTED THAT THE FOLEY CATHETERS WERE STOPPED UP AND PATIENT HAD TO CHANGE THEM SOONER THAN THEY WERE SUPPOSED TO. CUSTOMER STATED THAT FIRST ONE LASTED FOR THREE WEEKS, THE SECOND ONE LASTED FOR TWO WEEKS AND THEY HAD 3RD ONE IN NOW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02651 | 2023-04-19 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI14 | 1758SI14 |  | IT WAS REPORTED THAT THE BALLOONS WERE NOT INFLATING EVENLY SO THAT THE FOLEY CATHETER WAS LOPSIDED (LOT #NGGT3581) AND WAS PRESSING INTO THE BLADDER WALL THUS NOT DRAINING PROPERLY. PATIENT HAD THEM INFLATE IT FIRST TO TEST THE BALLOONS BEFORE USAGE AND NO PATIENT HARM. PATIENT HAD KEPT 6-8 OCCURRENCES (LOT #NGCZ0518) (LOT #NGCW0564) IN THE LAST YEAR WITH THE PACKAGES SO THERE MAY BE MULTIPLE LOTS BUT AS PATIENT WAS IN A WHEELCHAIR, THEY DID NOT WANT TO OPEN THEM UNTIL THEY HAD TO SHIP THEM. AS PER INFORMATION RECEIVED VIA EMAIL ON 05APR2023, IT WAS STATED THAT 3 FOLEY CATHETER WAS RECEIVED UNDER LOT# NGGT3581, 2 FOLEY CATHETER RECEIVED UNDER THE LOT #NGCZ0518, ONE FOLEY CATHETER RECEIVED UNDER THE LOT #NGCW0564. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE BALLOON MOLDING. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01425 | 2023-04-19 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001029 |  | IT WAS REPORTED BY THE CUSTOMER THAT "INTRODUCER WILL NOT INSERT EASILY INTO VESSEL DURING PICC PROCEDURE. CAUSING MAJOR TRAUMA TO VASCULATURE". ADDITIONAL INFORMATION RECEIVED 04/07/2023: OPENED ANOTHER INTRODUCER KIT AND PLACED THE LINE. THERE WERE NO COMPLICATIONS SECONDARY TO THE VESSEL TRAUMA. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16768361 | 2023-04-19 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | PORT-A-CATH | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | BARD ACCESS SYSTEMS, INC. |  |  |  | RETAINED PORT-A-CATH. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01427 | 2023-04-19 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER BASIC KIT (4F) (SINGLE-LUMEN) (20CM) | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P6154118 |  | IT WAS REPORTED THAT "POWER MIDLINE NOTICED TO BE LEAKING AROUND THE 0CM MARK. WE HAVE REMOVED ANOTHER MIDLINE WITH A VERY SIMILAR STORY. LEAKING ON FLUSHING, NO MEDICATION ENTERING THE MIDLINE. ALL PROBLEM SOLVING COMPLETED; ASPIRATION, FLUSH, DRESSING CHANGE, BUNG CHANGE, PATIENT POSITION CHANGE. ON INSPECTION BY ONE OF MY STAFF, THE MIDLINE APPEARED TO HAVE A HOLE/SPLIT NEAR THE 0CM MARKER. INSERTED (B)(6) 2023. REMOVED (B)(6) 2023 IN WARD. 4FR, 20CM SINGLE LUMEN, POWER MIDLINE." ADDITIONAL INFORMATION RECEIVED: THERE WERE NO ERRONEOUS TEST RESULTS AND NO CHANGE IN THE COURSE OF TREATMENT. ALTERNATE ACCESS WAS REQUIRED. THERE WAS A MEDICATION DELAY IN TREATMENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK IN THE CATHETER IS UNCONFIRMED AS THE FAILURE COULD NOT BE REPRODUCED. ONE 4 FR S/L POWERMIDLINE CATHETER WAS RETURNED FOR EVALUATION. DURING AN INITIAL VISUAL OBSERVATION, NO OBVIOUS DAMAGE WAS FOUND THROUGHOUT THE RETURNED DEVICE. FUNCTIONAL TESTING WAS PERFORMED BY PRESSURIZING WATER INTO THE RETURNED CATHETER USING A 12 ML SYRINGE FILLED WITH WATER. THE FUNCTIONAL TEST DID NOT REVEAL A LEAK IN THE DEVICE, AS TESTING WAS PERFORMED WITH AND WITHOUT THE NEEDLELESS INJECTION CAP. THE COMPLAINT OF A LEAKING CATHETER AT THE 0 CM DEPTH MARKER WAS UNCONFIRMED, AS THE FAILURE WAS NOT ABLE TO BE REPRODUCED IN THE LAB. THE DISTAL END OF THE DEVICE WAS CLOSED OFF AND THE DEVICE WAS PRESSURIZED WITH WATER TO ATTEMPT TO IDENTIFY THE FAILURE; HOWEVER, A LEAK COULD NOT BE FOUND DURING THIS TESTING. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01423 | 2023-04-19 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC SOLO2 5F 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | IT WAS REPORTED BY THE CUSTOMER THAT "INTRODUCER WILL NOT INSERT INTO VASCULATURE" NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DIFFICULTY INSERTING THE MICROINTRODUCER IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE 5 FR X 7 CM MICROEZ MICROINTRODUCER WAS RETURNED FOR EVALUATION. A PRODUCT LABEL INDICATING LOT: REGW2650. A PHOTOGRAPH OF THE IMPLICATED 5FR MICROINTRODUCER/DILATOR/SHEATH WAS ALSO PROVIDED AND APPEARED TO CORRESPOND WITH THE RETURNED PHYSICAL SAMPLE. DRIED BLOOD RESIDUE AND EVIDENCE OF USE WAS NOTED ON THE DEVICE. THE DILATOR WAS PRESENT WITHIN THE SHEATH WHICH HAD NOT BEEN PEELED. THE SHAFT OF THE SHEATH AND DILATOR WAS OBSERVED TO BE BENT. MICROSCOPIC INSPECTION OF THE SHEATH TIP DID NOT REVEAL ANY DEFORMATION OR DAMAGE. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE ADVANCEMENT AGAINST RESISTANCE, INADEQUATE SKIN NICK, AND SHEATH TIP TRANSITION. A REVIEW OF THE MANUFACTURING RECORDS DID NOT REVEAL EVIDENCE TO INDICATE A MANUFACTURING RELATED ROOT CAUSE. SINCE THE DEFORMATION ON THE MICROINTRODUCER SHAFT SUGGEST THAT RESISTANCE WAS LIKELY EXPERIENCED DURING INSERTION, THE COMPLAINT IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02674 | 2023-04-19 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589 , NGDY1685 AND 2120842164. AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON 29MAR 2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022 AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON (B)(6) 2022. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED . ON (B)(6) 2022 THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED . ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC, CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01418 | 2023-04-19 | 2023-03-30 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT 5F DUAL-LUMEN REVERSE TAPER PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7857505 |  | IT WAS REPORTED BY THE CUSTOMER "THE GROSHONG 5FR DOUBLE LUMEN CATHETER BEGAN TO LEAK AT THE CONNECTION BETWEEN THE DISTAL AND PROXIMAL ROUTES (LUMENS). PATIENT WAS ON A CONTINUOUS INFUSION OF MEDICATION AND HAD NO RESISTANCE, IT STARTED IN ONE ROUTE AND THEN HAPPENED IN THE OTHER. THE PICC LINE WAS REMOVED." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02673 | 2023-04-19 | 2022-12-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589, NGDY1685 AND 2120842164. AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022 AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON (B)(6) 2022. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED. ON (B)(6) 2022 THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED AS THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC . ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01426 | 2023-04-19 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC SOLO2 5F 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | IT WAS REPORTED BY THE CUSTOMER THAT "INTRODUCER WILL NOT SLIDE IN AS EXPECTED. CUSTOMER SAYS IT IS CAUSING TRAUMA TO THE SKIN AND VESSEL BECAUSE OF FORCE NEEDED TO INSERT THE DEVICE. ADDITIONAL INFORMATION RECEIVED 04/07/2023: OPENED ANOTHER INTRODUCER KIT AND PLACED THE LINE. THERE WERE NO COMPLICATIONS SECONDARY TO THE VESSEL TRAUMA. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DIFFICULTY INSERTING THE INTRODUCER SHEATH WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. ONE 5 FR X 7 CM MICROEZ MICROINTRODUCER WAS RETURNED FOR EVALUATION. THE SLIGHT BEND IN THE SHAFT OF THE DILATOR AND SHEATH WAS NOTED. DEFORMATION AT THE DISTAL END OF THE SHEATH WAS OBSERVED. A SINGLE PHOTOGRAPH OF THE IMPLICATED 5FR X 7CM MICROEZ MICROINTRODUCER WAS ALSO PROVIDED. THE PHOTO SAMPLE CORRESPONDED WITH THE RETURNED PHYSICAL SAMPLE. MICROSCOPIC INSPECTION OF THE DEFORMED SHEATH TIP REGION REVEALED A SECTION TO BE BUNCHED INWARDS. BLOOD RESIDUE WAS PRESENT AT THE DAMAGED SECTION. EVIDENCE OF A SHEATH TIP TAPER WAS OBSERVED. A REVIEW OF THE MANUFACTURING RECORDS DID NOT REVEAL EVIDENCE OF A MANUFACTURING RELATED ROOT CAUSE. POSSIBLE CONTRIBUTING FACTORS COULD INCLUDE TOO SMALL OF AN INCISION, A STEEP INSERTION ANGLE, A POOR TRANSITION OF THE INTRODUCER SHEATH TO THE DILATOR, OR DAMAGE TO THE SHEATH PRIOR TO USE. SHEATH DAMAGE CAN BE MINIMIZED BY SIMULTANEOUSLY ADVANCING THE SHEATH AND DILATOR AS A SINGLE UNIT USING A ROTATIONAL MOTION. IF NECESSARY, A SMALL INCISION MAY BE MADE ADJACENT TO THE GUIDEWIRE TO FACILITATE INSERTION OF THE SHEATH AND DILATOR. THE DAMAGE PRESENT ON THE MICRINTRODUCER SHEATH IS INDICATIVE OF A DIFFICULT INSERTION; THEREFORE, THE COMPLAINT IS CONFIRMED BUT THE EXACT CONTRIBUTING FACTORS REMAIN UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02667 | 2023-04-19 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILIZED WATER WAS UNABLE TO DRAIN DURING PRETEST. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 2 MINUTES WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE IP7603444, REVISION 0, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02671 | 2023-04-19 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT A FEW DAYS AFTER THE OPERATION, THE CATHETER FELL OUT SPONTANEOUSLY. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 16MAY2023, STATED THAT THE DETAILS REGARDING ANY DAMAGE ON THE BALLOON WAS NOT PROVIDED FROM THE CUSTOMER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND SYRINGE. VISUAL INSPECTION NOTED THAT THE TEMPERATURE WIRE WAS CUT AND THE BALLOON WAS INFLATED UPON RECEIVAL OF SAMPLE. DEFLATED BALLOON PASSIVELY WITH NO CUFFING OR LEAKS NOTED. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02670 | 2023-04-19 | 2022-12-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE EYE AT THE TIP OF THE CATHETER WAS DIRTY WHEN OPENED AND WAS NOT USED TO THE PATIENT. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. BASED ON THE EVALUATION, OBSERVED THERE WAS BROWN STAIN ON THE SURFACE OF TIP CATHETER. HOWEVER, THE ORIGIN OF THE STAIN COULD NOT BE DETERMINED. A POTENTIAL ROOT CAUSE FOR THIS POTENTIAL FAILURE MODE COULD BE DUE TO UNCLEAN MACHINE OR DIRT ON CATHETER DUE TO MACHINE FRICTION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02654 | 2023-04-19 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER BLEW UP TUBE RATHER THAN THE END, TWICE IN THE PAST WEEK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02653 | 2023-04-19 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | URETHRAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE ISSUES WITH THE INTERMITTENT CATHETERS. THE URINE WOULD NOT COME OUT, AND IT WAS CLOGGED. NO IODINE CAME IN THE PACKAGE AND ONLY ONE GLOVE CAME IN THE PACKAGE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INSUFFICIENT EYELET AREA BLOCKED EYE OR INNER DIAMETER OF DRAINAGE EYE SMALL ENOUGH TO ALLOW FOR OCCLUSION DURING USE". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INSTRUCTIONS FOR USE: PREPARING FOR CATHETERIZATION: OPEN KIT AND REMOVE CONTENTS. THE GLOVES ARE NOT NECESSARY TO MAINTAIN ASEPTIC CONDITIONS DURING THE PROCEDURE. THE GLOVES ARE FOR THE OPERATOR'S PROTECTION. MALE/FEMALE CATHETERIZATION: REMOVE CAP FROM INSERTION TIP WHILE SQUEEZING TABS. SAVE CAP FOR CLOSING THE BAG. SLIDE THE CATHETER HALFWAY INTO THE INSERTION TIP. MALE: PREPARE THE MALE URETHRAL MEATUS AND THE SURROUNDING AREA WITH POVIDONE-IODINE SWABS PROVIDED. FEMALE: PREPARE THE FEMALE URETHRAL MEATUS BY HOLDING THE OUTER LABIA APART AND PREP THE URETHRAL MEATUS AND SURROUNDING AREA WITH POVIDONE IODINE SWABS PROVIDED. MALE: HOLDING THE PENIS, ADVANCE THE INSERTION TIP INTO THE URETHRA NO FURTHER THAN THE FLANGE BASE. FEMALE: SPREAD THE INNER LABIA, ADVANCE THE INSERTION TIP INTO THE URETHRA NO FURTHER THAN THE FLANGE BASE. RELEASE THE INNER LABIA. PLACE YOUR NONDOMINANT HAND ON THE FINGER CONTROL GUIDE TO STABILIZE CATHETER IN URETHRA. WITH YOUR DOMINANT HAND, GRASP CATHETER THROUGH BAG APPROXIMATELY 1" BELOW THE FINGER CONTROL GUIDE AND PUSH CATHETER INTO URETHRA. THE CATHETER SHOULD BE INTRODUCED BY SHORT, REPETITIVE PUSHING MOTIONS. REPEAT MOTIONS UNTIL CATHETER REACHES BLADDER AND URINE STARTS TO FLOW. ALLOW URINE TO FLOW FREELY, MAKING CERTAIN THE CATHETER GUIDE IS ELEVATED AT LEAST 4" ABOVE LOWER PORTION OF THE BAG. ALLOW FLOW UNTIL BLADDER IS EMPTY OR UNTIL BAG IS FILLED. PRECAUTION: IT IS RECOMMENDED THAT THE COLLECTION BAG BE HELD. THE CATHETER COULD POSSIBLY SEPARATE FROM THE COLLECTION BAG WHEN URINE INCREASES WEIGHT OF THE BAG. WITHDRAW THE CATHETER FROM THE URETHRA. REMOVE THE REMAINING PORTION OF THE CATHETER FROM THE COLLECTION BAG BY PULLING IT THROUGH THE INSERTION TIP. NOTE: THE CATHETER IS DESIGNED TO PASS THROUGH THE INSERTION TIP. THE FILLED COLLECTION BAG MAY BE CLOSED BY REPLACING THE CAP OVER THE INSERTION TIP. MAKE SURE THE CAP SNAPS ON SECURELY. SPECIMEN COLLECTION: TO COLLECT A SPECIMEN, OBTAIN A SAMPLE CUP/TUBE AND ALCOHOL WIPE. REMOVE GUIDE TIP BY PULLING TAB(S) UPWARD FROM BAG. WIPE PORT WITH ALCOHOL. POUR SPECIMEN THROUGH PORT AT TOP OF BAG INTO CUP/TUBE. DRAINING BAG: REMOVE INSERTION TIP BY PULLING TAB(S) UPWARD FROM BAG. DRAIN URINE THROUGH PORT AT TOP OF BAG.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
2020394-2023-00264 | 2023-04-19 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U4150230RX | U4150230RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE OF A CHRONIC TOTAL OCCLUSION IN THE SUPERFICIAL FEMORAL ARTERY VIA A CONTRALATERAL APPROACH, THE CATHETER ALLEGEDLY BENT EASILY AND IT FAILED TO BE ADVANCED THROUGH THE SHEATH. IT WAS FURTHER REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A MALFUNCTION. H10: D4 (EXPIRATION DATE: 11/2024), G3. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02669 | 2023-04-19 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE ITEM WAS RECEIVED AND EXPECTED TO BE PREFILLED, BUT IT WASN'T, AND IN RELATION TO 'THE ITEM NOT PREFILLED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 09MAY2023, IT WAS NOT PREFILLED. | THE REPORTED EVENT IS INCONCLUSIVE DUE TO UNABLE TO CONFIRM REPORTED EVENT BASED ON ATTACHED PHOTO. NO FUNCTIONAL TESTING COULD BE PERFORMED AS ONLY PHOTOS ARE PROVIDED AS SAMPLE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO OPERATOR HANDLING ISSUE OR DEFECTIVE PART GENERATED AT SUPPLIER SITE OR DURING TRANSIT TO BARD. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. ¿ USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR ¿ FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02655 | 2023-04-19 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT AFTER INSERTION IN THE TREATMENT ROOM, THE CATHETER WAS SPONTANEOUSLY REMOVED IN THE WARD AND THE NEW CATHETER WAS REPLACED BECAUSE WATER LEAKAGE FROM THE TIP OF THE BALLOON WAS SUSPECTED. THE EVENT OCCURRED 12 HOURS AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BALLOON. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT CAUSED THE REPORTED FAILURE. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED A 0.041" PINHOLE FOUND ON THE SHAFT AND LOCATED 0.4850" FROM THE BIFURCATION. THIS IS OUT OF SPECIFICATION, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02719 | 2023-04-20 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE PATIENT WAS HAVING A TROUBLE IN INSERTING THESE MAGIC 3 HYDROPHILIC INTERMITTENT COUDE CATHETER. PATIENT SAID THE STERILE WATER PACKET DID NOT HAVE ENOUGH LUBRICANT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "AIR IN LINES". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE. INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01461 | 2023-04-20 | 2023-04-04 | I, n, j, u, r, y | Manufacturer report | BARD POWERPICC SOLO | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED "THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT: REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. THE EVENT INVOLVED A SL POWERPICC SOLO IN THE UPPER ARM." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02718 | 2023-04-20 | 2023-03-30 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT DEVELOPED A FUNGAL INFECTION RIGHT WHERE THE PUREWICK FEMALE EXTERNAL CATHETERS WERE PLACED AND ALSO STATED THAT THE PUREWICK FEMALE EXTERNAL CATHETER TURNED BLACK BETWEEN THE SEPARATION OF LABIA AND DESCRIBED IT AS CHARCOAL LOOKING. REPRESENTATIVE ADVISED TO REACH OUT TO DOCTOR FOR MEDICAL ADVICE. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR YEAST INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. WARNINGS: ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: ¿ AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER ¿ HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. ¿ EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. ¿ PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01460 | 2023-04-20 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT 4FR POWERMIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P4154108D |  | IT WAS REPORTED BY THE CUSTOMER THAT "EXTENSION LEG CAME OUT OF THE SUTURE WING. HAD TO USE ANOTHER CATHETER." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01457 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER "AN ISSUE WHERE THE WIRE THAT COMES WITH THE MICRO-INTRODUCER KIT WOULD NOT ADVANCE THROUGH THE PEEL-AWAY SHEATH. SHE STATES THE END OF THE WIRE 'FELT BUCKLED'. THE ISSUE HAPPENED TWICE WITH MICRO-INTRODUCER KITS FROM THE SAME BOX." THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02716 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALLS OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 06APR2023, STATED THAT THEY HAD SEEN SKIN LESIONS OFTEN ON MEN WHEN ASKED REGARDING NEGATIVE CONSEQUENCES FOR THE PATIENT. IT WAS UNCLEAR WHETHER IT IS THE FAULT OF THE BARDIA CATHETERS, BUT THEY FELT THAT IT HAD BECOME MORE SO IN RECENT YEARS. THIS APPLIED TO SEVERAL DIFFERENT PATIENTS BUT THEY LIVE IN NURSING HOMES. SOME WERE SLIM AND SOME HAD HIGHER BODY MASS INDEX. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 06APR2023, STATED THAT THEY HAD SEEN SKIN LESIONS OFTEN ON MEN WHEN ASKED REGARDING NEGATIVE CONSEQUENCES FOR THE PATIENT. IT WAS UNCLEAR WHETHER IT IS THE FAULT OF THE BARDIA CATHETERS, BUT THEY FELT THAT IT HAD BECOME MORE SO IN RECENT YEARS. THIS APPLIED TO SEVERAL DIFFERENT PATIENTS BUT THEY LIVE IN NURSING HOMES. SOME WERE SLIM AND SOME HAD HIGHER BODY MASS INDEX. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 03MAY2023, STATED THAT IT WAS UNEVEN IN THE NUMBER OF CATHETERS. THEY USUALLY JUST DID NOT LAST 12 WEEKS. THEY USUALLY WERE CHANGED MORE FREQUENTLY. MOST OFTEN DUE TO BLOCKAGE IN THE CATHETERS. MANY OF THE MEN HAD ALSO HAD SKIN ABRASIONS FROM THESE CATHETERS, AND THUS PERMANENT SKIN DAMAGE. THE CUSTOMER COULD NOT ANSWER THE NUMBER OF CATHETERS AS THEY GENERALLY SAW THESE PROBLEMS IN ALMOST ALL BUT ONE PATIENT. THE PROBLEM HAD BEEN THERE FOR A LONG TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USE FOR UROLOGICAL CARE. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02715 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER INSIDE THE FOLEY CATHETER BALLOON WAS UNABLE TO REMOVE DURING A PRETEST. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. VISUAL EVALUATION NOTED RECEIVED A 3-WAY TEMP SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. REPLACED RETURNED INFLATION VALVE WITH IN-HOUSE VALVE AND RE-ATTEMPTED DEFLATION WITH BOTH SYRINGES. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE FOREIGN MATTER DURING PROCESSES (FLAP OR FLASH DURING PUNCHING). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿WARNINGS 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. DIRECTIONS FOR USE: 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDO-ELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR. (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02711 | 2023-04-20 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT WHEN THE NURSE FLUSHED THE LUMEN OF CATHETER FOR THE PATIENT, IT WAS FOUND THAT THE LUMEN WAS OCCLUDED. THEN A NEW CATHETER WAS USED IMMEDIATELY, AND THE CATHETERIZATION WAS PERFORMED SUCCESSFULLY. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 04APR2023, AND NO PATIENT INJURY AND THE IRRIGATION LUMEN WAS BLOCKED AND SO THEY WERE UNABLE TO IRRIGATE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INCORRECT PARAMETER SETTING - BLOCKED DRAINAGE EYE/LUMEN OR NO DRAINAGE EYE OR USER RELATED (SALT ACCUMULATION). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATER ALLERGY". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02720 | 2023-04-20 | 2023-03-30 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT WAS PRONE TO GETTING URINARY TRACT INFECTIONS, BUT THEY SAID THEY WERE FINDING THAT THE PATIENT WAS GETTING LESS URINARY TRACT INFECTIONS WHEN USING BRIEFS INSTEAD OF THE PUREWICK. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. WARNINGS: ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: ¿ AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER ¿ HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. ¿ EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. ¿ PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | Y | N | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01458 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER "AN ISSUE WHERE THE WIRE THAT COMES WITH THE MICRO-INTRODUCER KIT WOULD NOT ADVANCE THROUGH THE PEEL-AWAY SHEATH. SHE STATES THE END OF THE WIRE 'FELT BUCKLED'. THE ISSUE HAPPENED TWICE WITH MICRO-INTRODUCER KITS FROM THE SAME BOX." THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01462 | 2023-04-20 | 2023-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER, "2 DAYS AFTER THE INSERTION, THE CATHETER BEGAN TO LEAK CLOSE TO THE INSERTION, REQUIRING REMOVAL AND PASSAGE OF A NEW CATHETER". NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02729 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02723 | 2023-04-20 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02726 | 2023-04-20 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT FEW OF THE INTERMITTENT CATHETERS DID NOT HAVE LIQUID IN THE PACKETS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02724 | 2023-04-20 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED, THAT FOLEY CATHETER SURFACE WAS BUMPY. THERE WAS A BUMPY PROTRUSION JUST ABOVE THE FOLEY CATHETER BIFURCATION AND ON THE SURFACE OF THE SHAFT. THE BIFURCATION WHERE THE INFLATION LUMEN WAS CONNECTED TO THE SHAFT WAS SLIGHTLY RAISED. BUT, IT WAS UNCLEAR WHERE THE SHAFT WAS BUMPY. IT WAS UNABLE TO DETERMINE IF THE BUMPINESS AT THE BRANCH WAS ALSO A DEFECT OR NOT. | H10.THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02725 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED, THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED, THAT THE CATHETER FALLS OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED ,THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED, THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED, THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02728 | 2023-04-20 | 2023-03-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT FOLEY CATHETER HAD LEAK ALONG SEAM OF THE CATHETER TUBE WHEN INSERTED IN PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02712 | 2023-04-20 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT FOREIGN OBJECTS WERE ADHERING TO THE FOLEY CATHETER. AS PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 06APR2023, STATED THAT FOREIGN MATERIAL WAS FOUND ON THE SURFACE OF CATHETER. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR EVENT IS NOT REPORTABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02732 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND TO BE A CASCADING FAILURE OF AN EVENT PREVIOUSLY REPORTED. NOTE: NOTE: THE CORRECT ZIP CODE UNDER TAB E IS 182 33. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02731 | 2023-04-20 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALLS OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02736 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALLS OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 06APR2023, STATED THAT THEY HAD SEEN SKIN LESIONS OFTEN ON MEN WHEN ASKED REGARDING NEGATIVE CONSEQUENCES FOR THE PATIENT. IT WAS UNCLEAR WHETHER IT WAS THE FAULT OF THE BARDIA CATHETERS, BUT THEY FELT THAT IT HAD BECOME MORE SO IN RECENT YEARS. THIS APPLIED TO SEVERAL DIFFERENT PATIENTS BUT THEY LIVE IN NURSING HOMES. SOME WERE SLIM AND SOME HAD HIGHER BODY MASS INDEX. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 03MAY2023, STATED THAT IT WAS UNEVEN IN THE NUMBER OF CATHETERS. THEY USUALLY JUST DID NOT LAST 12 WEEKS. THEY WERE USUALLY CHANGED MORE FREQUENTLY. MOST OFTEN DUE TO BLOCKAGE IN THE CATHETERS. MANY OF THE MEN ALSO HAD SKIN ABRASIONS FROM THESE CATHETERS, AND THUS PERMANENT SKIN DAMAGE. THE CUSTOMER COULD NOT ANSWER THE NUMBER OF CATHETERS AS THEY GENERALLY SAW THESE PROBLEMS IN ALMOST ALL BUT ONE PATIENT. THE PROBLEM HAD BEEN THERE FOR A LONG TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO POOR COATING THAT CAN LEADING CATHETER TO PERFORATION OR EROSION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE THE LATEX. VISUALLY INSPECTED THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATE 10ML OF STERILE WATER OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. ALLOW THE PRESSURE WITHIN TO BALLOON FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02735 | 2023-04-20 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALLS OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNKNOWN). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNKNOWN). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNKNOWN). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNKNOWN). BUT THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNKNOWN). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNKNOWN). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02734 | 2023-04-20 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02721 | 2023-04-20 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE BALLOON WAS DAMAGED DURING PRETEST. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 06APR2023, STATED THAT THE BALLOON WAS NOT DAMAGED BY APPEARANCE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. AS THE REPORTED EVENT IS UNCONFIRMED, LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02710 | 2023-04-20 | 2023-03-30 | I, n, j, u, r, y | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT LATEX ALLERGY PATIENT ACCIDENTALLY HAD A LATEX FOLEY INSERTED IN THE PACU AND DEVELOPED A SEVERE LATEX REACTION. FOLEY WAS REMOVED AND REPLACED WITH A SILICONE CATHETER. MEDICAL INTERVENTION IS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE MISHANDLING OF DEVICE BY USER. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE PRODUCT CATALOG NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02699 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 06APR2023, STATED THAT THEY HAD SEEN SKIN LESIONS OFTEN ON MEN WHEN ASKED REGARDING NEGATIVE CONSEQUENCES FOR THE PATIENT. IT WAS UNCLEAR WHETHER IT IS THE FAULT OF THE BARDIA CATHETERS, BUT THEY FELT THAT IT HAD BECOME MORE SO IN RECENT YEARS. THIS APPLIED TO SEVERAL DIFFERENT PATIENTS BUT THEY LIVE IN NURSING HOMES. SOME WERE SLIM AND SOME HAD HIGHER BODY MASS INDEX. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND TO BE A CASCADING FAILURE OF AN EVENT PREVIOUSLY REPORTED. NOTE: THE CORRECT ZIP CODE UNDER TAB E IS 182 33. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02708 | 2023-04-20 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE BALLOON OF CATHETER WAS BURST AND THE CATHETER FELL OUT. THEN A NEW CATHETER WAS REPLACED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 04APR2023 NO MISSING PIECES OF BALLOON LEFT IN PATIENT AND NO PATIENT INJURY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01454 | 2023-04-20 | 2023-01-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. HARD BASE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0604520CE |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE INTRODUCER NEEDLE WAS ALLEGEDLY DIFFICULT TO REMOVE. IT WAS FURTHER REPORTED THAT, THE WIRE ALLEGEDLY STUCK. REPORTEDLY, THE WIRE WAS REMOVED ALONG WITH NEEDLE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE M.R.I. HARD BASE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE M.R.I. HARD BASE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED. MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE BARDPORT KIT. FOLLOWING COMPONENTS WERE RECEIVED: ONE GROSHONG CATHETER WITH A LOADED CATHETER STYLET, ONE INTRODUCER NEEDLE WITH A LOADED GUIDEWIRE, ONE 8.0FR PEEL-APART SHEATH AND VESSEL DILATOR, ONE VEIN PICK, ONE STRAIGHT NON-CORING NEEDLE, ONE TUNNELER, ONE GUIDEWIRE HOOP, ONE CATH-LOCK, ONE FLUSHING CONNECTOR, ONE SEALED SAFETY INFUSION SET, AND ONE RIGHT-ANGLE NEEDLE LOADED TO A BARDPORT MRI IMPLANTABLE PORT WERE RETURNED FOR EVALUATION. GROSS VISUAL, FUNCTIONAL, MICROSCOPIC AND DIMENSIONAL EVALUATION WERE PERFORMED ON THE RETURNED DEVICE. THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED STRETCHED, MATERIAL SEPARATION, DEFORMATION AND FRACTURE ISSUE AS BENDS WERE NOTED THROUGHOUT THE GUIDEWIRE AND COMPLETE BREAKS WERE NOTED TO BOTH ROUND AND FLAT CORE WIRES WITHIN THE UNCOILED PORTION OF THE GUIDEWIRE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED DIFFICULT TO REMOVE AND PHYSICAL RESISTANCE ISSUES AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT ARE UNKNOWN. BASED ON THE MEASUREMENTS RECORDED DURING SAMPLE EVALUATION AND APPLICABLE DRAWINGS, NO MEASUREMENTS RECORDED COULD BE CONFIRMED TO BE OUT OF TOLERANCE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. (EXPIRY DATE: 12/2030). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01455 | 2023-04-20 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT TWO YEARS POST PORT PLACEMENT PROCEDURE, A LONGITUDINAL CRACK WAS ALLEGEDLY FOUND ON THE CATHETER. REPORTEDLY, THE CRACKED CATHETER WILL BE REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 12/2021). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00265 | 2023-04-20 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE OF CHRONIC TOTAL OCCLUSION OF A HIGHLY CALCIFIED TARGET LESION BELOW THE RIGHT KNEE VIA AN IPSILATERAL APPROACH, THE DEVICE WAS ALLEGEDLY UNABLE TO ADVANCE AROUND THE ANKLE. IT WAS FURTHER REPORTED THAT THE TIP OF THE CATHETER WAS ALLEGEDLY FOUND TO COME OFF FROM THE SHAFT OF THE CATHETER. REPORTEDLY, DETACHED COMPONENT WAS NOT FULLY SEPARATED FROM THE CATHETER. THERE WAS NO REPORTED PATIENT INJURY. | .H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE CROSSER CTO RECANALIZATION CATHETER WAS RECEIVED AND EVALUATED. UPON VISUAL EVALUATION, MARKER BAND WAS NOTED TO BE PINCHED WITH DELAMINATION. NO FUNCTIONAL TESTING DUE TO THE CONDITION OF THE DEVICE. THEREFORE, THE INVESTIGATION IS DETERMINED TO BE CONFIRMED FOR THE REPORTED MATERIAL SEPARATION ISSUE AS DISTAL TIP APPEARED TO BE SEPARATED FROM THE OUTER CATHETER WITH A TEAR, BUT STILL ATTACHED TO THE INNER CORE WIRE AND THE INVESTIGATION IS ALSO CONFIRMED FOR THE IDENTIFIED DEFORMATION DUE TO COMPRESSIVE STRESS AND PEELED ISSUE AS MARKER BAND WAS NOTED TO BE PINCHED WITH DELAMINATION. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE TO ADVANCE AS THE REPORTED EVENT CAN NOT BE REPRODUCED. A DEFINITIVE ROOT CAUSE FOR THE REPORTED FAILURE TO ADVANCE, MATERIAL SEPARATION AND THE IDENTIFIED DEFORMATION DUE TO COMPRESSIVE STRESS, AND PEELED ISSUES COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 10/2024), G3, H6 (DEVICE) H11: H6 (METHOD, RESULT, CONCLUSION) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02682 | 2023-04-20 | 2023-01-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589, NGDY1685 AND 2120842164. AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022, AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON (B)(6) 2022. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED. ON (B)(6) 2022 THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC. ON (B)(6) 2023. THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02687 | 2023-04-20 | 2023-03-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589, NGDY1685 AND 2120842164. AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON 29MAR2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022 AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON (B)(6) 2022. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED. ON (B)(6) 2022 THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED AS THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC. ON 24JAN2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02683 | 2023-04-20 | 2023-01-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589, NGDY1685 AND 2120842164. AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022 AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON (B)(6) 2022. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED. ON (B)(6) 2022 THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED AS THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02686 | 2023-04-20 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589, NGDY1685 AND 2120842164. AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022 AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON (B)(6) 2022. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED. ON (B)(6) 2022 THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED AS THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02685 | 2023-04-20 | 2023-02-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589, NGDY1685 AND 2120842164. AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022 AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON (B)(6) 2022. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED. ON (B)(6) 2022 THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED AS THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02684 | 2023-04-20 | 2023-01-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE URINARY CATHETERS WERE IN CLOSE PROXIMITY BECAUSE URINE WAS PRESENT IN THE PROTECTION. THE IMPLANTATION DATE WAS (B)(6) 2022, AND THE OCCURRENCE DATE WAS (B)(6) 2023. REPLACEMENT OF THE MEDICAL DEVICE WAS DONE. THE LOT# NUMBERS REPORTED WERE NGGQ0589, NGDY1685 AND 2120842164. AS PER THE ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION ON (B)(6) 2022 AND URINARY CATHETER WITH LOT NUMBER LOT 2120842164 WAS CHANGED ON 06DEC2022. ON 25JAN2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGGQ0589 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER 2120842164 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION BUT URINARY CATHETER WAS NOT CHANGED. ON (B)(6) 2022 THERE WAS MANIFOLD LEAK AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2022, THE USER SCHEDULED CHANGE FOR CATHETER WITH PRODUCT CATALOG NUMBER SCC226514. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION, AND CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED AS THERE WAS LEAK IN PROTECTION AND THE BALLOON DEFLATED TO 7CC. ON (B)(6) 2023, THE CATHETER WITH LOT NUMBER NGDY1685 WAS INSTALLED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THERE WAS LEAK IN PROTECTION AND URINARY CATHETER WITH LOT NUMBER NGDY1685 WAS CHANGED. ON (B)(6) 2023, THE CATHETER WITH 10ML BALLOON WAS INSTALLED. ON (B)(6) 2023, THE CATHETER REMOVED BY DOCTOR AS BALLOON DEFLATED TO 5CC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02688 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE ISSUES WITH THE INTERMITTENT CATHETERS. THE URINE WOULD NOT COME OUT AND IT WAS CLOGGED. NO IODINE CAME IN PACKAGE AND ONLY ONE GLOVE CAME IN PACKAGE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT OPERATION". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A LABELING REVIEW IS NOT REQUIRED BECAUSE A REVIEW OF THE LABEL COULD NOT HAVE PREVENTED THIS ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02693 | 2023-04-20 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WAS A HOLE IN THE FOLEY CATHETER BALLOON. AFTER THE CATHETER INSERTION TO THE PATIENT, STERILE WATER WAS INJECTED AND THE WATER CAME OUT EASILY WHILE THE USER WAS CHECKING THE CATHETER TO SEE WHETHER WATER WOULD COME OUT. THE EVENT OCCURRED JUST AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BALLOON. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: ¿[WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. MATERIAL: BALLOON CATHETER: SILICONE; SILVER COATING; THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO 22 EVERY 2FR: BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER: ACCESSORIES: CLOSED DRAINAGE BAG; (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.); SYRINGE PREFILLED WITH STERILE WATER; WATER SOLUBLE LUBRICANT; ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION; TWEEZERS; GAUZE PADS; WATERPROOF SHEET; COTTON BALLS; GLOVES; STATLOCK FOLEY. [INTENDED USE & EFFECT- EFFICACY]: THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE]: METHOD OF USE: CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. PRECAUTIONS FOR USE: TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT: OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING: CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING: REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING: DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS]: PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES: NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.); BALLOON-RUPTURE METHOD. WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST.NON-RUPTURE METHOD: ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) BALLOON-RUPTURE METHOD: INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) MRI SAFETY: BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. MALFUNCTION AND ADVERSE EVENTS: MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE; DIFFICULTY OR FAILURE TO REMOVE THE DEVICE; OCCLUSION OF CATHETER INNER LUMENS; ENCRUSTATION; ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE; DEVICE DAMAGE DUE TO INAPPROPRIATE USE; FAILURE TO MEASURE TEMPERATURE; IMPROPER TEMPERATURE INDICATION. ADVERSE EVENTS: URINARY-TRACT INFECTION; HEMORRHAGE, HEMATURIA; ALLERGY REACTION TO THE DEVICE; CALCULUS FORMATION; EDEMA; PAIN; DISCOMFORT; INJURY OF BLADDER OR URETHRAL; URETHRITIS, URINARY INCONTINENCE; RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]: STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02692 | 2023-04-20 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT WHEN THE FOLEY CATHETER WAS REPLACED DUE TO OBSTRUCTION, THE PATIENT EXPERIENCED DIFFICULTY DRAINING THE CATHETER. ALTHOUGH THE CATHETER WAS CUT AND OTHER MEASURES WERE TAKEN, IT COULD NOT BE DRAINED, AND THE BALLOON WAS REMOVED WHILE INFLATED. STATED THAT THERE WAS SLIGHT BLEEDING. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. SAMPLE HAS BEEN EVALUATED AND OBSERVED CATHETER WAS CUT INTO 2 PIECES AND THE INFLATION FUNNEL WAS NOT RETURNED. THE BALLOON WAS STILL SLIGHTLY INFLATED. THE CATHETER WAS UNABLE TO BE INFLATED AND DEFLATED USING SYRINGE AND UNABLE TO PERFORM FLOW RATE TEST SINCE SAMPLE HAVE BEEN CUT. SAMPLE WAS DISSECTED AND OBSERVED SALT ACCUMULATION THAT MIGHT HAVE CAUSE BLOCKAGE INSIDE THE DRAINAGE LUMEN. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE BLOCKED DRAINAGE EYE/LUMEN OR NO DRAINAGE EYE OR USER RELATED (SALT ACCUMULATION). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02690 | 2023-04-20 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI20 | 0165SI20 |  | IT WAS REPORTED THAT ON CHECKING THE PRODUCT, THE CATHETER WAS FOUND BENT. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 10APR2023, THE INFORMATION IF THE PRODUCT WAS USED ON THE PATIENT WAS NOT PROVIDED FROM CUSTOMER. | THE REPORTED EVENT WAS CONFIRMED ¿ UNKNOWN CAUSE. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. BASED ON THE EVALUATION, IT WAS OBSERVED DENT ON MIDDLE OF CATHETER¿S SHAFT. THE CATHETER WAS ABLE TO INFLATE AND DEFLATE WITHOUT DIFFICULTY. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO OPERATOR'S HANDLING METHOD. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02689 | 2023-04-20 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | WIDEBAND¿ SELF-ADHERING MALE EXTERNAL CATHETER | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 36303 | 36303 |  | IT WAS REPORTED THAT THE MALE EXTERNAL CATHETER HAD TOO MUCH ADHESIVE AND THE PATIENT HAD ISSUES TRYING TO REMOVE CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: 1) VERIFY CORRECT SIZE PRIOR TO USE. 2) TRIM PUBIC HAIR IF NECESSARY. 3) WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. 4) OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. 5) PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. 6) UNROLL THE CATHETER OVER PENIS. 7) GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. 8) CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: 1) ENSURE DRAINAGE BAG IS EMPTY. 2) DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. 3) GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER.¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02695 | 2023-04-20 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWKIT03 | PWKIT03 |  | IT WAS REPORTED THAT EXTENSION TUBE IN PUREWICK URINE COLLECTION SYSTEM GOING FROM PUMP TO BOTTLE WAS NOT LONG ENOUGH AND WAS TOLD BY A REP THAT IT WAS 18IN, BUT CUSTOMER SAID IT WAS ONLY 6 IN, VERY UPSET PATIENT GOT A URINARY TRACT INFECTION BECAUSE OF THE ISSUE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS GIVEN TO URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE COMPONENT SELECTION / COMPONENT DETERIORATION". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PLACE THE PUREWICK¿ URINE COLLECTION SYSTEM (A) NEXT TO THE BED OR CHAIR WHERE IT WILL BE USED. IT MUST BE CLOSE ENOUGH FOR THE COLLECTOR TUBING WITH ELBOW CONNECTOR (G) TO EASILY REACH THE USER AND SHOULD BE PLACED ON A STABLE, EVEN SURFACE. FOR OPTIMAL RESULTS, POSITION THE PUREWICK¿ URINE COLLECTION SYSTEM EITHER ON THE FLOOR OR AS CLOSE TO THE FLOOR AS YOUR FACILITY¿S PROTOCOL WILL ALLOW.[NOTE THE LETTERS CORRELATE TO A DIAGRAM WITHIN THE IFU]. NOTE: THE PUREWICK¿ URINE COLLECTION SYSTEM SHOULD BE STATIONARY WHILE IN USE. DEVICE IS NOT DESIGNED TO BE USED WHILE IN TRANSPORT. 1. INSPECT THE PACKAGE FOR DAMAGE. IF ANY DAMAGE IS NOTICED, CONTACT CUSTOMER SUPPORT IMMEDIATELY. 2. CAREFULLY REMOVE ALL CONTENTS FROM THE PACKAGE, INCLUDING THE PACKING MATERIAL AND INSPECT PRIOR TO USE. IF ANY DAMAGE IS NOTICED, CONTACT CUSTOMER SUPPORT IMMEDIATELY. CAUTION: AVOID CREATING A TRIPPING HAZARD WITH THE COLLECTOR TUBING OR POWER CORD. ATTACH THE PUMP TUBING (D) TO THE PUREWICK¿ URINE COLLECTION SYSTEM CONNECTOR PORT (F) AND THE CONNECTOR PORT (E) ON THE COLLECTION CANISTER LID. ENSURE TUBING CONNECTIONS ARE CONNECTED PROPERLY. CHECK COLLECTOR TUBING FOR BLOCKAGE OR FLOW RESTRICTION SUCH AS PINCHED OR KINKED TUBING." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02696 | 2023-04-20 | 2023-03-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119416M | 119416M |  | IT WAS REPORTED THAT THE PATIENT HAD A SUBJECT LINE FOLEY CATHETER INSERTED ON (B)(6) 2023, AND WHEN NURSE TURNED THE PATIENT ON (B)(6) 2023, THEY NOTICED THE WIRE OF THE INDWELLING TEMP SENSING FOLEY CATHETER ALMOST BREAKING AND COMING THROUGH OUTSIDE OF THE CATHETER AND THEN FOLEY CATHETER WAS REMOVED, AND THERE WAS NO HARM TO THE PATIENT DUE TO THE EXPERT NURSING CARE. THEY DO NOT HAVE THE ORIGINAL PACKAGE, BUT THEY DO HAVE THE FAULTY FOLEY CATHETER. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT. THE PRODUCT CAUSED THE REPORTED FAILURE. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED A USED THREE-WAY FOLEY CATHETER. VISUAL INSPECTION NOTED THE THERMISTOR WIRE SNAKING THROUGH THE LUMEN ON THE CATHETER AND WAS VERY CLOSE TO PROTRUDING OUT THE SURFACE OF THE CATHETER. SOME ENCRUSTATION WAS NOTED WITHIN THE EYELETS AND ON THE BALLOON, HOWEVER, NO ADDITIONAL DEFECTS WERE NOTED. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE WHICH STATES, "THE THERMISTOR CABLE AND WIRE MUST NOT BE STRETCHED, WRINKLED, OR DEFORMED". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED DIRECTIONS FOR USE WASH HANDS AND DON CLEAN GLOVES EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERIURETHRAL AREA 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL 5. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP 6. DON STERILE GLOVES 7. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE 8. POSITION FENESTRATED DRAPE ON PATIENT 9. SATURATE 3 FOAM SWAB STICKS IN POVIDONE IODINE 10.ATTACH THE WATER FILLED SYRINGE TO THE INFLATION PORT NOTE: IT IS NOT NECESSARY TO PRE-TEST THE FOLEY CATHETER BALLOON 11. REMOVE FOLEY CATHETER FROM WRAP AND LUBRICATE CATHETER 12. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA INORA MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD 13. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND A. WHEN CATHETER TIP HAS NTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE B. INSERT CATHETER TWO MORE INCHES AND INFLATE CATHETER BALLOON 14. INFLATE CATHETER BALLOON USING ENTIRE 10CC OF STERILE WATER PROVIDED IN THE PREFILLED SYRINGE NOTE: USE OF LESS THAN 10CC CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON 15. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK 16. SECURE THE FOLEY CATHETER TO THE PATIENT USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED (SEE STATLOCK® FOLEY STABILIZATION DEVICE IFU) NOTE: PLEASE MAKE SURE PATIENT IS APPROPRIATE FOR USE OF STATLOCK® STABILIZATION DEVICE 17. POSITION HANGER ON BED RAIL AT THE FOOT OF THE BED NOTE: EXERCISE CARE TO KEEP BAG OFF THE FLOOR 18. USE GREEN SHEETING CLIP TO SECURE DRAINAGE TUBE TO THE SHEET. MAKE SURE TUBE IS NOT KINKED 19. INDICATE TIME AND DATE OF CATHETER INSERTION ON PROVIDED LABELS. PLACE DESIGNATED LABELS ON PATIENT CHART AND DRAINAGE SYSTEM 20. DOCUMENT PROCEDURE ACCORDING TO HOSPITAL PROTOCOL FOLEY CATHETER REMOVAL 1. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER LOCK OR SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE 2. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY 3. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION 4. IF THE BALLOON WILL NOT DEFLATE AND IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL 5. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT¿. CORRECTIONS: D, H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02698 | 2023-04-20 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED DURING THE OPERATION AND THE INTRAOPERATIVE URINE FLOW WAS GOOD. AT THE END OF THE OPERATION, THE CATHETER WAS PULLED OUT, AND WHEN THE CUFF WAS CHECKED, IT WAS DEFLATED. WHEN THE USER POURED WATER INTO THE CATHETER THAT CAME OUT AND CHECKED THE SITUATION, THERE WAS A WATER LEAK. IT WAS WILTING THE NEXT MORNING. THE EVENT OCCURRED FOUR HOURS AFTER PLACEMENT. THERE WAS NO BALLOON RUPTURE OR TEAR OBSERVED ON THE FOLEY CATHETER. THE WATER LEAKED FROM THE BALLOON. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE LUMEN COMPROMISED BY A PUNCTURE OR CUT. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: [WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02700 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE DUE TO SAMPLE WAS NOT AVAILABLE FOR INVESTIGATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USE FOR UROLOGICAL CARE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, INCORRECT WIRE PRESSING TECHNIQUE, INCORRECT STRIPPING TECHNIQUE AND ETC. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02706 | 2023-04-20 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM FOLEY CATHETER BALLOON DURING PRETEST. STATED THAT IT WAS DIFFICULT TO POUR WATER. THE BALLOON RUPTURE OR TEAR WAS OBSERVED ON THE FOLEY CATHETER. IT WAS STATED THAT THE CATHETER WAS INSERTED BY HAND AND NO MATERIALS POSSIBLY CAME IN CONTACT TO THE CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. A DHR REVIEW IS NOT REQUIRED FOR THIS COMPLAINT. A RISK REVIEW AND LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02701 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02704 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALLS OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02702 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALLS OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02703 | 2023-04-20 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT ON INJECTING FIXATION WATER DURING PRETEST, WATER LEAKAGE FROM A PINHOLE IN THE BALLOON WAS CONFIRMED. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS. PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER" CORRECTION: H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02707 | 2023-04-20 | 2023-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED FOR BLADDER IRRIGATION DURING THE OPERATION. AFTER THE OPERATION, THE CATHETER FELL OUT AND BALLOON BURST PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023,NO MISSING PIECES OF BALLOON LEFT IN PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02691 | 2023-04-20 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE ONE OF OUR DAC PATIENT HAS REPORTED THIS MORNING THAT THE PATIENT HAD ISSUES WITH 3 OF D23695S14 AND THE PATIENT WAS REPORTED THAT WHEN A NURSE CAME OVER TO DO AN EMERGENCY CHANGE OVER THE WEEKEND, THEY HAVE DISCOVERED THAT THE CATHETER WAS OUT OF IT¿S FOIL STERILE PACKAGING (FOIL BAG WAS LOOSE WITHIN BOX AS WAS CATHETER). NURSE REFUSED TO USE THEM AS THEY WEREN'T IN STERILE CONDITION. THEY HAVE TRIED ALL 3 AND THEY HAD THE SAME PROBLEM IN EACH OF THE BOXES. LUCKILY NURSE HAD A SPARE AND MANAGED TO CHANGE IT FOR OUR PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02737 | 2023-04-20 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01465 | 2023-04-20 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR CE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770362 |  | PER CUSTOMER: THE WAVEFORM ANALYSIS WAS INACCURATE. OF NOTE, I DID HAVE THE LEADS PLACED APPROPRIATELY AND DIDN¿T HAVE ANYTHING ELECTRICAL (CELLPHONE, WATCH, CALL LIGHT, TELE) NEAR PATIENT. ADDITIONAL INFORMATION RECEIVED 4/11/2023: ADVANCED 1CM GOT DEFLECTION AND SMALLER P WAVE, PULLED BACK TO 0, GOT THE ¿GREEN DIAMOND¿ AND MAXIMUM P WAVE HEIGHT. CXRAY TIP IN PROXIMAL SVC, HAPPENED TO GET A CHEST CT LATER AND SHOWED TIP IN MID SVC. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01464 | 2023-04-20 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR CE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770362 |  | PER CUSTOMER: THE WAVEFORM ANALYSIS WAS INACCURATE. OF NOTE, I DID HAVE THE LEADS PLACED APPROPRIATELY AND DIDN¿T HAVE ANYTHING ELECTRICAL (CELLPHONE, WATCH, CALL LIGHT, TELE) NEAR PATIENT. ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: ¿GREEN DIAMOND¿, MAXIMUM P WAVE HEIGHT. I¿M ASSUMING THIS IS JUST THE ¿RABBIT EARS¿. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02733 | 2023-04-20 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO POOR COATING THAT CAN LEADING CATHETER TO PERFORATION OR EROSION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE THE LATEX. VISUALLY INSPECTED THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATE 10ML OF STERILE WATER OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. ALLOW THE PRESSURE WITHIN TO BALLOON FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01466 | 2023-04-20 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR CE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770362 |  | PER CUSTOMER: THE WAVEFORM ANALYSIS WAS INACCURATE. OF NOTE, I DID HAVE THE LEADS PLACED APPROPRIATELY AND DIDN¿T HAVE ANYTHING ELECTRICAL (CELLPHONE, WATCH, CALL LIGHT, TELE) NEAR PATIENT. ADDITIONAL INFORMATION RECEIVED 4/11/2023: ADVANCED CATHETER AND WOULD GET WHAT APPEARED TO BE DEFLECTION AND SMALLER P WAVE SO LEFT OUT 5CM AND GOT ¿GREEN DIAMOND¿ WITH MAXIMUM P WAVE HEIGHT. CHEST XRAY - PICC LINE WITH TIP AT THE LEVEL OF THE MID SVC. RADIOLOGIST SAID TO ADVANCE 4CM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01456 | 2023-04-20 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER PORT TI ISP, 8 FR. GROSHONG, INT W SP, SINGLE LUMEN | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8708560 |  | IT WAS REPORTED THAT FOURTEEN DAYS POST PORT PLACEMENT, THE PORT ALLEGEDLY FLIPPED. IT WAS FURTHER REPORTED THAT THE PORT WAS ALLEGEDLY INACCESSIBLE. REPORTEDLY, THE PATIENT WILL UNDERGO REMOVAL OF THE FLIPPED PORT AND REPLACEMENT OF A NEW PORT. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURES AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 02/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
MW5116889 | 2023-04-20 | 2023-04-14 | I, n, j, u, r, y | Voluntary report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT SLIM MR 8F - LOG15341 | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | BARD ACCESS SYSTEMS, INC. | LOG1534175 | 1618000 |  | PORT PLACED ON (B)(6) 2023 AND POST-PROCEDURE XRAY SHOWED RIGHT-SIDED POWER-INJECTABLE SINGLE-LUMEN TUNNELED CHEST PORT, WITH CATHETER TIP IN THE EXPECTED LOCATION OF THE CAVOATRIAL JUNCTION. ACCESS OF THE PORT WAS ATTEMPTED ON (B)(6) 2023 AND WAS UNSUCCESSFUL; PORT WAS ABLE TO BE FLUSHED BUT NO BLOOD RETURN. CHEST XRAY ON (B)(6) 2023 SHOWED THE RIGHT CHEST INFUSION PORT WAS DISPLACED WITH ITS TIP NOTED ABOUT THE RIGHT LOWER NECK. PORT WAS REMOVED AND REPLACED ON (B)(6) 2023. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02772 | 2023-04-21 | 2023-04-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53814G | 53814G |  | IT WAS REPORTED THAT THE CUSTOMER ORDERED THROUGH LIBERATOR AND RECEIVED A TOTAL OF 12 BOXES OF INTERMITTENT CATHETER DID NOT HAVE SHIPPING SHEET WITH THE ORDER NUMBER. ALSO STATED THAT FOR 7 OF THE 12 BOXES RECEIVED, NONE OF THE CATHETERS HAD SLEEVES ON THEM AND THEY WOULD NEED REPLACEMENTS. PER FOLLOW UP VIA PHONE ON (B)(6)2023, IT WAS REPORTED THAT CUSTOMER UNDERSTANDS THAT THEY WERE SENT A SUBSTITUTE CATHETER BECAUSE THE ONES THEY NORMALLY USED WERE ON BACKORDER WITHOUT A RELEASE DATE. THEY ALSO STATED THAT ITEM #53814G LOT # JUGV1681 HAD NO HOLES IN THE TIP, AND THEY HAD SAMPLES OF THAT ONE AVAILABLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01480 | 2023-04-21 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7957505J |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, CLEAR LIQUID LEAKED FROM THE INSERTION SITE. THERE WERE TIMES WHEN THE INFUSION BAG DID NOT COLLAPSE FROM THE BEGINNING OF USE, AND UNCOMFORTABLE WAS FELT FROM THE BEGINNING. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING CATHETER WAS UNCONFIRMED, AS THE FAILURE COULD NOT BE REPRODUCED. ONE 5 FR D/L GROSHONG CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED SOME USE RESIDUES THROUGHOUT THE DEVICE, AND A STATLOCK STABILIZATION DEVICE WAS ATTACHED TO THE RETURNED GROSHONG CATHETER. NON-BD EXTENSION TUBING WAS RETURNED ATTACHED TO BOTH LUERS OF THE COMPLAINANT CATHETER. FUNCTIONAL TESTING OF PRESSURIZING THE CATHETER BY CLAMPING THE END CLOSED AND ATTEMPTING TO INFUSE WATER INTO THE DEVICE USING A 12 ML SYRINGE DID NOT REVEAL ANY LEAKS. FUNCTIONAL TESTING WAS ATTEMPTED WITH AND WITHOUT THE RETURNED EXTENSION TUBING, AND A LEAK COULD NOT BE IDENTIFIED FROM THE TESTING. BOTH LUMENS WERE PATENT TO INFUSION. BECAUSE THE LEAK COULD NOT BE IDENTIFIED, THE COMPLAINT OF A LEAKING CATHETER IS UNCONFIRMED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01335 | 2023-04-21 | 2023-01-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH, ACE 18GX2.25 | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1182250 |  | IT WAS REPORTED BY THE CUSTOMER, ACCUCATH CATHETER KINKING, PATIENT LINE DID NOT LAST AND HAD TO BE STUCK AGAIN. NO OTHER INFORMATION WAS PROVIDED. TWO DEVICES WERE RETURNED FOR EVALUATION. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE ANALYSIS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF KINKED IV CATHETERS WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS TWO 18GA X 2.25¿ ACCUCATH PERIPHERAL IV CATHETERS. USAGE RESIDUES WERE OBSERVED THROUGHOUT BOTH SAMPLES. THE FIRST SAMPLE (SAMPLE 1) EXHIBITED THREE SHARP KINKS OCCURRING FROM JUST DISTAL OF THE MOLDED JOINT TO APPROXIMATELY 2CM FROM THE MOLDED JOINT. THE SECOND SAMPLE (SAMPLE 2) EXHIBITED ONE SHARP KINK APPROXIMATELY 2.5CM FROM THE MOLDED JOINT. MICROSCOPIC INSPECTION OF BOTH SAMPLES CONFIRMED THE PRESENCE OF SHARP KINKS IN THE CATHETER SHAFTS. INSPECTION OF THE DISTAL ENDS OF THE CATHETERS REVEALED DEFORMATION OF BOTH TIPS. THE EDGES OF THE CATHETER TIPS WERE CURVED INWARD AND BUCKLED. THE CATHETER KINKING AND TIP DEFORMATIONS WERE CONSISTENT WITH CATHETER DAMAGE CAUSED BY ADVANCEMENT AGAINST RESISTANCE, SUCH AS INTO TISSUE. THE USE RESIDUES SUGGESTED THAT RESISTANCE WAS ENCOUNTERED DURING DEVICE USE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02776 | 2023-04-21 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE OPERATOR ERROR, USER RELATED (EG: INCORRECT CATHETER SIZE USED, INSUFFICIENT LUBRICANT APPLIED). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02786 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02739 | 2023-04-21 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01481 | 2023-04-21 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED "OUR SHERLOCK HAS TO BE HELD ONCE YOU CONNECT THE STYLIST. IF YOU DON'T HOLD THE TWO CLIPS TOGETHER WITH PRESSURE THAT'S THE ONLY WAY YOU CAN GET THE WHITE & YELLOW LINE OF THE EKG. 2 OF OUR PICC'S TODAY WE HAD TO LEAVE 2-3 CM OUT THE 2 PATIENTS WE HAD TO DO DRESSINGS ON HAD TO BE PULLED BACK BECAUSE OF HEART PALPITATIONS." ADDITIONAL INFORMATION RECEIVED 04/11/2023: IT WAS STATED THERE WAS NO PATIENT HARM. THIS REPORT ADDRESSES THE SECOND DEVICE AND PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01482 | 2023-04-21 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED "OUR SHERLOCK HAS TO BE HELD ONCE YOU CONNECT THE STYLIST. IF YOU DON'T HOLD THE TWO CLIPS TOGETHER WITH PRESSURE THAT'S THE ONLY WAY YOU CAN GET THE WHITE & YELLOW LINE OF THE EKG. 2 OF OUR PICC'S TODAY WE HAD TO LEAVE 2-3 CM OUT THE 2 PATIENTS WE HAD TO DO DRESSINGS ON HAD TO BE PULLED BACK BECAUSE OF HEART PALPITATIONS." ADDITIONAL INFORMATION RECEIVED 04/11/2023: IT WAS STATED THERE WAS NO PATIENT HARM. THIS REPORT ADDRESSES THE FIRST DEVICE AND PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01473 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 3FR SL PROVENA MIDLINE MAX BARRIER KIT W/BIOPATCH AND PROBE COVER | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | S4153108BDP |  | IT WAS REPORTED BY THE CUSTOMER, PLACED A MIDLINE AND IT WAS LEAKING THIS MORNING, SO IT WAS REPLACED. AFTER PULLING THE LINE THIS MORNING, SHE FLUSHED IT AND IT SEEMED TO BE WORKING WITH NO LEAKING THROUGHOUT THE LINE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01484 | 2023-04-21 | 2023-01-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK SLIDING POST TRICOT CRESCENT PAD INCLUDES SKIN PREP | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | PIC0220CE |  | IT WAS REPORTED THAT THE BRACKET HAS COME LOOSE ON THE PLATE. THE ATTACHMENT LOOSENS UPON PROVOCATION WITHOUT MUCH FORCE. NO OTHER INFORMATION WAS PROVIDED. (B)(6) 2023 - THREE DEVICES WERE RETURNED OR EVALUATION. THE THREE DEVICES EXHIBITED PLASTIC HOUSING THAT WAS EASILY PULLED BACK FROM THE BASE PAD OF THE STATLOCK. THIS REPORT ADDRESSES THE THIRD RETURNED DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF BRACKETS COMING LOOSE FROM THE BASE WITH LITTLE FORCE IS CONFIRMED AND WAS DETERMINED TO BE RELATED TO MANUFACTURING. THREE UNOPENED STATLOCK PICC PLUS CATHETER STABILIZATION DEVICES WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED NO USE RESIDUES THROUGHOUT THE DEVICES. A TACTILE EVALUATION OF THE DEVICES REVEALED THE PLASTIC HOUSING WAS EASILY PULLED BACK FROM THE BASE PAD OF THE STATLOCK DEVICE. A MICROSCOPIC OBSERVATION REVEALED SMALL AMOUNTS OF ADHESIVE ON THE THREE RETURNED SAMPLES. THE COMPLAINT OF THE BRACKETS COMING UP FROM THE BASE PAD OF THE STATLOCK DEVICE IS CONFIRMED AND LIKELY RELATED TO MANUFACTURING. THIS FAILURE MAY OCCUR DUE TO INSUFFICIENT ADHESIVE APPLIED DURING DEVICE ASSEMBLY. THE INVESTIGATION WAS FORWARDED TO THE MANUFACTURING FACILITY FOR FURTHER EVALUATION, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01483 | 2023-04-21 | 2023-01-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK SLIDING POST TRICOT CRESCENT PAD INCLUDES SKIN PREP | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | PIC0220CE |  | IT WAS REPORTED THAT THE BRACKET HAS COME LOOSE ON THE PLATE. THE ATTACHMENT LOOSENS UPON PROVOCATION WITHOUT MUCH FORCE. NO OTHER INFORMATION WAS PROVIDED. 04/13/2023 - THREE DEVICES WERE RETURNED OR EVALUATION. THE THREE DEVICES EXHIBITED PLASTIC HOUSING THAT WAS EASILY PULLED BACK FROM THE BASE PAD OF THE STATLOCK. THIS REPORT ADDRESSES THE SECOND RETURNED DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF BRACKETS COMING LOOSE FROM THE BASE WITH LITTLE FORCE IS CONFIRMED AND WAS DETERMINED TO BE RELATED TO MANUFACTURING. THREE UNOPENED STATLOCK PICC PLUS CATHETER STABILIZATION DEVICES WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED NO USE RESIDUES THROUGHOUT THE DEVICES. A TACTILE EVALUATION OF THE DEVICES REVEALED THE PLASTIC HOUSING WAS EASILY PULLED BACK FROM THE BASE PAD OF THE STATLOCK DEVICE. A MICROSCOPIC OBSERVATION REVEALED SMALL AMOUNTS OF ADHESIVE ON THE THREE RETURNED SAMPLES. THE COMPLAINT OF THE BRACKETS COMING UP FROM THE BASE PAD OF THE STATLOCK DEVICE IS CONFIRMED AND LIKELY RELATED TO MANUFACTURING. THIS FAILURE MAY OCCUR DUE TO INSUFFICIENT ADHESIVE APPLIED DURING DEVICE ASSEMBLY. THE INVESTIGATION WAS FORWARDED TO THE MANUFACTURING FACILITY FOR FURTHER EVALUATION, AND BD IS WORKING CLOSELY WITH THE MANUFACTURING FACILITY TO PREVENT RECURRENCE OF THE REPORTED EVENT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01485 | 2023-04-21 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D2 |  | IT WAS REPORTED BY THE CUSTOMER THAT "PICC LEAKING AT INSERTION SITE. THE NURSES STATED THE ARM LOOKED SLIGHTLY SWOLLEN IN THE MORNING SO THEY ORDERED A VASCULAR ULTRASOUND TO EVALUATE FOR A CLOT. THE ULTRASOUND WAS NEGATIVE. THE NURSE EVALUATED THE PATIENT AND FELT THERE MAY BE A BREAK IN THE PICC. ADDITIONAL INFORMATION RECEIVED 4/17/2023: IT WAS REPORTED BY THE CUSTOMER "PICC WAS REPORTED TO BE LEAKING AT SITE. WHEN FURTHER EVALUATED AFTER DISCONTINUING PICC FROM PATIENT, IT WAS NOTED TO HAVE A BREAK IN THE PICC AT 5/6 CM MARK." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01495 | 2023-04-21 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG+ SENSOR | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770358 |  | IT WAS REPORTED THE ECG TIP LOCATION IS NOT ACCURATE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01496 | 2023-04-21 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROVPPSOLO 4FDL MBP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1294108D4 |  | IT WAS REPORTED BY THE CUSTOMER THAT PICC COULD NOT BE USED FOR POWER INJECTION. ADDITIONAL INFORMATION RECEIVED 4/19/2023: THERE WAS NO HARM BUT PATIENT NEEDED TO GET STUCK WITH AN IV. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01500 | 2023-04-21 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405J |  | IT WAS REPORTED THAT WHEN THE SAFETY INTRODUCER NEEDLE WAS REMOVED, THE LOCK COVER WAS BROKEN AND DETACHED. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF SAFETY MECHANISM FAILURE TO ACTIVATE WAS CONFIRMED AND THE CAUSE APPEARED TO BE SUPPLIER RELATED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 21GA SAFETY INTRODUCER NEEDLE. THE PLASTIC SAFETY COLLAR HAD BEEN ADVANCED AND WAS RECEIVED DETACHED FROM THE NEEDLE. THE METAL SAFETY CLIP WAS PARTIALLY ADVANCED ALONG THE NEEDLE SHAFT BUT DID NOT COVER THE NEEDLE TIP. THE CLIP WAS MISALIGNED ON THE NEEDLE SHAFT. MICROSCOPIC INSPECTION OF THE SAFETY CONFIRMED THAT IT WAS MISALIGNED ON THE NEEDLE SHAFT. THE SAFETY MECHANISM FUNCTIONED AS INTENDED WHEN PROPERLY ASSEMBLED ON THE NEEDLE SHAFT. THE FAILURE OF THE SAFETY MECHANISM TO CAPTURE THE NEEDLE TIP WAS CAUSED BY THE MISALIGNED CLIP, WHICH APPEARED TO HAVE BEEN ASSEMBLED INCORRECTLY. THE BD SUPPLIER HAS BEEN NOTIFIED OF THIS EVENT AND COMPLAINTS WILL CONTINUE TO BE TRENDED AND MONITORED. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01499 | 2023-04-21 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | CK000821 |  | IT WAS REPORTED BY THE CUSTOMER THAT A 3 FR CATHETER PLACED IN FEBRUARY SPRUNG A LEAK AT THE CONNECTION BETWEEN THE HUB AND THE EXTERNAL LUMEN. THE LINE HAD TO BE REPLACED. ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023: IT WAS REPORTED THAT ON (B)(6) AT 2030 FLUID WAS NOTED BY BEDSIDE NURSE AND CONFIRMED BY VAT. CATHETER WAS REMOVED. NEW PICC PLACED (B)(6) 2023. PATIENT UNABLE TO RECEIVE CPN FOR 2 DAYS UNTIL NEW LINE PLACED. CATHETER WAS SECURED WITH SECUREACATH. NO PATIENT HARM OR SYMPTOMS RELATED TO THE FLUID LEAK. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAK IN THE EXTENSION TUBING IS CONFIRMED. ONE 3 FR S/L POWERPICC SV WAS RETURNED FOR EVALUATION. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE TUBING. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUES THROUGHOUT THE RETURNED CATHETER. A MICROSCOPIC OBSERVATION REVEALED A SPLIT IN THE TUBING JUST DISTAL OF THE PURPLE LUER. THE SPLIT EXHIBITED ¿C¿ SHAPED IMPRESSIONS LEADING TO THE FRACTURE SITE, AND BEACH MARKS WERE OBSERVED ALONG THE FRACTURE SURFACE. THE COMPLAINT OF A LEAKING CATHETER WHERE THE TUBING MEETS THE HUB IS CONFIRMED. AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01507 | 2023-04-21 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET BASIC TRAY WITH MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1174118 |  | IT WAS REPORTED BY THE CUSTOMER THAT "PICC TEAM WAS CALLED TO TROUBLE SHOOT A PICC LINE THAT WAS PLACED ON 4/1 BY PICC TEAM FOR LONG TERM ANTIBIOTICS. VAT RN WENT TO THE BEDSIDE TO ASSESS PICC LINE AND PICC LINE LOOKED FINE FROM THE OUTSIDE. BEDSIDE RN WAS ABLE TO ADMINISTER TPA BUT WAS NOT ABLE TO DRAW ANY OF IT BACK. THEY HAD ALSO CHANGED THE NEEDLESS CONNECTOR. PICC RN ORDERED A CHEST X-RAY AND A RIGHT HUMERUS X-RAY. RIGHT HUMERUS X-RAY SHOWS A KINKED PICC LINE AND THE PICC LINE WILL NEED TO BE REPLACED. TWO KINKS WERE FOUND-- ONE AT APPROXIMATELY 7 CM AND A SECOND AT ~8 CM. PICC WAS PLACED IN UPPER RIGHT BASILIC VEIN. CATHETER TRIM LENGTH WAS 38.5 AND 2.5 CM OF THE CATHETER WAS LEFT EXTERNAL." ADDITIONAL INFORMATION RECEIVED 5/23/2023: CATHETER PLACED RIGHT UPPER BASILIC VEIN. CATHETER TRIMMED TO 38.5 CM AND 2.5 CM LEFT EXTERNAL TO THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER KINK WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. AN AP RADIOGRAPH OF THE PATIENT¿S RIGHT HUMERUS WAS RETURNED FOR EVALUATION. A RIGHT-SIDED PICC, CONSISTENT WITH THE IMPLICATED 4FR S/L POWERPICC CATHETER, WAS SEEN IN THE RADIOGRAPH. TWO POSSIBLE KINKED AREAS WERE FOUND NEAR THE VEIN INSERTION SITE. POSSIBLE CONTRIBUTING FACTORS FOR A KINK IN THE CATHETER COULD INCLUDE THE ANGLE OF THE CATHETER AS IT ENTERS THE VEIN, CATHETER PLACEMENT, OR CATHETER SECUREMENT. THE IFU STATES, ¿AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT.¿ AND ¿AVOID SHARP OR ACUTE ANGLES DURING IMPLANTATION WHICH COULD COMPROMISE THE PATENCY OF THE CATHETER LUMEN.¿ A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE REPORT OF A KINKED CATHETER HAS BEEN DOCUMENTED AND WILL BE MONITORED AS PART OF COMPLAINT TRENDING. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01508 | 2023-04-21 | 2023-04-04 | I, n, j, u, r, y | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. DL POWERPICC INSERTED IN THE UPPER ARM." IT WAS REPORTED THIS OCCURRED WITH 2 DEVICES. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01505 | 2023-04-21 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT CAME TO THE HOSPITAL FOR TREATMENT OF BREAST CANCER. AT 11:00 AM ON (B)(6)2023, IT WAS FOUND THAT THE CENTRAL VENOUS CATHETER OF PERIPHERAL INTUBATION WAS PERMEATED 2CM AWAY FROM THE PUNCTURE POINT, AND A SMALL HOLE WAS FOUND ON THE CATHETER, WHICH WAS IMMEDIATELY STOPPED AND REPLACED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01509 | 2023-04-21 | 2023-04-04 | I, n, j, u, r, y | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE; DL POWERPICC INSERTED IN THE UPPER ARM." IT WAS REPORTED THIS OCCURRED WITH 2 DEVICES. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02748 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02751 | 2023-04-21 | 2023-03-30 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE PATIENT ALWAYS GETTING A URINARY TRACT INFECTION, ONCE A MONTH. THE PATIENT THOUGHT THEY CURRENTLY HAD ONE. THEY JUST WAITING ON RESULTS FROM DOCTOR OFFICE. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION IS UNKNOWN. AS PER FOLLOW UP VIA PHONE ON 18APR2023, IT WAS REPORTED THAT CUSTOMER WENT TO DOCTOR ON (B)(6) AND URINE SAMPLE WAS GIVEN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED WAS RETURNED FOR EVALUATION. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD NOT BE CHOSEN DUE TO INSUFFICIENT INFORMATION. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: [WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BEINJURED.] APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBERLATEX PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES]: BARDEX®SILVER LUBRICATH®FOLEYCATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD®SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO26EVERY 2FR. BALLOON CATHETER: ACCESSORIES: SYRINGE PREFILLED WITH STERILE WATER; WATER-SOLUBLE LUBRICANT; (NEITHER OF THEM MAYBE INCLUDED IN SOME PRODUCT VARIATIONS.); [INTENDED USE & EFFECT-EFFICACY]; THIS DEVICE IS AN INDWELLING CATHETER IN THEBLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE]: METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUMEOF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. TO DEFLATE BALLOON AND REMOVE CATHETER, INSERTA LUERTIP(NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. PRECAUTIONS FOR USE: WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE ANDREMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGEBY HANDS.[THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETERCANNOT BE REMOVED.] WHEN INSERTING CATHETER WITH ASTYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON.[IF CONTRAST MEDIUM IS USED,BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY.] DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS]: PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OFTHEBALLOON AND/OR THE CATHETERIS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDINGTO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.); BALLOON-RUPTURE METHOD: WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAYREMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. NON-RUPTURE METHOD: ATTACH LUER TIP SYRINGE TO THE INFLATIONVALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATIONWOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. F SITUATION WOULDN'T BE IMPROVED WITH3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5)IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. BALLOON-RUPTURE METHOD INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2)IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A)UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUMUNDER ULTRASONOGRAPHIC GUIDANCE. IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. MALFUNCTION AND ADVERSE EVENTS: MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE; DIFFICULTY OR FAILURE TO REMOVE THE DEVICE; OCCLUSION OF CATHETER INNER LUMENS; ENCRUSTATION; ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE; DEVICE DAMAGE DUE TO INAPPROPRIATE USE. ADVERSE EVENTS: URINARY-TRACT INFECTION; HEMORRHAGE, HEMATURIA; ALLERGY REACTION TO THE DEVICE; CALCULUS FORMATION; EDEMA; PAIN; DISCOMFORT; INJURY OF BLADDER OR URETHRAL; URETHRITIS, URINARY INCONTINENCE; RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]: STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02783 | 2023-04-21 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT NORMALLY USED THE INTERMITTENT CATHETER. PATIENT STATED THAT THEY DID NOT LIKE TO USE THE BURST PACKET AND THEY HAD TO CUT THEM OPEN. PATIENT ALSO SAID THEY FINDING THAT SOMETHING IN THEIR BODY WAS RESTRICTING URINE FLOW AFTER THE CATHETER WAS IN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02777 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01478 | 2023-04-21 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERLOC SAFETY INFUSION SET 20G X 0.75IN WITH Y-SITE | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | NA | 0672034 |  | IT WAS REPORTED BY THE CUSTOMER THEY HAVE A Y- SITE LEAKING FROM A POWERLOC NEEDLE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01479 | 2023-04-21 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1274108D5 |  | THE CUSTOMER REPORTED LEAKING FROM CLEAR T-PORT WHERE STYLET GOES THROUGH, NO PATIENT HARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01470 | 2023-04-21 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0607540 | 0607540 |  | IT WAS REPORTED THAT SHORTLY AFTER PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED A REACTION TO THE PORT. IT WAS FURTHER REPORTED THE PATIENT'S TUNGSTEN LEVELS WERE HIGH. REPORTEDLY, THERE IS NO ADDITIONAL INTERVENTION OR MEDICATION PRESCRIBED FOR TREATMENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: ADDITIONAL INFORMATION WAS RECEIVED, AND THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE REPORTABLE AS MALFUNCTION. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 04/2024). H11: H6 (PATIENT, RESULT, CONCLUSION). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02778 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALLS OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02744 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02759 | 2023-04-21 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT WATER LEAKED FROM THE INFLATION VALVE OF THE FOLEY CATHETER. THOUGH INJECTING DISTILLED WATER INTO THE BALLOON, THE WATER DID NOT ENTER AND CAME BACK AND LEAKED FROM THE SYRINGE SIDE OF THE VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02758 | 2023-04-21 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, FIXATION WATER IN THE BALLOON WAS DIFFICULT TO REMOVE FROM THE FOLEY CATHETER. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED UNDER 2 MINUTES WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS S THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02769 | 2023-04-21 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT PRIOR TO USE, A FOREIGN MATERIAL WAS FOUND ADHERE TO THE BALLOON AREA OF THE FOLEY TRAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02746 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND TO BE A CASCADING FAILURE OF AN EVENT PREVIOUSLY REPORTED. NOTE: THE CORRECT ZIP CODE UNDER TAB E IS (B)(6). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02785 | 2023-04-21 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THESE MAGIC 3 GO HYDROPHILIC MALE INTERMITTENT CATHETER DID NOT HAVE ENOUGH LUBE ON THEM AND DID NOT KNOW HOW TO USE IT AND THE PATIENT WAS REQUESTING ANOTHER PRODUCT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT/MISSING TRANSLATION; MISSING INSTRUCTIONS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON'T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02741 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT WATER DID NOT ENTER THE FOLEY CATHETER THROUGH THE INFLATION VALVE. ON ATTEMPTING TO INJECT WATER THROUGH THE INFLATION VALVE USING THE SYRINGE FROM THE TRAY, THE WATER LEAKED AND DID NOT ENTER THE CATHETER. HOWEVER, WHEN ANOTHER SYRINGE STORED IN THE HOSPITAL WAS USED, WATER WAS ABLE TO BE INJECTED WITHOUT ANY ISSUES. ONLY THE SYRINGE WAS RECOVERED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02756 | 2023-04-21 | 2023-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER WAS UNABLE TO REMOVE FROM THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02760 | 2023-04-21 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT FOLEY CATHETER WAS DAMAGED AND URINE CAME OUT OF THE CATHETER AFTER INDWELLING, NO BALLOON RUPTURE OR TEAR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02747 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, INCORRECT WIRE PRESSING TECHNIQUE, INCORRECT STRIPPING TECHNIQUE ETC. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02753 | 2023-04-21 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119416M | 119416M |  | IT WAS REPORTED THAT THE SAME LOT NUMBER FOLEY CATHETER FAILED TWICE DURING USE. THERE WAS PROBLEM WITH FLUSHING. ALSO STATED THAT THEY HAD ISOLATED 5 TRAYS OF THIS LOT NUMBER AND 2 WAS WASTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02749 | 2023-04-21 | 2023-03-29 | I, n, j, u, r, y | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE OPERATOR ERROR, USER RELATED (EG: INCORRECT CATHETER SIZE USED, INSUFFICIENT LUBRICANT APPLIED). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02750 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02781 | 2023-04-21 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THESE INTERMITTENT CATHETER CAUSED THE PATIENT PAIN, FELT SCRATCHY, AND AND DID NOT COMPLETELY EMPTY THE PATIENT BLADDER. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿INNER DIAMETER OF DRAINAGE LUMEN TOO SMALL/KINKED OR PARTIALLY COLLAPSED DRAINAGE LUMEN". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02762 | 2023-04-21 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT AFTER INSERTION OF THE CATHETER, NO URINE OUTFLOW WAS OBSERVED. WHEN THE PRODUCT WAS INSPECTED USING WATER AFTER REMOVAL, THE WATER DIDN'T FLOW. THE EVENT OCCURRED JUST AFTER THE INSERTION AND THE URINE FLOWED AFTER REPLACEMENT. THERE WAS NO KINK OR BENT ON THE CATHETER AND NO PATIENT'S URINE STONE. NO ABNORMALITY OBSERVED IN THE CONDITION OF PATIENT'S BLADDER. TO RESOLVE THE PROBLEM, COMPRESSED THE BLADDER FROM THE BODY SURFACE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02767 | 2023-04-21 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123518A |  | IT WAS REPORTED THAT THE BALLOON DID NOT DEFLATE DURING THE BALLOON CHECK PRIOR TO PLACEMENT. AS PER THE INQUIRY SHEET ATTACHED, 10ML OF WATER WAS ASPIRATED BY THE SYRINGE. USUALLY BALLOON CAN BE DEFLATED WITHIN FEW SECONDS. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARDIA® SILICONE-COATED FOLEY CATHETER IS A BALLOON CATHETER. THERE ARE DIFFERENT TYPE VARIATIONS, FOR EXAMPLE, 2WAY AND 3WAY PRODUCTS. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILICONE COATING <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 30 EVERY 2FR 1. BALLOON CATHETER [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (9) WHEN IRRIGATING, OPEN THE CAP OF IRRIGATION FUNNEL AND ATTACH IRRIGATION LINE OR TUBE UNDER ASEPTIC TECHNIQUE. AFTER USE, CLOSE THE CAP. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02755 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514CE |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02765 | 2023-04-21 | 2023-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE CATHETER BALLOON WAS FOUND LEAKING WHILE CHANGING THE CYSTOSTOMY TUBE FOR THE PATIENT AS PER THE FOLLOW UP INFORMATION RECEIVED ON 10APR2023, STATED THAT IT WAS A WATER LEAK AND THERE WAS A DAMAGE ON THE BALLOON BUT OF WHICH KIND COULD NOT BE DETERMINED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY". THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02779 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02740 | 2023-04-21 | 2023-03-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT AT 6:36 AM ON (B)(6) 2023 AS THE BALLOON BURST, THE CATHETER FELL OUT AT THE URETHRAL ORIFICE AND THEN THE CATHETER WAS PULLED OUT ACCORDING TO THE DOCTOR'S ADVICE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01476 | 2023-04-21 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER, 5F DL, 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | CK000590A |  | IT WAS REPORTED BY THE CUSTOMER THAT THE CLINICIAN WAS PREPARING PICC, DREW BACK GUIDEWIRE TO TRIM AND NOTICED A RED STRING/ THREAD ATTACHED WITH THE GUIDEWIRE FROM INSIDE THE PICC. (FOREIGN OBJECT). NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02780 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02784 | 2023-04-21 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THESE MAGIC 3 GO HYDROPHILIC MALE INTERMITTENT CATHETER DID NOT HAVE ENOUGH LUBE ON THEM AND DID NOT KNOW HOW TO USE IT AND THE PATIENT WAS REQUESTING ANOTHER PRODUCT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE. INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02768 | 2023-04-21 | 2023-04-01 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD VAGINAL DRYNESS AND A URINARY TRACT INFECTION FROM THE PUREWICK. LMS REPRESENTATIVE ADVISED THAT PATIENT NEED TO SEE THEIR DOCTOR. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS: ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER ¿ HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. ¿ EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. ¿ PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02743 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETERS WERE STOPPED UP AND PATIENT HAD TO CHANGE THEM SOONER THAN THEY WERE SUPPOSED TO. CUSTOMER STATED THAT FIRST ONE LASTED FOR THREE WEEKS, THE SECOND ONE LASTED FOR TWO WEEKS AND THEY HAD 3RD ONE IN NOW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02745 | 2023-04-21 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 06APR2023, STATED THAT THE COMPLETE LOT NUMBER SHOULD BE MYGN2859. STATED THAT THEY HAD SEEN SKIN LESIONS OFTEN ON MEN WHEN ASKED REGARDING NEGATIVE CONSEQUENCES FOR THE PATIENT. IT WAS UNCLEAR WHETHER IT IS THE FAULT OF THE BARDIA CATHETERS, BUT THEY FELT THAT IT HAD BECOME MORE SO IN RECENT YEARS. THIS APPLIED TO SEVERAL DIFFERENT PATIENTS BUT THEY LIVE IN NURSING HOMES. SOME WERE SLIM AND SOME HAD HIGHER BODY MASS INDEX. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND TO BE A CASCADING FAILURE OF AN EVENT PREVIOUSLY REPORTED. NOTE: THE CORRECT ZIP CODE UNDER TAB E IS 182 33. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16784948 | 2023-04-21 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | THREE DIFFERENT PICCS WERE REQUIRED TO BE ATTEMPTED TO BE PLACED ON A PATIENT SECONDARY TO THE LINES LEAKING AT THE RUBBER STOPPER WHILE BEING FLUSHED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01469 | 2023-04-21 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393270 | 5393270 |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT PROCEDURE, THE CATHETER ALLEGEDLY CHANGED TO MILKY COLOUR. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL SAMPLE WAS NOT RETURNED FOR EVALUATION. THREE ELECTRONIC PHOTOS WERE PROVIDED FOR REVIEW. BOTH THE EXTENSION LEGS WERE NOTED TO BE DISCOLORED, DEFORMED AND BULGED. MILKY WHITE OPACIFICATION WAS OBSERVED ON BOTH OF THE EXTENSION LEG. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED CATHETER OPACIFICATION AND IDENTIFIED DISCOLORATION AND STRETCHED ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01467 | 2023-04-21 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 7707540 | 7707540 |  | IT WAS REPORTED THAT SOMETIMES POST PORT PLACEMENT VIA INTERNAL JUGULAR VEIN, THE PORT CHAMBER ALLEGEDLY HAD A COMPLETE BLOCKAGE DURING SECOND CYCLE OF CHEMOTHERAPY. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 01/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16785579 | 2023-04-21 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 1174108D5 | 1174108D5 |  | WHEN INSERTING PICC RUBBER STOPPER LEAKED SALINE WHILE FLUSHING. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16785533 | 2023-04-21 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 1174108D5 | 1174108D5 |  | RUBBER TIP ON PICC WHERE WIRE ENTERS LEAKS SALINE WHEN FLUSHED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16785076 | 2023-04-21 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | RUBBER LUMEN LEAKING WHERE THE WIRE FOR THE PICC IS. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16784895 | 2023-04-21 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | SEVERAL REPORTS OF PICC LINES LEAKING AT THE RUBBER-LIKE STOPPER UPON ATTEMPTS TO FLUSH. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16784739 | 2023-04-21 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 1174108D5 | 1174108D5 |  | RUBBER STOPPER ON SINGLE-LUMEN PICC LEAKS SALINE WHEN FLUSHED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16784738 | 2023-04-21 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | 1174108D5 | 1174108D5 |  | SALINE LEAKS FROM RUBBER STOPPER IN PICC LINE DURING INSERTION PROCEDURE. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16784352 | 2023-04-21 | 2012-05-29 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SOLO POWER PICC-DOUBLE LUMEN | CATHETER, INTRAVASCULAR, THERAPEUTIC, LONG - TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. |  |  |  | DURING PLACEMENT OF POWER PICC (DOUBLE LUMEN), THE GUIDEWIRE DID NOT WANT TO COME OUT OF LINE WITH GENTLE PULLING BACK AND COULD NOT BE RETHREAD. NURSE HAD ATTEMPTED TO PULL BACK OUT OF THE INTRODUCER TO ENSURE THAT IT WAS NOT POSITIONAL AND IT WAS NOT. CATHETER WAS ALSO FLUSHED WITH 10-20 ML OF NORMAL SALINE TO ENSURE THAT FRICTION WASN'T INVOLVED AND IT WAS NOT. THE DEVICE WAS REMOVED FROM THE PATIENT AND REPLACED WITH A NEW PICC. THERE WAS NO PATIENT INJURY. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02742 | 2023-04-21 | 2023-03-28 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175818 |  | IT WAS REPORTED THAT THE GP HAVE SENT PRESCRIPTION FOR FOLEY CATHETER. PATIENT WANTED TO MAKE SURE THE PRESCRIPTION WAS NOT FOR THE BARD LUB SILICON CATHETER AS THEY HAVE STATED THAT THIS CATHETER CAUSED THEM AN INFECTION RESULTING IN THEM HAVING TO TAKE ANTIBIOTICS. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 30MAR2023, STATED THAT THE EXACT INFECTION EXPERIENCED BY THE CUSTOMER WAS NOT REPORTED. ANTIBIOTICS WERE GIVEN TO THE PATIENT. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 03APR2023, STATED THAT THE INFECTION WAS CAUSED BY CATHETER AND MEDICATION WAS PRESCRIBED BY DOCTOR. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
16783780 | 2023-04-21 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET LOT#REGZ0135 NOTED SMALL LEAK DURING FLUSHING. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00272 | 2023-04-21 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DORADO | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | DR4054 | DR4054 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON WAS ALLEGEDLY PULLED OFF. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. ONE PHOTO WAS PROVIDED AND REVIEWED. THE PHOTO SHOWS THE BALLOON OF THE DORADO DEVICE. IT IS SEEN BROKEN AND PULLED AWAY FROM THE CATHETER SHAFT AT THE BOND JOINT WITH THE INNER GUIDEWIRE LUMEN EXPOSED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED BREAK AS A BREAK AT THE BOND JOINT OF THE DEVICE IS SEEN FROM THE PHOTO REVIEW. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED BREAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRY DATE: 08/2024). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01475 | 2023-04-21 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT SOME TIME POST PORT PLACEMENT, THE CATHETER ALLEGEDLY CRACKED. IT WAS FURTHER REPORTED THAT THE MEDICINE ALLEGEDLY LEAKED SUBCUTANEOUSLY. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED ON THE ATTACHED CATHETER. THE EDGES OF THE BREAK ON THE CATHETER WERE NOTED TO BE UNEVEN. UPON INFUSION, A LEAK FROM THE PARTIAL CIRCUMFERENTIAL BREAK WAS OBSERVED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND LEAK ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00271 | 2023-04-21 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE OF A CALCIFIED TARGET LESION IN THE POPLITEAL ARTERY, THE TIP OF THE CATHETER ALLEGEDLY SEPARATED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 10/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02775 | 2023-04-21 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT CATHETER SPONTANEOUSLY WITHDREW AND UPON CHECKING IT, THERE WAS A PINHOLE AROUND 20 CM FROM THE TIP. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02816 | 2023-04-24 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO POOR COATING THAT CAN LEADING CATHETER TO PERFORATION OR EROSION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE THE LATEX. VISUALLY INSPECTED THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. -USE LUER TIP SYRINGE TO INFLATE WITH STATE 10ML OF STERILE WATER OR - FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. ALLOW THE PRESSURE WITHIN TO BALLOON FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02818 | 2023-04-24 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02814 | 2023-04-24 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO POOR COATING THAT CAN LEADING CATHETER TO PERFORATION OR EROSION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE THE LATEX. VISUALLY INSPECTED THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. -USE LUER TIP SYRINGE TO INFLATE WITH STATE 10ML OF STERILE WATER OR - FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. ALLOW THE PRESSURE WITHIN TO BALLOON FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02815 | 2023-04-24 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123512A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12, 14 AND 16. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGU1645, BATCH# MYFX4463 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLEUM BASE. THEY WILL DAMAGE THE LATEX. VISUALLY INSPECTED THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. -USE LUER TIP SYRINGE TO INFLATE WITH STATE 10ML OF STERILE WATER OR - FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE "STICK" IN THE VALVE. ALLOW THE PRESSURE WITHIN TO BALLOON FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02817 | 2023-04-24 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123514A |  | IT WAS REPORTED THAT THE CUSTOMER HAD SEEN A DECREASED QUALITY OF THE BARDIA SILICONE ELASTOMER FOLEY CATHETERS FOR A LONG TIME. STATED THAT THE CATHETER FALL OUT (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THE INCIDENT OCCURRED DURING USE. THE CUSTOMER EXPERIENCED MAJOR PROBLEMS WITH BARDIA CATHETERS IN BOTH SIZES 12 AND 14. THE CATHETERS COULD RARELY STAY FOR 12 WEEKS. THEY OFTEN GET STUCK (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). ALSO STATED THAT THE CATHETERS SLIPPED OUT. THE CUSTOMER ALSO EXPERIENCE THAT THE CATHETERS OFTEN CAUSE DAMAGE (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). SOMETIMES THE CUFF WAS BROKEN, WHICH COULD INDICATE SOMETHING ELSE (BATCH# MYGQ5399,BATCH# MYGQ5394 AND BATCH# UNK). BUT, THE CATHETERS ALSO CAME OUT WITH A CUFF (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THIS APPLIES TO ALMOST EVERYONE WITH BARDIA CATHETERS. IT WAS NOT A SPECIAL BATCH, IT APPLIED TO EVERYONE AND FOR A LONG TIME. THE CUSTOMER STATED THAT THEY SHOULD BE ABLE TO TAKE ANY AS A PRODUCT SAMPLE AS THIS HAD BEEN A PROBLEM FOR A VERY LONG TIME. THERE WAS NO SPECIAL CARTON OR BATCH NUMBER. ALSO STATED THAT THEY COULD SEE A GENERAL DETERIORATION IN BARDIA CATHETERS (BATCH# MYGQ5399, BATCH# MYGQ5394 AND BATCH# UNK). THEY ALSO DON'T SEE THE SAME PROBLEM WHEN THEY HAD SILICONE CATHETERS OR TIEMANN CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00280 | 2023-04-24 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQF75124 | CQF75124 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE CONTRAST WAS ALLEGEDLY LEAKING OUT OF THE TIP OF THE BALLOON FURTHEST FROM THE SHEATH. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE CONQUEST 40 PTA DILATATION CATHETER WAS RETURNED FOR EVALUATION. NO SPECIFIC ANOMALIES WERE NOTED DURING THE VISUAL EVALUATION. ON THE FUNCTIONAL TESTING THE RETURNED BALLOON WAS INFLATED WITH IN-HOUSE PRESTO INFLATION DEVICE AND THE DEVICE WAS LEAKING FROM THE DISTAL TIP OF THE BALLOON, FURTHER THE BALLOON FIBERS WERE STRIPPED AND UNDER MICROSCOPIC OBSERVATION A PINHOLE RUPTURE WAS NOTED. NO OTHER FUNCTIONAL TESTING PERFORMED. AS THE DEVICE LEAKS DURING THE FUNCTIONAL TESTING AND PIN-HOLE BALLOON RUPTURE WAS NOTED UNDER MICROSCOPIC OBSERVATION. THEREFORE, INVESTIGATION WAS CONFIRMED FOR THE REPORTED LEAKS AND IDENTIFIED PIN-HOLE BALLOON RUPTURE. A DEFINITIVE ROOT CAUSE FOR THE REPORTED LEAKS AND IDENTIFIED PIN-HOLE BALLOON RUPTURE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 07/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00279 | 2023-04-24 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8088 | VA8088 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE VIA DIALYSIS ACCESS GRAFT IN LEFT ARM, THE DEVICE ALLEGEDLY GOT CAUGHT IN THE TIP OF THE SHEATH. IT WAS FURTHER REPORTED THAT THE DEVICE ALLEGEDLY RIPPED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 11/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02811 | 2023-04-24 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53818G |  | IT WAS REPORTED THAT THE HYDRO SIL GO CATHETERS, SOME OF WHICH HAD NO LUBRICATION OR LITTLE TO NONE, WERE VERY UNCOMFORTABLE TO USE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "MECHANICAL FAILURE". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02813 | 2023-04-24 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT BEFORE USE, THERE WAS A CRACK AT THE TIP OF THE CATHETER. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 03APR2023, STATED THAT THE SAMPLES RECEIVED WITH DIFFERENT SERIAL# ((B)(6)). | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND TO BE A DUPLICATE OF AN EVENT PREVIOUSLY REPORTED. H11: THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02821 | 2023-04-24 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CATHETER WOULD NOT BALLOON AND BUMP OUTSIDE OF CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02823 | 2023-04-24 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CATHETER STABILIZATION DEVICE | CATHETER STABILIZATION DEVICE | C.R. BARD, INC. (COVINGTON) -1018233 | FOL0102 | FOL0102 |  | IT WAS REPORTED THAT THE PATIENT WAS HAVING TROUBLE WITH A 16FR TELEFLEX CATHETER AND FOLEY STATLOCK. REPRESENTATIVE RESPONDED THAT THE STATLOCK FOLEY STABILIZATION DEVICE ACCOMMODATES LATEX 8-22 FR. AND SILICONE 8-26FR. FOLEY CATHETERS. SPOKE TO HOME NURSE AND THEY STATED THAT TRIED MULTIPLE FOLEY STATLOCK FROM THE SAME BOX, SO ASSUMED THEY WERE ALL ONE LOT. THEY WOULD NOT CLAMP SHUT. ADVISED IT WAS POSSIBLE THE WALL THICKNESS OF THE COMPETITOR'S CATHETER WAS SUCH THEIR STATLOCK WOULD NOT ACCOMMODATE IT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02820 | 2023-04-24 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | WIDEBAND¿ SELF-ADHERING MALE EXTERNAL CATHETER | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 36302 | 36302 |  | IT WAS REPORTED THAT THE MALE EXTERNAL CATHETER WOULD HOLD FOR UP TO ONE HOUR AND LATER WOULD FALL OFF. ALSO STATED THAT THE CATHETERS WERE STICKING. THE CUSTOMER REPLACED THE CATHETERS WITH A DIFFERENT LOT. PER NOTIFICATION RECEIVED FROM INVESTIGATOR ON 03APR2023, IT WAS STATED THAT THE STRONG ADHESIVE THAT WAS FOUND DURING THE SAMPLE EVALUATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02825 | 2023-04-24 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT IT WAS DIFFICULT TO DEFLATE DURING THE PRETEST BEFORE USE IN PATIENTS. ABOUT A 7ML OF WATER WAS ASPIRATED BY SYRINGE, THE BALLOON WAS DEFLATED COMPLETELY. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) WITH THE RETURNED SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 2 MINUTES WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01514 | 2023-04-24 | 2022-12-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA SLIM-CATH DIALYSIS CATHETER KIT 12F X 20CM | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5553200 |  | IT WAS REPORTED THAT GUIDEWIRE GOT STUCK WHEN INSERTING CATHETER. A NEW CATHETER WAS USED AND INSERTED. (B)(6) 2023 - THE RETURNED GUIDEWIRE WAS FOUND ELONGATED WITH THE INNER CORE WIRES PROTRUDING FROM BETWEEN THE OUTER COIL WIRES. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, COMPLAINT AND BATCH HISTORY, APPLICABLE PREVIOUS INVESTIGATION(S), APPLICABLE MANUFACTURING RECORDS, PHOTO ANALYSIS AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GUIDEWIRE WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. A PHOTOGRAPH OF THE IMPLICATED 70CM GUIDEWIRE WAS RETURNED FOR EVALUATION OF THIS COMPLAINT. THE OUTER CORE WIRE ON THE GUIDEWIRE WAS ELONGATED NEAR THE DISTAL J-TIP END. THE INNER CORE WIRE WAS SEEN PROTRUDING FROM BETWEEN THE COILS OF THE OUTER COIL WIRE. ELONGATION OF THE COIL WIRE CAN OCCUR WHEN THE INNER CORE WIRE IS EITHER BROKEN OR SHEARED. THE EXACT ROOT CAUSE OF THE DAMAGE COULD NOT BE DETERMINED FROM THE PHOTOGRAPH ALONE. MICROSCOPIC EXAMINATION AND DIMENSIONAL ANALYSIS COULD NOT BE CONDUCTED WITHOUT THE PHYSICAL SAMPLE. POSSIBLE CONTRIBUTING FACTORS COULD INCLUDE THE INSERTION TECHNIQUE, DIMENSIONAL INCOMPATIBILITY BETWEEN INTERFACING DEVICES, OR THE GUIDEWIRE BEING RETRACTED THROUGH THE NEEDLE. NORMAL MOVEMENT OF THE GUIDEWIRE IS AWAY FROM THE SHARPENED BEVEL AND WILL NOT DAMAGE THE WIRE. IF THE GUIDEWIRE DIRECTION IS REVERSED, THE GUIDEWIRE IS THEN PULLED AGAINST THE SHARPENED EDGE OF THE NEEDLE BEVEL AND CAN CAUSE SHEARING DAMAGE OF THE WIRE AND/OR CAUSE THE WIRE TO BECOME STICK WITHIN THE NEEDLE. IF THE GUIDEWIRE MUST BE WITHDRAWN WHILE THE NEEDLE IS INSERTED, REMOVE BOTH THE NEEDLE AND WIRE AS A UNIT TO PREVENT THE NEEDLE FROM DAMAGING OR SHEARING THE GUIDEWIRE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11 | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02827 | 2023-04-24 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI12 |  | IT WAS REPORTED THAT THERE WAS A PINHOLE IN THE FOLEY CATHETER AND WATER LEAKED FROM THERE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02828 | 2023-04-24 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166L14 |  | IT WAS REPORTED THAT THE BALLOON BURST. WHEN ATTEMPTING TO REMOVE THE FOLEY CATHETER, IT WAS FOUND WITH THE BALLOON TORN. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 12APR2023, NO MISSING PIECES WERE REPORTED. HOWEVER, SINCE THE SAMPLE HAS NOT YET REACHED US, WE HAVE NOT YET CONFIRMED IF THERE ARE ANY MISSING PIECES OR NOT. IF WE OBTAIN ADDITIONAL INFORMATION OF THE MISSING PIECES, WE WILL LET YOU KNOW. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02829 | 2023-04-24 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT STERILIZED WATER DID NOT COME OUT DURING THE PRETEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01515 | 2023-04-24 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG (TPS) STYLET 5F BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1275118 |  | IT WAS REPORTED, PICC WAS CALLED TO TROUBLE SHOOT PICC LINE. BEDSIDE RN WAS HAVING A DIFFICULT INSTILLING THE TPA, SO AN X-RAY OF THE CHEST AND THE ARM WAS ORDERED. THE X-RAY OF THE ARM SHOWED A KINK IN THE LINE AND THAT WAS THE REASON WHY THE BEDSIDE RN COULD NOT INSTILL THE TPA. PICC RN PLACED A STILETTO AND AN ACCUCATH IN HER RIGHT UPPER EXTREMITY. I SPOKE WITH NP, AND TOLD HER THAT IF CENTRAL ACCESS IS NEEDED, WE COULD PLACE A PICC LINE ON THE LEFT SIDE FOR HER. PICC WAS PLACED IN RIGHT UPPER BASILIC ON (B)(6) 2023 AND REMOVED ON (B)(6) 2023. CATHETER WAS 49 CM LONG AND 2CM OF THE CATHETER WAS LEFT EXTERNAL. KINK OCCURRED AT ~5 CM MARK. NO POWER INJECTION SCANS WERE PERFORMED ON THIS PATIENT. NEEDED TO REPLACE LINE WITH PIV X2, EXTRA UNNECESSARY POKES AND PAIN. STAFF ATTEMPTED TPA X2 ON (B)(6) 2023 AT 0340 AND 1230, UNNECESSARY MEDICATIONS GIVEN. PATIENT ANXIETY AND WORRY REGARDING SOMETHING HAPPENING AGAIN TO NEW LINE.". | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02808 | 2023-04-24 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT WATER LEAKED FROM EITHER THE DRAIN BAG OR THE CATHETER. DETAILS ARE IN THE PROCESS OF CONFIRMING. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 07APR2023, THERE WERE NO OBVIOUS VISIBLE DEFECTS TO THE RETURNED SAMPLE (BOTH CATHETER AND BAG). WE COULD NOT DETERMINE THE LEAK LOCATION. THERE COULD HAVE BEEN A LEAK FROM EITHER, SO BOTH THE DRAIN BAG AND THE CATHETER SHOULD BE INSPECTED, THE LEAK LOCATION IS ONE.HOWEVER, WE CANNOT CONFIRM THE EXACT LOCATION OF THE LEAK.PLEASE INVESTIGATE BOTH THE DRAIN BAG AND THE CATHETER AND DETERMINE THE EXACT LOCATION AND CAUSE OF THE LEAK BY THE INSPECTION. | THE REPORTED EVENT WAS NOT CONFIRMED BECAUSE DEVICE MEETS SPECIFICATIONS AND THE REPORTED FAILURE COULD NOT BE REPRODUCED. VISUAL AND FUNCTIONAL INSPECTION OBSERVED THE EXTERIOR AND FUNCTIONALITY OF THE SAMPLE AND DID NOT FIND ANY EVIDENCE OF A FAILURE THAT WOULD SUPPORT THE REPORTED EVENT. NO ABNORMALITIES ON BOTH CATHETER AND DRAINAGE BAG. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS NOT CONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]: METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] APPLICABLE PATIENTS: · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS]: METHOD FOR USE: DO NOT REUSE. DO NOT RESTERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] APPLICABLE PATIENTS: PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX; PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES]: BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL: BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING; THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 14FR, 16FR. BALLOON CATHETER WITH TEMPERATURE-SENSING. ACCESSORIES: CLOSED DRAINAGE BAG; (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.); SYRINGE PREFILLED WITH STERILE WATER; WATER SOLUBLE LUBRICANT; ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION; TWEEZERS; GAUZE PADS; WATERPROOF SHEET; COTTON BALLS; GLOVES; STATLOCK® FOLEY. [INTENDED USE & EFFECT- EFFICACY]: THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE]: METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. PRECAUTIONS FOR USE: TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT: OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. UNKINK TUBING. HOW TO DISCONNECT CATHETER FROM TUBING: CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY.] DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING: REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING: DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS]: PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS): EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. IMPORTANT PRECAUTIONS: WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.); BALLOON-RUPTURE METHOD. WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. NON-RUPTURE METHOD: ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) BALLOON-RUPTURE METHOD: INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) MRI SAFETY: BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. MALFUNCTION AND ADVERSE EVENTS: MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE; DIFFICULTY OR FAILURE TO REMOVE THE DEVICE; OCCLUSION OF CATHETER INNER LUMENS; ENCRUSTATION; ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE; DEVICE DAMAGE DUE TO INAPPROPRIATE USE; FAILURE TO MEASURE TEMPERATURE; IMPROPER TEMPERATURE INDICATION. ADVERSE EVENTS: URINARY-TRACT INFECTION; HEMORRHAGE, HEMATURIA; ALLERGY REACTION TO THE DEVICE; CALCULUS FORMATION; EDEMA; PAIN; DISCOMFORT; INJURY OF BLADDER OR URETHRAL; URETHRITIS, URINARY INCONTINENCE; RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]: STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02807 | 2023-04-24 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT IMMEDIATELY AFTER PLACEMENT, THE CATHETER FELL OFF ALTHOUGH NO ISSUE DURING A PRETEST. AS PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL ON (B)(6) 2023, THE BALLOON RUPTURE WAS OBSERVED TO THE RETURNED SAMPLE. | THE REPORTED EVENT WAS CONFIRMED - CAUSE UNKNOWN. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. VISUAL INSPECTION NOTED IRREGULAR BURST WITHOUT MISSING PIECES. INTRODUCED WATER INTO THE INFLATION LUMEN AND DID NOT EXPERIENCE ANY OBSTRUCTIONS TO IMPEDE FLOW. THE EXACT CAUSE ON HOW AND WHEN PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, WATERPROOF SHEET, TWEEZERS, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE TWO TYPES OF CATHETER, TWO WAY AND THREE WAY, AND SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE CATHETER AND/OR BAG INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. ALSO, THE STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES AN INDWELLING BLADDER CATHETER AND A URINARY DRAINAGE BAG THAT ARE USED FOR URINARY DRAINAGE AND BLADDER IRRIGATION. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) USING 3WAY CATHETER, CAPPED ADAPTER ATTACHED TO IRRIGATION LUMEN IS DESIGNED TO CONNECT CATHETER WITH BLADDER IRRIGATION LINE OR TUBE. WHILE IN USE, OPEN THE CAP AND ATTACH IRRIGATION LINE OR TUBE TO IT USING ASEPTIC TECHNIQUE. AFTER USE, CLEANSE THE OPENING OF THE LUMEN AND CLOSE THE CAP. (10) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (11) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (12) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTIONS: D,F,H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02805 | 2023-04-24 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CATHETER WOULD NOT BALLOON AND BUMP OUTSIDE OF CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02803 | 2023-04-24 | 2023-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT AFTER INSERTING, THE LUMEN WAS OCCLUDED, AND THE FOLEY CATHETER WAS CUT OFF AND THE DRAINAGE BECAME FLUENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02810 | 2023-04-24 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE TOO FLIMSY AND TOO SHORT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INDICATIONS FOR USE THE MAGIC3 GO¿ INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT FEMALE PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. ADVERSE REACTIONS: URINARY TRACT INFECTION, BLEEDING FROM THE URETHRA, IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS.¿. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02789 | 2023-04-24 | 2023-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HEMATURIA FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 2551H22 |  | IT WAS REPORTED THAT FOLEY CATHETER WAS DIFFICULT TO INFLATE DUE TO THE VALVE DEFECT. THE CATHETER WAS REPLACED WITH 2551HT24. PER NOTIFICATION RECEIVED FROM INVESTIGATOR ON 01APR2023, IT WAS STATED THAT VISUAL EVALUATION NOTED A HOLE IN THE BIFURCATION BETWEEN THE DRAINAGE AND IRRIGATION ARM. | THE REPORTED EVENT IS CONFIRMED, MANUFACTURING RELATED. THE USED THREE-WAY FOLEY CATHETER WAS RETURNED WITHOUT THE ORIGINAL PACKAGING. PHOTO SAMPLES WERE SUBMITTED SHOWING THE BALLOON DEFLATED, HOWEVER, COULD NOT BE EVALUATED FOR THE REPORTED FAILURE. VISUAL EVALUATION NOTED A HOLE IN THE BIFURCATION BETWEEN THE DRAINAGE AND IRRIGATION ARM DUE TO WEBBING; THE DEVICE DOES NOT MEET SPECIFICATION, AS IT DOES NOT MEET THE WEBBING - MAXIMUM ACCEPTABLE SECTION OF IP 25.1 WHICH STATES " IF THE WEBBING HAS A PINHOLE IT IS NOT ACCEPTABLE", RESIDUE WAS ALSO NOTED ON THE CATHETER FROM BEING USED ON THE PATIENT. NO VISIBLE DEFECTS WERE NOTED ON THE INFLATION VALVE. THE BALLOON WAS SUCCESSFULLY INFLATED WITH 50CC DI WATER USING A LUER LOCK SYRINGE. THE BALLOON WAS INFLATED CONCENTRICALLY AND PASSIVELY DEFLATED WITHOUT ISSUES. A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE, "PROGRAM ERROR". AS NO PATIENT INVOLVEMENT WAS REPORTED THE DEVICE WAS NOT USED, HOWEVER IT IS INTENDED FOR TREATMENT PURPOSES. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABEL/PACK REVIEW IS NOT REQUIRED AS A REVIEW OF THE LABEL COULD NOT HAVE PREVENTED THE REPORTED EVENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02790 | 2023-04-24 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS INFLATED DURING TESTING, BUT COULD NOT BE DEFLATED AFTERWARDS. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN 3 MINUTES AND 35 SECONDS WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02791 | 2023-04-24 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 119314 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON WAS INFLATED DURING TESTING, BUT COULD NOT BE DEFLATED AFTERWARDS. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INADEQUATE PRESSURE ON THE MACHINE TO PUNCH. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16794579 | 2023-04-24 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | DURING PICC PROCEDURE, BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLETS S1275108FD5 LOT# REHN2437 2 KITS NOTED LEAKAGE AT RUBBER-LIKE STOPPER AND STYLET BEFORE INSERTION AND REPLACED. 3RD ONLY LEAKED AFTER INSERTION AND RESOLVED WITH STYLET REMOVAL. PICC SUCCESSFULLY PLACED. 2 PICCS RETAINED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01513 | 2023-04-24 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT THIRTY DAYS POST PORT PLACEMENT, PATIENT ALLEGEDLY EXPERIENCED SWELLING AND PAIN. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING TREATMENT OR MEDICATIONS FOR PAIN AND SWELLING. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02797 | 2023-04-24 | 2023-04-01 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR LESS THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE " DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. " " REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER." "ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS." BAW2456229 REV 0 IS FOUND TO BE ADEQUATE TO FULFILL S03FA211 REVISION 17 AS IT STATES : PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE:- EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE "¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS." "6. REPLACE THE PUREWICK¿ FEMALE EXTERNAL CATHETER AT LEAST EVERY 8-12 HOURS OR IF SOILED WITH FECES OR BLOOD. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICK¿ FEMALE EXTERNAL CATHETER." " ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS."" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02795 | 2023-04-24 | 2023-03-30 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT CUSTOMER JUST FOUND THIS TWITTER THREAD ABOUT UROSEPSIS, AND IT SEEMS A CLINICIAN WAS REFERENCING PUREWICK IN THAT A PATIENT WAS RETAINING URINE BECAUSE THEY WERE USING AN EXTERNAL CATHETER INSTEAD OF AN INTERNAL ONE, AND THEY FELT THIS CONTRIBUTED TO THE CONDITION. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02804 | 2023-04-24 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE EMERGENCY CARE WHO WAS HAVING ISSUES WITH THIS CATHETER. THE CATHETER WAS BLOCKED AND THE LUMEN FOR THE INFLOW/OUTFLOW DID NOT WORK PROPERLY. ON (B)(6) 2023, A PATIENT CAME TO EMERGENCY ROOM USING THE BLADDER CATHETER. THIS BLADDER CATHETER HAD BECOME CLOGGED AND ALSO, THE LUMEN FOR INFLOW AND OUTFLOW OF THE BALLOON WAS DEFECTIVE. WITH BLADDER IRRIGATION ONLY INFLOW WAS POSSIBLE, OUTFLOW WAS NOT POSSIBLE. ALSO, BALLOON CONTENTS CANNOT BE DRAINED, IT IMMEDIATELY VACUUMS, WITH DIFFICULTY 2CC EXTRACTION. FINALLY TOGETHER WITH DOCTOR LABORIOUSLY ABLE TO DRAIN 5CC BALLOON CONTENTS, TURNED OUT TO BE FULL CONTENTS. CAD REMOVED AND NEW ONE PLACED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02802 | 2023-04-24 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE EMERGENCY CARE WHO WAS HAVING ISSUES WITH THIS CATHETER. THE CATHETER WAS BLOCKED AND THE LUMEN FOR THE INFLOW/OUTFLOW DID NOT WORK PROPERLY. ON (B)(6) 2023, A PATIENT CAME TO EMERGENCY ROOM USING THE BLADDER CATHETER. THIS BLADDER CATHETER HAD BECOME CLOGGED AND ALSO, THE LUMEN FOR INFLOW AND OUTFLOW OF THE BALLOON WAS DEFECTIVE. WITH BLADDER IRRIGATION ONLY INFLOW WAS POSSIBLE, OUTFLOW WAS NOT POSSIBLE. ALSO, BALLOON CONTENTS CANNOT BE DRAINED, IT IMMEDIATELY VACUUMS, WITH DIFFICULTY 2CC EXTRACTION. FINALLY TOGETHER WITH DOCTOR LABORIOUSLY ABLE TO DRAIN 5CC BALLOON CONTENTS, TURNED OUT TO BE FULL CONTENTS. CAD REMOVED AND NEW ONE PLACED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16794614 | 2023-04-24 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | POWERPICC PROVENA | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | S1275108FD5 | S1275108FD5 |  | DURING PICC PROCEDURE, BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLETS S1275108FD5 LOT# REHN2437 2 KITS NOTED LEAKAGE AT RUBBER-LIKE STOPPER AND STYLET BEFORE INSERTION AND REPLACED. 3RD ONLY LEAKED AFTER INSERTION AND RESOLVED WITH STYLET REMOVAL. PICC SUCCESSFULLY PLACED. 2 PICCS RETAINED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01522 | 2023-04-25 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT ONE YEAR SIX MONTHS AND SIXTEEN DAYS POST PORT PLACEMENT, CRACK WAS ALLEGEDLY FOUND SIXTEEN CENTIMETERS FROM THE TIP OF THE REMOVED CATHETER. IT WAS FURTHER REPORTED THAT THE PATIENT COMPLAINED OF PAIN. HOWEVER, THERE IS NO INFORMATION PROVIDED REGARDING ADDITIONAL INTERVENTION OR MEDICATION TO TREAT PAIN. REPORTEDLY, THE PORT SYSTEM WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE X-PORT IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL, TACTILE AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. TWO SPLITS WERE NOTED ON THE PORT SEPTUM. A COMPOUND BREAK WAS NOTED APPROXIMATELY 12.6CM FROM THE DISTAL END OF THE CATH-LOCK AND A PARTIAL CIRCUMFERENTIAL BREAK WAS NOTED APPROXIMATELY 13.1CM FROM THE DISTAL END OF THE CATH-LOCK. UPON INFUSION, LEAKS FROM THE BREAKS ON THE ATTACHED CATHETER WERE OBSERVED WHILE WATER EXITED THE DISTAL END OF THE CATHETER. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND IDENTIFIED WEAR AND DEFORMATION ISSUES. THE IDENTIFIED BREAK IS TYPICAL OF FLEXURAL FATIGUE, WHICH IS DUE TO THE REPETITIVE, CYCLIC KINKING OF THE CATHETER. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 05/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02842 | 2023-04-25 | 2023-04-04 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THESE INTERMITTENT CATHETERS CAUSED PATIENT TO GET A URINARY TRACT INFECTION AND WAS HOSPITALIZED WITH SEPSIS. CAREGIVER LOOKING FOR OTHER BRANDS AND SENT SAMPLE OF OTHER PRE-LUBED CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION AND SEPSIS. PER ADDITIONAL INFORMATION FROM CARDINAL HEALTH, IT WAS REPORTED THAT PATIENT GOT A URINARY TRACT INFECTION WHILE USING THE MAGIC 3 INTERMITTENT CATHETER. PATIENT WAS IN THE HOSPITAL BECAUSE OF THE ISSUE. THEY RECEIVED SAMPLES AND WANTED THE CURE ULTRA CATHETER. THIS WAS THE FIRST TIME THE PATIENT HAD USED THIS PRODUCT. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿ PATIENT SENSITIVITY TO MATERIALS, MECHANICAL AND/OR CHEMICAL RESPONSES FROM THE PATIENT¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSE BUT IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON'T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02840 | 2023-04-25 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE BALLOON CANNOT BE INFLATED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, THERE WAS RESISTANCE DURING INFLATING THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02843 | 2023-04-25 | 2023-03-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT URINE LEAKED FROM THE FOLEY CATHETER SHAFT DURING USE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02844 | 2023-04-25 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119110 | 119110 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS UNABLE TO DRAIN URINE. FOLEY CATHETER WAS INSERTED FOR THE PURPOSE OF URINARY DRAINAGE DURING SURGERY. HOWEVER, FOLEY CATHETER WAS REMOVED BECAUSE THERE WAS NO URINE DRAINAGE AT ALL. FLUSHED SALINE SOLUTION INTO THE DRAINAGE LUMEN OF THE REMOVED FOLEY CATHETER, BUT SALINE SOLUTION COULD NOT BE DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02846 | 2023-04-25 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT CUSTOMER HAS HEARD MULTIPLE COMPLAINTS ON POST PROCEDURE AND UPON REMOVAL OF THE CATHETER, MORE SPECIFICALLY THAT PATIENTS WERE COMPLAINING OF BURNING WHILE THE FOLEY WAS ON PLACE AND DIFFICULTY WITH THE REMOVAL AND THE MATERIAL WAS NOT AS SLICK AS IT ONCE WAS AND ASKED IF THERE WERE ANY CHANGES HAVE BEEN MADE TO THE PRODUCT OR IF ANY OTHER COMPLAINTS SIMILAR TO THIS HAVE BEEN MADE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02850 | 2023-04-25 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0102L16 | 0102L16 |  | IT WAS REPORTED THAT THE PATIENT HAD NOTICED A DIFFERENCE IN THE DIAMETER AT THE BASE OF THE FOLEY CATHETER WHERE IT JOINED TO THE FUNNEL AREA, BUT UP FROM THERE A COUPLE INCHES. NO ADDITIONAL INFORMATION PROVIDED . | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02852 | 2023-04-25 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53618G |  | IT WAS REPORTED THAT THE HYDRO SIL GO CATHETERS, SOME OF WHICH HAD NO LUBRICATION OR LITTLE TO NONE, WERE VERY UNCOMFORTABLE TO USE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, USE THE ADHESIVE LABEL(S) TO HANG THE PACKAGE ON A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. 6. KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL. 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02851 | 2023-04-25 | 2023-02-27 | I, n, j, u, r, y | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT DURING CYSTOURETHROSCOPY AND URETHROPLASTY, THE FOLEY CATHETER BULB WAS DEFLATED AND THERE WAS RESISTANCE MET DURING ATTEMPT TO REMOVE AND THEN BULB WAS FURTHER DEFLATED, RESISTANCE PERSISTED. THE BULB WAS INFLATED AND DEFLATED AGAIN WITH SALINE AND DESPITE THE USE OF LUBRICATE, IT WAS DIFFICULT TO REMOVE. THIS RESULTED IN A THIN RIM NOTED TO BE PRESENT AFTER REMOVAL AND SUPERFICIAL LACERATION ALONG POSTERIOR EDGE OF THE URETHRAL MEATUS. THIS AREA REQUIRED REPAIR USING INTERRUPTED STITCHES PARALLEL TO THE LUMEN OF THE URETHRA. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | OTHER |  | 2023-04
3006260740-2023-01527 | 2023-04-25 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED "HUB/OUTER BODY BREAKAGE UNDER PRESSURE DURING MOC INJECTION." ADDITIONAL INFORMATION 04/12/2023: CONFIRMED THE EVENT OCCURRED (B)(6) 2023 IN ONCOLOGY. A FLUID LEAK WAS OBSERVED BUT THERE WAS NO NEGATIVE CONSEQUENCES TO THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02855 | 2023-04-25 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER INSERTION TRAY | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 802110 | 802110 |  | IT WAS REPORTED THAT THE BARDIA FOLEY CATHETER INSERTION TRAY WAS EMPTY, ONLY HAD THE PACKAGING MATERIAL AND NO MEDICAL PRODUCTS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01528 | 2023-04-25 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2174108 |  | IT WAS REPORTED THAT "DURING THE LAST FLUSHING THE HCP NOTED LEAKAGE OF SALINE SOLUTION FROM THE INSERTION POINT. WHEN REMOVED, THE CATHETER WAS FOUND TO BE RUPTURED AT APPROXIMATELY 10/15 CM. THE INCIDENT OCCURRED AFTER USE" ADDITIONAL INFORMATION RECEIVED 04/17/2023: THERE WERE NO REPORTED CONSEQUENCES TO THE PATIENT. THE PICC WAS IMPLANTED ON (B)(6)2023 AND HAS BEEN IN CONTACT WITH CHEMOTHERAPY DRUGS. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01525 | 2023-04-25 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER "CATHETER PLACEMENT OPERATOR FOUND THAT THE SPIRAL METAL WIRE WAS DETACHED FROM THE DISTAL END OF THE GUIDE WIRE. THE CATHETER WAS INSERTED SUCCESSFULLY AFTER THE GUIDE WIRE WAS REPLACED." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01521 | 2023-04-25 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5393190 | 6393190 |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT, THE CLAMP ALLEGEDLY BROKE DURING USE. REPORTEDLY, THE DEVICE WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01524 | 2023-04-25 | 2023-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED VIA MEDWATCH "MALFUNCTION IN PICC LINE WIRE, STOPCOCK PERMITTED AIR INTRUSION WHEN ASPIRATED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02856 | 2023-04-25 | 2023-04-04 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT HAD URINARY TRACT INFECTION BUT HAD NOW RECOVERED. REPRESENTATIVE DISCUSSED THE PATIENT CLEANING AND USAGE THE PUREWICK FEMALE EXTERNAL CATHETERS. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS MORE THAN 90 DAYS. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION WAS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02863 | 2023-04-25 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THEY WERE TRYING TO REWARM PATIENT TO 36C ON THE ARCTIC SUN DEVICE. THE DEVICE WAS REGISTERING PT1 (PATIENT TEMPERATURE) WAS BELOW 31C. TEMPERATURE SENSING FOLEY AND RECTAL PROBE IN PLACE. SECONDARY PROBE WAS REGISTERING PATIENT TEMPERATURE WAS 35.4C. THEY DRAINED PADS AND RECONNECTED AND REPLACED CABLES AND PROBES AND NOW THE TEMPERATURE WAS REGISTERING CORRECTLY BUT NURSE WANTED TO KNOW WHAT TO DO IF IT HAPPENED AGAIN. NURSE NOTED WHEN THIS OCCURS TO VERIFY PLACEMENT OF PROBE. NURSE PLUG PROBE INTO BEDSIDE MONITOR TO SEE IF TEMPERATURE READING WAS INCONSISTENT. USE ALTERNATE SOURCE TO CHECK PATIENT TEMPERATURE. UNPLUG AND PLUG BACK IN CABLE FROM PROBE AND DEVICE. FLEX CABLE TO SEE IF TEMPERATURE JUMPS AROUND. IF TEMPERATURE CHANGES NOTED, THEN REPLACE THE CABLE. MIS EXPLAINED SINCE NURSE CHANGED THE CABLES AND PROBES IT WOULD LIKELY NOT RETURN BUT THEY COULD NOT IDENTIFY WHICH ONE WAS THE ISSUE WITHOUT TROUBLE SHOOTING IN THE MOMENT. MIS ADVISED TO CALL BACK WITH ANY ALERTS OR QUESTIONS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02864 | 2023-04-25 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TOUCHLESS PLUS CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 4A5142 |  | IT WAS REPORTED THAT CUSTOMER RECENTLY SWITCHED THEIR CATHETERS FROM THE RUBBER TO THE VINYL ONES ON THIS ORDER. THEY STATED THAT A PATIENT WAS BLEEDING WHEN THEY TRIED TO INSERT THE CATHETER. NO MEDICAL INTERVENTION REQUIRED. MATERIAL# (B)(4). MATERIAL# (B)(4). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED.\ | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02858 | 2023-04-25 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 165814UK |  | IT WAS REPORTED THAT THE SILICONE FOLEY CATHETER HAD A HOLE IN THE TUBING THAT HAD COME ACROSS A RADAR FROM 14.2.23. IT WAS A BARDIA AQUAFIL WITH MATERIAL REFERENCE (B)(4), LOT MYGR5472 AND USE BY (B)(6) 2026. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02853 | 2023-04-25 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123516A | 123516A |  | IT WAS REPORTED THAT THE HEAD OF NURSES INDICATED THAT THE URINARY CATHETER WAS PRESENTING PROBLEMS WHEN INFLATING THE RETENTION BALLOON, THE STERILE WATER DID NOT PASS TO THE BALLOON AND THE CATHETER PRESENTED PROBLEMS WHEN INJECTING THE STERILE WATER TO INFLATE THE BALLOON. IF USED ON THE PATIENT, THEY REMOVED THE PROBE. THE PROBE WAS USED IN THE AREA OF TOCOSURGERY. THE PROBE WAS REMOVED FROM THE PATIENT, IT DID NOT CAUSE PROBLEMS IN THE PATIENT BECAUSE IT WAS REMOVED ON TIME. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01523 | 2023-04-25 | 2023-01-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT PICC LINE THAT HAS BEEN ISSUED WAS DAMAGED. DOCTOR IS FINDING IT DIFFICULT TO CONNECT PROPERLY, ALSO THEY ARE NOT ABLE TO PUNCTURE FROM ITS TIP. NO OTHER INFORMATION WAS PROVIDED. 04/21/2023 - THE RETURNED SAMPLE EXHIBITED A NEARLY COMPLETE SPLIT. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, PHOTO SAMPLE, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED CATHETER WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS THREE PHOTOGRAPHS WHICH DEPICTED A 4FR S/L GROSHONG NXT CATHETER. THE DEPICTED PORTION OF THE DEVICE INCLUDED THE LUER ADAPTER, EXTENSION TUBING, AND A PORTION OF CATHETER SHAFT. THE CATHETER WAS INSERTED INTO THE ARM OF A PATIENT. A NEARLY COMPLETE SPLIT WAS OBSERVED AT THE PROXIMAL END OF THE EXTENSION TUBING. THE SPLIT EXHIBITED A DIAGONALLY ALIGNED PROFILE. DAMAGE WAS EVIDENT IN ALL THREE SUBMITTED PHOTOGRAPHS; HOWEVER, INSPECTION OF THE PHOTOGRAPHS WAS INSUFFICIENT TO IDENTIFY THE CAUSE OF THE DAMAGE. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE BURST DAMAGE CAUSED BY OVER-PRESSURIZATION AND SHARP INSTRUMENT CONTACT. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02867 | 2023-04-25 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE 23 OF THESE INTERMITTENT CATHETER HAD EMPTY BURST PACKS. CUSTOMER STATED THIS WAS THE 1ST TIME THIS HAD EVER HAPPENED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02837 | 2023-04-25 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THESE MAGIC 3 HYDROPHILIC INTERMITTENT COUDE CATHETER HURTS WHILE REMOVING. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿CATHETER SWELLS, SOFTENS, ETC., DUE TO INCOMPATIBILITY WITH COATING¿. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01519 | 2023-04-25 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT VUE, ISP, MR, 8FR, MICRO | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608052 | 1608052 |  | IT WAS REPORTED THAT DURING PORT DEVICE IMPLANT PROCEDURE, THE PORT STEM WAS ALLEGEDLY BENT. IT WAS FURTHER REPORTED THAT THE PORT STEM WAS ALLEGEDLY BROKEN. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01520 | 2023-04-25 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1616000 | 1616000 |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, NEEDLE ALLEGEDLY CRACKED DURING USE. IT WAS FURTHER REPORTED THAT ASPIRATION OF THE VEIN WAS ALLEGEDLY MORE DIFFICULT DUE TO THE CRACKED NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 02/2024) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02836 | 2023-04-25 | 2023-03-22 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 1758SI22 | 1758SI22 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS NOT DRAINING. DOCTOR OF MEDICINE ATTEMPTED TO REMOVE CATHETER AFTER BALLOON DEFLATION BUT WAS UNABLE. DOCTOR OF MEDICINE CUT BALLOON FILLING PORT AND BALLOON STAYED INFLATED. DOCTOR OF MEDICINE REMOVED CATHETER WITH INFLATED BALLOON. DOCTOR OF MEDICINE WENT BACK IN WITH CYSTOSCOPE TO ASSURE HOMEOSTASIS AND THEN REPLACED NEW CATHETER WITHOUT DIFFICULTY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00282 | 2023-04-25 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG120126 | ATG120126 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN A CALCIFIED LESION, THE PTA BALLOON ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT BLOOD WAS ALLEGEDLY FOUND IN THE SHAFT AFTER DEFLATING THE BALLOON. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01516 | 2023-04-25 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716000J |  | IT WAS REPORTED THAT FOUR YEARS FOUR MONTHS TWENTY SIX DAYS POST PORT PLACEMENT, THE PORT SYSTEM WAS ATTEMPTED TO BE REMOVED VIA THE UPPER RIGHT ARM, THE CATHETER ALLEGEDLY COULD NOT BE REMOVED. IT WAS FURTHER REPORTED THAT THE BLOOD RETURN COULD NOT BE CONFIRMED THROUGH THE CATHETER. REPORTEDLY THE CATHETER WAS CUT AND THE PORT BODY AND THE PROXIMAL CATHETER SEGMENT WERE REMOVED. THE HCP DECIDED THAT THE DISTAL CATHETER SEGMENT REMAINS INSIDE OF THE PATIENT¿S BODY CONSIDERING THE RISK OF REMOVAL AND CONTINUES THE FOLLOW-UP OBSERVATION. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWER PORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00284 | 2023-04-25 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U415064RX | U415064RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE IN THE COMMON FEMORAL ARTERY VIA A CONTRALATERAL APPROACH, THE PTA BALLOON ALLEGEDLY RUPTURED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 02/2025). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01517 | 2023-04-25 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0606150J |  | IT WAS REPORTED THAT THIRTEEN YEARS ELEVEN MONTHS AND FOURTEEN DAYS POST PORT PLACEMENT VIA THE SUBCLAVIAN VEIN, THE PORT BODY AND THE VICINITY OF THE CATH LOCK WERE DETACHED AND THE CATHETER WAS GRASPED. IT WAS FURTHER REPORTED THAT WHEN THE CATHETER WAS PULLED, THE CATHETER WAS ALLEGEDLY BROKEN. REPORTEDLY, THE DISTAL CATHETER SEGMENT AND THE PORT BODY WERE REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE TITANIUM LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00283 | 2023-04-25 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80166 | ATG80166 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, BALLOON ALLEGEDLY RUPTURED AT 4 ATM. IT WAS FURTHER REPORTED THAT BLOOD WAS ALLEGEDLY FOUND IN THE SHAFT AFTER DEFLATING THE BALLOON. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED BALLOON RUPTURE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 04/2024), G3. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01518 | 2023-04-25 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT ONE WEEK POST DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY CAME OUT FROM THE INSERTION SITE AND COULD NOT CONTINUE THE DIALYSIS PROCEDURE. REPORTEDLY, NEEDED NEW DEVICE FOR INSERTION. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00286 | 2023-04-25 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80144 | ATG80144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED. IT WAS FURTHER REPORTED THAT BLOOD WAS ALLEGEDLY FOUND IN THE SHAFT AFTER DEFLATING THE BALLOON. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 12/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16806269 | 2023-04-25 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDEX LUBRI-SIL ALL-SILICONE TEMP SENSING FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | 119216M | 119216M |  | LEAKING AT CONNECTION POINT OF SPOUT TO BAG. |  | N | Y |  |  | 2023-04
2020394-2023-00285 | 2023-04-25 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U4150230RX | U4150230RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE BELOW THE KNEE, THE PTA BALLOON WAS ALLEGEDLY UNABLE TO CROSS THE LESION. IT WAS FURTHER REPORTED THAT THE TIP ALLEGEDLY BROKE INSIDE THE PATIENT. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 04/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02831 | 2023-04-25 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | URETERAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE ISSUES WITH THE INTERMITTENT CATHETERS. THE URINE WOULD NOT COME OUT AND IT WAS CLOGGED. NO IODINE CAME IN PACKAGE AND ONLY ONE GLOVE CAME IN PACKAGE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-04
16806505 | 2023-04-25 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARDEX LUBRI-SIL ALL-SILICONE TEMP SENSING FOLEY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. | 119216M | 119216M |  | SMALL HOLE IN DRAINAGE BAG. URINARY CATHETER REMOVED AND REPLACED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02835 | 2023-04-25 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0102L12 | 0102L12 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS PLACED AND BALLOON INFLATED PRIOR TO PROCEDURE WITHOUT DIFFICULTY. POST OPERATIVELY, THE REGISTER NURSE DEFLATED BALLOON TIP WITH EMPTY SYRINGE. WHEN UNABLE TO REMOVE ANY MORE FLUID FROM TIP, REGISTER NURSE ATTEMPTED TO PULL CATHETER OUT BUT RESISTANCE MET CAUSING PATIENT DISCOMFORT. CHARGE REGISTER NURSE TRIED TO TROUBLE SHOOT BY TRYING TO SAW IF ANYMORE FLUID COULD BE REMOVED FROM TIP. NOTHING CAME OUT. FOLEY LEFT IN PLACE. DURING SHIFT, NIGHT REGISTER NURSE AND DAY REGISTER NURSE MADE ANOTHER ATTEMPT TO PULL FOLEY AND THIS TIME IT CAME OUT. TIP OF CATHETER HAD BALLOON STILL SLIGHTLY INFLATED. PATIENT IN STABLE CONDITION WITH NO INJURY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
16806112 | 2023-04-25 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | INDWELLING URINARY CATHETER | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. |  |  |  | WHILE ATTACHING THE SYRINGE TO OBTAIN THE URINE FROM THE SAMPLE ATTACHED TO THE URINARY CATHETER CAME OFF AND URINE WAS LEAKING. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02888 | 2023-04-26 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51612 | 51612 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE NOT WORKING. CUSTOMER SAID WHEN USING THE URINE WENT EVERYWHERE AND NOT EASY TO LUBRICATE. REPRESENTATIVE OFFERED TO SEND A SAMPLE OF DIFFERENT STYLE CATHETERS AND PATIENT APPROVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02884 | 2023-04-26 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175812 |  | IT WAS REPORTED THAT THE CATHETER BALLOON DID NOT INFLATE AND THEREFORE THE NURSED WEREN'T ABLE TO USE THE CATHETER AND USED A SPARE IN STOCK RESULTED IN EARLY ORDERING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02883 | 2023-04-26 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT DURING A PRETEST FOR CHECKING THE BALLOON, A CRACK WAS FOUND IN THE FOLEY CATHETER BALLOON. IT WAS STATED THAT, SINCE IT IS INFORMATION FROM A DISTRIBUTOR, NO ADDITIONAL INFORMATION WAS OBTAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02869 | 2023-04-26 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123514A | 123514A |  | IT WAS REPORTED THAT THE HEAD OF NURSES INDICATED THAT THE URINARY CATHETER WAS PRESENTING PROBLEMS WHEN INFLATING THE RETENTION BALLOON, THE STERILE WATER DID NOT PASS TO THE BALLOON AND THE CATHETER PRESENTED PROBLEMS WHEN INJECTING THE STERILE WATER TO INFLATE THE BALLOON. IT WAS NOT USED ON THE PATIENT AND IT DID NOT CAUSE PROBLEMS IN THE PATIENT. IN TOTAL THERE WERE 5 PIECES OF THIS SAME BATCH. THE DEFICIENT PROBES WERE NOT USED WITH PATIENTS, SINCE WHEN PERFORMING THE TEST (INFLATING THE BALLOON) IT WAS NOT POSSIBLE, SO THEY WERE DISCARDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02871 | 2023-04-26 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE HEAD OF NURSES INDICATED THAT THE URINARY CATHETER WAS PRESENTING PROBLEMS WHEN INFLATING THE RETENTION BALLOON, THE STERILE WATER DID NOT PASS TO THE BALLOON AND THE CATHETER PRESENTED PROBLEMS WHEN INJECTING THE STERILE WATER TO INFLATE THE BALLOON. IT WAS NOT USED ON THE PATIENT AND IT DID NOT CAUSE PROBLEMS IN THE PATIENT. IN TOTAL THERE WERE 5 PIECES OF THIS SAME BATCH MYDX4908. THE DEFICIENT PROBES WERE NOT USED WITH PATIENTS, SINCE WHEN PERFORMING THE TEST (INFLATING THE BALLOON) IT WAS NOT POSSIBLE, SO THEY WERE DISCARDED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02872 | 2023-04-26 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998004 | 998004 |  | IT WAS REPORTED THAT MANOMETER SECTION OF INFLATION DEVICE WAS BROKEN OFF. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01543 | 2023-04-26 | 2023-04-02 | I, n, j, u, r, y | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT SOMETIMES POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN, SWELLING AND BRUISING AROUND THE SEPTUM. REPORTEDLY PAIN MEDICATIONS AND HYPOALLERGENIC DRESSING WERE PROVIDED FOR PAIN, SWELLING AND BRUISING. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 10/2024). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01530 | 2023-04-26 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 7F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0603870CE |  | IT WAS REPORTED THAT UPON PULMONARY ARTERIAL CATHETER REMOVAL, A LARGE PART OF THE CATHETER REMAINED IN THE VEIN. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 7F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 7F PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 01/2025). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01529 | 2023-04-26 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED PAIN. IT WAS FURTHER REPORTED THAT THE PAIN WAS TREATED WITH PAIN RELIEVING MEDICATION. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 04/2024) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01531 | 2023-04-26 | 2023-04-06 | D, e, a, t, h | Manufacturer report | POWERPORT IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1759601 | 1759601 |  | IT WAS REPORTED THAT THE PATIENT ALLEGEDLY EXPIRED WHILE USING A PORT DEVICE. NO FURTHER INFORMATION WAS PROVIDED. THE PATIENTS CAUSE OF DEATH IS UNKNOWN. | H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01532 | 2023-04-26 | 2023-01-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 1716070J |  | IT WAS REPORTED THAT DURING A PORT PLACEMENT PROCEDURE, THE GUIDEWIRE WAS ALLEGEDLY STUCK INTO THE INTRODUCER NEEDLE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 6F PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE INTRODUCER NEEDLE WITH A LOADED GUIDEWIRE WERE RETURNED FOR EVALUATION. FUNCTIONAL, GROSS VISUAL AND MICROSCOPIC VISUAL EVALUATIONS WERE PERFORMED. THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED STRETCHED ISSUE AS THE LOADED GUIDEWIRE WAS NOTED TO BE STUCK WITHIN THE INTRODUCER NEEDLE AND THE GUIDEWIRE WAS NOTED TO BE UNCOILED AND STRETCHED. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE TO ADVANCE AND PHYSICAL RESISTANCE ISSUES, AS THE EXACT CIRCUMSTANCES AT THE TIME OF THE REPORTED EVENT WAS UNKNOWN. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01544 | 2023-04-26 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY EXPERIENCED PAIN AND BRUISING AROUND THE PORT. IT WAS FURTHER REPORTED THAT DEVICE ALLEGEDLY HAD DIFFICULTY IN DRAWING BLOOD FROM THE PORT AND REPOSITION BY LIFTING THE ARM TO GET THE PORT TO WORK. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 02/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02889 | 2023-04-26 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123614A | 123614A |  | IT WAS REPORTED THAT THE PATIENT STATED THAT THESE CATHETERS HAD BEEN GETTING CLOGGED. REPRESENTATIVE ASKED PATIENT IF BLOCKAGE WAS FROM SEDIMENT AND PATIENT ONLY SAW A POWDER OR SAND ONE TIME . PATIENT SAID THAT THE DOCTOR'S OFFICE WAS ABLE TO UNCLOG IT ONE TIME. ADVISED HOW TO CLEAN OUT THE CATHETER PATIENT DID NOT DO IT AND PATIENT WOULD DO THAT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
2020394-2023-00289 | 2023-04-26 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG120144 | ATG120144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY BURST DURING DILATION. IT WAS FURTHER REPORTED THAT THE BALLOON ALLEGEDLY DETACHED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 10/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
16815218 | 2023-04-26 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | LUTONIX® DRUG COATED BALLOON | DRUG-ELUTING PERIPHERAL TRANSLUMINAL ANGIOPLASTY CATHETER | C.R. BARD, INC. | LX1813053005F | LX1813053005F |  | DEVICE WOULDN'T ADVANCE THROUGH SHEATH FOR A SECOND PASS. NEW CATHETER OPENED AND WORKED. NO PATIENT INJURY. |  | N | Y |  |  | 2023-04
1018233-2023-02905 | 2023-04-26 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, THE FIXATION WATER LEAKED FROM THE BALLOON. IT WAS A BALLOON CHECK EVENT AND A FORCEPS WAS NOT USED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01533 | 2023-04-26 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PROV 4F DL IR 135CM NI | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S3274335 |  | IT WAS REPORTED "CALLER STATES CLOTTING IN THE POWERPICC PROVENA CATHETER PLACED IN PEDIATRIC PATIENT. PRODUCT CODE S3274335 LOT# REGR1311 WAS PLACED AND CLOTTED AND EXCHANGED OVER A WIRE. RESPONSE: EMAILED POWERPICC PROVENA INSTRUCTIONS AND DISCUSSED FLUSHING RECOMMENDATIONS.". ADDITIONAL INFO RECEIVED ON 9 APRIL 2023: PATIENT HAS CRITICAL MEDICATIONS RUNNING IN THEIR POWER PICC CONTINUOUSLY AND WHEN THEY DON'T INFUSE PROPERLY THEY MISS THEIR MEDICATIONS AND DECOMPENSATE. PATIENT IS FEMALE UNDER A YEAR OF AGE AND THE LINE IS PLACED IN THEIR FEMORAL VEIN. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01548 | 2023-04-26 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER 4F RADIOLOGY (DL) (W/ HYDRO-GLIDE INSERTION STYLET) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3274355 |  | IT WAS REPORTED THAT THE CUSTOMER IS HAVING AN ISSUE WITH ITEM 3274355. PER THE RN, BLOOD NOTED TO BE LEAKING FROM THE CATHETER RIGHT WHERE IT MEETS THE SECUREMENT WINGS. CLEAR AND OBVIOUS LEAK OBSERVED. ADDITIONAL INFORMATION: THE EVENT DID NOT OCCUR DURING AN URGENT OR LIFE-THREATENING MEDICAL SITUATION. TPN INFUSION WAS INTERRUPTED BUT NO ADVERSE OUTCOME WAS REPORTED. THE CATHETER WAS SECURED WITH STATLOCK. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01550 | 2023-04-26 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROVPPSOLO 4FDL MBP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1294108D4 |  | IT WAS REPORTED BY CUSTOMER THAT THE IV CONTRAST WAS INJECTED VIA MULTIPLE INJECTORS/MACHINE AND THE MACHINE GAVE A DEFAULT PRESSURE LIMITING ALARM AND SO IT PREVENTED FROM INFUSING IV CONTRAST TO THE PATIENT VIA THEIR POWER PICC LINES. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01551 | 2023-04-26 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE INTRODUCER WILL NOT INSERT INTO VASCULATURE. DATE OF EVENT (B)(6) 2023. ADDITIONAL INFO RECEIVED ON 14-APRIL-2023 CUSTOMER REPORTED THEY HAD TO DROP ANOTHER INTRODUCER KIT TO ACCESS THE PATIENT. THE EVENT CAUSED TRAUMA TO THE SKIN AND VESSEL UPON INSERTION BUT DID NOT REQUIRE TREATMENT. THERE WAS NO PATIENT HARM OR NEGATIVE OUTCOME. THE EVENT DID NOT INVOLVE AN URGENT OR LIFE-THREATENING MEDICAL SITUATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01552 | 2023-04-26 | 2023-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F DL POLYURETHANE CATHETER RADIOLOGY (135 CM NITINOL GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275335 |  | IT WAS REPORTED BY THE CUSTOMER "RED LUMEN ON 5F DUAL LUMEN PICC BROKE OFF. THE PICC WAS PLACED IN IR ON (B)(6) 2023, CATHETER STARTED LEAKING ON (B)(6) 2023 AND COMPLETELY BROKE OFF ON (B)(6).". | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01553 | 2023-04-26 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE INTRODUCER WILL NOT INSERT INTO VASCULATURE. DATE OF EVENT: (B)(6) 2023. ADDITIONAL INFORMATION RECEIVED 04/14/2023. THE CLINICIAN HAD TO OPEN ANOTHER INTRODUCER KIT TO ACCESS THE PATIENT. THE REPORTED EVENT CAUSED TRAUMA TO THE SKIN AND VESSEL UPON INSERTION BUT NO TREATMENT WAS NEEDED. THE WAS NO HARM OR NEGATIVE OUTCOME TO THE PATIENT. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01555 | 2023-04-26 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE CATHETER PLACEMENT OPERATOR INSPECTED THE CATHETER BY FLUSHING PRIOR TO INSERTION ON (B)(6) 2023. WHEN ATTACHING THE EXTENSION TUBING FOLLOWING CATHETER INSERTION, THE CONNECTION BETWEEN THE CATHETER AND THE EXTENSION TUBING LEAKED LIQUID. THE TWO PARTS COULD NOT BE ENGAGED FIRMLY OR SECURED. SIX CATHETERS FROM THE SAME BATCH EXPERIENCED THE SAME PROBLEM. THE OPERATOR HAD TO REPLACE THE EXTENSION TUBING. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01556 | 2023-04-26 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE CATHETER PLACEMENT OPERATOR INSPECTED THE CATHETER BY FLUSHING PRIOR TO INSERTION ON (B)(6) 2023. WHEN ATTACHING THE EXTENSION TUBING FOLLOWING CATHETER INSERTION, THE CONNECTION BETWEEN THE CATHETER AND THE EXTENSION TUBING LEAKED LIQUID. THE TWO PARTS COULD NOT BE ENGAGED FIRMLY OR SECURED. SIX CATHETERS FROM THE SAME BATCH EXPERIENCED THE SAME PROBLEM. THE OPERATOR HAD TO REPLACE THE EXTENSION TUBING. THIS REPORT ADDRESSES THE THIRD DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01557 | 2023-04-26 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE CATHETER PLACEMENT OPERATOR INSPECTED THE CATHETER BY FLUSHING PRIOR TO INSERTION ON (B)(6) 2023. WHEN ATTACHING THE EXTENSION TUBING FOLLOWING CATHETER INSERTION, THE CONNECTION BETWEEN THE CATHETER AND THE EXTENSION TUBING LEAKED LIQUID. THE TWO PARTS COULD NOT BE ENGAGED FIRMLY OR SECURED. SIX CATHETERS FROM THE SAME BATCH EXPERIENCED THE SAME PROBLEM. THE OPERATOR HAD TO REPLACE THE EXTENSION TUBING. THIS REPORT ADDRESSES THE FOURTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01559 | 2023-04-26 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE CATHETER PLACEMENT OPERATOR INSPECTED THE CATHETER BY FLUSHING PRIOR TO INSERTION ON (B)(6) 2023. WHEN ATTACHING THE EXTENSION TUBING FOLLOWING CATHETER INSERTION, THE CONNECTION BETWEEN THE CATHETER AND THE EXTENSION TUBING LEAKED LIQUID. THE TWO PARTS COULD NOT BE ENGAGED FIRMLY OR SECURED. SIX CATHETERS FROM THE SAME BATCH EXPERIENCED THE SAME PROBLEM. THE OPERATOR HAD TO REPLACE THE EXTENSION TUBING. THIS REPORT ADDRESSES THE SIXTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01560 | 2023-04-26 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER "THERE WAS A HOLE IN THE PICC LINE. PATIENT WAS RECEIVING TREATMENT WITH EPIRUBICIN AND CYCLOPHOSPHAMIDE. CX CALLS THE DISTRICT NURSE FROM HEALTH CARE CENTER AS BLOOD COMES FROM THE NEEDLE WHEN SHE FLUSHES. THE PATIENT COMES INTO THE ONCOLOGY DEPARTMENT WHERE A HOLE IS FOUND ON THE PICC LINE A FEW CM INSIDE THE SKIN. THE PICC LINE IS BEING PHASED OUT." NO OTHER INFORMATION WAS PROVIDED. | HH11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01562 | 2023-04-26 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 3F SL MBPLUS (3CG, BIOPATCH, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1173108D4 |  | CUSTOMER REPORTED TO BD VASCULAR ACCESS DEVICES SALES REPRESENTATIVE THAT THEY HAVE BEEN EXPERIENCING AIR WITHIN THE NEW SEPTUM OF THE PICC LINES THAT HOLDS THE WIRE THAT IS WITHIN OUR PICCS. IF THE NURSE PULLS BACK FOR BLOOD IT ALLOWS AIR AND THE SAME GOES FOR FLUSHING. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01561 | 2023-04-26 | 2023-03-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | CUSTOMER REPORTED THERE WAS A HOLE IN THE PICC LINE. DURING BLOOD TRANSFUSION ON (B)(6) 2023, BLOOD WAS VISIBLE AT THE PICC LINE CONNECTION. A HEMATOMA DEVELOPED ABOVE THE PUNCTURE SITE. THE PICC LINE WAS FLUSHED AND ASPIRATED AND THERE WAS NO LEAKAGE. ON (B)(6) 2023 THERE WAS NO BACKFLOW AND RESISTANCE DURING FLUSHING. ON (B)(6) 2023, THE PICC WAS BEING DISMANTLED AND HOLES WERE VISIBLE BUT UNCLEAR WHERE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01558 | 2023-04-26 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE CATHETER PLACEMENT OPERATOR INSPECTED THE CATHETER BY FLUSHING PRIOR TO INSERTION ON (B)(6) 2023. WHEN ATTACHING THE EXTENSION TUBING FOLLOWING CATHETER INSERTION, THE CONNECTION BETWEEN THE CATHETER AND THE EXTENSION TUBING LEAKED LIQUID. THE TWO PARTS COULD NOT BE ENGAGED FIRMLY OR SECURED. SIX CATHETERS FROM THE SAME BATCH EXPERIENCED THE SAME PROBLEM. THE OPERATOR HAD TO REPLACE THE EXTENSION TUBING. THIS REPORT ADDRESSES THE FIFTH DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01563 | 2023-04-26 | 2023-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "RED HUB ON 5F DUAL LUMEN POWERPICC BROKE. " NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01565 | 2023-04-26 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHERLOCK 3CG SENSOR (RADEL) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770352 |  | IT WAS REPORTED THE SENSOR, WHEN IT IS ON THE PATIENT, THE LOLLIPOP MOVES AROUND BY ITSELF AND RETURNS TO ITS ORIGINAL PLACE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01567 | 2023-04-26 | 2023-04-07 | I, n, j, u, r, y | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT, BIOPATCH, PROBE COVER, BELDAR NEEDLE GUIDE | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1174108D5 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE RN WAS ATTEMPTING TO FLUSH PICC LINE WHEN SUDDENLY A SPRAY OF SOMETHING WET WAS FELT ON HER FACE, FOREHEAD AND EYE AREA. AFTER LOOKING DOWN, THE RN REALIZED THE "SPRAY" WAS A MIX OF BLOOD PRODUCT AND SALINE FROM THE PICC LINE. UPON FURTHER INSPECTION, IT WAS DISCOVERED THAT THE RUBBER STOPPER AT THE END OF THE PICC WHICH HOLDS THE GUIDE WIRE DISLODGED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02893 | 2023-04-26 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT UPON RETURNING TO THE WARD AFTER SURGERY, THE CATHETER WAS FOUND FALLEN OUT OF THE PATIENT. ACCORDING TO THE INQUIRY SHEET ATTACHED, THERE WAS A TEAR OR RUPTURE OBSERVED ON THE BALLOON. SIMILAR EVENT HAS OCCURRED TO THE SAME PATIENT. THE EVENT OCCURRED AFTER 1 HOUR OF PLACEMENT AND THE WATER LEAKED FROM THE BALLOON. NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE BALLOON AND THE CATHETER WAS NOT PULLED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01549 | 2023-04-26 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROVPPSOLO 4FDL MBP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1294108D4 |  | IT WAS REPORTED BY CUSTOMER THAT THE IV CONTRAST WAS INJECTED VIA MULTIPLE INJECTORS/MACHINE AND THE MACHINE GAVE A DEFAULT PRESSURE LIMITING ALARM AND SO IT PREVENTED FROM INFUSING IV CONTRAST TO THE PATIENT VIA THEIR POWER PICC LINES. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01547 | 2023-04-26 | 2023-04-06 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO CATHETER 5F BASIC TRAY (DUAL-LUMEN) (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6295118 |  | IT WAS REPORTED THAT, ON TRYING TO REMOVE THE GUIDEWIRE WHILE INSERTING THE PICC, THE GUIDEWIRE UNRAVELLED AND HAD TO BE FISHED OUT OF THE PATIENT. THE PICC WAS REMOVED AND THE PATIENT XRAYED TO ENSURE NO FRAGMENTS LEFT INSITU. CUSTOMER CONFIRMED THAT THE PICC HAD BEEN TRIMMED BUT THAT THE GUIDEIWRE WAS PULLED BACK. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01554 | 2023-04-26 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE CATHETER PLACEMENT OPERATOR INSPECTED THE CATHETER BY FLUSHING PRIOR TO INSERTION ON (B)(6) 2023. WHEN ATTACHING THE EXTENSION TUBING FOLLOWING CATHETER INSERTION, THE CONNECTION BETWEEN THE CATHETER AND THE EXTENSION TUBING LEAKED LIQUID. THE TWO PARTS COULD NOT BE ENGAGED FIRMLY OR SECURED. SIX CATHETERS FROM THE SAME BATCH EXPERIENCED THE SAME PROBLEM. THE OPERATOR HAD TO REPLACE THE EXTENSION TUBING. THIS REPORT ADDRESSES THE FIRST DEVICE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01545 | 2023-04-26 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY (W/ MICROINTRODUCER) (55CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174118 |  | IT WAS REPORTED, PICC THAT WE IMPLANTED ON (B)(6) HAD TO BE REMOVED ON WEDNESDAY ON (B)(6) DUE TO CATHETER BREAKAGE ABOUT 20CM FROM THE TIP. IN A FEW WEEKS ANOTHER ONE MAY BE PLACED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01535 | 2023-04-26 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT WHEN CONDUCTING PUNCTURING FOR THE GROSHONG CATHETER, THE NURSE FOUND THAT RESISTANCE WAS FELT AT THE PROXIMAL END OF THE GUIDE WIRE, UNABLE TO REMOVE THE GUIDE WIRE. THE GUIDE WIRE WAS THEREFORE REMOVED TOGETHER WITH THE MICRO-INTRODUCER. A CAREFUL EXAMINATION REVEALED THAT THE PROXIMAL GUIDE WIRE WAS NOT SMOOTH AND PROTRUDED, WHICH MADE THE REMOVAL IMPOSSIBLE. ANOTHER MST WAS USED FOR RE-PUNCTURE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01534 | 2023-04-26 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, GUARDIVA, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D4 |  | IT WAS REPORTED BY THE CUSTOMER THAT NEW GEL SEPTUM USED TO HOLD PICC WIRES DURING INSERTIONS IS NOT SEALED AND ALLOWS AIR DURING CHECKING FOR BLOOD RETURN AND DURING FLUSHING. ADDITIONAL INFORMATION ON 4/13: "CUSTOMER REPORTED THAT IF THE PICC HAS ANY SORT OF COMPLICATION SUCH AS NOT DRAWING OR DIFFICULTY FLUSHING, AIR WILL COME INTO THE SYRINGE FROM AROUND THE WIRE, OR SALINE WILL SQUIRT OUT OF IT. IT'S NOTICEABLE IF THE PICC GOES IN AND DROPS RIGHT AWAY. THE BIGGEST PROBLEM IS THE AIR IN THE SYRINGE BECAUSE YOU HAVE TO TAKE OFF THE SYRINGE AND PUSH THE AIR OUT, AND THEN YOU ARE OPENING UP THE CLOSED SYSTEM MORE THAT SHOULD BE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01536 | 2023-04-26 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED "THE PATIENT WAS ADMITTED WITH LEUKEMIA AND RECEIVED DRUG CHEMOTHERAPY. AT 13:56 ON (B)(6) 2023, WHEN THE NURSE WAS PREPARING TO USE THE CENTRAL VENOUS CATHETER WITH PERIPHERAL CANNULA FOR DEEP VEIN CATHETERS FOR CHEMOTHERAPY TREATMENT, SHE FOUND THAT THE FRONT SECTION OF THE GUIDE WIRE WAS DISINTEGRATED AND CATHETERS COULD NOT BE INSERTED. SHE IMMEDIATELY REPLACED ANOTHER CENTRAL VENOUS CATHETER OF THE SAME MODEL AND COMPLETED THE OPERATION WITHOUT CAUSING ADVERSE CONSEQUENCES TO THE PATIENT." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01546 | 2023-04-26 | 2023-04-03 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 |  | 2194108 |  | IT WAS REPORTED THAT THE PICC FLUSHED AND CHECKED FOR VENOUS RETURN PRE DRUG ADMINISTRATION. NO ISSUES WITH PICC. BEVCIZUMAB COMMENCED. 11 MINUTES INTO TREATEMENT PATIENT PRESSED BUZZER TO ALERT NURSE THAT UPPER ARM WAS WET FROM PICC SITE. ON INSPECTION DRESSING VISIBLE WET. CHEST X RAY SHOWED NO OBVIOUS FRACTURE. INFUSION STOPPED. PICC DRESSING CHANGED. SITE TREATED AS EXTRAVASATION AS PER MEDICAL TEAM FOLLOWING EXTRAVASATION POLICY. PICC REMOVED AND INFUSIONAL SERVICES INFORMED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01537 | 2023-04-26 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED THERE IS LEAKAGE FROM PICC LINE WHICH COMES FROM HOLES IN THE CATHETER PART TRANSITION TO FIXATION PART. THE PATIENT IS DOING WELL, AND HAS HAD A NEW PICC. THE EVENT OCCURRED DURING USE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01539 | 2023-04-26 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED BY THE CUSTOMER THAT "WITH THE NEW DEVICE (LEFT) ITS DIFFICULT TO ASPIRATE BLOOD RETURN AND WHEN FLUSHING WITH SALINE, BUBBLES OUT OF THE END." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02894 | 2023-04-26 | 2023-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0166V16S |  | IT WAS REPORTED THAT THE DEFLATION LUMEN WAS OCCLUDED AND AFTER INSERTING, THE CATHETER WAS CUT OFF AND BECAME FLUENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01540 | 2023-04-26 | 2023-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER 4F BASIC TRAY (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6194118 |  | IT WAS REPORTED THAT, STAFF NOTED PICC APPEARED TO BE 'CUT' AND LEAKING UNABLE TO BE USED. ADDITIONAL INFORMATION RECEIVED 12 APRIL 2023: INFUSATE (POSSIBLY WITH SMALL VOLUME OF BLOOD) EXPOSED TO PATIENT¿S SKIN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02899 | 2023-04-26 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTERMITTENT ROCHESTER CATHETERS | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT CUSTOMER USES THE INTERMITTENT CATHETER, BUT THE LOT THEY WERE USING WAS NOT WORKING THERE WAS NO FLOW. PATIENT HAD DISCARD MANY OF THEM, BECAUSE THEY WERE NOT USEABLE, AND THEY CALLED SEVERAL TIMES AND HAD NOT RECEIVED A RESOLUTION TO THIS ISSUE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿KINKED OR PARTIALLY COLLAPSED DRAINAGE LUMEN COLLAPSED DRAINAGE LUMEN DUE TO WALL THICKNESS AND DUROMETER". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02898 | 2023-04-26 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 175808 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON INVERTED OVER THE CATHETER FORMING A RING THAT MADE FOLEY REMOVAL DIFFICULT FOR THE NURSE AND UNCOMFORTABLE FOR THE PATIENT. NO MEDICAL INTERVENTION WAS REPORTED. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INCORRECT BALLOON DESIGN (BALLOON WALL THICKNESS EXCESSIVE)". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE, AND FEDERAL LAWS AND REGULATIONS. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING PETROLATUM BASE. THEY WILL DAMAGE SILICONE AND MAY CAUSE THE BALLOON TO BURST. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. UNLESS PACKAGE IS OPENED OR DAMAGED. CAUTION: DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. SINGLE USE ONLY. DO NOT REUSE. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SLIP TIP SYRINGE. DO NOT USE NEEDLE. RECOMMENDED INFLATION CAPACITIES 3CC BALLOON: USE 5CC STERILE WATER : 5CC BALLOON: USE 10CC STERILE WATER, 30CC BALLOON: USE 35CC STERILE WATER. DO NOT EXCEED RECOMMENDED CAPACITIES. NOTE: AGGRESSIVE TRACTION, PARTICULARLY IN THE PRESENCE OF SUTURING, IS NOT RECOMMENDED FOR SILICONE FOLEY CATHETERS. BARD, BARDEX AND LUBRI-SIL ARE REGISTERED TRADEMARKS OF C. R. BARD, INC. OR AN AFFILIATE. U.S. PATENT NUMBER (B)(6) AND PATENT PENDING. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. ALLOW THE BALLOON TO DEFLATE SLOWLY ON ITS OWN. DO NOT ASPIRATE OR MANUALLY ACCELERATE THE DEFLATION OF THE BALLOON. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02901 | 2023-04-26 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 39302 | 39302 |  | IT WAS REPORTED THAT THEY WERE HARD TO REMOVE THE MALE EXTERNAL CATHETER. NO ADDITIONAL INFORMATION PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: 1) VERIFY CORRECT SIZE PRIOR TO USE. 2) TRIM PUBIC HAIR IF NECESSARY. 3) WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. 4) OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. 5) PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. 6) UNROLL THE CATHETER OVER PENIS. 7) GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. 8) CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: 1) ENSURE DRAINAGE BAG IS EMPTY. 2) DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. 3) GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER." | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01538 | 2023-04-26 | 2023-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC, IR, 135 CM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275335J |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS PRIMED, SALINE SOLUTION LEAKED FROM THE CATHETER. THERE WAS NO REPORTED PATIENT INVOLVEMENT. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01542 | 2023-04-26 | 2023-04-04 | I, n, j, u, r, y | Manufacturer report | BARD TRIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED "WE HAVE A PATIENT IN THE ICU WHO HAD A 15 CM TRIALYSIS CATHETER PLACED BY THE CRITICAL CARE TEAM. AFTER PLACEMENT THE PATIENT WAS STARTED ON DIALYSIS AND DEVELOPED SIGNIFICANT ECTOPY INCLUDING FREQUENT RUNS OF NON-SUSTAINED VT. THE X-RAY SHOWED THAT THE DEVICE WAS IN TOO FAR AND SO IT WAS PULLED BACK SEVERAL CENTIMETERS AND DIALYSIS WAS CONTINUED. INITIAL XRAY AFTER PLACEMENT PROVIDED AND AFTER 3CM REMOVAL. THE ECTOPY CONTINUED THROUGHOUT THE NIGHT DESPITE THE WITHDRAWAL AND CESSATION OF HEMODIALYSIS. THE NEXT DAY THE X-RAY WAS REVIEWED AND NOTED TO STILL BE SIGNIFICANTLY TOO FAR ADVANCED. THE '15 CM CATHETER' WAS REMOVED AN ADDITIONAL 7 CM TO A TOTAL OF 10CM OUT OF THE PATIENT AND THE FOLLOW UP X-RAY WAS NOTED IN PROPER LOCATION. THE CATHETER CLEARLY SAYS THAT IT IS A 15 CM CATHETER, BUT IT IS CLEARLY A 24 CM CATHETER AS THE X-RAY SHOWS A RULER MEASURING THE FULL LENGTH OF THE CATHETER. YOU CAN SEE HOW MUCH OF THE '15 CM CATHETER' IS OUTSIDE THE PATIENT WHILE THE TIP IS LOCATED IN THE CAVOATRIAL JUNCTION (FROM A RIGHT IJ INSERTION SITE)." ADDITIONAL INFO RECEIVED ON 10 APRIL 2023: NO FURTHER ADVERSE EVENTS BEYOND THE TEMPORARY NON-SUSTAINED VT THAT REPEATEDLY OCCURRED WHILE THE CATHETER WAS INSERTED TO ITS FULL LENGTH. THE DEVICE IS STILL IN THE PATIENT BUT A NOTE HAS BEEN CREATED TO ENSURE THAT THE CATHETER IS SAVED IN THE EVENT THAT IT IS REMOVED PRIOR TO DISCHARGE FROM FACILITY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01510 | 2023-04-26 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. BD POWERPICC TRIPLE LUMEN IN UPPER ARM. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02903 | 2023-04-26 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI12 | 0165SI12 |  | IT WAS REPORTED, THAT ON THE NEXT DAY AFTER PLACEMENT. THE CATHETER BALLOON FELL OUT OF THE PATIENT DURING THE MANAGEMENT OF THE CATHETER IN AN INDWELLING PATIENT. ACCORDING TO THE INQUIRY SHEET ATTACHED, THERE WAS NO BALLOON TEAR OR RUPTURE. NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02904 | 2023-04-26 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165L18 | 0165L18 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WOULD NOT STAY INFLATED. THE PATIENT PUSHED MORE SOLUTION IN BUT IT WAS NOT FULLY INFLATED, BUT THEY HAD NO LEAKS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01585 | 2023-04-27 | 2023-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | IT WAS REPORTED THAT FLUID LEAK FROM THE SIDE OF THE CATHETER. THERE WAS NO REPORTED IMPACT ON PATIENT. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01573 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 4F SINGLE-LUMEN CATHETER BASIC TRAY W/135CM NITINOL GW | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6194335 |  | IT WAS REPORTED THAT "FLOW RATES ARE TOO LOW. WITH HIGH PRESSURE INFUSION, TOO LOW A FLOW RATE WAS ACHIEVED." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02944 | 2023-04-27 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | MALE COUDE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50814G | 50814G |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER DID NOT HAD A HOLE IN THE TIP. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "MECHANICAL FAILURE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02939 | 2023-04-27 | 2023-04-07 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE CUSTOMER THOUGHT THEY HAD AN INFECTION AND WANTED THE OTHER CATHETERS AND STATED THEY DID NOT LIKE THESE. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PATIENT SENSITIVITY TO MATERIALS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02948 | 2023-04-27 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI18 |  | IT WAS REPORTED THAT URINE LEAKED FROM THE CATHETER. AS PER THE ADDITIONAL INFORMATION RECEIVED VIA IBC ON (B)(6) 2023, URINE LEAKED FROM THE GAP BETWEEN THE CATHETER AND THE URETHRA. NO SAMPLE WAS AVAILABLE AND WAS DISCARDED AT THE FACILITY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02953 | 2023-04-27 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE CATHETER SHAFT DURING THE TEST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02954 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI16 |  | IT WAS REPORTED THAT WATER WAS SPONTANEOUSLY REMOVED FROM THE CATHETER BALLOON. AFTER THE FOLEY CATHETER WAS PLACED IN THE PATIENT, STERILE DISTILLED WATER WAS INJECTED INTO THE CATHETER TO CHECK THE BALLOON FIXATION, BUT THE WATER SPONTANEOUSLY REMOVED, AND THE CATHETER FELL OUT OF THE PATIENT. THERE WAS NO DAMAGE TO THE OUTSIDE OF THE BALLOON AND NO SIGNS OF WATER SEEPING OUT. ALTHOUGH IT HAS NOT BEEN CONFIRMED, THE AGENCY SAID, THE NURSES WHO REPORTED IT SEEMED TO BE CONCERNED THAT THE WATER IN THE BALLOON WAS LEAKING INTO THE DRAINAGE FUNNEL. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 11APR2023, THE CATHETER FELL OUT DUE TO STERILE WATER LEAK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02956 | 2023-04-27 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE PATIENT INDWELLED A THREE-LUMEN URINARY CATHETER AFTER THE OPERATION AND WAS CONNECTED TO THE BLADDER FOR IRRIGATION ON (B)(6) 2023. THE BALLOON WAS INJECTED WITH 30ML OF WATER. THE URINARY CATHETER WAS BLOCKED IN THE EARLY MORNING OF THE FIRST DAY AFTER THE OPERATION. AFTER THE EXAMINATION, THE BALLOON WAS FOUND RUPTURED AND THE URINARY CATHETER FELL OUT. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, THE BALLOON WAS BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02966 | 2023-04-27 | 2023-04-07 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETER IRRITATED THE PATIENT. CUSTOMER STATED THAT THE PATIENT PLACED WICK THEMSELF AND SPOKE WITH PATIENT WHO CONFIRMED THEY HAD HISTORY OF BEDSORES BUT NO PERINEAL AREA ISSUES, CHANGED WICK EVERY 8 TO 12 HOURS OR AFTER EACH BOWEL MOVEMENT, NO DIARRHEA. THEY STATED THEY INFORMED NURSE OF IRRITATION BUT HAD NOT SPOKE TO THEIR DOCTOR AND THEY USING BARRIER CREAM PROVIDED BY NURSE. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE OR MATERIAL SURFACE IS ROUGH, ABRASIVE OR UNCOMFORTABLE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "WARNINGS: ¿ TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. ¿ NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. ¿ DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: ¿ NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER ¿ HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE ¿ EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ¿ EXPERIENCING MODERATE/HEAVY MENSTRUATION AND CANNOT USE A TAMPON ¿ ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: ¿ PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ¿ ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. ¿ REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. ¿ CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01577 | 2023-04-27 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PRISTINE LONG-TERM HEMODIALYSIS CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5403190 | 5403190 |  | IT WAS REPORTED THAT SIX DAYS POST DIALYSIS CATHETER PLACEMENT, THE CATHETER ALLEGEDLY CLOTTED. IT WAS FURTHER REPORTED THAT THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 04/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01580 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED BY THE CUSTOMER RN WAS ATTEMPTING TO FLUSH PICC LINE WHEN SPRAYED WITH SALINE. RUBBER STOPPER AT THE END OF THE PICC WHICH HOLDS THE GUIDE WIRE WAS LEAKING AND WAS SPRAYING SALINE ONTO THE NURSE WHEN ATTEMPTING TO FLUSH. RISK OF BLOOD EXPOSURE IF PORT ISN'T SEALED. IT WAS STATED THIS OCCURRED WITH TWO PICCS. THIS REPORT ADDRESSES THE SECOND DEVICE. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01582 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SL CV FOR MAQUET IAB | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | NA | 0684000472 |  | IT WAS REPORTED BY THE CUSTOMER THAT POUCHES FAILED FOR INCOMPLETE SEALS. NO OTHER INFORMATION WAS PROVIDED. | THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02964 | 2023-04-27 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT UPON RETURNING TO THE WARD AFTER SURGERY, THE CATHETER WAS FOUND FALLEN OUT OF THE PATIENT. ACCORDING TO THE INQUIRY SHEET ATTACHED, THERE WAS A TEAR OR RUPTURE OBSERVED ON THE BALLOON. SIMILAR EVENT HAS OCCURRED TO THE SAME PATIENT. THE EVENT OCCURRED AFTER 0.5 HOUR OF PLACEMENT AND THE WATER LEAKED FROM THE BALLOON. NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE BALLOON AND THE CATHETER WAS NOT PULLED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 11APR2023, NO MISSING PIECES WERE REPORTED, SPONTANEOUS CATHETER REMOVAL DUE TO THE BALLOON BURST WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01578 | 2023-04-27 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D5 |  | IT WAS REPORTED BY THE CUSTOMER "RN WAS ATTEMPTING TO FLUSH PICC LINE WHEN SPRAYED WITH SALINE. RUBBER STOPPER AT THE END OF THE PICC WHICH HOLDS THE GUIDE WIRE WAS LEAKING AND WAS SPRAYING SALINE ONTO THE NURSE WHEN ATTEMPTING TO FLUSH. RISK OF BLOOD EXPOSURE IF PORT ISN'T SEALED." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01584 | 2023-04-27 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE 18G 2.25¿ INTERMEDIATE KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | AC1182252 |  | CUSTOMER REPORTED THAT THE GUIDEWIRE MECHANISM IS ROUGH AND MORE DIFFICULT TO DEPLOY. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01589 | 2023-04-27 | 2023-04-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT WAS ADMITTED WITH GASTRIC CANCER, AND A PICC CATHETER WAS INSERTED FOR INFUSION. AT 10:30 ON (B)(6) 2023 WHEN THE NURSE USED THE CENTRAL VENOUS CATHETER WITH PERIPHERAL CANNULA, IT WAS FOUND THAT THE PLACEMENT WAS NOT SMOOTH. AFTER WITHDRAWING A LITTLE, THE PLACEMENT WAS STILL NOT SMOOTH. AFTER THE COMPLETE WITHDRAWAL, IT WAS FOUND THAT THE FRONT END OF THE CATHETER SHEATH SPLIT AND COULD NOT CONTINUE TO BE USED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01590 | 2023-04-27 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CATHETER 5F BASIC TRAY (DUAL-LUMEN) (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7957505J |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT HAD A GROSHONG CATHETER IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023, LEAKAGE OCCURRED FROM THE CATHETER PART WHICH WAS OUTSIDE OF THE PATIENT¿S BODY DURING THE ADMINISTRATION OF ANTIBIOTICS. THE DEVICE WAS REMOVED FROM THE PATIENT¿S BODY. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01588 | 2023-04-27 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | IT WAS REPORTED BY THE CUSTOMER THAT ON (B)(6) PICC CATHETERIZATION WAS GIVEN UNDER STANDARD OPERATION IN NEONATAL SURGERY WARD, THE PROCESS WAS SMOOTH, BLOOD RETURN WAS GOOD, AND THE POSITION WAS FIXED. AFTER THE CHEST RADIOGRAPH EXAMINATION, THE RESULTS SHOWED THAT THE CATHETERIZATION TIP WAS LOCATED AT THE LEVEL OF THE 5TH THORACIC VERTEBRAE, THE NIGHT INFUSION WAS SMOOTH, BLOOD RETURN WAS GOOD, A LITTLE EXUDATION AND BLOOD EXUDATION COULD BE SEEN, AND THE REPLACEMENT APPLICATION WAS GIVEN (B)(6) IN THE MORNING, THERE WAS MORE SEEPAGE UNDER THE FILM, AND THE DRESSING WAS REPLACED. AT 10:00 ON (B)(6), THE FILM WAS REPLACED DURING THE CONTINUOUS INFUSION PROCESS, AND LIQUID SEEPAGE WAS FOUND AT THE NEEDLE HANDLE, SO THE INFUSION WAS STOPPED IMMEDIATELY. AFTER INFORMING THE HEAD NURSE, B ULTRASOUND EXAMINATION RESULTS SHOWED THAT THERE WAS NO THROMBOSIS, AND EXTUBATION WAS GIVEN AFTER COMMUNICATING WITH THE PARENTS ABOUT THE RISK. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01574 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SL CV FOR MAQUET IAB | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0684000472 |  | IT WAS REPORTED BY THE CUSTOMER THAT "POUCHES EXHIBITED CUTS AND NICKS." NO OTHER INFORMATION WAS PROVIDED. THE CUSTOMER STATED THIS OCCURRED WITH MULTIPLE POUCHES HOWEVER THE EXACT QUANTITY INVOLVED WAS NOT PROVIDED TO BD VASCULAR ACCESS DEVICES FIELD ASSURANCE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01586 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (22G) (8CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | F122080T |  | IT WAS REPORTED BY THE CUSTOMER "PLACING POWERGLIDE PRO RIGHT BASILIC UNDER ULTRASOUND. DURING INSERTION PATIENT ACTIVE, TWISTING ARM. CANNULATED VEIN WITH NEEDLE, FULLY ADVANCED NEEDLE, RELEASED GUIDE WIRE, ADVANCED CATHETER. ATTEMPTED TO REMOVE GUIDEWIRE AND RESISTANCE NOTED. ATTEMPTED TO REMOVE CATHETER WITH GUIDE WIRE AND RESISTANCE NOTED. PATIENT BROUGHT TO IGT. UNDER FLUOROSCOPY IT WAS NOTED WIRE AND CATHETER HAD COILED BACK. CATHETER AND GUIDEWIRE REMOVED, INTACT, UNDER FLUOROSCOPY. NEW POWERGLIDE PRO INSERTED AFTER 2 MORE ATTEMPT BY RADIOLOGIST TO RIGHT BRACHIAL VEIN. NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01579 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED BY THE CUSTOMER RN WAS ATTEMPTING TO FLUSH PICC LINE WHEN SPRAYED WITH SALINE. RUBBER STOPPER AT THE END OF THE PICC WHICH HOLDS THE GUIDE WIRE WAS LEAKING AND WAS SPRAYING SALINE ONTO THE NURSE WHEN ATTEMPTING TO FLUSH. RISK OF BLOOD EXPOSURE IF PORT ISN'T SEALED. IT WAS STATED THIS OCCURRED WITH TWO PICCS. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01572 | 2023-04-27 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PROVPPSOLO 4FDL MBP | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1294108D4 |  | IT WAS REPORTED BY THE CUSTOMER "THE PICC CANNOT BE USED FOR POWER INJECTION. THERE WAS NO HARM BUT THE PATIENT NEEDED TO GET STUCK WITH AN IV. PATIENT HAS CANCER." IT WAS STATED THE PATIENT DOES NOT HAVE AN INFECTIONS DISEASE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01575 | 2023-04-27 | 2023-03-31 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THE GROSHONG CATHETER WAS RUPTURED IN THE BODY OF THIS PATIENT AND THE BROKEN END WAS REMOVED USING THE INTERVENTIONAL SURGERY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02930 | 2023-04-27 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI12 |  | IT WAS REPORTED THAT TWO BALLOON RUPTURES OCCURRED IN THE SAME PATIENT AND IN THE SAME LOT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02931 | 2023-04-27 | 2023-04-07 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT CAUGHT URINARY TRACT INFECTION DUE TO THEIR INCONTINENCE. IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS REPORTED FOR URINARY TRACT INFECTION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNINGS: TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL.. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
2020394-2023-00293 | 2023-04-27 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE OF A CHRONIC TOTAL OCCLUSION IN THE SUPERFICIAL FEMORAL ARTERY VIA AN IPSILATERAL APPROACH, THE CATHETER WAS ALLEGEDLY UNABLE TO CROSS THE LESION AFTER FIVE MINUTES ATTEMPT. IT WAS FURTHER REPORTED THAT THE CATHETER WAS ALLEGEDLY SEPARATED. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETER PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETER PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED FAILURE TO ADVANCE AND MATERIAL SEPARATION ISSUES COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02906 | 2023-04-27 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998004 | 998004 |  | IT WAS REPORTED THAT PRESSURE COULD NOT BE APPLIED. PRESSURE WAS GIVEN TO THE DEVICE FOR THE DILATION OF URETHRA, BUT NO PRESSURE WAS APPLIED. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 09APR2023, THE BASE OF THE METER PART WAS CRACKED AND DETACHED FROM THE MAIN BODY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02910 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT WHEN 30 HOURS HAD PASSED AFTER THE FOLEY CATHETER WAS PLACED, THE STERILE WATER HAD BEEN DRAINED AND THE BALLOON SPONTANEOUSLY WITHDREW. AFTER THAT, WHEN 10 ML OF STERILIZED WATER WAS INJECTED AGAIN INTO THE MALFUNCTIONING BALLOON FOR VERIFICATION BY MEDICAL PERSONNEL, THE FIXED WATER WAS COMPLETELY REMOVED AFTER 36 HOURS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00295 | 2023-04-27 | 2023-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST 40 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQF5064 | CQF5064 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PROXIMAL TIP OF THE PTA BALLOON ALLEGEDLY HAD CRACK. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 02/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00296 | 2023-04-27 | 2023-03-31 | I, n, j, u, r, y | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG120164 | ATG120164 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE VIA SUBCLAVIAN VEIN AND RIGHT INNOMINATE VEIN, THE PTA BALLOON ALLEGEDLY SEPARATED FROM THE CATHETER INSIDE THE PATIENT'S VEIN. IT WAS FURTHER REPORTED THAT SNARE WAS USED TO RETRIEVE THE BALLOON. REPORTEDLY NEW PUNCTURE WAS MADE THROUGH PATIENT'S FEMORAL VEIN TO REMOVE THE BALLOON. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, PHOTOS AND IMAGES WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 03/2025). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02912 | 2023-04-27 | 2023-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS INSERTED AND BALLOON INFLATED. CLINICIANS DID NOT GET ANY URINE RETURN SO PHYSICIAN ASSUMED IT WAS INSERTED INCORRECTLY. THEY DEFLATED THE BALLOON AND REMOVED THE CATHETER BUT THEN NOTICED THERE WAS NO TIP OR BALLOON ATTACHED. THEY COULD NOT LOCATE ANY REMNANTS IN THE BLADDER VIA CYSTOSCOPE SO THEY WERE SURE IF IT WAS MISSING IN THE BEGINNING OR NOT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02913 | 2023-04-27 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123614A | 123614A |  | IT WAS REPORTED THAT THE PATIENT STATED THAT THESE CATHETERS HAD BEEN GETTING CLOGGED. REPRESENTATIVE ASKED PATIENT IF BLOCKAGE WAS FROM SEDIMENT AND PATIENT ONLY SAW A POWDER OR SAND ONE TIME. PATIENT SAID THAT THE DOCTOR'S OFFICE WAS ABLE TO UNCLOG IT ONE TIME. ADVISED HOW TO CLEAN OUT THE CATHETER PATIENT DID NOT DO IT AND PATIENT WOULD DO THAT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
2020394-2023-00294 | 2023-04-27 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE RX | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U4150530RX | U4150530RX |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED AT NOMINAL PRESSURE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 11/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02921 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI20 | 0165SI20 |  | IT WAS REPORTED THAT THE FOLEY CATHETERS WERE LEAKING AT THE INSERTION POINT. THE CAREGIVER WANTED TO KNOW IF THERE WAS A 22FR CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02918 | 2023-04-27 | 2023-04-06 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE PATIENT HAD BLEEDING, PAINFUL WITHDRAWAL, AND URINARY TRACT INFECTION WHILE USING THIS MAGIC 3 INTERMITTENT CATHETER. UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR URINARY TRACT INFECTION. PER ADDITIONAL INFORMATION RECEIVED VIA SAMPLE FORM ON (B)(6) 2023, IT WAS REPORTED THAT THE INTERMITTENT CATHETER WAS VERY PAINFUL INSERTING AND MORE SO WHEN REMOVING IT AND IT HAS BEEN USING CATHETERS OVER 2 YEARS AND HAD VERY LITTLE PROBLEMS, THE ONES THAT REPLACED THESE WERE SAME THEY HAVE BEEN USING ALL ALONG. STATED THAT THE CUSTOMER HAD URINARY TRACT INFECTION TWICE, URGENT CARE TWICE AND ANTIBIOTIC TWICE. PER FOLLOW-UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT THE CUSTOMER RECEIVED A SAMPLE PACK FROM LIBERATOR AND WAS ABLE TO FIND ANOTHER CATHETER THAT WORKED WITHOUT ANY PAIN OR DISCOMFORT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02920 | 2023-04-27 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 63618 | 63618 |  | IT WAS REPORTED THAT THE CUSTOMER STATED, A PATIENT WAS UNABLE TO USE THE INTERMITTENT CATHETER AS THERE WAS NO ENOUGH LUBRICANT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02926 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L20 | 0119L20 |  | IT WAS REPORTED THAT THE LAST TIME THIS OCCURRED THE FOLEY BALLOON WAS OVER INFLATED, AS TOO LARGE A SYRINGE WAS USED CAUSING EXTRA PRESSURE, THIS HAPPENED THREE TIMES ON THE SAME PATIENT. IT WAS A DUAL LUMEN FOLEY FOR IRRIGATION AND IT WAS A YELLOW RUBBER MATERIAL, NOT SILICONE LIKE MOST OF THEIR FOLEY CATHETERS. A 10ML SYRINGE WAS USED BUT ONLY 6 ML STERILE WATER WAS DRAWN UP IN THE SYRINGE . THE FIRST CATHETER HELD WATER ON TEST BUT RUPTURED ON USE BUT HAD ONLY BEEN INFLATED WITH APPROXIMATELY 3ML OF WATER WHEN IT RUPTURED. THE SECOND FOLEY BALLOON RUPTURED WHEN TESTED AND SOME PRESSURE WAS APPLIED BEFORE USING ON THE PATIENT. THE THIRD FOLEY BALLOON WAS INSPECTED VISUALLY AND DUE TO LOOKING CRACKED OR DRY-ROTTED WAS NOT EVEN TESTED BASED ON THE OTHER FAILURES. THE PACKAGING FOR THE PRODUCT INDICATED TO INFLATE WITH 10ML IN THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02925 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L20 | 0119L20 |  | IT WAS REPORTED THAT THE LAST TIME THIS OCCURRED THE FOLEY BALLOON WAS OVER INFLATED, AS TOO LARGE A SYRINGE WAS USED CAUSING EXTRA PRESSURE, THIS HAPPENED THREE TIMES ON THE SAME PATIENT. IT WAS A DUAL LUMEN FOLEY FOR IRRIGATION AND IT WAS A YELLOW RUBBER MATERIAL, NOT SILICONE LIKE MOST OF THEIR FOLEY CATHETERS. A 10ML SYRINGE WAS USED BUT ONLY 6 ML STERILE WATER WAS DRAWN UP IN THE SYRINGE . THE FIRST CATHETER HELD WATER ON TEST BUT RUPTURED ON USE BUT HAD ONLY BEEN INFLATED WITH APPROXIMATELY 3ML OF WATER WHEN IT RUPTURED . THE SECOND FOLEY BALLOON RUPTURED WHEN TESTED AND SOME PRESSURE WAS APPLIED BEFORE USING ON THE PATIENT. THE THIRD FOLEY BALLOON WAS INSPECTED VISUALLY AND DUE TO LOOKING CRACKED OR DRY-ROTTED WAS NOT EVEN TESTED BASED ON THE OTHER FAILURES . THE PACKAGING FOR THE PRODUCT INDICATED TO INFLATE WITH 10ML IN THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02927 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L20 | 0119L20 |  | IT WAS REPORTED THAT THE LAST TIME THIS OCCURRED THE FOLEY BALLOON WAS OVER INFLATED, AS TOO LARGE A SYRINGE WAS USED CAUSING EXTRA PRESSURE, THIS HAPPENED THREE TIMES ON THE SAME PATIENT. IT WAS A DUAL LUMEN FOLEY FOR IRRIGATION AND IT WAS A YELLOW RUBBER MATERIAL, NOT SILICONE LIKE MOST OF THEIR FOLEY CATHETERS. A 10ML SYRINGE WAS USED BUT ONLY 6 ML STERILE WATER WAS DRAWN UP IN THE SYRINGE. THE FIRST CATHETER HELD WATER ON TEST BUT RUPTURED ON USE BUT HAD ONLY BEEN INFLATED WITH APPROXIMATELY 3ML OF WATER WHEN IT RUPTURED. THE SECOND FOLEY BALLOON RUPTURED WHEN TESTED AND SOME PRESSURE WAS APPLIED BEFORE USING ON THE PATIENT. THE THIRD FOLEY BALLOON WAS INSPECTED VISUALLY AND DUE TO LOOKING CRACKED OR DRY-ROTTED WAS NOT EVEN TESTED BASED ON THE OTHER FAILURES . THE PACKAGING FOR THE PRODUCT INDICATED TO INFLATE WITH 10ML IN THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02916 | 2023-04-27 | 2023-04-05 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT PATIENT GOT AN INFECTION FROM THESE INTERMITTENT CATHETERS. IT WAS UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "PATIENT SENSITIVITY TO MATERIALS, MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT IMPROPER SURFACE TREATMENT VIA CHLORINATION IMPROPER LEACHING OF THE PRODUCT IMPROPER". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DONT REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02928 | 2023-04-27 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0119L20 | 0119L20 |  | IT WAS REPORTED THAT THE LAST TIME THIS OCCURRED THE FOLEY BALLOON WAS OVER INFLATED, AS TOO LARGE A SYRINGE WAS USED CAUSING EXTRA PRESSURE. THIS HAPPENED THREE TIMES ON THE SAME PATIENT. IT WAS A DUAL LUMEN FOLEY FOR IRRIGATION AND IT WAS A YELLOW RUBBER MATERIAL, NOT SILICONE LIKE MOST OF THEIR FOLEY CATHETERS. A 10ML SYRINGE WAS USED BUT ONLY 6 ML STERILE WATER WAS DRAWN UP IN THE SYRINGE. THE FIRST CATHETER HELD WATER ON TEST BUT RUPTURED ON USE BUT HAD ONLY BEEN INFLATED WITH APPROXIMATELY 3ML OF WATER WHEN IT RUPTURED. THE SECOND FOLEY BALLOON RUPTURED WHEN TESTED AND SOME PRESSURE WAS APPLIED BEFORE USING ON THE PATIENT. THE THIRD FOLEY BALLOON WAS INSPECTED VISUALLY AND DUE TO LOOKING CRACKED OR DRY-ROTTED WAS NOT EVEN TESTED BASED ON THE OTHER FAILURES. THE PACKAGING FOR THE PRODUCT INDICATED TO INFLATE WITH 10ML IN THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02971 | 2023-04-28 | 2023-04-06 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETER CAUSED SOME REDNESS IN PATIENT. CUSTOMER DID USE A MEDICATED CREAM. PATIENT HAS BEEN USING MORE THAN 90 DAYS. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS REPORTED THAT PATIENT IS STILL USING THE CATHETERS BUT APPLIES MEDICATED CREAM BEFORE EACH USE TO ASSIST WITH IRRITATION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | Y | N | LAY USER/PATIENT |  | 2023-04
1018233-2023-02969 | 2023-04-28 | 2023-04-06 | I, n, j, u, r, y | Manufacturer report | BARDIA® COUDE OLIVE TIP URETHRAL FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 802516 | 802516 |  | IT WAS REPORTED THAT PATIENT HAD BEEN USING THE INTERMITTENT CATHETER FOR YEARS AND CALLED AND STATED THAT THEY PUT THE CATHETER IN, FELT RESISTANCE AND HAD SOME BLEEDING. REPRESENTATIVE NOT SURE IF IT WAS THE CATHETER AND IT COULD BE SOMETHING ELSE CAUSING THE PATIENT TO BLEED. THE PATIENT SAID WHEN THEY ROTATE IT, THEY FEEL IT HAD A HARD TIME MOVING INSIDE OF THEM AS WELL. ADVISED THE PATIENT TO NOT EVER USE FORCE. THE CATHETER IN THE PACKAGE SEEMS TO BE WAVY, NOT STRAIGHT. STATED THAT THEY DID NOT HAVE THE EXACT DATE OF THE INCIDENT, BUT IT WAS SOMETIME BETWEEN (B)(6) 2023 AND (B)(6) 2023. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW UP VIA MAIL ON 14APR2023, PATIENT HAD A HARD TIME INSERTING CATHETER, BUT THEY DID GET IT IN THE BLADDER FOLLOWED UP BY BLEEDING. GAVE THEIR SOMETHING ELSE THAT HELPED PASS IT SOMETHING THROUGH, PASSED A BLOOD CLOT AND THEN GOT A URINARY TRACT INFECTION. PATIENT WAS ON ANTI-BIOTIC. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02968 | 2023-04-28 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER CREEVY MODEL | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 01267520 | 1267520 |  | IT WAS REPORTED THAT THE WATER INJECTED INTO THE BALLOON HAS HIGH RESISTANCE DUE TO A DUST FOUND. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 14APR2023,WATER INJECTED INTO THE BALLOON HAS HIGH RESISTANCE ('DUST WAS FOUND'). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00301 | 2023-04-28 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQ5054 | CQ5054 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON WAS ALLEGEDLY RUPTURED AT 30 ATM DURING THE THIRD INFLATION. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A VIDEO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 01/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01595 | 2023-04-28 | 2023-04-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT SOMETIME POST DIALYSIS CATHETER PLACEMENT, THE DIALYSIS CATHETER ALLEGEDLY BROKE IN EXTENSION PART. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY LEAKED. REPORTEDLY, THE CATHETER WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 02/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01596 | 2023-04-28 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT PRIOR TO PORT PLACEMENT PROCEDURE, THE INSTRUCTIONS IN IFUS TO USE HEPARIN ALLEGEDLY DISAGREES WITH THE RECOMMENDATIONS OF CERTAIN AUTHORITIES. IT WAS FURTHER REPORTED THAT THE IFUS ARE ALLEGEDLY INADEQUATE AND SHOULD STATE DIFFERENT INSTRUCTIONS. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01583 | 2023-04-28 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT ON (B)(6)2023, THE PATIENT WHO HAD A PICC CATHETER SINCE (B)(6)2023 CAME FOR MEDICATION VIA THE EMERGENCY ROOM, ROUTINE MEDICATION WAS ADMINISTERED (MORPHINE, TILETHYL, DIAZEPAM, CIMITIDINE, INTRAVENOUS DIGESAN), THE SALINIZATION BY WHIRLING TECHNIQUE WITHOUT RESISTANCE, AFTER THE PICC WAS DRESSED BY THE TRAINED NURSE WITHOUT INTERCURRENCE, ACCOMPANIED BY THE CONSCIOUS AND ORIENTED PATIENT, THEN A NEW SALINIZATION WAS PERFORMED AND PRESENTED RESISTANCE. THE UN OBSTRUCTION TECHNIQUE WAS PERFORMED ON THE SAME DAY, THE NEXT DAY AND THE PICC WAS PULLED 2 CM WITHOUT SUCCESS IN UN OBSTRUCTING. WE PERFORMED THE PASSAGE OF A NEW PICC CATHETER, WHEN REMOVING THE PICC IT WAS VERIFIED THAT IT WAS ¿OBSTRUCTED¿, SALINIZATION WAS PERFORMED, WITH RESISTANCE WE TRIED TO PASS THE GUIDE WIRE USED IN THE OTHER PICC, WITHOUT SUCCESS. IT WAS OBSERVED THAT AT THE JUNCTION OF THE CATHETER WITH THE PART THAT IS CLAMPED, THE CATHETER APPARENTLY ¿IT CLOSED." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00298 | 2023-04-28 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONQUEST | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | CQ5064 | CQ5064 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE FOR VASCULAR STENOSIS OF CEPHALIC VEIN, THE PTA BALLOON ALLEGEDLY RUPTURED. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A VIDEO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 07/2025). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02974 | 2023-04-28 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE CATHETER WAS WITHDRAWN DUE TO BALLOON RUPTURE 7 DAYS AFTER PLACEMENT IN THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
2020394-2023-00302 | 2023-04-28 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U35130104 | U35130104 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE VIA RIGHT FEMORAL VEIN, THE PTA BALLOON ALLEGEDLY RUPTURED AT 6 ATM. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 04/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01597 | 2023-04-28 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 4808060 | 4808060 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PATIENT ALLEGEDLY EXPERIENCED ITCHING OVER THE PORT SITE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02975 | 2023-04-28 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | UNKNOWN FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNKNOWN |  | IT WAS REPORTED THAT FOLEY CATHETERS THAT WERE INSTALLED IN THE PATIENT, THE DEFECT SHOWED ITSELF UPON REMOVAL AND IT TOOK A LONG TIME TO EMPTY THE BALLOON. SO THE USER MADE A CUT WITH SCISSORS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01594 | 2023-04-28 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT PRIOR TO PORT PLACEMENT PROCEDURE, THE PATIENT PAPERWORK AND IFU WERE ALLEGEDLY MISSING. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 08/2024). DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01592 | 2023-04-28 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT TI ISP, 8 FR CHRONOFLEX, SUTURE PLUG, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1708060 | 1708060 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE DEVICE ALLEGEDLY HAD ISSUES WITH FLUSHING. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 06/2022). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02973 | 2023-04-28 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT ONLY GENERALLY THE PATIENT STATED THE FOLEY CATHETER WAS DEFECTIVE AND THEY LEAKED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02967 | 2023-04-28 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | TEMPERATURE SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 119316M |  | IT WAS REPORTED THAT THE USER FOUND THE FOLEY CATHETER OUT OF PATIENT, LYING IN THE BED AND THE BALLOON OF THE FOLEY WAS NOT INFLATED. THE USER INFLATED THE BALLOON WITH AIR AND AFTER 30 MINUTES FOUND THAT NO AIR REMAINED IN THE BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02978 | 2023-04-28 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS CONTAMINATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02980 | 2023-04-28 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | FEMALE EXTERNAL CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT PATIENT WAS EXPERIENCED LEAKING FROM THE PUREWICK FEMALE EXTERNAL CATHETER. STATED THAT THEY DID TROUBLESHOOT WITH WICKS PLACEMENT, AND THEY WOULD TRY IF STILL EXPERIENCING LEAKING WOULD CALL BACK. CUSTOMER STATED THAT THEY NEVER RECEIVED INSTRUCTION BOOK AND SENT REQUEST FOR NEW INSTRUCTION BOOK TO BE SENT . IT WAS NOTED THAT THE PATIENT HAD BEEN USING THE PUREWICK PRODUCTS FOR MORE THAN 90 DAYS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02987 | 2023-04-28 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS REMOVED, WATER WAS WITHDRAWN FROM THE BALLOON, BUT THE CATHETER WAS STUCK AND COULD NOT BE REMOVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02977 | 2023-04-28 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE LUBE IN THE INTERMITTENT CATHETER WAS HARD TO OPEN AND THE LUBE DRIED OUT TOO QUICK. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02979 | 2023-04-28 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE LUBE IN THE INTERMITTENT CATHETER WAS HARD TO OPEN AND THE LUBE DRIED OUT TOO QUICK. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02970 | 2023-04-28 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51612 | 51612 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE NOT WORKING. CUSTOMER SAID WHEN USING THE URINE WENT EVERYWHERE AND NOT EASY TO LUBRICATE. REPRESENTATIVE OFFERED TO SEND A SAMPLE OF DIFFERENT STYLE CATHETERS AND PATIENT APPROVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02972 | 2023-04-28 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165816 | 165816 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BROKE AT THE END CAUSED THEM TO USE ANOTHER INSERTION TRAY NOW THE PATIENT WAS SHORT ONE TRAY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01593 | 2023-04-28 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT SOME TIME POST PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT THE PORT BODY AND THE DISTAL CATHETER SEGMENT WERE REMOVED. REPORTEDLY, THE PORT WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02991 | 2023-04-28 | 2023-04-10 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THE PATIENT RECEIVED A URINARY TRACT INFECTION FROM THE INTERMITTENT CATHETER AND PREFERRED PRE LUBED CATHETERS. THE REPRESENTATIVE SENT MORE SAMPLES FOR PATIENT TO TRY NEW CATHETERS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO " PATIENT SENSITIVITY TO MATERIALS MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE. INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-04
MW5117160 | 2023-04-28 | 2023-03-29 | I, n, j, u, r, y | Voluntary report | ROTAREX | CATHETER, PERIPHERAL, ATHERECTOMY | BARD PERIPHERAL VASCULAR, INC. | 302272 | 80237 |  | DURING A LEFT LOWER EXTREMITY ANGIOGRAM, THE ROTAREX DEVICE WIRE BROKE LEAVING A SEGMENT IN THE VESSEL. PATIENT HAD TO RETURN TO THE OPERATING ROOM FOR REMOVAL OF LEFT LOWER EXTREMITY INTRA-ARTERIAL FOREIGN BODY AND A LEFT COMMON FEMORAL ENDARTERECTOMY WITH COMMON FEMORAL TO BELOW-KNEE POPLITEAL BYPASS WITH POLYTETRAFLUOROETHYLENE. |  | Y | N | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02989 | 2023-04-28 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE LUBE IN THE INTERMITTENT CATHETER WAS HARD TO OPEN AND THE LUBE DRIED OUT TOO QUICK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02990 | 2023-04-28 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THE LUBE IN THE INTERMITTENT CATHETER WAS HARD TO OPEN AND THE LUBE DRIED OUT TOO QUICK. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "SEAL STRENGTH TOO HIGH". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-02996 | 2023-04-28 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® HEMATURIA FOLEY CATHETER | HEMATURIA FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 2551H22 |  | IT WAS REPORTED THAT THE FOLEY CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT AFTER 3 HOURS OF PLACEMENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02999 | 2023-04-28 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX TEMP-SENSING CATHETER | UNKNOWN LATEX TEMP-SENSING CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE ARCTIC SUN DEVICE HAD LOW FLOW BANNER ON SCREEN, AND THEY HAVE HAD SOME ALERTS, BUT IT HAPPENED WHEN THE PATIENT HAD A BOWEL MOVEMENT. THE PATIENT TEMPERATURE WAS 37.5C, TARGET TEMPERATURE WAS 37C, WATER TEMPERATURE WAS 12.2C AND WATER FLOW RATE WAS 1 L/M. MIS DISCUSSED IMPORTANCE OF COVERAGE AND WRAPPING PADS AROUND THE PATIENT. NURSE STATED THAT THEY WERE NOT ACTIVELY COOLING BUT JUST USING FOR TEMPERATURE REGULATION. MIS WALKED NURSE THROUGH ACCESSING EVENT LOG. ONLY ALERTS WERE ALERT 50 (PATIENT TEMPERATURE 1 ERRATIC) AND ALARM 15 (UNABLE TO OBTAIN A STABLE PATIENT TEMPERATURE). NURSE CONFIRMED THESE ALERTS OR ALARMS OCCURRED WHEN TEMPERATURE PROBE WAS DISPLACED. MIS WALKED NURSE THROUGH STOP THERAPY, DRAIN, DISCONNECT AND RECONNECT. NURSE VERIFIED ALL LINES STRAIGHT WITH NO BENDS OR KINKS. THE FLOW RATE STABILIZED AT 1.1 L/M . MIS EXPLAINED MOST LIKELY BASED ON THE SIZE OF THE SINGLE PAD BEING USED AND MOVEMENT FROM THE INFANT THAT THE FLOW RATE WAS OK TO PROCEED WITH THERAPY. NOTED THAT IF FLOW RATE DROPS BELOW 1 L/M NURSE SHOULD CALL BACK. NURSE COULD NOT CONFIRM WHAT SIZE PAD WAS BEING USED. ONLY ONE PAD CONNECTED, AND IT WAS LYING FLAT. PATIENT WAS ROLLING AROUND ON IT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-03009 | 2023-04-28 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT IN A DAY AFTER PLACEMENT, WATER IN THE BALLOON LEAKED OUT. THE USER SUSPECTED THERE MIGHT BE A PINHOLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-03000 | 2023-04-28 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 175816 |  | IT WAS REPORTED THAT END OF CATHETER WAS MISSING AND THERE WAS NO BALLOON. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-03006 | 2023-04-28 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE NOT EMPTYING BLADDER COMPLETELY. THE PATIENT WANTED TO GO BACK TO CURE ULTRA FOR NEXT ORDER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿INSUFFICIENT TIP STRENGTH DUE TO WALL THICKNESS AND DUROMETER". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE REVIEW THE SELF CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW (APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS.". H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
1018233-2023-03014 | 2023-04-28 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT UPON INFLATING THE BALLOON DURING PRETEST, IT INFLATED ASYMMETRICALLY. THEN, WHEN ATTEMPTED TO DEFLATE THE BALLOON, IT WAS DIFFICULT TO REMOVE THE WATER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02995 | 2023-04-28 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI12 |  | IT WAS REPORTED THAT FOLEY CATHETER BALLOON WAS NOT INFLATED DURING PRETEST. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 17APR2023, CATHETER WAS NOT USED ON PATIENT. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT WAS UNCONFIRMED A LABELING REVIEW WAS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-02983 | 2023-04-28 | 2023-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53616G | 53616G |  | IT WAS REPORTED THAT THE PATIENT HAD A HARD TIME WITH MAGIC 3 INTERMITTENT CATHETER AND EXPERIENCED LEAKAGE AND THE CATHETER PACKAGING WAS HARD TO OPEN. REPRESENTATIVE OFFERED TO CHANGE CATHETER, BUT THE PATIENT DECLINED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "POOR INTERFACE BETWEEN THE CATHETER SHAFT AND THE DRAINAGE FUNNEL/HANDLE. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-04
3006260740-2023-01598 | 2023-04-28 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 7F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 0603870CE |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR AND TWO MONTHS POST PORT CATHETER PLACEMENT, DURING REMOVAL A LARGE PART OF THE CATHETER ALLEGEDLY BROKE AND REMAINED IN THE VEIN. IT WAS FURTHER REPORTED THAT THE CATHETER WAS REMOVED. REPORTEDLY, THE PATIENT EXPERIENCED PAIN IN THE RIGHT CLAVICLE BUT THERE IS NO INTERVENTION OR MEDICATION PRESCRIBED FOR THE TREATMENT FOR PAIN. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US, BUT IS SIMILAR TO THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 7F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510K NUMBER FOR THE M.R.I. LOW-PROFILE IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 7F PRODUCTS IS IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 01/2025). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01603 | 2023-04-30 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT "AFTER XRAY, THE ADDRESSING WAS OPENED AND THE PORTION FROM THE SCALE OF 29 ~ 30 CM WAS FOUND CRACKED. A NEW PRODUCT, INTRODUCER KIT AND INTRODUCER WERE ACUDUCTOR FOR CATHETER PLACEMENT." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED CATHETER IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. TWO PHOTOGRAPHS OF A 4 FR GROSHONG NXT CLEARVUE CATHETER WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPHS SHOWED A LARGE CIRCUMFERENTIAL SPLIT IN THE CATHETER TUBING NEAR THE 30 CM DEPTH MARKER. ONE PHOTOGRAPH SHOWED THE CATHETER INSERTED INTO A PATIENT WITH THE SPLIT ABOUT 2.5 CM PROXIMAL TO THE INSERTION SITE. THE OTHER PHOTOGRAPH SHOWED THE ENTIRE CATHETER ASSEMBLY AND DEMONSTRATED THE SPLIT WENT AROUND NEARLY THE ENTIRE CIRCUMFERENCE OF THE CATHETER TUBING. THE SPLIT WAS OBSERVED TO BE SLIGHTLY ANGLED WITH MOSTLY STRAIGHT AND EVEN EDGES; HOWEVER, FURTHER DETAILS OF THE FRACTURE SURFACES COULD NOT BE DISCERNED FROM THE RETURNED PHOTOGRAPHS. WHILE THE EXACT CAUSE OF THE OBSERVED DAMAGE COULD NOT BE DETERMINED FROM THE RETURNED PHOTOGRAPHS, POSSIBLE CAUSES OF THE SPLIT INCLUDE OVER-PRESSURIZATION, EXCESSIVE TENSILE FORCES, MATERIAL FATIGUE, AND SHARP INSTRUMENT DAMAGE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
3006260740-2023-01604 | 2023-04-30 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED, "THE HEAD NURSE OF ONCOLOGY DEPARTMENT PERFORMED THE CATHETER PLACEMENT. DURING PLACEMENT, THE GUIDE WIRE WAS INSERTED AFTER SUCCESSFUL PUNCTURE. THE OPERATOR PLANNED TO CROSS THE DISTAL END OF THE GUIDE WIRE INTO THE DILATOR OF MICRO-INTRODUCER SHEATH, BUT FAILED. AFTER THE DISTAL END OF THE GUIDE WIRE WAS CUT OFF, THE CATHETER WAS PLACED SUCCESSFULLY.". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-04
1018233-2023-03021 | 2023-05-01 | 2023-03-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT THE PATIENT USED A DISPOSABLE STERILE URINARY CATHETER AFTER SURGERY ON (B)(6) 2023. THE BALLOON WAS INTACT DURING USE, NO PETROLEUM-BASED LUBRICATING OIL WAS USED, AND THE WATER INJECTION DID NOT EXCEED 30ML. HOWEVER, THE BALLOON RUPTURED AFTER ABOUT A FEW HOURS OF USE. NO MISSING BALLOON PIECES INSIDE THE PATIENT'S BODY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03017 | 2023-05-01 | 2023-04-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 1758SI16 | 1758SI16 |  | IT WAS REPORTED THAT THE REGISTERED NURSE STATED THAT THEY INFLATED THE BALLOON WITH NO ISSUES, AFTER IT WAS INSERTED INTO THE PATIENT, THEY DID GET URINE BACKFLOW. HOWEVER, IT IMMEDIATELY FELL OUT WHEN THEY WERE SETTLING THE PATIENT FOR SURGERY. THEY REPORTED THAT THERE WAS NO TENSION WHEN FILLING THE BALLOON, IT WAS FILLED VERY EASILY AND LAST WEEK THEY HAD AN ISSUE WITH A FOLEY CATHETER USED ON A PATIENT IN L&D. THE DEVICE WAS INSERTED INTO THE PATIENT, THE BALLOON WAS FILLED WITH WATER, AND IT IMMEDIATELY FELL OUT. IT WAS UNCLEAR IF THE BALLOON BURST, OR IT WAS DEFECT IN THE TUBE, AND THEY INITIALLY THOUGHT THE DEVICE BROKE OFF OR BURST INSIDE THE PATIENT. HOWEVER, AFTER CLOSER INSPECTION AND MULTIPLE CONFIRMATORY TESTS IT WAS FOUND THAT THERE WAS NOTHING RETAINED INSIDE THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03024 | 2023-05-01 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123518A |  | IT WAS REPORTED THAT THE NURSE TRIED TO PUMP OUT THE BALLOON 5 TIMES AND WAS STILL UNABLE TO REMOVE THE FOLEY CATHETER. ABOUT 7CC OF DISTILLED WATER HAD BEEN PUMPED OUT. AFTER APPLYING FORCE PULLING , THE FOLEY CATHETER WAS FINALLY REMOVED. AFTER REMOVAL, IT WAS FOUND THAT THE BALLOON WAS NOT DEFLATED COMPLETELY, AND STILL SOME DISTILLED WATER AND AIR WERE INSIDE THE BALLOON. THE NURSE NEEDED TO APPLY FORCE FOR PULLING THE FOLEY CATHETER TO REMOVE IT FROM PATIENT. NURSE TAKE MORE ATTENTION ON FUTURE INFECTION CONTROL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01601 | 2023-05-01 | 2023-04-10 | I, n, j, u, r, y | Manufacturer report | BARD ACCUCATH | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT THE GUIDEWIRE BROKE OFF IN PATIENT WHEN NEEDLE WAS RETRACTED. THE GUIDEWIRE WAS COILED IN THE TISSUE. PATIENT HAD TO HAVE GUIDEWIRE REMOVED FROM TISSUE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01599 | 2023-05-01 | 2023-04-07 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT COMPLAINED OF CHEST DISTRESS THE DAY FOLLOWING PLACEMENT OF THE CATHETER. THE CATHETER WAS PULLED OUT BY 5 CM. AFTER THE PATIENT WAS SENT BACK TO THE WARD FOLLOWING POSITIONING BY X-RAY, IT WAS FOUND THAT 2 CM OF THE CATHETER GOT OUT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03026 | 2023-05-01 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119112 | 119112 |  | IT WAS REPORTED THAT WATER WAS DIFFICULT TO REMOVED FROM THE FOLEY CATHETER DURING THE PRETEST. THE CATHETER WAS NOT USED ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03022 | 2023-05-01 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT ORIGINAL BARD FOLEY CATHETERS (14 G) COULD BE INSERTED INTO 16 G TRANSLUMINAL ACCESS CATHETERS SMOOTHLY. HOWEVER, THE BALLOON OF THIS BATCH OF GOODS RECEIVED BY THE HOSPITAL WAS TOO LARGE TO INSERT INTO THE 16 G ACCESS CATHETERS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01600 | 2023-05-01 | 2023-04-09 | I, n, j, u, r, y | Manufacturer report | PROV 4F DL IR 135CM NI | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S3274335 |  | IT WAS REPORTED BY THE CUSTOMER THE POWER PICC THAT WAS PLACED AND NOW, THE POWER PICC IS HAVING TROUBLE INFUSING AND DRAWING BACK. PATIENT HAS CRITICAL MEDICATIONS RUNNING IN THEIR POWER PICC CONTINUOUSLY AND WHEN THEY DON'T INFUSE PROPERLY, THEY MISS THEIR MEDICATIONS AND DECOMPENSATE. PATIENT IS UNDER A YEAR OF AGE AND THE LINE IS PLACED IN THEIR FEMORAL VEIN. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01612 | 2023-05-02 | 2023-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERHICKMAN CV CATHETER, POLYURETHANE SINGLE-LUMEN, 8F | CHRONIC CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 0805810 | 0805810 |  | IT WAS REPORTED THAT SOMETIME POST CHRONIC CATHETER PLACEMENT, DURING REMOVAL OF THE CATHETER, THE CUFF ALLEGEDLY REMAINED IN PLACE WHICH REQUIRED SURGICAL REMOVAL. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00304 | 2023-05-02 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8088 | VA8088 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON WAS ALLEGEDLY DIFFICULT TO REMOVE THROUGH THE SHEATH AFTER INFLATION. IT WAS FURTHER REPORTED THAT THE BALLOON ALLEGEDLY GOT STUCK IN SHEATH. REPORTEDLY, THE SHEATH HAD TO BE REMOVED FROM THE PATIENT IN ORDER TO GET BALLOON OUT OF PATIENT ARM. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE VACCESS PTA DILATATION CATHETER WITH AN UNKNOWN SHEATH LOADED OVER THE BALLOON WAS RECEIVED FOR EVALUATION. THE SHEATH WAS NOTED TO HAVE SIGNS OF BUCKLING AT ITS DISTAL END. DURING FUNCTIONAL EVALUATION, NEGATIVE PRESSURE WAS APPLIED USING AN IN-HOUSE PRESTO INFLATION DEVICE. AN ATTEMPT WAS MADE TO REMOVE BALLOON FROM THE SHEATH AND SOME RESISTANCE WAS FELT BUT THE REMOVAL WAS SUCCESSFUL. NO FURTHER FUNCTIONAL TESTING WAS PERFORMED. ALSO, ONE PHOTO WAS PROVIDED AND REVIEWED. THE PHOTO SHOWS THE VACCESS DEVICE WITH AN UNKNOWN SHEATH LOADED OVER THE BALLOON. NO OTHER ANOMALIES NOTED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED SHEATH REMOVAL DIFFICULTY AS THE DEVICE WAS RETURNED WITH AN UNKNOWN SHEATH LOADED OVER THE BALLOON, WITH THE SHEATH ALSO HAVING SIGNS OF BUCKLING AT ITS DISTAL END. THE PHOTO REVIEW ALSO CONFIRMS THE FAILURE. A DEFINITIVE ROOT CAUSE FOR THE REPORTED SHEATH REMOVAL DIFFICULTY COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 11/2025), G3. H11: B5, H6 (METHOD, RESULT, CONCLUSION). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00305 | 2023-05-02 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VACCESS | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | VA8088 | VA8088 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON WAS ALLEGEDLY FAILED TO BE REMOVED THROUGH SHEATH AFTER INFLATION. IT WAS FURTHER REPORTED THAT THE BALLOON HAD ALLEGEDLY STUCK IN THE SHEATH. REPORTEDLY, SHEATH HAD TO BE REMOVED FROM PATIENT IN ORDER TO GET BALLOON OUT OF PATIENT ARM. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED SHEATH REMOVAL ISSUE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 11/2025), G3, H6 (METHOD) H11: E1 H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
16852611 | 2023-05-02 | 2022-12-22 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | FOLEY SIL 2-WAY 16 FR 5CC | CATHETER, RETENTION TYPE, BALLOON | C. R. BARD, INC. |  |  |  | RN (REGISTERED NURSE) PLACED THE FOLEY AND THEN THE CHEMO DRUG, GEMZAR, WAS INSTILLED IN THE BLADDER AT 1132. THERE WAS A CLAMP PLACED ON THE FOLEY TUBING. RN RETURNED FROM LUNCH AT 1230, ANOTHER RN REPORTED THAT THE CLAMP NEEDED TO BE REMOVED FROM THE FOLEY TUBING AT 1232. RN REMOVED THE CLAMP FROM THE FOLEY AND SAW URINE PASSING THROUGH THE FOLEY TUBING. LATER, THE RESIDENT ASKED THE RN WHERE WAS THE FOLEY. RN FOUND THE FOLEY BY THE HEAD OF THE BED AND SHOWED RESIDENT. HE NOTED THAT THERE WAS ABOUT 175 IN THE FOLEY BAG. RN REMINDED HIM THAT 50CC OF GEMZAR HAD BEEN INSTILLED IN THE BLADDER, SO HE DOCUMENTED 125 OF URINE OUTPUT. AROUND 1500 THE ROBOT WAS UNDOCKED, AND THE PATIENT WAS REPOSITIONED, AND AT THAT TIME IT WAS NOTED THAT THE FOLEY WAS ON THE GROUND. AT 1505, RESIDENT REPLACED THE FOLEY AND IMMEDIATELY 500CC OF URINE CAME OUT OF THE BLADDER INTO THE NEWLY PLACED BAG. SURGEON WAS PRESENT AND DID NOT STATE THAT ANYTHING ELSE NEEDED TO BE DONE. RN ASKED RESIDENT HOW MUCH URINE OUTPUT HE HAD DOCUMENTED IN TOTAL. HE HAD CHARTED THE PREVIOUS 125 FOR THE 1130 OUTPUT. BECAUSE OF THE CHANCE THAT GEMZAR MIGHT HAVE BEEN IN THE BLADDER LONGER THAN THE RECOMMENDED ONE-HOUR INSTILLATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03036 | 2023-05-02 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER COULD NOT BE WITHDRAWN FROM THE FOLEY CATHETER DURING PRETEST. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT WAS UNCONFIRMED A LABELING REVIEW WAS NOT REQUIRED. THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03048 | 2023-05-02 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 123516A | 123516A |  | IT WAS REPORTED THAT THE FOLEY CATHETER DEFECTIVE, WHEN THEY PUT WATER, IT JUST LEAKED OUT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE INSUFFICIENT LATEX STRENGTH, LOW/NON-UNIFORM RUBBERIZE THICKNESS, WIRE PRESSING TECHNIQUE, STRIPPING TECHNIQUE ETC. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THEREFORE, NO ADDITIONAL ACTION IS REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE LATEX AND MAY CAUSE BALLOON TO BURST. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. SINGLE PATIENT USE ONLY. DO NOT REUSE. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. VALVE TYPE: USE LUER SYRINGE. DO NOT USE NEEDLE. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. DO NOT USE IF PACKAGE IS OPENED OR DAMAGED. RECOMMENDED INFLATION CAPACITIES 5CC BALLOON: USE 10CC STERILE WATER 30CC BALLOON: USE 35CC STERILE WATER DO NOT EXCEED RECOMMENDED CAPACITIES. NOTE: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LAWS AND REGULATIONS. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT." THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03037 | 2023-05-02 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V18S |  | IT WAS REPORTED THAT THE PATIENT GOT CATHETERIZATION ON (B)(6) AND AFTER ABOUT 1 HOUR, THE PATIENT FELT THAT THERE WAS SOMETHING BURST IN THE BLADDER. THEN THE CATHETER WAS REMOVED AND FOUND THAT THE BALLOON WAS BURST BUT NO HARM CAUSED ON THE PATIENT PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 23APR2023, DOA WAS (B)(6) 2023, THE DATE WE KNEW THE CASE FROM THE AUTHORITY PLATFORM. NO MISSING PIECES OF BALLOON FOUND IN THE BODY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE THIN SAC ISSUE DUE TO ¿SHORT LATEX DWELL TIME" OR "LATEX DIP SPEED OUT TOO SLOW" WHICH CONTRIBUTE TO BURST BALLOON DEFECT. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01616 | 2023-05-02 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PUMPS ALLEGEDLY STARTED RINGING OCCLUDED IN THE MIDDLE OF THE TREATMENT. IT WAS FURTHER REPORTED THAT BLOOD RETURN AND FLUSH DID NOT WORKED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01668 | 2023-05-02 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "MICROINTRODUCER BENDING/KINKING". NO OTHER INFORMATION WAS PROVIDED. | H11: THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01676 | 2023-05-02 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5608300 |  | IT WAS REPORTED BY THE CUSTOMER "ON APRIL 6, THE COLLEAGUE WILL INSERT A BARD BRAND DIALYSIS CATHETER, POWER-TRIALYSIS 30 CM. THE GUIDE IS INSERTED INTO THE FEMORAL VEIN. THEN THE USUAL DILATATION WITH THE ACCOMPANYING DILATORS. IT IS NOT POSSIBLE TO INSERT THE DIALYSIS CATHETER ON THE GUIDE. THE DISTAL END OF THE GUIDE IS VERY SOFT, BUT WE SNAP IT OFF WITH A STERILE CUTTER AND THEN THE DIALYSIS CATHETER CAN BE THREADED OVER THE GUIDE WITH RESISTANCE. THE CATHETER IS VERY SLOW TO PASS OVER THE GUIDE. THE GUIDE CANNOT THEN BE BACKED OUT OF THE DIALYSIS CATHETER BUT ONLY GIVES A FAILING RESPONSE, SUCH AS THE GUIDE BEING STRETCHED. THE GUIDE AND CATHETER ARE THEN PULLED OUT AS ONE UNIT. THE PROCEDURE HAD TO BE STOPPED." ADDITIONAL INFORMATION RECEIVED 17-APR-2023: IT WAS REPORTED BY BD REP, "THERE WAS NO PATIENT HARM." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DIFFICULT CATHETER ADVANCEMENT OVER THE GUIDEWIRE WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01675 | 2023-05-02 | 2023-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SINGLE-LUMEN CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6174355 |  | IT WAS REPORTED BY THE CUSTOMER "DURING AN INFUSION OF MOGAMULIZUMAB A PERIPHERAL CENTRAL LINE LEAKAGE IS OBSERVED. A RADIOLOGICAL CHECK TO ENSURE ITS PERMEABILITY IS CARRIED OUT AND SHOWS EXTRAVASATION INTO THE PATIENT'S ARM. THE CATHETER IS THEN REMOVED AND A CRACK IS OBSERVED AT THE LEVEL OF THE EIGHTH CENTIMETER OF THE BODY OF THE CATHETER. NO CLINICAL CONSEQUENCES FOR THE PATIENT. REMOVAL OF THE CATHETER AND TRANSFER OF THE INFUSION TO THE PERIPHERAL ROUTE AFTER APPROVAL BY THE PHARMACIST RESPONSIBLE FOR THE CHEMOTHERAPY." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A LEAKING CATHETER WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS FOUR PHOTOGRAPHS WHICH DEPICTED A 4FR S/L POWERPICC CATHETER. TWO PHOTOGRAPHS DEPICTED THE ENTIRE DEVICE. THE OTHER TWO PHOTOGRAPHS DEPICTED A CLOSER VIEW OF A PORTION OF CATHETER SHAFT. A PARTIALLY CIRCUMFERENTIAL SPLIT WAS EVIDENT IN THE PHOTOGRAPHS. A LONGITUDINALLY ALIGNED SPLIT WAS VISIBLE AT ONE CORNER OF THE CIRCUMFERENTIAL SPLIT. MATERIAL BUCKLING WAS VISIBLE IN THE VICINITY OF THE SPLIT. CATHETER DAMAGE WAS EVIDENT IN THE SUBMITTED PHOTOGRAPHS; HOWEVER, INSPECTION OF THE PHOTOGRAPHS WAS INSUFFICIENT TO IDENTIFY THE CAUSE OF THE DAMAGE. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE MATERIAL FATIGUE FROM CYCLIC KINKING AND CATHETER COMPRESSION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01673 | 2023-05-02 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7274118 |  | CUSTOMER REPORTED A 4FR DOUBLE LUMEN CATHETER WAS INSERTED IN THE RIGHT ARM (B)(6) 2023 IN A MINOR PATIENT. THE DEVICE WAS REVIEWED DUE TO IT BEING OBSTRUCTED AND THE CHILD PRESENTING WITH DISCOMFORT AT THE MOMENT OF IRRIGATING (PAIN AT THE LEVEL OF THE ARM). CLINICIAN FOUND THE CATHETER UNUSABLE, "COVERED WITH AN INTEGRAL TRANSPARENT DRESSING, LAST HEALED ON (B)(6) 2023, DEVICE AT ZERO LEVEL, INSERTION POINT WITH HUMIDITY." AT THE TIME OF IRRIGATING THROUGH THE LUMEN WHITE IN COLOR, A SLIGHT DISCHARGE OF LIQUID IS OBSERVED FROM THE INSERTION POINT. GIVEN THIS AND A HISTORY OF DISCOMFORT, CLINICIAN DECIDE TO WITHDRAW THE CATHETER. ADDITIONAL INFORMATION RECEIVED ON APRIL 20, 2023: "IN THE ANALYSIS OF THE REPORT, THE POWER PICC CATHETER 7274118 WAS NOT FOUND TO BE BROKEN AND IT IS POSSIBLE THAT THERE WAS AN ANATOMICAL CONDITION OF THE PATIENT THAT CONTRIBUTED TO THE LEAKAGE OF FLUID FROM THE INSERTION SITE." | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01665 | 2023-05-02 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | IT WAS REPORTED THAT THE STYLET WAS REMOVED DURING THE LAST STEP OF THE CATHETER PLACEMENT PROCEDURE. DISCRETE YELLOW DOTS WERE FOUND WITH THE STYLET INSIDE THE POWERPICC CATHETER (3174118). | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF MATERIAL ON THE STYLET WAS CONFIRMED; HOWEVER, THE EXACT ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE STYLET. IT HAD BEEN FULLY REMOVED FROM THE CATHETER, WHICH WAS NOT RETURNED FOR EVALUATION. BIOLOGICAL RESIDUES WERE IN SEVERAL REGIONS ALONG THE LENGTH OF THE STYLET. SEVERAL BENDS WERE OBSERVED IN THE STYLET. MATERIAL APPEARED TO BE FLAKING OFF OF THE WIRE. MICROSCOPIC INSPECTION OF THE SAMPLE CONFIRMED THE PRESENCE OF WHITE MATERIAL ON THE STYLET. THE MATERIAL APPEARED BUNCHED AND PEELING. THE MATERIAL APPEARED TO PARTIALLY COAT THE STYLET WIRE. IT APPEARED THAT THE HYDROPHILIC COATING APPLIED TO THE STYLET WIRE DURING PRODUCT MANUFACTURE WAS BUNCHING AND PEELING. IT COULD NOT BE DETERMINED WHAT CAUSED THE DELAMINATION OF THE COATING. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE EXPOSURE TO INCOMPATIBLE CHEMICALS AND UNIDENTIFIED ENVIRONMENTAL FACTORS AFFECTING THE PROPERTIES OF THE COATING. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01672 | 2023-05-02 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000375 |  | IT WAS REPORTED BY THE CUSTOMER, PICC LEAKING, WAS NOTICED ON (B)(6) 2023 BUT POSSIBLY HAD SIGNS BEFORE AS PATIENT SAID THERE WAS BLOOD RUNNING DOWN HIS ARM. A CHEST XRAY WAS DONE. PICC FRACTURE AT 3.5CM, CLEAR FRACTURE ALMOST SEVERED. PICC WAS PLACED ON (B)(6) 2023 WITH GOOD PLACEMENT AND VESSEL FILL CAPACITY 4% IN THIS VESSEL, PLACED AT 1CM EXIT AND 44CM TRIMMED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING PICC WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR POWERPICC SOLO. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED BETWEEN THE 3 CM AND 4 CM DEPTH MARKERS. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED.FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR. 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER. OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING). AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01670 | 2023-05-02 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/ SHERLOCK 3CG (TPS) STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, WHEN FLUSHING THE CATHETER IT WAS LEAKING, IT IS A HOLE IN IT. PICC HAVE BEEN USED WITH DOXRUBICIN. NOW THEY FLUSHED IT AND IT WAS LEAKING CLOSE TO THE HUB. THEY PULLED IT OUT. THE PATIENT HAVE WORKED HARD WITH HIS ARM. OCCURRED DURING USE. ADDITIONAL INFORMATION 04/18/2023: THE PATIENT DID NOT GET ANY HARM BUT NEEDED A NEW PICC SO IT WAS ONE DISPLACEMENT MORE. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01671 | 2023-05-02 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7274118 |  | MINOR FEMALE PATIENT, HOSPITALIZED IN THE BURN ROOM IN THE COMPANY OF HER MOTHER, IS A CARRIER OF A 4-FR BILUMEN CATHETER INSERTED ON (B)(6) 2023 IN THE LEFT ARM, WITH THE LAST HEALING DATE OF (B)(6) 2023. OBSERVE WITH FULL DRESSING, INSERTION POINT WITH SLIGHT HUMIDITY, DEVICE 3 CENTIMETERS FROM ZERO LEVEL, IT IS DECIDED TO CURE, WITH STERILE TECHNIQUE, AT THE TIME OF FINISHING THE PROCEDURE AND PERFORMING IRRIGATION THROUGH A WHITE LUMEN, A LARGE AMOUNT OF FLUID IS OBSERVED FROM THE UNION OF THE BODY OF THE CATHETER (BIFURCATION) FOR WHICH IT IS DECIDED TO WITHDRAW. __ ADDITIONAL INFORMATION RECEIVED ON 26.APR.2023: TREATMENT WAS CONTINUED, PATIENT REQUIRED MORE PUNCTURES TO ACHIEVE VASCULAR ACCESS. THERE IS NO PATIENT HARM FROM THE INCIDENT, NO NEED FOR MEDICAL INTERVENTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED AND WAS DETERMINED TO BE CAUSED BY KINKING OF THE CATHETER. ONE 4 FR D/L POWERPICC SV CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUES THROUGHOUT THE RETURNED DEVICE, AND THE CATHETER TERMINATED AT THE 20 CM DEPTH MARKER. A MICROSCOPIC OBSERVATION REVEALED A PARTIALLY CIRCUMFERENTIAL SPLIT JUST DISTAL OF THE MOLDED JOINT. THE SPLIT HAD A GRANULAR FRACTURE SURFACE, A SUBSTANTIAL KINK IMPRESSION, AND ROUNDED EDGES FROM MATERIAL FATIGUE. BASED ON THE OBSERVATIONS OF THE RETURNED POWERPICC SV CATHETER, THE COMPLAINT OF A LEAKING CATHETER WAS CONFIRMED AND IS LIKELY DUE TO EXCESSIVE KINKING OF THE CATHETER. THIS CAN OCCUR IF THE SECUREMENT OF THE CATHETER CAUSES A SHARP KINK, AND OVER A PERIOD OF TIME THE CATHETER MAY FAIL OPPOSITE OF THE KINK IMPRESSION. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01669 | 2023-05-02 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY W/ SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT N/S COULD NOT BE PUSHED IN THE CATHETER INITIALLY. WHEN THE CLINICIAN TRIED TO ADD MORE FORCE, THE SOFT FLANGE AND GREY TAPER CONNECTOR ARE DISCONNECTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GROSHONG CONNECTOR WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR SINGLE LUMEN GROSHONG CATHETER. THE CATHETER WAS RECEIVED CUT INTO TWO SEGMENTS. THE PROXIMAL SEGMENT WAS INSERTED THROUGH THE DISTAL CONNECTOR SEGMENT AND OVERLAID THE METAL CANNULA OF THE PROXIMAL SEGMENT. THE TWO SEGMENTS WERE NOT CONNECTED. THE CATHETER WAS BUNCHED ON THE CANNULA. GREY PORTION OF THE PROXIMAL CONNECTOR SEGMENT WAS COMPLETELY DETACHED FROM THE SUTURE WING AND EXTENSION TUBING. THE SUTURE WING REGION EXHIBITED DEFORMATION. MICROSCOPIC INSPECTION OF THE SAMPLE REVEALED THE CONNECTOR COMPONENTS TO BE WELL-FORMED. MECHANICAL DAMAGE WAS OBSERVED ON THE LOCKING ARMS OF THE PROXIMAL CONNECTOR. TACTILE INSPECTION OF THE SAMPLE REVEALED TENSILE WEAKNESS THROUGHOUT THE EXTENSION TUBE. ASSEMBLY OF THE TWO CONNECTOR SEGMENTS WAS SUCCESSFUL AND UNREMARKABLE. THE RESULTANT ASSEMBLY WAS PATENT TO INFUSION AND ASPIRATION WITH NO OBSERVED LEAKS. THE SUTURE WING DEFORMATION AND TENSILE WEAKNESS WERE CONSISTENT WITH SEPARATION OF THE GREY PORTION OF THE CONNECTOR DUE TO PULLING STRESS. IT APPEARED THAT PULLING OCCURRED AS PART OF MANIPULATION OF THE CONNECTOR DURING ATTEMPTED ASSEMBLY. THE BUNCHING OF THE CATHETER SUGGESTED THAT ADVANCING THE CATHETER TOO FAR ONTO THE CANNULA MAY HAVE CONTRIBUTED TO DIFFICULTY ASSEMBLING THE CONNECTOR. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01622 | 2023-05-02 | 2023-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | 0668945 |  | IT WAS REPORTED THAT UPON UNPACKING OF THE SELDINGER, A BREACH WAS FOUND WITH THE INTRODUCER SHEATH. THIS MADE IT IMPOSSIBLE FOR THE INTRODUCER SHEATH TO PASS THROUGH THE GUIDE WIRE. A NEW SELDINGER KIT WAS THEREFORE USED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A SPLIT MICROINTRODUCER SHEATH WAS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE 4.5 FR X 5 CM MICROEZ MICROINTRODUCER WAS RETURNED FOR EVALUATION. THE SHAFT WAS OBSERVED TO BE CURVED AND THE DILATOR CONTAINED A SHARP BEND. BLOOD RESIDUE AND EVIDENCE OF USE WAS NOTED. ONE PHOTO SAMPLE OF THE DISTAL END OF A MICROINTRODUCER WAS ALSO PROVIDED. THE SHEATH APPEARED TO CONTAIN A LONGITUDINAL SPLIT EXTENDING FROM THE DISTAL END. BUNCHING AT THE TIP ALSO APPEARED TO BE PRESENT AND CORRESPONDED WITH THE RETURNED PHYSICAL SAMPLE. MICROSCOPIC INSPECTION OF THE DISTAL END OF THE SHEATH REVEALED A LONGITUDINAL SPLIT. ADDITIONAL DEFORMATION AND INWARD BUNCHING WAS NOTED AT THE TIP. EVIDENCE OF A SHEATH TIP TAPER WAS OBSERVED. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE DAMAGED SHEATH, ADVANCEMENT AGAINST RESISTANCE, AND INADEQUATE SKIN NICK. A REVIEW OF THE MANUFACTURING RECORDS DID NOT REVEAL EVIDENCE TO SUGGEST A MANUFACTURING RELATED ROOT CAUSE. SINCE BUNCHING AND SPLITTING OF THE SHEATH WAS NOTED ON THE RETURNED SAMPLE, THE COMPLAINT IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01664 | 2023-05-02 | 2023-01-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM CURVED DIALYSIS CATHETER (13F) (15CM) (TL) | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5615150 |  | PER MEDWATCH: GUIDEWIRE BECAME STUCK IN NEEDLE DURING SHORT TERM TRIAL CATHETER INSERTION. ADDITIONAL INFORMATION RECEIVED 02/28/2023: GUIDEWIRE WAS USED DURING AN INSERTION OF A TRIAL CATHETER FOR A PATIENT, IT BECAME STUCK IN NEEDLE DURING CATHETER INSERTION. NO PATIENT HARM REPORTED. (B)(6) 2023 ONE OF THE COILS OF THE COILED WIRE OF THE GUIDEWIRE WAS FOUND TO BE UNRAVELED NEAR THE DISTAL END OF THE GUIDEWIRE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DIFFICULTY INSERTING A GUIDEWIRE IS CONFIRMED. ONE 0.035 IN. J-TIP GUIDEWIRE IN A PLASTIC HOOP AND ONE 18 G SECURELOC INTRODUCER NEEDLE WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLE. WHILE THE GUIDEWIRE WAS RETURNED IN THE INTRODUCER NEEDLE, IT WAS NOT FOUND TO BE STUCK. THE CORE WIRE OF THE GUIDEWIRE WAS FOUND TO NOT BE BROKEN DURING TACTILE EVALUATION. A MICROSCOPIC OBSERVATION REVEALED THE BEVEL OF THE INTRODUCER NEEDLE WAS DAMAGED. A RELATIVELY LARGE AMOUNT OF DRIED BLOOD RESIDUE WAS OBSERVED ON THE RETURNED GUIDEWIRE, ESPECIALLY ON THE SECTION OF GUIDEWIRE THAT WAS WITHIN THE INTRODUCER NEEDLE. THIS BIOLOGICAL MATERIAL MAY HAVE CONTRIBUTED TO THE ALLEGED EVENT OF THE GUIDEWIRE BECOMING STUCK WITHIN THE INTRODUCER NEEDLE. ONE OF THE COILS OF THE COILED WIRE OF THE GUIDEWIRE WAS FOUND TO BE UNRAVELED NEAR THE DISTAL END OF THE GUIDEWIRE; HOWEVER, NO OTHER DAMAGE WAS OBSERVED ON THE RETURNED GUIDEWIRE. NORMAL MOVEMENT OF THE GUIDEWIRE IS AWAY FROM THE SHARPENED BEVEL AND WILL NOT DAMAGE THE WIRE. IF THE GUIDEWIRE DIRECTION IS REVERSED, THE GUIDEWIRE IS THEN PULLED AGAINST THE SHARPENED EDGE OF THE NEEDLE BEVEL AND CAN CAUSE SHEARING DAMAGE OF THE WIRE AND/OR CAUSE THE WIRE TO BECOME STUCK WITHIN THE NEEDLE. THE NEEDLE DEFORMATION WAS CONSISTENT WITH GUIDEWIRE RETRACTION AGAINST THE NEEDLE BEVEL. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01662 | 2023-05-02 | 2023-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SL PICC BASIC TRAY (W/ MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405J |  | IT WAS REPORTED THAT PRIOR TO PUNCTURING, THE PHYSICIAN CHECKED THE GUIDEWIRE WAS ABLE TO INSERT THOUGHT THE INTRODUCER NEEDLE USING THE SOFT PART OF THE GUIDEWIRE, HOWEVER AFTER PUNCTURING THE VESSEL WITH THE INTRODUCER NEEDLE, ONLY THE SOFT PART OF THE GUIDEWIRE WAS ALLEGEDLY INSERTED THROUGH THE INTRODUCER NEEDLE BUT UNABLE TO BE ADVANCED MORE. THERE WAS NO REPORTED PATIENT INJURY. (B)(6)2023 A MICROSCOPIC OBSERVATION REVEALED THE COILS OF THE RETURNED GUIDEWIRE WERE DAMAGED AND SLIGHTLY MISALIGNED AT ITS DISTAL END. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: SAMPLE ANALYSIS, PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), COMPLAINT AND LOT HISTORY REVIEW, APPLICABLE MANUFACTURE RECORDS, AND APPLICABLE FMEA DOCUMENTS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A GUIDEWIRE UNABLE TO PASS THROUGH AN INTRODUCER NEEDLE IS INCONCLUSIVE DUE TO UNKNOWN CLINICAL CONDITIONS. ONE 21 G SAFECAN INTRODUCER NEEDLE AND ONE 0.018 IN. GUIDEWIRE IN A PLASTIC HOOP WERE RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED USE RESIDUE ON THE RETURNED SAMPLES. KINKS WERE OBSERVED IN THE GUIDEWIRE NEAR ITS DISTAL END. NO BREAK IN THE CORE WIRE OF THE GUIDEWIRE WAS FOUND DURING TACTILE EVALUATION. AN ATTEMPT WAS MADE TO PASS THE GUIDEWIRE THROUGH THE RETURNED INTRODUCER NEEDLE AND THE GUIDEWIRE WAS FOUND TO PASS COMPLETELY THROUGH THE NEEDLE WITH EASE. SOME SLIGHT RESISTANCE WAS OBSERVED AS THE DAMAGED PORTION OF THE GUIDEWIRE ENTERED THE NEEDLE CANNULA. A MICROSCOPIC OBSERVATION REVEALED THE COILS OF THE RETURNED GUIDEWIRE WERE DAMAGED AND SLIGHTLY MISALIGNED AT ITS DISTAL END. NO OCCLUSION WAS OBSERVED WITHIN THE HUB OF THE RETURNED INTRODUCER NEEDLE; HOWEVER, THE INNER DIAMETER OF THE JUNCTION OF THE NEEDLE CANNULA AND HUB WAS OBSERVED TO BE SOMEWHAT ROUGH AND UNEVEN. WHILE THE RETURNED GUIDEWIRE WAS FOUND TO BE ABLE TO PASS THROUGH THE RETURNED INTRODUCER WITH EASE, UNKNOWN CLINICAL CONDITIONS SUCH AS INSERTION TECHNIQUE AND/OR ANGLE, PATIENT ANATOMY, AND POSITION OF INTRODUCER NEEDLE COULD HAVE AFFECTED THE PASSAGE OF THE GUIDEWIRE THROUGH THE INTRODUCER NEEDLE. CONSEQUENTLY THIS COMPLAINT IS INCONCLUSIVE AT THIS TIME. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01659 | 2023-05-02 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO 5F SINGLE-LUMEN CATHETER BASIC TRAY W/70CM NITINOL GW | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 6195355 |  | IT WAS REPORTED, "PICC LINE BENT AT THE VALVE, WITH RISK OF CRACKING." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01657 | 2023-05-02 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | 4FR DL POWERMIDLINE FULL KIT | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P4254108 |  | IT WAS REPORTED THAT THE MIDLINE WAS PLACED WITH NO INCIDENT, MIDLINE PROVIDED GOOD BLOOD RETURN WHEN ASSESSED BUT THEN WHEN FLUSHED BEGAN TO SPRAY SALINE AROUND THE SUTURE WINGS WHERE THEY MET THE CATHETER. MIDLINE WAS REMOVED AND REPLACED. ADDITIONAL INFORMATION RECEIVED MAY 1, 2023: PATIENT ADMITTED TO THE ICU DUE TO RESPIRATORY FAILURE, ON MULTIPLE DRIPS, CONNECTED TO A VENTILATOR, VERY DIFFICULT VENOUS ACCESS. EVENT DID NOT INVOLVE AN URGENT/LIFE THREATENING MEDICAL SITUATION. EVENT DID NOT INTERRUPT TREATMENT; CLINICIAN WAS ABLE TO FIX THE PROBLEM IMMEDIATELY WITHOUT INTERRUPTING PATIENT CARE OR PATIENT TREATMENT. PLANNED MEDICATIONS TO BE INFUSED, TYPICAL IV FLUIDS FOR HYDRATION, IV ANTIBIOTICS, AND IV SEDATIVES. THERE WAS NO PATIENT HARM, INJURY, OR NEGATIVE OUTCOME AS THE PROBLEM WAS CORRECTED IMMEDIATELY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01658 | 2023-05-02 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED, "WHEN THEY SHOULD USE THE CATHETER ON A CT SCAN IT WAS LEAKING. PATIENT HAVE THE CATHETER FOR CHEMO. NOW THEY SHOULD USE IT FOR CT SCAN AND WHEN FLUSHING IT WAS LEAKING. WHEN THEY PULLED IT OUT THEY SAW 2 HOLES IN IT. PATIENT NEEDED A NEW PICC. NO IMPACT ON THE OPERATOR, THE NURSE USED PROTECTION, IT WAS SALINE THAT WAS RUNNING." IT WAS STATED THE DEVICE IS CHEMO CONTAMINATED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR POWERPICC SOLO. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND TWO SPLITS WERE OBSERVED BETWEEN THE 5 CM AND 6 CM DEPTH MARKERS AND BETWEEN THE 6 CM AND 7 CM DEPTH MARKERS. BOTH CATHETER SPLITS CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: ¿ DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED ¿ FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR ¿ 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER ¿ OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING) AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01656 | 2023-05-02 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 6F RADIOLOGY BASIC TRAY (DUAL-LUMEN) (135 CM GUIDEWIRE) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3276135 |  | IT WAS REPORTED HAIR FOUND IN THE PACKAGING OF THE ITEM. IT WAS DISCOVERED FOUND BEFORE USE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A HAIR IN THE PACKAGING WAS CONFIRMED AND THE CAUSE WAS MANUFACTURING-RELATED. THE PRODUCT RETURNED FOR EVALUATION WAS 6FR D/L RAD POWERPICC CATHETER TRAY. THE TRAY WAS RECEIVED SEALED. A FIBER WAS OBSERVED CAUGHT WITHIN THE FLANGE SEAL OF THE TRAY. MICROSCOPIC INSPECTION OF THE SAMPLE CONFIRMED THE PRESENCE OF A FIBER CAUGHT WITHIN THE FLANGE SEAL. THE FIBER APPEARED CONSISTENT WITH A HUMAN HAIR. THE PACKAGING EXHIBITED A HAIR CAUGHT WITHIN THE SEAL, INDICATING THAT THE HAIR BECAME DEPOSITED ON THE TRAY PRIOR TO SEALING DURING KIT ASSEMBLY. PHOTOGRAPHS HAVE BEEN FORWARDED TO THE MANUFACTURING SITE FOR FURTHER EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01655 | 2023-05-02 | 2022-12-28 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT, "THE PICC BROKE AND WENT INSIDE THE BODY DURING FLUSHING, IT WAS PULLED WITH THE HELP OF INTERVENTION RADIOLOGIST. NO EXPOSURE TO BLOOD OR BODILY FLUIDS. TREATMENT WAS NOT ALTERED BECAUSE OF THE EVENT. PATIENT WAS DISCHARGED." IT WAS STATED THE INCIDENT INVOLVED INCORRECT TEST RESULTS. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01653 | 2023-05-02 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC CATHETER W/ SHERLOCK 3CG TPS BASIC TRAY (4F) (SINGLE-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174118 |  | IT WAS REPORTED BY THE CUSTOMER THAT "POWERPICC KINKED WITHIN 24 HOURS OF INSERTION. PICC HAD TO BE REMOVED AND REPLACED. ADDITIONAL INFORMATION PROVIDED 4/19/2023: DATE OF INSERTION: (B)(6) 2023 INTO VESSEL SELECTION: RIGHT UPPER BASILIC, DATE OF REMOVAL: (B)(6) 2023, CATHETER TRIM LENGTH: 41, CATHETER EXTERNAL (EXIT SITE) LENGTH: 1.5CM. NO POWER INJECTION SCAN WAS PERFORMED. PATIENT NEEDED TO REPLACE THE LINE IN THE OPPOSITE ARM, WHICH RESULTED IN EXTRA UNNECESSARY POKES AND PAIN, ADDITIONAL COSTS TO THE PATIENT, AND ANXIETY THAT ISSUE WILL REPEAT WITH THE NEW DEVICE. ADDITIONAL INFORMATION RECEIVED ON 5/23/2023: CATHETER PLACED RIGHT UPPER BASILIC VEIN. CATHETER TRIMMED TO 41CM WITH 1.5 CM EXTERNAL TO THE PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A KINK IN THE CATHETER WAS INCONCLUSIVE AS THE REPORTED KINK COULD NOT BE CONCLUSIVELY SEEN IN THE RETURNED IMAGES. TWO CONED DOWN RADIOGRAPHIC IMAGES OF THE RIGHT HUMERUS WERE RETURNED FOR EVALUATION. A PICC, CONSISTENT WITH THE IMPLICATED 4FR S/L POWERPICC CATHETER, WAS SEEN IN THE IMAGES. THE PICC CONTAINED GENTLE CURVES AS IT ENTERED THE MUSCLE AND VEINS. NO DEFINITIVE KINKS OR COILS WERE SEEN IN THE CATHETER SHAFT. IT IS POSSIBLE THAT THE CATHETER WAS KINKED, BUT IT COULD NOT BE INDEPENDENTLY CONFIRMED FROM THE PROVIDED IMAGES. POSSIBLE CONTRIBUTING FACTORS FOR A KINK IN THE CATHETER COULD INCLUDE CATHETER PLACEMENT, PATIENT ANATOMY, OR CATHETER SECUREMENT. THE IFU STATES, ¿AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT.¿ AND ¿AVOID SHARP OR ACUTE ANGLES DURING IMPLANTATION WHICH COULD COMPROMISE THE PATENCY OF THE CATHETER LUMEN.¿ A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE REPORT OF A KINKED CATHETER HAS BEEN DOCUMENTED AND WILL BE MONITORED AS PART OF COMPLAINT TRENDING. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01623 | 2023-05-02 | 2023-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWER-TRIALYSIS SHORT-TERM STRAIGHT DIALYSIS CATHETER (13 F) (20CM) (TL) | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5603200 |  | IT WAS REPORTED BY THE CUSTOMER " THE CLINICIAN PLACED THE LINE, CATHETER WENT OVER THE WIRE FINE BUT WHEN SHE ATTEMPTED TO REMOVE THE WIRE IT STUCK INSIDE THE DISTAL PORT. THEY TRIED MULTIPLE TIMES TO REMOVE THE WIRE AND ENDED UP HAVING TO PULL THE LINE AND WIRE AS 1 UNIT FOR FEAR OF SHEARING THE WIRE.". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01619 | 2023-05-02 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | CUSTOMER REPORTED THE FOLLOWING: RN WAS ATTEMPTING TO FLUSH PICC LINE WHEN SPRAYED WITH SALINE. RUBBER STOPPER AT THE END OF THE PICC WHICH HOLDS THE GUIDE WIRE WAS LEAKING AND WAS SPRAYING SALINE ONTO THE NURSE WHEN ATTEMPTING TO FLUSH. RISK OF BLOOD EXPOSURE IF PORT ISN'T SEALED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING T-LOCK ASSEMBLY WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE SAMPLE RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03058 | 2023-05-02 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER WAS DIFFICULT TO GET IN AND OUT DURING THE TEST. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT WAS UNCONFIRMED A LABELING REVIEW WAS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03055 | 2023-05-02 | 2023-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | TEMP SENSING FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119316M | 119316M |  | IT WAS REPORTED THAT CUSTOMER FOUND FOLEY CATHETER WAS OUT OF PATIENT, LYING IN THE BED, THE BALLOON OF THE FOLEY CATHETER WAS NOT INFLATED. THE CUSTOMER INFLATED THE BALLOON WITH AIR AND AFTER 30 MINS NO AIR REMAINED IN THE BALLOON. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "IMPERFECTION IN BALLOON DUE TO PROCESS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DIRECTIONS FOR USE: 1. WASH HANDS AND DON CLEAN GLOVES 2. EXPLAIN PROCEDURE TO PATIENT AND OPEN PERI-CARE KIT 3. USE THE PROVIDED PACKET OF WIPES TO CLEANSE PATIENT¿S PERIURETHRAL AREA 4. REMOVE GLOVES AND PERFORM HAND HYGIENE WITH PROVIDED ALCOHOL HAND SANITIZER GEL 5. USING PROPER ASEPTIC TECHNIQUE OPEN CSR WRAP 6. DON STERILE GLOVES 7. PLACE UNDERPAD BENEATH PATIENT, PLASTIC/¿SHINY¿ SIDE DOWN NOTE: USE CAUTION TO MAINTAIN ASEPTIC TECHNIQUE 8. POSITION FENESTRATED DRAPE ON PATIENT 9. SATURATE 3 FOAM SWAB STICKS IN POVIDONE IODINE 10.ATTACH THE WATER FILLED SYRINGE TO THE INFLATION PORT NOTE: IT IS NOT NECESSARY TO PRE-TEST THE FOLEY CATHETER BALLOON 11. REMOVE FOLEY CATHETER FROM WRAP AND 12. PREPARE PATIENT WITH 3 FOAM SWAB STICKS SATURATED IN POVIDONE IODINE. USE THE NONDOMINANT HAND FOR THE GENITALIA AND THE DOMINANT HAND FOR THE SWABS NOTE: USE EACH SWAB STICK FOR ONE SWIPE ONLY FEMALE PATIENT: WITH A DOWNWARD STROKE CLEANSE THE RIGHT LABIA MINORA AND DISCARD THE SWAB. DO THE SAME FOR THE LEFT LABIA MINORA. WITH THE LAST SWABSTICK CLEANSE THE MIDDLE AREA BETWEEN THE LABIA MINORA MALE PATIENT: CLEANSE THE PENIS IN A CIRCULAR MOTION STARTING AT THE URETHRAL MEATUS AND WORKING OUTWARD LUBRICATE CATHETER 13. PROCEED WITH CATHETERIZATION IN USUAL MANNER USING THE DOMINANT HAND A. WHEN CATHETER TIP HAS ENTERED BLADDER, URINE WILL BE VISIBLE IN THE DRAINAGE TUBE B. INSERT CATHETER TWO MORE INCHES AND INFLATE 14. INFLATE CATHETER BALLOON USING ENTIRE 10CC OF STERILE WATER PROVIDED IN THE PREFILLED SYRINGE NOTE: USE OF LESS THAN 10CC CAN RESULT IN ASYMMETRICALLY INFLATED BALLOON 15. ONCE INFLATED, GENTLY PULL CATHETER UNTIL THE INFLATED BALLOON IS SNUG AGAINST THE BLADDER NECK" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-02922 | 2023-05-02 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 1758SI14 | 1758SI14 |  | IT WAS REPORTED THAT THE BALLOONS WERE NOT INFLATING EVENLY SO THAT THE FOLEY CATHETER WAS LOPSIDED (LOT# NGGT3581) AND WAS PRESSING INTO THE BLADDER WALL THUS NOT DRAINING PROPERLY. PATIENT HAD THEM INFLATE IT FIRST TO TEST THE BALLOONS BEFORE USAGE AND NO PATIENT HARM. PATIENT HAD KEPT 6-8 OCCURRENCES (LOT# NGCZ0518) (LOT# NGCW0564) IN THE LAST YEAR WITH THE PACKAGES SO THERE MAY BE MULTIPLE LOTS BUT AS PATIENT WAS IN A WHEELCHAIR, THEY DID NOT WANT TO OPEN THEM UNTIL THEY HAD TO SHIP THEM. AS PER INFORMATION RECEIVED VIA EMAIL ON (B)(6) 2023, IT WAS STATED THAT 3 FOLEY CATHETER WAS RECEIVED UNDER LOT# NGGT3581, 2 FOLEY CATHETER RECEIVED UNDER THE LOT#NGCZ0518, ONE FOLEY CATHETER RECEIVED UNDER THE LOT#NGCW0564. SMASTAN(B)(6) 2023. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03052 | 2023-05-02 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53816G | 53816G |  | IT WAS REPORTED THAT THE SURE SLEEVE DID NOT WORK OR SLIDE CORRECTLY ON 6 CATHETERS. NO ADDITIONAL INFORMATION PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "NOT FOLLOWING DOCUMENTED PROCEDURE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS NONE KNOWN WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS URINARY TRACT INFECTION BLEEDING FROM THE URETHRA IRRITATION OF THE URETHRA REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8.WASH HANDS. " THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03054 | 2023-05-02 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WHEN PRETESTED, STERILE WATER WAS INJECTED AND INFLATED, BUT THEN THE WATER TRIED TO DRAIN, BUT THE WATER DID NOT RETURN AND THE BALLOON REMAINED INFLATED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) WITH THE RETURNED SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 2 MINUTES WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03053 | 2023-05-02 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS WITHDRAWN SPONTANEOUSLY. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03050 | 2023-05-02 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT ON THE DAY OF USE, THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03051 | 2023-05-02 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 53616G |  | IT WAS REPORTED THAT MAGIC 3 MALE INTERMITTENT CATHETERS WERE NOT LUBED ENOUGH TO USE. THE PATIENT ASKED IF THEY COULD USE THEIR OWN LUBE AS WELL. REPRESENTATIVE ADVISED YES, THEY WERE ABLE TO. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON'T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01615 | 2023-05-02 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN POWER PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNKNOWN POWER PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PUMPS ALLEGEDLY STARTED RINGING OCCLUDED IN THE MIDDLE OF THE TREATMENT. IT WAS FURTHER REPORTED THAT BLOOD RETURN AND FLUSH DID NOT WORKED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03038 | 2023-05-02 | 2023-04-11 | I, n, j, u, r, y | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI14 |  | IT WAS REPORTED THAT THE PATIENT'S CATHETER WAS FOUND TO BE BROKEN INTO TWO. UPON INSPECTION, ONE PART OF THE CATHETER WAS INSIDE OF THE PATIENT¿S PENIS WITH AT LEAST 2CM IN LENGTH. THE OTHER HALF OF THE CATHETER WAS ON THE FLOOR. THE DOCTOR WAS INFORMED AND REMOVED THE PART OF THE CATHETER WITHIN THE PATIENT BY DRAINING THE CATHETER BALLOON WITH A SYRINGE. THE PATIENTS WERE GIVEN A BLADDER SCAN, ASSESSED FOR TRAUMA TO THE PENIS, AND MONITORED TO CHECK THEY COULD PASS URINE WITHOUT PROBLEM. NO HARM WAS REPORTED TO THE PATIENT. THE BROKEN CATHETER HAS BEEN RETAINED FOR FURTHER INVESTIGATION. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE DESIGN". THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DIRECTIONS FOR USE: 1. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. 2. USING PROPER ASEPTIC METHODS, REMOVE CATHETER FROM PACKAGE. 3. PREPARE PATIENT PER HOSPITAL/NURSING RECOMMENDED PROCEDURE. 4. PROCEED WITH CATHETERIZATION USING STANDARD TECHNIQUES. 5. INFLATE THE 10ML BALLOON WITH A MAXIMUM OF 10ML STERILE WATER BY USING A WATER-FILLED LUER-TIP SYRINGE. DO NOT USE A NEEDLE TIP SYRINGE TO INFLATE BALLOON. 6. CONNECT CATHETER TO COLLECTION CONTAINER. 7. TO DEFLATE THE 10ML BALLOON, GENTLY REINSERT THE LUER TIP SYRINGE INTO THE VALVE AND ASPIRATE. CAUTION: TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE IN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF PERMITTED BY HOSPITAL PROTOCOL, THE VALVE ARM MAY BE SEVERED. IF THIS FAILS, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. VALVE TYPE: USE LUER SLIP TIP SYRINGE. DO NOT USE NEEDLE. RECOMMENDED INFLATION CAPACITIES. 5ML BALLOON: USE 5ML STERILE WATER. 10ML BALLOON: USE 10ML STERILE WATER. 30ML BALLOON: USE 35ML STERILE WATER. DO NOT EXCEED RECOMMENDED CAPACITIES." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03039 | 2023-05-02 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT 21 DAYS AFTER PLACEMENT, THE CATHETER WAS UNABLE TO REMOVE. EVEN AFTER INJECTING THE FIXATION WATER, THE CATHETER WAS UNABLE TO BE REMOVED. EVENTUALLY, THE PATIENT WAS TRANSPORTED TO A HOSPITAL WITH A UROLOGIST ((B)(6)), WHERE THE CATHETER WAS REMOVED. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01610 | 2023-05-02 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT PROCEDURE, THE PUMPS ALLEGEDLY STARTED RINGING OCCLUDED IN THE MIDDLE OF THE TREATMENT. IT WAS FURTHER REPORTED THAT THE DEVICE ALLEGEDLY HAD A SUCTION AND FLUSHING ISSUE. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. D4 (EXPIRY DATE: 03/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01611 | 2023-05-02 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1808060 | 1808060 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PUMPS ALLEGEDLY STARTED RINGING OCCLUDED IN THE MIDDLE OF THE TREATMENT. IT WAS FURTHER REPORTED THAT BLOOD RETURN AND FLUSH DID NOT WORKED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. (EXPIRATION DATE: 11/2023). THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01613 | 2023-05-02 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PUMPS ALLEGEDLY STARTED RINGING OCCLUDED IN THE MIDDLE OF THE TREATMENT. IT WAS FURTHER REPORTED THAT BLOOD RETURN AND FLUSH DID NOT WORKED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 03/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01614 | 2023-05-02 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PUMPS ALLEGEDLY STARTED RINGING OCCLUDED IN THE MIDDLE OF THE TREATMENT. IT WAS FURTHER REPORTED THAT BLOOD RETURN AND FLUSH DID NOT WORKED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 01/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01617 | 2023-05-02 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE PUMPS ALLEGEDLY STARTED RINGING OCCLUDED IN THE MIDDLE OF THE TREATMENT. IT WAS FURTHER REPORTED THAT BLOOD RETURN AND FLUSH DID NOT WORKED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 03/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01618 | 2023-05-02 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6F CHRONW/OS | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806061 |  | IT WAS REPORTED THAT APPROXIMATELY ONE YEAR AND ONE MONTH POST PORT PLACEMENT, AMOUNT OF BLOOD RETURNED WAS ALLEGEDLY SMALL. IT WAS FURTHER REPORTED THAT THE LIQUID DRIP INJECTION WAS NOT SMOOTH. FURTHERMORE, THE CATHETER WAS ALLEGEDLY FOUND TO BE BROKEN. REPORTEDLY, THE PORT WAS REMOVED AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRATION DATE: 04/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03089 | 2023-05-03 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | TEMPERATURE-SENSING FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT A SYRINGE WAS CONNECTED TO THE INFLATION VALVE AND STERILE WATER WAS INJECTED, BUT STERILE WATER LEAKED FROM THE CONNECTION AREA. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. IT IS UNKNOWN IF THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED THAT THE BALLOON WAS PARTIALLY INFLATED UPON RECEIVAL OF SAMPLE. DEFLATED BALLOON PASSIVELY USING RETURNED SYRINGE AND THE CATHETER RESTED WITH NO CUFFING OR LEAKS NOTED. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED AROUND THE INFLATION VALVE OR THE CATHETER ITSELF. THE BALLOON PASSIVELY DEFLATED WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE IP7603444, REVISION 0, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03091 | 2023-05-03 | 2023-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE URINE LEAKED FROM NEAR CATHETER SHAFT. | THE REPORTED EVENT WAS CONFIRMED MANUFACTURING RELATED. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT CAUSED THE REPORTED FAILURE. IT IS UNKNOWN IF THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED A HOLE MEASURING 0.1685" FOUND ON THE SHAFT OF THE CATHETER AND LOCATED 0.4915" FROM THE BIFURCATION. THIS IS OUT OF SPECIFICATION PER INSPECTION PROCEDURE (B)(4), REVISION 0, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER BARD® SILVER LUBRI-SIL® FOLEY TRAY CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK FOLEY. SOME CATHETER TYPES OF THE DEVICE MAY HAVE A TEMPERATURE SENSOR FOR MEASURING PATIENT¿S CORE BODY TEMPERATURE AND THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. THE SURFACES OF THE CATHETER ARE COATED WITH A MINUTE AMOUNT OF METALLIC SILVER AND FURTHER COATED WITH POLYURETHANE, HAVING ANTIPROLIFERATIVE EFFECTS ON MICROORGANISMS ON THE CATHETER. <MATERIAL> BALLOON CATHETER: SILICONE; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD®* SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 22 EVERY 2FR 1. BALLOON CATHETER FOLEY CATHETER TEMPERATURE-SENSING CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE COMBINES A DISPOSABLE CATHETER THAT IS DESIGNED TO BE PLACED IN THE BLADDER FOR THE PURPOSE OF URINARY DRAINAGE AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE PACKAGED COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE CATHETER SHAFT WITH THE LUBRICANT JELLY. (3) CAREFULLY INSERT THE CATHETER INTO THE URETHRAL MEATUS. AFTER THE BALLOON ADVANCED IN THE BLADDER, ATTACH THE NEEDLELESS SYRINGE, AND GENTLY INFUSE THE SPECIFIED VOLUME OF STERILE WATER TO INFLATE THE BALLOON. (4) PULL THE CATHETER SLIGHTLY TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK. (5) TO DEFLATE AND REMOVE THE BALLOON, ATTACH A NEEDLELESS SYRINGE TO LET STERILE WATER IN THE BALLOON COME OUT SPONTANEOUSLY THROUGH BALLOON DEFLATION WITHOUT ASPIRATION. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER SLOWLY WHILE CONFIRMING THAT NO ABNORMAL RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 (9) WHEN CATHETER WITH TEMPERATURE-SENSING IS USED, CONNECT LEAD WIRE TO MONITOR WITH RELAY CABLE. (10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 (11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 (12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.]. (14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. (15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. (16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. (17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. (18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE WAS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (26) WHEN USING STATLOCK FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE CATHETER CONNECTOR CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.¿ CORRECTIONS: D, F, H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03092 | 2023-05-03 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123612A |  | IT WAS REPORTED THAT THE FOLEY CATHETER 12F WAS UNABLE TO DRAIN URINE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03095 | 2023-05-03 | 2023-04-12 | I, n, j, u, r, y | Manufacturer report | BARDEX® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 165814 | 165814 |  | IT WAS REPORTED THAT A PHYSICIAN STATED IN AN ONLINE SURVEY THAT THE PATIENT EXPERIENCED URINARY TRACT INFECTION PRIOR TO THE DEVICE PLACEMENT AND URINARY TRACT INFECTION COMPLICATIONS WERE DIRECTLY ATTRIBUTABLE TO THE DEVICE WHICH OCCURRED 7 DAYS AFTER THE DATE OF INSERTION AND THE INFECTION WAS NOT SYMPTOMATIC IN RELATION TO BARDIA ALL-SILICONE MALE LENGTH CATHETERS, PRE-CONNECTED TO A 350ML CENTER ENTRY URINE METER WITH 2500ML DRAINAGE BAG, FRENCH SIZE 14. MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE IS AVAILABLE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: "STERILE UNLESS PACKAGE IS OPENED OR DAMAGED, EXCEPT FOR ANY INDIVIDUALLY PACKAGED COMPONENTS WITHIN THE TRAY WHICH ARE NOT LABELED AS STERILE. THESE COMPONENTS ARE NOT TERMINALLY STERILIZED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE. PROPER TECHNIQUES FOR URINARY CATHETER INSERTION PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION. INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT. USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE. DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD. PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE: SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED. MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS. MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING. KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES. EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT. ROUTINE HYGIENE (E.G., CLEANSING OF THE METAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE. LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01680 | 2023-05-03 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TPS STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 2194108 |  | IT WAS REPORTED "HOLE IN CATHETER. THE CATHETER WAS USED FOR A WHILE BUT WHEN THEY FLUSHED IT BEFORE GIVEN CHEMO, THE CATHETER WAS LEAKING. THEN THEY PULLED IT OUT AND SAW A HOLE. NO INJURY TO THE PATIENT.". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03098 | 2023-05-03 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 129614 |  | IT WAS REPORTED THAT WATER LEAK OCCURRED FROM THE FOLEY CATHETER BALLOON. AT FIRST IT WAS INFLATED, BUT AFTER ABOUT 30 MINUTES, THE WATER LEAKED OUT. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS EVALUATED AND OBSERVED THERE WAS A TEAR AT THE BOTTOM PART OF SAC COLLAR THAT ALLOWED WATER TO LEAK OUT. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO BUBBLE DURING DIPPING PROCESS/UNDER SIZE AND LOW SAC WEIGHT-TOO SHORT PICK-UP. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03099 | 2023-05-03 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT LEAKAGE OF STERILE WATER WAS OBSERVED DURING PRE-TEST. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. ONE SAMPLE EXHIBITED THE REPORTED FAILURE. THE DEVICE HAD NOT MET SPECIFICATIONS. THE REPORTED FAILURE IS CONSIDERED OUT OF SPECIFICATION AS THE REPORTED FAILURE WAS REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT CAUSED THE REPORTED FAILURE. ALTHOUGH A SPECIFIC CAUSE CANNOT BE DETERMINED, BASED ON THE RISK DOCUMENT A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE IMPERFECTION IN BALLOON DUE TO PROCESS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION NOTED A HOLE ON THE BALLOON MEASURING 0.031" WITH THE PIN GAUGE. THIS IS OUT OF SPECIFICATION WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01683 | 2023-05-03 | 2023-04-12 | I, n, j, u, r, y | Manufacturer report | BARD POWERGLIDE PRO | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT: "REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE MIDLINE INSERTION SITE. THE EVENT INVOLVED AN 18G POWERGLIDE PRO RT INSERTED IN THE RIGHT UPPER ARM." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03101 | 2023-05-03 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED THAT WHEN THE PATIENT WENT TO POP THE WATER SACHET FOR THIS LOT OF INTERMITTENT CATHETER, IT WAS FOUND THAT THE SACHETS WERE FILLED WITH AIR AND THERE WAS NO LIQUID. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿FOR UROLOGICAL USE ONLY. INTENDED FOR USE BY PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICES ARE PASSED TO THE URINARY BLADDER VIA THE URETHRA. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. PRIOR TO OPENING THE SEALED CATHETER POUCH, APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. WET THE CATHETER BY HOLDING THE PACKAGE WITH THE PRINTED SIDE UP AND TIP THE PACKAGE END-TO-END THREE TO SIX TIMES TO WET THE CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. PEEL OPEN THE PACK AT THE FUNNEL/HANDLE END JUST ENOUGH TO EXPOSE 5 CM (2¿) OF THE CATHETER, OR FOR PRODUCTS WITH AN INSERTION SLEEVE, EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. USING THE CATHETER WITH INSERTION SLEEVE (GRIPPER): HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. WITHOUT INSERTION SLEEVE (GRIPPER): HOLD THE HANDLE/FUNNEL END AND REMOVE THE CATHETER FROM THE PACKAGING. GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA AND ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03104 | 2023-05-03 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT AFTER CATHETER INSERTION IN THE EMERGENCY ROOM, UPON ADMISSION TO THE EICU, URINE OUTFLOW WAS OBSERVED, BUT THE CATHETER WAS EASILY REMOVED. IT WAS MENTIONED THAT THERE WAS NO FIXED WATER AND THE USER SUSPECTED A PINHOLE ON THE BALLOON. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) FIG. 6 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.7 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG.8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03088 | 2023-05-03 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L12 |  | IT WAS REPORTED THAT THE SYRINGE DID NOT CONTAIN 10ML OF STERILE WATER. THE SYRINGE INCLUDED IN THE PACKAGE CONTAINED ONLY 9ML OF STERILE WATER AND THE EVENT OCCURRED BEFORE USE TO THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01686 | 2023-05-03 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH, | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED VIA MEDWATCH THAT "DURING PICC PLACEMENT RUBBER-LIKE STOPPER DISLODGED FROM EXTENSION". NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01685 | 2023-05-03 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED BY THE CUSTOMER THAT THERE WERE PROTRUSIONS ON THE SURFACE OF THE INTRODUCER SHEATH, MAKING THE DEVICE UNUSABLE. THE INTRODUCER SHEATH WAS REPLACED WITH A NEW ONE TO CONDUCT RE-PUNCTURE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED MICROINTRODUCER SHEATH IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE PHOTO SAMPLE OF A 4.5 FR MICROEZ MICROINTRODUCER WAS PROVIDED FOR EVALUATION. THE IMAGE SHOWED THE DEVICE BEING HELD OUTSIDE OF PATIENT USE. THE DILATOR IS PRESENT WITHIN THE SHEATH AND HAS NOT BEEN PEELED. DAMAGE AT THE DISTAL END OF THE SHEATH WAS NOTED. THE CHARACTERISTICS OF THE COMPONENT CANNOT BE CLEARLY INSPECTED FROM THE IMAGE. BASED ON THE INFORMATION PROVIDED, POSSIBLE CONTRIBUTING FACTORS INCLUDE ADVANCEMENT AGAINST RESISTANCE AND INADEQUATE SKIN INCISION. A REVIEW OF THE MANUFACTURING RECORDS DID NOT REVEAL EVIDENCE TO SUGGEST A MANUFACTURING RELATED ROOT CAUSE. SINCE DEFORMATION AT THE SHEATH TIP WAS NOTED IN THE IMAGE, THE COMPLAINT IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01689 | 2023-05-03 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PEDIATRIC PATIENT UNDERWENT THE PICC PLACEMENT FROM THE RIGHT UPPER LIMB UNDER ULTRASOUND ON MARCH 16. UPON UNPACKING OF THE PICC CATHETER, FOREIGN MATTERS (HAIR) WERE FOUND INSIDE THE PACKAGE. THE OPERATION WAS SUSPENDED AND A NEW PRODUCT WAS UNPACKED TO COMPLETE THE CATHETER PLACEMENT. NO OTHER INFORMATION WAS PROVIDED. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A HAIR INSIDE THE PACKAGING OF A GROSHONG CATHETER KIT IS INCONCLUSIVE BECAUSE THE KIT WAS RETURNED OPENED AND A HAIR COULD NOT BE FOUND. ONE 4 FR S/L GROSHONG CATHETER KIT WAS RETURNED OPENED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED THE KIT HAD BEEN RETURNED OPENED WITH NO USE RESIDUES WITHIN THE PACKAGING. RETURNED INSIDE THE KIT WAS THE GROSHONG CATHETER, THE STYLET, A NEEDLELESS INJECTION CAP, A GUIDEWIRE, A NEEDLE, AND A GROSHONG CONNECTOR. THE PACKAGING WAS RETURNED OPENED, AND A HAIR INSIDE THE PACKAGING COULD NOT BE IDENTIFIED. THE PACKAGING WAS ALSO OBSERVED TO HAVE KINKS IN THE PLASTIC, POSSIBLY FROM TRANSPORTATION. THE COMPLAINT OF CONTAMINATION CAUSED BY A HAIR INSIDE THE PACKAGING OF A GROSHONG CATHETER KIT IS INCONCLUSIVE BECAUSE THE SAMPLE WAS RETURNED OPENED AND A HAIR WAS NOT IDENTIFIED. ATTEMPTS TO MITIGATE CONTAMINATION ARE IN PLACE AND WILL CONTINUE TO BE MONITORED AT THE MANUFACTURING FACILITIES AT BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03100 | 2023-05-03 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE WATER PACKET ONLY HAD AIR IN IT AND THE PATIENT WAS UNABLE TO USE THE MAGIC 3 INTERMITTENT COUDE CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03087 | 2023-05-03 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER WAS NOT LUBRICATED ENOUGH, AND PATIENT HAD TROUBLE GETTING THE CATHETER TO PASS THROUGH THE PROSTATE WITHOUT BENDING. PER FOLLOW UP VIA PHONE ON (B)(6)2023, IT WAS STATED THAT THEIR SUPPLIER WAS LIBERATOR AND SINCE THEY SWITCHED FROM THE LIQUID CATHETERS TO THE GEL ALTERNATIVES, CUSTOMER HAD SEEN IMPROVEMENTS DURING USAGE. THEIR UROLOGIST RECOMMENDED THE CURE ULTRA CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR, VISCOMETER FAILURE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03068 | 2023-05-03 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST IN THE OPERATION ROOM, STERILIZED DISTILLED WATER COULD BE INJECTED, BUT THE WATER COULD NOT BE DRAINED. WHEN APPLYING PRESSURE TO THE BALLOON AREA WITH A FINGER, THE STERILE WATER RETURNED, BUT THE USER DID NOT USE IT BECAUSE THE USER WAS AFRAID TO USE THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03071 | 2023-05-03 | 2023-04-12 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PUREWICK URINE COLLECTION SYSTEM WAS NOT SUCTIONING PROPERLY. THE PATIENT NEED TO DO TROUBLESHOOT BECAUSE THEY COULD NOT RETURN USED SYSTEM. THE PATIENT HAD BEEN USING THIS PRODUCT FOR LESS THAN 90 DAYS. AS PER FOLLOW-UP VIA PHONE ON 16DEC2022, IT WAS REPORTED THAT TROUBLESHOOTING WAS COMPLETED WITH LMS BUT CUSTOMER WAS HAVING ISSUES WITH PLACEMENT. AS PER ADDITIONAL INFORMATION ON 28JAN2023, PATIENT SAID THAT THE MACHINE WOULD NOT SUCTION PATIENTS SMALL URINE STREAM. AS PER ADDITIONAL INFORMATION RECEIVED VIA FORM WITH SAMPLE RETURNED ON 12APR2023, IT WAS REPORTED THAT THE DESPITE FOLLOWING THE INSTRUCTIONS, THE UNIT DID NOT WORK AS DESCRIBED. THE URINE WAS SUCTIONED AWAY ONLY WHEN THE URINE STREAM WAS STRONG BUT FOR A WEAK URINE STREAM IT BACKED UP BACK TO THE WICK WHICH CAUSING BURNING AND PAIN DURING ITS USE . ALSO PATIENT EXPERIENCED URINARY TRACT INFECTION AND PRESCRIBED ANTIBIOTICS. | THE REPORTED EVENT IS INCONCLUSIVE, DUE TO PATIENT ANATOMY UNKNOWN. A POTENTIAL ROOT CAUSES INCLUDE INADEQUATE MATERIAL SELECTION/MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE/MATERIAL SURFACE IS ROUGH, ABRASIVE OR UNCOMFORTABLE. THE DHR REVIEW DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: DO NOT USE THE PUREWICKTM FEMALE EXTERNAL CATHETER WITH BEDPAN OR ANY MATERIAL THAT DOES NOT ALLOW FOR SUFFICIENT AIRFLOW. TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03086 | 2023-05-03 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER WAS NOT LUBRICATED ENOUGH AND THEY HAD TROUBLE GETTING THE CATHETER TO PASS THROUGH THE PROSTATE WITHOUT BENDING . | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01678 | 2023-05-03 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT TWENTY DAYS POST DIALYSIS CATHETER PLACEMENT, THE CATHETER GOT EXPELLED OUT WITH CUFFS APPEARED PAPERY WITH NO FIBROSIS. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY GOT SLIPPED ITSELF AND SPONTANEOUSLY CAME OUT FROM PATIENT'S SKIN. REPORTEDLY, CATHETER WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE 23CM GLIDEPATH D/L CATHETER WAS RETURNED FOR EVALUATION. GROSS VISUAL, MICROSCOPIC VISUAL, TACTILE EVALUATION AND FUNCTIONAL TESTING WERE PERFORMED ON THE RETURNED DEVICE. UNDER MICROSCOPIC OBSERVATION, WHAT APPEARED TO BE FIBER DISTURBANCE WAS NOTED THROUGHOUT THE MATTED TISSUE CUFF AND THE TISSUE CUFF DID NOT MOVE FREELY THROUGHOUT THE CATHETER. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, ONE ELECTRONIC PHOTO WAS PROVIDED FOR REVIEW. THE PHOTO SHOWS ONE GLIDEPATH DIALYSIS CATHETER IMPLANTED IN A PATIENT BODY AND THE TISSUE CUFF WAS NOTED TO BE MOVED FROM THE CATHETER INSERTION SITE. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED CATHETER EXPULSION AND LOOSENING OF IMPLANT ISSUES AS THE CUFF WAS NOTED TO BE MOVED OUT FROM THE INSERTION SITE. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 05/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03064 | 2023-05-03 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE CUSTOMER HAD ORDERED 3-4 BOXES OF PUREWICK FEMALE EXTERNAL CATHETERS. THE NEW PUREWICK CATHETERS DID NOT HAVE BECTON DICKINSON (BD) LABELLED ON THE PACKAGE, WHEREAS THE FIRST TWO PACKAGES THEY ORDERED DID HAVE THE LABEL, AND THE PRODUCT FELT THICKER. THE CUSTOMER ALSO STATED THAT THE CATHETERS APPEARED TO HAVE LESS SUCTION. PER FOLLOW UP VIA PHONE ON 01-MAY-2023, A FAMILY MEMBER REPORTED THE FEMALE EXTERNAL CATHETERS WERE FOR HER MOTHER, BUT SHE HAD PASSED AWAY. NO ADDITIONAL INFORMATION WAS PROVIDED. PER CLINICAL FOLLOW UP VIA PHONE ON (B)(6) 2023 THE PATIENT¿S DAUGHTER WAS UNABLE TO PROVIDE ANY ADDITIONAL INFORMATION REGARDING THE REPORTED ISSUES. SHE REPORTED THE 82-YEAR-OLD FEMALE PATIENT PASSED AWAY ON (B)(6) 2023 DUE TO PNEUMONIA. SHE ADDITIONALLY STATED THE PATIENT¿S DEATH WAS UNRELATED TO THE PUREWICK DEVICE. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS THE EVENT WAS NOT REPORTABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03065 | 2023-05-03 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT CUSTOMER STATED THAT THE BURST PACKET IN 30 WAS DAMAGED WHEN RECEIVED AND THE CUSTOMER DISCARDED THE ENTIRE BOX OF CATHETERS AND THE CATHETER PACKET WAS DRY. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT PROCESS PARAMETERS". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. BASED ON PATIENT CODE 2645 THE PRODUCT WAS NOT USED ON THE PATIENT. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELLING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00307 | 2023-05-03 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ULTRAVERSE 035 | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | U357566 | U357556 |  | IT WAS REPORTED THAT DURING ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED AND CAME OFF THE CATHETER. IT WAS FURTHER REPORTED THAT THE DETACHED BALLOON WAS SNARED OUT OF THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY. | THE INITIAL MDR WAS INADVERTENTLY SUBMITTED WITH A G3 DATE OF 04/06/2023. THE CORRECT G3 DATE IS 04/05/2023. MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ULTRAVERSE 035 PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, A COMPLETE CIRCUMFERENTIAL RUPTURE WAS NOTED TO THE BALLOON, AND THE DETACHED PIECE WAS NOT RETURNED LEAVING THE INNER GUIDEWIRE LUMEN EXPOSED. NO ANOMALIES WERE NOTED TO THE Y-BODY/STRAIN RELIEF. FUNCTIONAL EVALUATION WAS NOT PERFORMED DUE TO THE CONDITION OF THE RETURNED DEVICE. ONE PHOTO WAS PROVIDED AND REVIEWED. THE PHOTO SHOWS THE ULTRAVERSE 035 WITHIN ITS INNER PACKAGING. THE LABEL CORRESPONDS WITH THE INFORMATION AVAILABLE IN TRACKWISE. COMPLETE CIRCUMFERENTIAL RUPTURE OF THE BALLOON IS SEEN. NO OTHER ANOMALIES ARE NOTED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED DETACHMENT AND COMPLETE CIRCUMFERENTIAL BALLOON RUPTURE AS THE DEVICE WAS RETURNED WITH A PORTION OF THE BALLOON COMPLETELY DETACHED. A DEFINITIVE ROOT CAUSE FOR THE REPORTED COMPLETE CIRCUMFERENTIAL BALLOON RUPTURE AND DETACHMENT COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. B5, D4 (EXPIRATION DATE: 07/2025). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03069 | 2023-05-03 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER WAS DIFFICULT TO GET IN AND OUT DURING THE TEST. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE FOREIGN MATTER DURING PROCESSES (FLAP OR FLASH DURING PUNCHING). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS ·PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01690 | 2023-05-03 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000375 |  | IT WAS REPORTED BY THE CUSTOMER, "ON (B)(6) 2023 CHEST XRAY DONE AND REPORTED AS: 'TIP APPEARS TO HAVE ADVANCED AND LIES LOW IN THE SVC/RIGHT ATRIUM AT THE LEVEL OF T 8/9. PERSISTENT PICC WAS PLACED ON (B)(6) 2023 WITH GOOD PLACEMENT AND VESSEL FILL CAPACITY 4% IN THIS VESSEL, PLACED AT 1CM EXIT AND 44CM TRIMMED. ON (B)(6) 2023 PICC WAS DOCUMENTED TO BE AT 1CM.". | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03066 | 2023-05-03 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50618 | 50618 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER DID NOT COME WITH ENOUGH LUBE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03150 | 2023-05-04 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0167L26 | 0167L26 |  | IT WAS REPORTED, THAT THE PHYSICIAN INSERTED 3-WAY FOLEY CATHETER AND UNABLE TO INFLATE BALLOON OR IRRIGATE CATHETER. NEW 3-WAY CATHETER OBTAINED AND INSERTED WITHOUT ISSUE. NO HARM TO PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03147 | 2023-05-04 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT WHITE DEPRESSED DAMAGES WITH THE BALLOON PART FOR A TOTAL OF 5350 FOLEY CATHETERS FROM TWO BATCHES (BATCH# MYGN5842 AND BATCH# MYGR5328) WERE FOUND DURING USE AND INSPECTION. THE QUALITY DEPARTMENT OF COMPANY HAD DETERMINED THAT THERE WERE RISKS OF BALLOON BURSTING AND THAT THESE PRODUCTS COULD NOT BE USED IN THE PRODUCTION. THE ISSUE WAS DETECTED DURING REUSE OF A SINGLE USE MEDICAL DEVICE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, STATED THAT CUSTOMER WAS A DISTRIBUTION CENTER. STATED THAT 5 PHYSICAL SAMPLES OF EACH BATCH WOULD BE PROVIDED. THE USER PURCHASED 5350 UNITS AND INSPECTED SOME OF THEM HAD THE ISSUE, THUS THE CUSTOMER WAS WORRIED ABOUT THE QUALITY OF THE WHOLE BATCHES AND REQUEST TO CHANGE AND RETURN 5350 UNITS WITH ANOTHER BATCH. PER SAMPLE RECEIVED BY THE BD REPRESENTATIVE ON (B)(6) 2023, STATED THAT THE MANUFACTURING LOT# OF SAMPLES RECEIVED FOR THE BATCH MYGN5842 WAS 2HJ031 AND FOR THE BATCH MYGR5328 WERE 2JD072, 2HH041, 2HU135 AND 2HL031 (BATCH# MYGN5842 AND SERIAL# 2HJ031, BATCH# MYGR5328 AND SERIAL# 2JD072,BATCH# MYGR5328 AND SERIAL# 2HH041, BATCH# MYGR5328 AND SERIAL# 2HU135 AND BATCH# MYGR5328 AND SERIAL# 2HL031). PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, STATED THAT THE CUSTOMER WAS AN OLD CUSTOMER, THEY HAVE USED BD FOLEY FOR ALMOST 18 YEARS SINCE 2006. THIS WAS THE FIRST INCIDENT EXPERIENCED BY THEM. THE MANUFACTURING LOT# OF SAMPLES FOR THE BATCH MYGN5842 WAS 2FD047 AND FOR THE BATCH MYGR5328 WERE 2HU136, 2FR083, 2JC013 AND 2HU134 (BATCH# MYGN5842 AND SERIAL# (B)(4) BATCH# MYGR5328 AND SERIAL# (B)(4) BATCH# MYGR5328 AND SERIAL# 2FR083, BATCH# MYGR5328 AND SERIAL# (B)(4) AND BATCH# MYGR5328 AND SERIAL# (B)(4) STATED THAT THIS CUSTOMER PER CHASE FOLEY FOR REPROCESSING. WHEN THEY RECEIVED FOLEY, THEY WOULD JUST CHECK PRODUCT MODEL, SIZE, QUANTITY AND EXPIRATION DATE. WHILE DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. THERE WAS NO SAMPLING INSPECTION PERFORMED, THEY TESTED THE BALLOON OF ALL CATHETERS. STATED THAT DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL FROM IBC ON (B)(6) 2023, STATED THAT THE SAMPLES REPORTED BY CUSTOMER WERE THOSE PICKED UP BY CUSTOMER FOR FUNCTION TESTING DURING THEIR REPRODUCTION PROCESS AND THESE PHYSICAL SAMPLES COULD NOT BE RETURNED FOR INVESTIGATION, BUT THE PHOTOS ATTACHED TO WERE BELONGING TO THEM. WHILE FOR THE SAMPLES RECEIVED, THESE WERE REPRESENTATIVE SAMPLES FROM THE SAME REPORTED BATCHES RETURNED FOR INVESTIGATION, THEY WERE NOT THE REAL AFFECTED SAMPLES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03148 | 2023-05-04 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED, THAT WHITE DEPRESSED DAMAGES WITH THE BALLOON PART FOR A TOTAL OF (B)(4) FOLEY CATHETERS FROM TWO BATCHES (BATCH#: MYGN5842 AND BATCH#: MYGR5328) WERE FOUND DURING USE AND INSPECTION. THE QUALITY DEPARTMENT OF COMPANY HAD DETERMINED, THAT THERE WERE RISKS OF BALLOON BURSTING AND THAT THESE PRODUCTS COULD NOT BE USED IN THE PRODUCTION. THE ISSUE WAS DETECTED DURING REUSE OF A SINGLE USE MEDICAL DEVICE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 13 APR 2023. STATED ,THAT CUSTOMER WAS A DISTRIBUTION CENTER. STATED, THAT 5 PHYSICAL SAMPLES OF EACH BATCH WOULD BE PROVIDED. THE USER PURCHASED (B)(4) UNITS AND INSPECTED SOME OF THEM HAD THE ISSUE, THUS THE CUSTOMER WAS WORRIED ABOUT THE QUALITY OF THE WHOLE BATCHES AND REQUEST TO CHANGE AND RETURN (B)(4) UNITS WITH ANOTHER BATCH. PER SAMPLE RECEIVED BY THE BD REPRESENTATIVE ON 14 APR 2023. STATED, THAT THE MANUFACTURING LOT# OF SAMPLES RECEIVED FOR THE BATCH#: MYGN5842 WAS (B)(4), AND FOR THE BATCH#: MYGR5328 WERE (B)(4), (BATCH#: MYGN5842 AND SERIAL#: (B)(4), BATCH#: MYGR5328 AND SERIAL#: (B)(4), BATCH#: MYGR5328 AND SERIAL#: B)(4), BATCH#: MYGR5328 AND SERIAL#: (B)(4), AND BATCH#: MYGR5328 AND SERIAL#: (B)(4)). PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 15 APR 2023. STATED, THAT THE CUSTOMER WAS AN OLD CUSTOMER. THEY HAVE USED BD FOLEY FOR ALMOST 18 YEARS SINCE 2006. THIS WAS THE FIRST INCIDENT EXPERIENCED BY THEM. THE MANUFACTURING LOT# OF SAMPLES FOR THE BATCH#: MYGN5842 WAS(B)(4). AND FOR THE BATCH#: MYGR5328 WERE (B)(4), (BATCH#: MYGN5842 AND SERIAL#: (B)(4), BATCH#: MYGR5328 AND SERIAL#: (B)(4), BATCH#: MYGR5328 AND SERIAL#: (B)(4), BATCH#: MYGR5328 AND SERIAL#: (B)(4), AND BATCH#: MYGR5328 AND SERIAL#: (B)(4)). STATED, THAT THIS CUSTOMER PER CHASE FOLEY FOR REPROCESSING. WHEN THEY RECEIVED FOLEY, THEY WOULD JUST CHECK PRODUCT MODEL, SIZE, QUANTITY AND EXPIRATION DATE. WHILE DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. THERE WAS NO SAMPLING INSPECTION PERFORMED, THEY TESTED THE BALLOON OF ALL CATHETERS. STATED, THAT DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL FROM IBC ON 21 APR 2023. STATED ,THAT THE, SAMPLES REPORTED BY CUSTOMER WERE THOSE PICKED UP BY CUSTOMER FOR FUNCTION TESTING. DURING THEIR REPRODUCTION PROCESS AND THESE PHYSICAL SAMPLES COULD NOT BE RETURNED FOR INVESTIGATION. BUT THE PHOTOS ATTACHED TO WERE BELONGING TO THEM. WHILE FOR THE SAMPLES RECEIVED. THESE WERE REPRESENTATIVE SAMPLES FROM THE SAME REPORTED BATCHES RETURNED FOR INVESTIGATION. THEY WERE NOT THE REAL AFFECTED SAMPLES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03146 | 2023-05-04 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT WATER FROM THE FOLEY CATHETER BALLOON WAS DIFFICULT TO REMOVE DURING THE PRETEST. WHEN THE ACTUAL PRODUCT WAS RECEIVED, THE SYRINGE WAS CONNECTED, AND THE WATER WAS COMPLETELY DRAINED. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE PRODUCT WAS NOT USED FOR PATIENT TREATMENT OR DIAGNOSIS. THE PRODUCT HAD NOT CAUSED THE REPORTED FAILURE. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL: VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING AND A STRAIGHT TIPPED SYRINGE. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) USING THE RETURNED SYRINGE AND THE CATHETER RESTED WITH NO LEAKS NOTED. THE BALLOON PASSIVELY DEFLATED IN UNDER 2 MINUTES WITH NO CUFFING OR LEAKS NOTED. THIS IS WITHIN SPECIFICATION PER INSPECTION PROCEDURE (B)(4), REVISION 0, WHICH STATES, CATHETER LEAKS, VALVE LEAKS, BALLOON PERFORATION, LUMEN TO LUMEN, PREMATURELY DEFLATED AND DISLODGED BALLOON ARE NOT ALLOWED, AND MUST INFLATE AND DEFLATE WITH THE USE OF A SYRINGE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT IS UNCONFIRMED NEITHER A RISK REVIEW NOR LABELING REVIEW IS REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03139 | 2023-05-04 | 2023-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE PATIENT'S TEMPERATURE CAME OUT IN THE LOW 34C EVEN THOUGH THE RECTOMETER WAS IN THE 36C, AND CHANGING THE CODE DID NOT IMPROVE THE PATIENT'S TEMPERATURE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ASSEMBLY TECHNIQUE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01696 | 2023-05-04 | 2023-03-13 | I, n, j, u, r, y | Manufacturer report | POWERPICC SOLO 4F SL 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000937A |  | IT WAS REPORTED BY THE CUSTOMER, "PICC LINE DAMAGE. FATHER WENT TO GIVE NIGHT INFUSION OF ANTIBIOTICS AND WITH SALINE FLUSH 'IT SQUIRTED EVERYWHERE'. THERE IS A VISIBLE HOLE IN THE CATHETER. PICC WAS HEPARIN LOCKED AT THE TIME OF THE EVENT. THERE WERE NO PATIENT SYMPTOMS RELATED TO THE FLUID LEAK. THE CATHETER WAS REMOVED AND A NEW PICC WAS REQUESTED BUT THE PATIENT HAS NO VESSELS LEFT FOR PLACEMENT. TEAM DECIDE TO STOP ERTAPENUM EARLY AND MONITOR FOR ISSUES. CATHETER HAD BEEN SECURED WITH STATLOCK.' | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A HOLE IN THE CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR POWER PICC CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH CATHETER DAMAGE CAUSED BY CONTACT WITH A SHARP EDGED INSTRUMENT SUCH AS A SCALPEL. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A BREAK WAS OBSERVED JUST DISTAL OF THE MOLDED SUTURE JOINT. MICROSCOPIC EXAMINATION OF THE CATHETER BREAK CONFIRMED IT WAS TYPICAL OF CONTACT WITH A SHARP BLADED INSTRUMENT, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: THE FRACTURE SURFACE WAS REFLECTIVE IN NATURE AND CONTAINED A STRIATED PATTERN. THIS CAN OCCUR DUE TO PATTERN TRANSFER OF THE SHARPENED EDGE TYPICALLY FOUND ON A SCALPEL-TYPE INSTRUMENT. SHARPLY FORMED FRACTURE EDGES PLANAR SHAPE TO THE FRACTURE SURFACE AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H.11. | Y | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01699 | 2023-05-04 | 2023-04-11 | I, n, j, u, r, y | Manufacturer report | BARD SL PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED DURING A CLINICAL ASSESSMENT "REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. SINGLE LUMEN UPPER ARM." NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION AND SPECIFIC PATIENT DETAILS WERE NOT DISCLOSED AS PART OF THE SURVEY. AS SUCH, ADDITIONAL INFORMATION AND THE SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03151 | 2023-05-04 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THERE WAS A CUFF ROLL FOUND TO THE FOLEY CATHETER. AFTER THE CATHETER PLACEMENT IN THE OPERATING ROOM, ATTEMPTING TO REMOVE THE CATHETER IN ICU, THE USER FELT RESISTANCE DURING PULLING IT. CHECKING THE CATHETER, A CUFF ROLL WAS FOUND TO IT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03149 | 2023-05-04 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT WHITE DEPRESSED DAMAGES WITH THE BALLOON PART FOR A TOTAL OF 5350 FOLEY CATHETERS FROM TWO BATCHES (BATCH# MYGN5842 AND BATCH# MYGR5328) WERE FOUND DURING USE AND INSPECTION. THE QUALITY DEPARTMENT OF COMPANY HAD DETERMINED THAT THERE WERE RISKS OF BALLOON BURSTING AND THAT THESE PRODUCTS COULD NOT BE USED IN THE PRODUCTION. THE ISSUE WAS DETECTED DURING REUSE OF A SINGLE USE MEDICAL DEVICE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 13APR2023, STATED THAT CUSTOMER WAS A DISTRIBUTION CENTER. STATED THAT 5 PHYSICAL SAMPLES OF EACH BATCH WOULD BE PROVIDED. THE USER PURCHASED 5350 UNITS AND INSPECTED SOME OF THEM HAD THE ISSUE, THUS THE CUSTOMER WAS WORRIED ABOUT THE QUALITY OF THE WHOLE BATCHES AND REQUEST TO CHANGE AND RETURN 5350 UNITS WITH ANOTHER BATCH. PER SAMPLE RECEIVED BY THE BD REPRESENTATIVE ON (B)(6) 2023, STATED THAT THE MANUFACTURING LOT# OF SAMPLES RECEIVED FOR THE BATCH MYGN5842 WAS 2HJ031 AND FOR THE BATCH MYGR5328 WERE 2JD072, 2HH041, 2HU135 AND 2HL031 (BATCH# MYGN5842 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6),BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6) AND BATCH# MYGR5328 AND SERIAL# (B)(6). PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 15APR2023, STATED THAT THE CUSTOMER WAS AN OLD CUSTOMER, THEY HAVE USED BD FOLEY FOR ALMOST 18 YEARS SINCE 2006. THIS WAS THE FIRST INCIDENT EXPERIENCED BY THEM. THE MANUFACTURING LOT# OF SAMPLES FOR THE BATCH MYGN5842 WAS 2FD047 AND FOR THE BATCH MYGR5328 WERE 2HU136, 2FR083, 2JC013 AND 2HU134 (BATCH# MYGN5842 AND SERIAL# 2FD047, BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6) AND BATCH# MYGR5328 AND SERIAL# (B)(6). STATED THAT THIS CUSTOMER PER CHASE FOLEY FOR REPROCESSING. WHEN THEY RECEIVED FOLEY, THEY WOULD JUST CHECK PRODUCT MODEL, SIZE, QUANTITY AND EXPIRATION DATE. WHILE DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. THERE WAS NO SAMPLING INSPECTION PERFORMED, THEY TESTED THE BALLOON OF ALL CATHETERS. STATED THAT DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL FROM IBC ON 21APR2023, STATED THAT THE SAMPLES REPORTED BY CUSTOMER WERE THOSE PICKED UP BY CUSTOMER FOR FUNCTION TESTING DURING THEIR REPRODUCTION PROCESS AND THESE PHYSICAL SAMPLES COULD NOT BE RETURNED FOR INVESTIGATION, BUT THE PHOTOS ATTACHED TO WERE BELONGING TO THEM. WHILE FOR THE SAMPLES RECEIVED, THESE WERE REPRESENTATIVE SAMPLES FROM THE SAME REPORTED BATCHES RETURNED FOR INVESTIGATION, THEY WERE NOT THE REAL AFFECTED SAMPLES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03155 | 2023-05-04 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI18 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE INFLATION VALVE WHEN INJECTING WATER WITH A SYRINGE. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AS THE REPORTED EVENT WAS UNCONFIRMED A LABELING REVIEW WAS NOT REQUIRED. THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01698 | 2023-05-04 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD TL PICC | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNK |  | IT WAS REPORTED DURING A CLINICAL ASSESSMENT ""FLUID LEAKING FROM CENTRAL LINE INSERTION SITE. TRIPLE LUMEN POWERPICC IN THE CHEST." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION AND SPECIFIC PATIENT DETAILS WERE NOT DISCLOSED AS PART OF THE SURVEY. AS SUCH, ADDITIONAL INFORMATION AND THE SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03157 | 2023-05-04 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER WAS UNABLE TO REMOVE. WHEN THE ACTUAL SAMPLE WAS RECOVERED, A SYRINGE WAS CONNECTED TO THE CATHETER AND THE WATER HAD BEEN COMPLETELY REMOVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01705 | 2023-05-04 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PROVENA MIDLINE 3F SL | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000879 |  | IT WAS REPORTED BY THE CUSTOMER THAT "I DID A 3 FR AND COULDN'T ADVANCE THE INTRODUCER BECAUSE A LIP FORMED ON THE END AND WOULDN'T PENETRATE THE VEIN." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01706 | 2023-05-04 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F NURSE BASIC TRAY WITH MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3174118 |  | IT WAS REPORTED, THAT THIS PATIENT FROM THE (B)(6) OF THIS HOSPITAL NEEDS TO HAVE THIS POWERPICC CATHETER PLACED. NO ABNORMALITY WAS FOUND WHEN THE OPERATOR TRIMMING THE CATHETER, AND THE CATHETER WAS INSERTED SMOOTHLY. RESISTANCE WAS MET WHEN REMOVING THE GUIDE WIRE. AND IT WAS REMOVED SUCCESSFULLY EVENTUALLY. AFTER REMOVAL, THE STYLET WAS FOUND BENT. THE CATHETER WAS FOUND OCCLUDED ONE HOUR LATER. X-RAY SHOWED THAT THE CATHETER WAS KINKED IN THE BODY. THE CATHETER WAS THEN REMOVED SUBSEQUENTLY. AND THE PATIENT SWITCHED TO A CVC. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT: DEVICE NOT RETURNED FOR EVALUATION. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01707 | 2023-05-04 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1385108D5 |  | IT WAS REPORTED BY THE CUSTOMER THAT "PICC KINKED." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
MW5117271 | 2023-05-04 | 2023-04-28 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BARD POWERPICC PROVENA CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, LONG-TERM GREATER THAN 30 DAYS | BARD ACCESS SYSTEMS, INC. | CK000951 | CK000951 |  | PATIENT ORDERED TO RECEIVE A DOUBLE LUMEN PICC LINE FOR TPN. THE IV TEAM NURSE WAS SETTING UP THE TRAY FOR INSERTION AND PRE-FLUSHING THE CATHETER WAS DONE PRIOR TO STARTING THE PROCEDURE. THE T-PORT THAT WAS ATTACHED TO THE PICC LINE WAS LEAKING FROM THE STOPPER WHEN FLUSHED IN THE LOCATION OF WHERE THE WIRE IS INSERTED. THE PRODUCT WAS NOT USED FOR INSERTION AND WAS DISPOSED OF. A TRIPLE LUMEN PICC WAS INSERTED INSTEAD WITH DIFFERENT TYPE OF T-PORT ATTACHMENT. NO COMPLICATIONS WERE OBSERVED. NO HARM TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
MW5117256 | 2023-05-04 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | POWER GLIDE PRO MIDLINE CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. | 20G | F220108PT |  | MIDLINE CATHETER FROM BARD BEING INSERTED, GUIDEWIRE HAD BECOME DISLODGED. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01702 | 2023-05-04 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERGLIDE PRO MIDLINE CATHETER (18G) (8CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | F118081T |  | IT WAS REPORTED BY THE CUSTOMER" GUIDEWIRE RECOILED WHEN HE WAS CHECKING IT PRIOR TO INSERTION." ADDITIONAL INFORMATION RECEIVED (B)(6) 2023: IT WAS REPORTED BY THE CUSTOMER "WHEN I WAS CHECKING THE GUIDEWIRE BEFORE USING THE MIDLINE CATHETER, I NOTICED THAT I COULD ADVANCE THE SLIDE, BUT THE GUIDEWIRE WOULD NOT ADVANCE. UPON FURTHER EVALUATION, IT LOOKS LIKE THE GUIDEWIRE RECOILED INTO THE DEVICE. I HAD NOT YET POKED THE PATIENT, AS WE WERE TAUGHT TO MAKE SURE THE GUIDEWIRE IS INTACT EVEN BEFORE INSERTION. THE PRODUCT I WAS USING WAS THE MIDLINE CATHETER KIT." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01700 | 2023-05-04 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER W/SHERLOCK 3CG TPS STYLET 4F | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1194108D |  | IT WAS REPORTED BY CUSTOMER THAT PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT WITH A PICC LINE THAT WAS LEAKING. UPON FURTHER INSPECTION, LEAKING WAS NOTED AT THE SEAM OF THE SOLO VALVE. ADDITIONAL INFORMATION RECEIVED APRIL 25, 2023: PICC LEAKING AT THE VALVE. FOUND DURING INFUSION BY HOMECARE NURSE. LINE REPLACED THROUGH EMERGENCY ROOM BY PICC NURSE. | THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03128 | 2023-05-04 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT WATER LEAKED FROM THE FOLEY CATHETER. AFTER CATHETER INSERTION, THE BALLOON WAS INFLATED WITH DISTILLED WATER AND THE SYRINGE WAS DETACHED FROM THE CATHETER. AFTERWARDS, DISTILLED WATER WAS FOUND TO BE LEAKING FROM THE DISTILLED WATER INSERTION SITE. WHEN ATTEMPTING TO CHECK THE CATHETER, THE CATHETER SPONTANEOUSLY FELL OFF AND THE BALLOON WAS FOUND TO HAVE BEEN COMPLETELY DEFLATED. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. EVALUATION FOUND INTERLUMEN LEAKAGE DUE TO TEAR ON RUBBERIZED LAYER. THE REPORTED EVENT IS CONFIRMED AS MANUFACTURING RELATED ISSUE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE(S) THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. MATERIAL BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDOELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 4) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: D H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS EVALUATED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01697 | 2023-05-04 | 2023-04-07 | I, n, j, u, r, y | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D1 |  | IT WAS REPORTED BY THE CUSTOMER THAT, "ON (B)(6) 2023 THE LINE WAS OCCLUDED AND WOULD NOT FLUSH OR DRAW. WHEN PATIENT'S ARM WAS LIFTED, THE LINE WOULD FLUSH AND DRAW. X-RAY SHOWED KINK NEAR INSERTION SITE. LINE RETRACTED 1 CM AND CATHETER DREW AND FLUSHED WITH EASE. ON (B)(6) 2023, PICC (SAME LINE AS BEFORE) OCCLUDED. PICC WAS RETRACTED 1 CM. RED LUMEN WITH BRISK BLOOD RETURN FLUSHES EASILY, PURPLE LUMEN FLUSHES WITH MILD RESISTANCE AND HAD NO BLOOD RETURN. PATIENT ALSO COMPLAINING OF PAIN IN LEFT SHOULDER. AXILLARY ARM CIRCUMFERENCE NOTED TO BE 3CM GREATER THAN ON DATE OF INSERTION. AN ULTRASOUND WAS TAKEN THE SAME DAY AND SHOWED L BRACHIAL VEIN THROMBUS - NO NOTED FLOW AROUND PICC IN LEFT CEPHALIC VEIN BUT PROXIMAL AND DISTAL PORTIONS OF LEFT CEPHALIC VEIN APPEAR PATENT. THE CATHETER WAS REMOVED AND REPLACED. PER CLINICIAN, THE EVENT DID NOT OCCUR DURING AN URGENT OR LIFE-THREATENING MEDICAL SITUATION AND DID NOT CAUSE A CLINICALLY SIGNIFICANT DELAY IN TREATMENT." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED A SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01694 | 2023-05-04 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 1608062 | 1608062 |  | IT WAS REPORTED THAT APPROXIMATELY FIVE MONTHS POST PORT PLACEMENT, THE DEVICE ALLEGEDLY WAS UNABLE TO BE FLUSHED. IT WAS FURTHER REPORTED THAT THE DEVICE ALLEGEDLY HAD A SUCTION PROBLEM. REPORTEDLY A SHORTER NEEDLE WAS USED BUT THE NEEDLE WAS ALLEGEDLY UNABLE TO ACCESS THE PORT AND THE PORT HAD ALLEGEDLY FLIPPED. THERE WAS NO REPORTED PATIENT INJURY. | MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03135 | 2023-05-04 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53612G | 53612G |  | IT WAS REPORTED THAT THE EYELETS ON SEVEN OF THE INTERMITTENT CATHETERS WERE NOT CUT OR PUNCHED PROPERLY. PATIENT DID NOT HAVE THE LOT NUMBER OR ANY SAMPLES TO SEND. PER FOLLOW UP VIA PHONE ON 14APR2023, CUSTOMER STATED THAT THEY WERE STILL HAVING WITH THE CATHETERS. CUSTOMER SAID EVERY OTHER DAY THEY HAD AT LEAST ONE THAT WILL NOT LET URINE FLOW THROUGH. CUSTOMER SAID IT WAS ALWAYS WITH THE FIRST ONE THEY TRY BUT THEY CAN ALWAYS INSERT THE SECOND ONE. CUSTOMER STATED THEY CAN FEEL THE CATHETER WHEN IN INSERTED AND THAT THEY KNOW THEY WERE INSERTING CORRECTLY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "BLOCKED EYE OR INNER DIAMETER OF DRAINAGE EYE SMALL ENOUGH TO ALLOW FOR OCCLUSION DURING USE". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON'T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
2020394-2023-00314 | 2023-05-04 | 2023-04-06 | I, n, j, u, r, y | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80144 | ATG80144 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE PTA BALLOON ALLEGEDLY RUPTURED IN THE CEPHALIC ARCH. IT WAS FURTHER REPORTED THAT THE BALLOON ALLEGEDLY SEPARATED FROM THE TRAILING EDGE OF THE SHAFT. REPORTEDLY, THE BALLOON WAS DIFFICULT TO REMOVE FROM THE INTRODUCER SHEATH AND WAS REMOVED TOGETHER AS ONE UNIT. FURTHER, A DISSECTION OF THE VENOTOMY SITE WAS DONE TO REMOVE THE BALLOON WITH SUBSEQUENT CLOSURE AND PROCEDURE WAS COMPLETED THROUGH NEW VENOTOMY SITE. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE ATLAS GOLD PTA DILATATION CATHETER WAS RECEIVED FOR EVALUATION. DURING VISUAL EVALUATION, BALLOON MATERIAL WAS NOTED ON THE PROXIMAL GLUE WELD, INDICATING THAT THE BALLOON HAD BEEN ATTACHED. THE INNER GUIDEWIRE LUMEN WAS STILL INTACT WITH THE BALLOON THAT WAS PARTIALLY INVERTED AT ITS DISTAL END. NO KINKS WERE NOTED TO THE CATHETER. AN UNKNOWN SHEATH WAS LOADED OVER THE INNER GUIDEWIRE LUMEN WITH GLUE BULLET SEATED INCORRECTLY. NO FUNCTIONAL TESTING WAS PERFORMED DUE TO THE CONDITION OF THE SAMPLE. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED BALLOON RUPTURE AS NO FUNCTIONAL TESTING COULD BE PERFORMED DUE TO THE CONDITION OF THE RETURNED DEVICE. THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED DETACHMENT AS NO COMPLETE DETACHMENT WAS NOTED ON THE DEVICE. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED SHEATH REMOVAL DIFFICULTY AS THE DEVICE WAS RETURNED WITH AN INTRODUCER SHEATH LOADED OVER IT. THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED BREAK AS THE BALLOON WAS SEEN STILL ATTACHED TO THE INNER GUIDEWIRE LUMEN AT ITS DISTAL END. A DEFINITIVE ROOT CAUSE FOR THE ALLEGED BALLOON RUPTURE, SHEATH REMOVAL DIFFICULTY, BALLOON DETACHMENT AND IDENTIFIED BREAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. D4 (EXPIRY DATE: 03/2026). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03117 | 2023-05-04 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ALL SILICONE CATHETER | SILICONE CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT IN AN ONLINE SURVEY A PHYSICIAN STATED THAT, "NO, THEY DO NOT AGREE THAT THE INSTRUCTIONS FOR USE (IFU) FOR THE FOLLOWING BARD/BECTON DICKINSON PRODUCT(S) PROVIDE(S) SUFFICIENT INFORMATION ABOUT THE PRODUCT(S) AND ITS/THEIR MEDICAL APPLICATION(S) BECAUSE 'NOT APPLICABLE OR NEVER USED' IN RELATION TO PRECONNECTED UNCOATED ALL-SILICONE FOLEY CATHETERS". | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INCORRECT/ MISSING TRANSLATION; MISSING INSTRUCTIONS; VENDOR/PRINTER ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03120 | 2023-05-04 | 2022-09-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS BEING USED FOR A CLINICAL PROCEDURE WHEN THE TIP OF THE CATHETER PARTIALLY SNAPPED OFF AND A NEW SET WAS OPENED TO CONTINUE WITH THE PROCEDURE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE, ¿OPERATOR ERROR OR MACHINE MALFUNCTION". THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER FOR THIS DEVICE IS UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. CORRECTION: D,H. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03119 | 2023-05-04 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53812G | 53812G |  | IT WAS REPORTED THAT THE PATIENT'S LAST ORDER WAS DRY AND NOT LUBRICATED WELL IN THE INTERMITTENT CATHETER. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS: THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. INSTRUCTIONS FOR USE: REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. 1. ALWAYS WASH HANDS PRIOR TO USE. 2. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. 3. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 4. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. 5. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). 6. WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL: 7. PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. 8. WASH HANDS. " H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01695 | 2023-05-04 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED VIA MEDWATCH "DURING PICC PLACEMENT RUBBER-LIKE STOPPER DISLODGED FROM EXTENSION." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03126 | 2023-05-04 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 129414M |  | IT WAS REPORTED THAT THE BALLOON WAS NOT INFLATED DURING PATIENT INDWELLING. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE DEVICE DID NOT FAIL TO MEET RELEVANT SPECIFICATIONS. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: BARD® SILVER TSC TRAY SINGLE USE ONLY WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS · PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® SILVER TSC TRAY CONSISTS OF A BALLOON CATHETER WITH TEMPERATURE-SENSING, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS ,GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 14FR, 16FR 1. BALLOON CATHETER WITH TEMPERATURE-SENSING 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY INTENDED USE & EFFECT- EFFICACY: THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND MEASUREMENT OF CORE BODY TEMPERATURE. IT COMBINES A BALLOON CATHETER WITH TEMPERATURE-SENSING DESIGNED TO BE PLACED IN THE BLADDER, AND A URINARY DRAINAGE BAG. DIRECTIONS FOR USE: 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 2. PRECAUTIONS FOR USE 1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. 2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. 3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. 4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. 6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) 9) CONNECT THE LEAD WIRE OF CATHETER TO TEMPERATURE MONITOR WITH EXTENSION CABLE. 10) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. 11) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK®, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 12) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 13) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. 14) DO NOT STRETCH CATHETER AS DAMAGE TO OR DISLODGEMENT OF LEAD WIRE AS TEMPERATURE PROBE MAY CAUSE IMPROPER TEMPERATURE MEASUREMENT. 15) WHEN ENDO-ELECTRIC SURGERY IS PERFORMED, CARE SHOULD BE TAKEN TO PREVENT BURNS IN THE LOCAL TISSUE. 16) DO NOT WET THE LEAD WIRE AND THE JUNCTION WITH EXTENSION CABLE. 17) THIS DEVICE IS COMPATIBLE ONLY WITH MONITORS REQUIRING YSI 400-SERIES TYPE TEMPERATURE PROBES. 18) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. 19) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 20) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 21) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 22) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. 23) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. 24) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG.THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR. 25) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. 26) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. <MRI SAFETY> 1) BASED ON NON-CLINICAL STUDIES, MRI SAFETY FOR THE DEVICE IS CONFIRMED UNDER THE FOLLOWING CONDITIONS: - STATIC MAGNETIC FIELD OF 3-TESLA OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - SPATIAL GRADIENT MAGNETIC FIELD OF 720-GAUSS/CM OR LESS WITH REGARD TO MAGNETIC FIELD INTERACTIONS. - MAXIMUM MR SYSTEM REPORTED WHOLE-BODY-AVERAGED SPECIFIC ABSORPTION RATE (SAR) OF 3.5-W/KG AT 1.5- OR 3-W/KG AT 3-TESLA FOR 15 MINUTES OF SCANNING. 2) THE POSITION OF THE WIRE OF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS AN EFFECT ON THE AMOUNT OF HEATING THAT MAY DEVELOP DURING AN MRI PROCEDURE. ACCORDINGLY, THE FOLEY CATHETER WITH TEMPERATURE SENSOR MUST BE POSITIONED IN A STRAIGHT CONFIGURATION DOWN THE CENTER OF THE PATIENT TABLE (I.E., DOWN THE CENTER OF THE MR SYSTEM WITHOUT ANY LOOP) TO PREVENT POSSIBLE EXCESSIVE HEATING ASSOCIATED WITH AN MRI PROCEDURE. 3) IMPORTANTLY, THE MRI PROCEDURE SHOULD BE PERFORMED USING AN MR SYSTEM OPERATING AT A STATIC MAGNETIC FIELD STRENGTH OF 1.5-TESLA OR 3-TESLA ONLY. THE SAFE USE OF AN MR SYSTEM OPERATING AT LOWER OR HIGHER FIELD STRENGTH FOR A PATIENT WITH A FOLEY CATHETER WITH TEMPERATURE SENSOR HAS NOT BEEN DETERMINED. 4) IF THE FOLEY CATHETER WITH TEMPERATURE SENSOR HAS A REMOVABLE EXTENSION CABLE, IT SHOULD BE DISCONNECTED PRIOR TO THE MRI PROCEDURE. 5) REMOVE ALL ELECTRICALLY CONDUCTIVE MATERIAL FROM THE BORE OF THE MR SYSTEM THAT IS NOT REQUIRED FOR THE PROCEDURE (I.E., UNUSED SURFACE COILS, CABLES, ETC.). 6) KEEP ELECTRICALLY CONDUCTIVE MATERIAL THAT MUST REMAIN IN THE BORE OF THE MR SYSTEM FROM DIRECTLY CONTACTING THE PATIENT BY PLACING THERMAL AND/OR ELECTRICAL INSULATION INCLUDING AIR BETWEEN THE CONDUCTIVE MATERIAL AND THE PATIENT. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE - FAILURE TO MEASURE TEMPERATURE - IMPROPER TEMPERATURE INDICATION 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03131 | 2023-05-04 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 37303 | 37303 |  | IT WAS REPORTED THAT THE PATIENT HAD ISSUES WITH FOUR MALE EXTERNAL CATHETER. THEY THAT HAD TOO MUCH ADHESIVE AND THE PATIENT COULD NOT GET THEM OFF. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿INTENDED USE: THE SELF-ADHERING MALE EXTERNAL CATHETER IS USED FOR THE DRAINAGE OF URINE. THE CATHETER IS APPLIED BY THE PATIENT OR CAREGIVER. DESCRIPTION / INDICATION: THE SELF-ADHERING MALE EXTERNAL CATHETER IS DESIGNED FOR THE MANAGEMENT OF ADULT MALE URINARY INCONTINENCE. CONTRAINDICATION: DO NOT USE ON IRRITATED OR COMPROMISED SKIN. WARNING: REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY OR ILLNESS OF THE PATIENT. PRECAUTION: DO NOT USE IF ALLERGIC REACTION OCCURS OR IF PATIENT HAS KNOWN ALLERGIES TO DEVICE COMPONENTS. FOR GOOD HYGIENE, CHANGE CATHETER DAILY. USE OF A SINGLE DEVICE FOR LONGER PERIODS THAN 24 HOURS MAY INCREASE THE RISK OF COMPLICATIONS. FOLLOW DIRECTIONS TO REMOVE IF EXPERIENCING SWELLING, NUMBNESS, DISCOMFORT, PAIN, DISCOLORATION, OR ABNORMAL APPEARANCE. NOTE: IF EXPERIENCING PROBLEMS WITH USE OF THE DEVICE, PLEASE CONSULT YOUR HEALTHCARE PROFESSIONAL FOR ASSISTANCE. DIRECTIONS TO APPLY: VERIFY CORRECT SIZE PRIOR TO USE. TRIM PUBIC HAIR IF NECESSARY. WASH PENIS WITH MILD SOAP AND WARM WATER. DRY THOROUGHLY. WEAR TIME MAY BE REDUCED IF SKIN IS NOT DRY OR CREAM/OIL IS USED. OPEN PACKAGE AT PERFORATION. REMOVE CATHETER FROM PLASTIC INSERT, IF PRESENT. PLACE ROLLED END OVER THE END OF THE PENIS, LEAVING A SMALL SPACE BETWEEN THE END OF THE PENIS AND THE CONE OF THE CATHETER. UNROLL THE CATHETER OVER PENIS. GENTLY SQUEEZE THE CATHETER TO PROPERLY SEAL ADHESIVE TO THE SKIN. IF POSSIBLE, AVOID LEAVING A ROLLED ¿COLLAR¿ AROUND THE BASE OF THE PENIS. IMPORTANT: WEAR TIME MAY BE REDUCED IF ADHESIVE IS NOT PROPERLY SEALED TO THE SKIN. CONNECT THE CATHETER TO A DRAINAGE SYSTEM. MAKE SURE TO CHECK THAT CONNECTIONS ARE SECURE BEFORE USE. DIRECTIONS TO REMOVE: ENSURE DRAINAGE BAG IS EMPTY. DISCONNECT CATHETER FROM THE DRAINAGE SYSTEM. GENTLY ROLL THE CATHETER FORWARD AND OFF THE PENIS. IF NECESSARY, APPLY A WARM WET COMPRESS (SUCH AS A WET WASHCLOTH) AROUND THE CATHETER TO HELP LOOSEN THE ADHESIVE. DISPOSAL: AFTER REMOVAL, DISPOSE BY PLACING THE USED CATHETER INTO A WASTE CONTAINER.¿ SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03132 | 2023-05-04 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | URETHRAL CATHETER KIT | C.R. BARD, INC. (COVINGTON) -1018233 | 4A5144 | 4A5144 |  | IT WAS REPORTED THAT THERE WAS LEAKAGE FROM THE TOUCHLESS INTERMITTENT CATHETER BECAUSE THE SEAM ON THE BOTTOM OF THE BAG WAS NOT SEALED WELL, AND THIS HAD BEEN HAPPENING MULTIPLE TIMES WITH THIS BATCH. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03130 | 2023-05-04 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L20 |  | IT WAS REPORTED THAT WATER LEAKAGE WAS FOUND DURING A PRETEST. | THE REPORTED EVENT WAS CONFIRMED ¿ MANUFACTURING RELATED. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT WAS USED FOR UROLOGICAL CARE. BASED ON THE EVALUATION, A TEAR WAS FOUND ON THE RUBBERIZED LAYER CAUSING INTERLUMEN LEAKAGE. THE REPORTED EVENT IS CONFIRMED AS MANUFACTURING RELATED ISSUE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE, INSUFFICIENT LATEX STRENGTH, LOW/NON UNIFORM RUBBERIZED THICKNESS ETC CAUSING TORN RUBBERIZE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND A POSSIBLE MANUFACTURING ISSUE THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER-CONTAINING CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARDEX® SILVER LUBRICATH® FOLEY CATHETER IS A BALLOON CATHETER. SOME MAY INCLUDE A SYRINGE PREFILLED WITH STERILE WATER AND A WATER-SOLUBLE LUBRICANT AS AN ACCESSORY. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX, POLYURETHANE, SILVER THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO26EVERY 2FR 1. BALLOON CATHETER 2. ACCESSORIES SYRINGE PREFILLED WITH STERILE WATER WATER-SOLUBLE LUBRICANT (NEITHER OF THEM MAY BE INCLUDED IN SOME PRODUCT VARIATIONS.) [INTENDED USE & EFFECT- EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 1) FIG. 1 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 2) FIG. 2 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.3) FIG. 3 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.4) FIG. 4 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 5) FIG. 5 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG.6) FIG. 6 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 7) FIG. 7 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG.8) FIG. 8 3. MALFUNCTION AND ADVERSE EVENTS (1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE (2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: F,H. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03136 | 2023-05-04 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0167L26 | 0167L26 |  | IT WAS REPORTED THAT THE PHYSICIAN INSERTED 3-WAY FOLEY CATHETER AND UNABLE TO INFLATE BALLOON OR IRRIGATE CATHETER. NEW 3-WAY CATHETER OBTAINED AND INSERTED WITHOUT ISSUE. NO HARM TO PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "PHYSIOLOGICAL INTERFERENCES ~LOW LATEX MODULUS." IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: ¿ WARNING: DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. USE LUER TIP SYRINGE TO INFLATE WITH STATED ML OF STERILE WATER. OR FOR PRE-FILLED PRODUCTS, REMOVE CLIP AND SQUEEZE RESERVOIR TO INFLATE WITH STATED ML OF STERILE WATER. STORAGE: STORE CATHETERS AT ROOM TEMPERATURE AWAY FROM DIRECT EXPOSURE TO LIGHT, PREFERABLY IN THE ORIGINAL BOX. FOR UROLOGICAL USE ONLY. TO DEFLATE CATHETER BALLOON: GENTLY INSERT A LUER SLIP TIP SYRINGE IN THE CATHETER VALVE. NEVER USE MORE FORCE THAN IS REQUIRED TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. ALLOW THE PRESSURE WITHIN THE BALLOON TO FORCE THE PLUNGER BACK AND FILL THE SYRINGE WITH WATER. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN, PREVENTING BALLOON DEFLATION. IF NECESSARY, CONTACT ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PROTOCOL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. CONTAINS OR PRESENCE OF NATURAL RUBBER LATEX. CAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTIONS. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LAWS AND REGULATIONS.¿ H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03186 | 2023-05-05 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI12 | 0165SI12 |  | IT WAS REPORTED THAT THERE WAS A LINEAR BLACKENING ON THE SIDE OF THE CATHETER. AS PER THE ADDITIONAL INFORMATION RECEIVED VIA IBC ON 24APR2023, IT WAS A FOREIGN MATERIAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03196 | 2023-05-05 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT WHITE DEPRESSED DAMAGES WITH THE BALLOON PART FOR A TOTAL OF 5350 FOLEY CATHETERS FROM TWO BATCHES (BATCH# MYGN5842 AND BATCH# MYGR5328) WERE FOUND DURING USE AND INSPECTION. THE QUALITY DEPARTMENT OF COMPANY HAD DETERMINED THAT THERE WERE RISKS OF BALLOON BURSTING AND THAT THESE PRODUCTS COULD NOT BE USED IN THE PRODUCTION. THE ISSUE WAS DETECTED DURING REUSE OF A SINGLE USE MEDICAL DEVICE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 13APR2023, STATED THAT CUSTOMER WAS A DISTRIBUTION CENTER. STATED THAT 5 PHYSICAL SAMPLES OF EACH BATCH WOULD BE PROVIDED. THE USER PURCHASED 5350 UNITS AND INSPECTED SOME OF THEM HAD THE ISSUE, THUS THE CUSTOMER WAS WORRIED ABOUT THE QUALITY OF THE WHOLE BATCHES AND REQUEST TO CHANGE AND RETURN 5350 UNITS WITH ANOTHER BATCH. PER SAMPLE RECEIVED BY THE BD REPRESENTATIVE ON 14APR2023, STATED THAT THE MANUFACTURING LOT# OF SAMPLES RECEIVED FOR THE BATCH MYGN5842 WAS 2HJ031 AND FOR THE BATCH MYGR5328 WERE 2JD072, 2HH041, 2HU135 AND 2HL031 (BATCH# MYGN5842 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6) AND BATCH# MYGR5328 AND SERIAL# (B)(6) ). PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 15APR2023, STATED THAT THE CUSTOMER WAS AN OLD CUSTOMER, THEY HAVE USED BD FOLEY FOR ALMOST 18 YEARS SINCE 2006. THIS WAS THE FIRST INCIDENT EXPERIENCED BY THEM. THE MANUFACTURING LOT# OF SAMPLES FOR THE BATCH MYGN5842 WAS 2FD047 AND FOR THE BATCH MYGR5328 WERE 2HU136, 2FR083, 2JC013 AND 2HU134 (BATCH# MYGN5842 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6) AND BATCH# MYGR5328 AND SERIAL# (B)(6) ). STATED THAT THIS CUSTOMER PER CHASE FOLEY FOR REPROCESSING. WHEN THEY RECEIVED FOLEY, THEY WOULD JUST CHECK PRODUCT MODEL, SIZE, QUANTITY AND EXPIRATION DATE. WHILE DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. THERE WAS NO SAMPLING INSPECTION PERFORMED, THEY TESTED THE BALLOON OF ALL CATHETERS. STATED THAT DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL FROM IBC ON 21APR2023, STATED THAT THE SAMPLES REPORTED BY CUSTOMER WERE THOSE PICKED UP BY CUSTOMER FOR FUNCTION TESTING DURING THEIR REPRODUCTION PROCESS AND THESE PHYSICAL SAMPLES COULD NOT BE RETURNED FOR INVESTIGATION, BUT THE PHOTOS ATTACHED TO THEM. WHILE FOR THE SAMPLES RECEIVED, THESE WERE REPRESENTATIVE SAMPLES FROM THE SAME REPORTED BATCHES RETURNED FOR INVESTIGATION, THEY WERE NOT THE REAL AFFECTED SAMPLES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03189 | 2023-05-05 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI14 |  | IT WAS REPORTED THAT THE BALLOON WAS RUPTURED AT THE TIME OF REMOVAL. AS PER THE NOTIFICATION RECEIVED ON 28APR2023, AFTER INSERTION , THE URINE DID NOT FLOW OUT. | THE REPORTED EVENT WAS NOT CONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS NOT CONFIRMED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE WRONG DURATION SETTING ON AIR PURGING THAT CAUSE BLOCKED DRAINAGE LUMEN/DRAINAGE EYE OCCLUSION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: WARNINGS: 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. THE URETHRA MAY BE INJURED. (2) DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. CONTRAINDICATIONS: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RE-STERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER SHAPE, CONFIGURATION AND PRINCIPLES: BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS. SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY INTENDED USE & EFFECT- EFFICACY: THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. DIRECTIONS FOR USE: 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED. (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR. (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. PRECAUTIONS: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRA-SONOGRAPHIC GUIDANCE. C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS STORAGE METHOD AND EXPIRATION DATE: 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03190 | 2023-05-05 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT WHILE TRANSFERRING FROM THE OPERATING ROOM TO THE WARD, THE CATHETER CAUGHT ON SOMETHING AND RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03192 | 2023-05-05 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER LEAKAGE WAS FOUND. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03193 | 2023-05-05 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI16 | 0165SI16 |  | IT WAS REPORTED THAT THE BALLOON WAS RUPTURED AT THE TIME OF REMOVAL. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03191 | 2023-05-05 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT DURING A PRETEST, STERILE WATER LEAKAGE WAS FOUND. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 24APR2023, ON CHECKING THE RETURNED SAMPLE, A BALLOON RUPTURE WAS CONFIRMED. | THE REPORTED EVENT WAS CONFIRMED ¿ UNKNOWN CAUSE. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. BASED ON THE EVALUATION, IT WAS OBSERVED IRREGULAR BURST WITHOUT MISSING PIECES. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01725 | 2023-05-05 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | IT WAS REPORTED THAT DURING PRE-FLUSHING PRIOR TO PLACEMENT, THE PICC CATHETER WAS FOUND LEAKING LIQUID FROM THE SCALE OF 51 CM TO 52 CM. THIS MADE THE PRODUCT UNUSABLE. A NEW PICC CATHETER WAS OPENED FOR INSERTION. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A LEAKING GROSHONG CATHETER IS CONFIRMED AND WAS DETERMINED TO BE DUE TO STYLET DAMAGE. ONE 4 FR GROSHONG CLEARVUE CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION SHOWED NO BLOOD USE RESIDUES THROUGHOUT THE DEVICE, AND THE CATHETER WAS RETURNED WITH THE INLAID STYLET. A MICROSCOPIC OBSERVATION REVEALED A DIAGONALLY ALIGNED SPLIT BETWEEN THE 52 CM AND 53 CM DEPTH MARKERS. MULTIPLE LONGITUDINAL SCORING MARKS WERE OBSERVED ON THE INSIDE SURFACE OF THE CATHETER IN THE VICINITY OF THE SPLIT, AND THE FRACTURE SURFACE APPEARED TO HAVE BRAIDED MARKS SIMILAR TO THE PATTERN OF THE STYLET. THE SPLIT APPEARED TO CONTINUE INSIDE THE CATHETER WHILE ONLY A PORTION OF THE SPLIT WENT COMPLETELY THROUGH THE CATHETER. THE FEATURES AND IMPRESSIONS ALONG THE SPLIT AND AROUND THE SPLIT ARE CONSISTENT WITH STYLET DAMAGE. THE COMPLAINT OF A LEAKING CATHETER IS CONFIRMED. ATTEMPTS OF REMOVAL AND INSERTION OF THE STYLET BEFORE FLUSHING MAY CAUSE DAMAGE TO THE CATHETER, AND MANIPULATION OF THE STYLET INSIDE THE CATHETER MAY CAUSE THE CATHETER TO SPLIT. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03200 | 2023-05-05 | 2023-04-15 | I, n, j, u, r, y | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE PATIENT GOT INFECTION ON LAST SHIPMENT AND THE REPRESENTATIVE WORKED WITH THEM ON TRYING READY TO GO CATHETERS. THE PATIENT DID NOT REMEMBER WHAT ORDER IT WAS OR WHAT ITEM IT WAS THEY THOUGHT MIGHT BE WITH DIRTY HANDS AND BURSTING AND INSERTING IT MIGHT HAVE CAUSED URINARY TRACT INFECTION. AFFECTED QUANTITY NOT REPORTED. IT WAS UNKNOWN IF THE DEVICE CAUSED THE INFECTION AND UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿ PATIENT SENSITIVITY TO MATERIALS, MECHANICAL AND / OR CHEMICAL RESPONSES FROM THE PATIENT¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSE BUT IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: ¿ THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. ¿ WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. ¿ RELEASE THE STERILE WATER FROM THE FOIL PACKET. ¿ TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. ¿ PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. ¿ WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. ¿ WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. ¿ NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. ¿ USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. ¿ TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. ¿ FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: ¿ THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. ¿ URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01719 | 2023-05-05 | 2023-04-17 | I, n, j, u, r, y | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED "REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE. DOUBLE LUMEN IN UPPER ARM." | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03197 | 2023-05-05 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 129416 | 129416 |  | IT WAS REPORTED, THAT THE CUSTOMER HAD ENCOUNTERED ISSUE WITH THE BALLOONS OF SILICONE INDWELLING CATHETERS LOSING FLUID AND DEFLATING EARLIER THIS WEEK. THE FOLEY CATHETER WAS PLACED ON (B)(6) 2023 IN SURGERY, AND IT WAS DOCUMENTED THAT THE BALLOON WAS INFLATED. HOWEVER, AT THE TIME OF CATHETER REMOVAL, THE BALLOON WAS FOUND TO BE COMPLETELY DEFLATED (LOT# NGGR3688). THIS WAS THE FIRST TIME THEY HAD HEARD OF THIS ISSUE. BUT THE REGISTERED NURSE ON THIS PATIENT CARE UNIT NOTED, THIS MULTIPLE TIMES RECENTLY (LOT# UNK). ALTHOUGH THEY DID NOT HAVE ANY SPECIFICS OR THE CATHETER. THEY DID RE-INFLATE THE BALLOON YESTERDAY AND PLACED IT IN A BAG TO SAVE IT. TODAY, THERE WAS SOME FLUID IN THE BAG, WHICH WOULD INDICATE THE BALLOON WAS SLOWLY LEAKING. PER ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2023, IT WAS REPORTED, THAT THE CUSTOMER WERE SHOWING THAT THEY HAD A TOTAL OF FOUR DIFFERENT FOLEYS IN THEIR OFFICE WAITING THAT THEY CAN RETURN FOR INVESTIGATION. THE 3 CATHETERS THAT SEEM TO HAVE AN ISSUE WITH DEFLATING BALLOONS, ONE FOLEY HAD SLIPPED OUT OF THE PATIENT, AND NURSE STATED, THAT THEY TESTED THE BALLOON AND FELT A LEAK (LOT# NGFX5338), OTHER FOLEY (LOT# NGGX0360) AND AN OTHER FOLEY (LOT# NGG20749). PER FOLLOW UP VIA PHONE 17 APR 2023, CUSTOMER HAD RETURNED FOUR EACH FOR TESTING. NO MEDICAL INTERVENTION WAS REPORTED FOR ANY OF THE PATIENTS INVOLVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01731 | 2023-05-05 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER "PATIENT HAD A SINGLE LUMEN 4 FR PICC PLACED ON (B)(6) 2023. PROCEDURE NOTES THAT THERE WAS NO DIFFICULTY WITH INSERTION. ON 4/17/23 VAT WAS CALLED TO ASSESS LINE DUE TO PT, WHO HAPPENS TO BE A NURSE, REPORTED PAIN/COOLNESS WITH FLUSHING ABOVE THE INSERTION SITE. WHEN CATHETER WAS REMOVED A PARTIAL FRACTURE WAS NOTED AT APPROXIMATELY 4 CM MARK." ADDITIONAL INFORMATION RECEIVED: PICC WAS REMOVED AND ANOTHER ONE PLACED IN OTHER ARM. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A FRACTURE IN THE CATHETER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4 FR S/L POWERPICC CATHETER. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE ACCUMULATED THROUGH FLEXURAL FATIGUE. FLEXURAL FATIGUE OCCURS DUE TO CYCLIC KINKING OF THE CATHETER TUBE IN WHICH PHYSIOLOGICAL, PLACEMENT, USAGE, AND MECHANICAL FACTORS MAY GRADUALLY FORM A CRACK(S) IN THE CATHETER. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SPLIT WAS OBSERVED JUST DISTAL OF THE 4 CM DEPTH MARKER. THE CATHETER SPLIT CONTAINED PHYSICAL FEATURES ASSOCIATED WITH MATERIAL FATIGUE, AND THE CHARACTERISTICS OBSERVED WHICH SUPPORTED THIS TYPE OF FAILURE INCLUDED: DAMAGE WHICH WAS CIRCUMFERENTIALLY ALIGNED. FRACTURE EDGES WHICH WERE ROUNDED AND POLISHED DUE TO REPEATED MATERIAL WEAR 'C' SHAPED IMPRESSIONS LEADING INTO THE FRACTURE SITE WHICH ARE CONSISTENT WITH MATERIAL BUCKLING DUE TO MOVEMENT WHICH CAUSED THE FRACTURE EDGES TO BE PRESSED TOGETHER OVERALL ELLIPTICAL SHAPE TO THE FRACTURE CROSS-SECTION (A RESULT OF REPEATED KINKING OF THE TUBING) AN EXAMINATION OF THE CATHETER STRUCTURE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. THE DAMAGE LOCATION SUGGESTED THAT CATHETER SECUREMENT, ACCESS AND MAINTENANCE TECHNIQUES MAY HAVE CONTRIBUTED. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01722 | 2023-05-05 | 2023-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER BASIC TRAY (5F) (DUAL-LUMEN) (MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 8295118 |  | IT WAS REPORTED BY THE CUSTOMER THAT AT 9:25 AM, (B)(6) 2023, THE PATIENT'S CATHETER WAS FOUND TO BE BROKEN WHEN THE INFUSION TUBE WAS OPENED AND FLUSHED. IMMEDIATELY CLAMP THE PROXIMAL HEART END WITH HEMOSTATIC FORCEPS, COMFORT THE PATIENT, INFORM THE HEAD NURSE AND THE LEADER OF THE STATIC THERAPY TEAM, AND AFTER BEDSIDE INSPECTION, PULL OUT THE TUBE, CHECK THE CATHETER, AND PROVIDE PSYCHOLOGICAL COMFORT TO THE PATIENT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03195 | 2023-05-05 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED, THAT THE INTERMITTENT CATHETER WAS NOT LUBRICATED ENOUGH, AND PATIENT HAD TROUBLE GETTING THE CATHETER TO PASS THROUGH THE PROSTATE WITHOUT BENDING. PER FOLLOW UP VIA PHONE ON (B)(6) 2023, IT WAS STATED, THAT THEIR SUPPLIER WAS (B)(4). AND SINCE THEY SWITCHED FROM THE LIQUID CATHETERS TO THE GEL ALTERNATIVES, CUSTOMER HAD SEEN IMPROVEMENTS DURING USAGE. THEIR UROLOGIST RECOMMENDED THE CURE ULTRA CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01733 | 2023-05-05 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT,4FR POWERMIDLINE | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | P4154108D |  | IT WAS REPORTED BY THE CUSTOMER THAT "CUSTOMER WAS BREAKING APART THE DILATOR OF THE MIDLINE WHEN THE GRIPS BROKE OFF. NURSE HAD TO TAKE 4X4S FROM KIT TO KEEP TRYING TO PEEL OFF THE ORANGE PART OFF OF THE DILATOR." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01734 | 2023-05-05 | 2023-04-18 | I, n, j, u, r, y | Manufacturer report | BARD TRIALYSIS CATHETER | CATHETER, HEMODIALYSIS, TRIPLE LUMEN, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | THE FOLLOWING WAS OBTAINED DURING CLINICAL ASSESSMENT. REDNESS OR SWELLING WAS IDENTIFIED AT OR AROUND THE CENTRAL LINE INSERTION SITE TRIPLE LUMEN IN THE GROIN. NO OTHER INFORMATION WAS PROVIDED. | 11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THIS COMPLAINT WAS RECEIVED AS PART OF A CLINICAL SURVEY. CONTACT INFORMATION WAS REQUESTED BUT NOT PROVIDED BY THE CUSTOMER IN THE SURVEY PROCESS. AS SUCH, ANY ADDITIONAL INFORMATION INCLUDING REGARDING THE PATIENT OR SUBJECT SAMPLE CANNOT BE OBTAINED. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01730 | 2023-05-05 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER, SHERLOCK 3CG TIP POSITIONING SYSTEM,STYLET 4F FULL TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT CAME TO THE EMERGENCY ROOM WHEN IT LEAKED UNDER THE DRESSING AT THE HEALTH CENTER WHEN THE PICC LINE WAS REPOSITIONED (CARE AND MAINTAINS). WHEN THE PICC LINE WAS FLUSHED FLUID CAME FROM THE INSERTION SITE. BLOOD WAS ALSO OBTAINED DURING ASPIRATION FROM THE INSERTION SITE. THE PICC LINE ARE DRAWN OUT FROM THE PATIENT WITHOUT PROBLEMS. WHEN FLUSHED OUTSIDE THE BODY, A SMALL HOLE IS VISIBLE THAT WAS INSIDE IN THE PATIENT'S ARM. THE PATIENT FELT WELL THE WHOLE TIME. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01727 | 2023-05-05 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PPICC SV 3FSL,3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 22173208 |  | IT WAS REPORTED BY THE CUSTOMER THAT LAST WEEK A PATIENT'S LINE SPLIT EXTERNALLY, IT WAS A 3FR. THE LINE WAS SECURED AT 5CM AT THE SKIN WITH A SECUR-A-CATH IN SITU. THE HOLE IS BETWEEN THE 2-3CM MARKING. THE PATIENT HAD A FEW DAYS WITH A CANNULA BEING INSERTED IN OUR ED DEPARTMENT, DOSES WERE LATE DUE TO THIS. THE TREATMENT WAS DELAYED, AND SHE RECEIVED MULTIPLE CANNULAS. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01726 | 2023-05-05 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 2194108 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT IS USING THE PICC FOR UNKNOWN CHEMO. NOW THE CATHETER IS LEAKING, AND THEY PULLED IT OUT. ADDITIONAL INFORMATION RECEIVED APRIL 22 2023: IT WAS REPORTED BY THE CUSTOMER "THE PATIENT WAS NOT HARMED BUT NEEDED A NEW PICC." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01724 | 2023-05-05 | 2023-04-15 | I, n, j, u, r, y | Manufacturer report | PPICC BASIC 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1385118QJ |  | IT WAS REPORTED BY THE CUSTOMER THAT THE CATHETER WAS INSERTED WITH THE STYLET TIP PROTRUDING FROM THE CATHETER TIP. BECAUSE THERE WAS VASCULAR TORTUOSITY NEAR THE SHEATH CENTER AND THE CATHETER COULD NOT BE ADVANCED, WHEN THE CATHETER WAS PULLED BACK, THE STYLET WAS BROKEN. CT EXAMINATION WAS PERFORMED AND REVEALED THAT THE STYLET SEGMENT MIGRATED TO THE PULMONARY ARTERY. THE DOCTOR CONTINUES THE FOLLOW-UP OBSERVATION. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF BROKEN STYLET WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 3CG STYLET WITH A T-LOCK ASSEMBLY. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND CONFIRMED TO BE BROKEN. THE BREAK OCCURRED ON THE STYLET TUBING OUTSIDE THE MAGNET STRING REGION. THE FOLLOWING OBSERVATIONS WERE MADE WHICH WERE CONSISTENT WITH THIS FAILURE TYPE: MICROSCOPIC EXAMINATION REVEALED DEFORMATION OF THE STYLET TUBING AT THE BREAK SITE. MICROSCOPIC EXAMINATION REVEALED LOCALIZED DELAMINATION OF THE STYLET TUBING AT THE DAMAGE SITE(S), INDICATIVE OF TUBING KINKING/FOLDING. KINKING OF THE SURROUNDING STYLET TUBING. MEASUREMENTS OF THE STYLET CORE WIRE WERE TAKEN AND CONFIRMED TO BE WITHIN SPECIFICATIONS. THE CONDITION OF THE SAMPLE PREVENTED ACCURATE MEASUREMENTS OF THE STYLET TUBING AND MAGNET FROM BEING TAKEN. THE OBSERVED FEATURES OF THE BREAK IN THE STYLET WERE INDICATIVE OF CIRCUMSTANCES WHICH INCLUDED STYLET KINKING, LIKELY DURING THE INSERTION PROCESS. ADDITIONALLY, THE DESCRIPTION OF THE EVENT INDICATED THE STYLET TIP WAS PROTRUDING FROM THE CATHETER TIP, WHICH LIKELY ALLOWED THE STYLET TO KINK AND ULTIMATELY BREAK DURING THE INSERTION PROCESS. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H11. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01723 | 2023-05-05 | 2023-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD PICC CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT THE INTRODUCER WIRE HAD A BURR. DISTAL TO CHANGE TOO FLOPPY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED GUIDEWIRE IS CONFIRMED BUT THE EXACT CAUSE REMAINS UNKNOWN. ONE 50 CM GUIDEWIRE WAS RETURNED FOR EVALUATION. THE WIRE AND WELD TIPS WERE FOUND TO BE FULLY INTACT. THE PROXIMAL END OF THE FLOPPY SECTION WAS OBSERVED TO CONTAIN SLIGHT BENDS. NO APPARENT KINKS WERE NOTED. A MICROSCOPIC INSPECTION OF THE DAMAGED GUIDEWIRE SECTION DID NOT REVEAL ANY BURRS OR PROTRUSIONS FROM THE COIL WIRE. THE OUTER DIAMETER WAS MEASURED AND WAS FOUND TO BE WITHIN MANUFACTURING SPECIFICATION. BASED ON THE DESCRIPTION OF THE REPORTED EVENT, POSSIBLE CONTRIBUTING FACTORS INCLUDE ADVANCEMENT AGAINST RESISTANCE AND PATIENT ANATOMY. SINCE THE GUIDEWIRE WAS OBSERVED TO BE BENT, THE COMPLAINT IS CONFIRMED; HOWEVER, THE EXACT CAUSE REMAINS UNKNOWN. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01732 | 2023-05-05 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | IT WAS REPORTED THAT WHEN THE VIP NEUROSURGERY CUSTOMER WAS PLACING THE CATHETER IN THIS PATIENT WITH CEREBRAL HEMORRHAGE, THE PUNCTURE WAS CONDUCTED SMOOTHLY AND BLOOD WAS SEEN JUST WITH ONE STICK, HAVING GOOD BLOOD REFLEX. BUT THE GUIDE WIRE COULD NOT BE INSERTED. MULTIPLE ATTEMPTS WERE MADE, INCLUDING ADJUSTING THE ANGLE OF INSERTION, BUT ALL FAILED. SINCE THE PATIENT WAS ALREADY ON THE OPERATING BED AND THERE WERE NO MST KITS IN THE HOSPITAL, THE CUSTOMER HAD TO OPENED A PICC CATHETER INTRODUCER KIT AND USED THE NEW INTRODUCER NEEDLE, THROUGH WHICH THE GUIDE WIRE PASSED. AFTER OPERATION, WE ATTEMPTED TO CROSS THE GUIDE WIRE THROUGH THE TWO ENDS OF THE PROBLEMATIC INTRODUCER NEEDLE OUT OF THE PATIENT AND FOUND THAT THE INTRODUCER NEEDLE WAS OCCLUDED. ADDITIONAL INFORMATION 5/4: IT WAS REPORTED BY THE CUSTOMER "PATIENT WAS STABLE AND NO NEGATIVE OUTCOMES DUE TO DELAY." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF INABILITY TO INSERT THE GUIDEWIRE INTO THE INTRODUCER NEEDLE WAS CONFIRMED. THE SAMPLE WAS EVALUATED AND THIS EVENT WAS DETERMINED TO BE RELATED TO A MANUFACTURE CONDITION. BD IS WORKING TO PREVENT FUTURE OCCURRENCES OF SIMILAR EVENTS AND VALUES YOUR SUPPORT AND PARTNERSHIP IN COMMUNICATING THIS EVENT AND COORDINATING THE RETURN OF THE SAMPLE. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03188 | 2023-05-05 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT WHITE DEPRESSED DAMAGES WITH THE BALLOON PART FOR A TOTAL OF 5350 FOLEY CATHETERS FROM TWO BATCHES (BATCH# MYGN5842 AND BATCH# MYGR5328) WERE FOUND DURING USE AND INSPECTION. THE QUALITY DEPARTMENT OF COMPANY HAD DETERMINED THAT THERE WERE RISKS OF BALLOON BURSTING AND THAT THESE PRODUCTS COULD NOT BE USED IN THE PRODUCTION. THE ISSUE WAS DETECTED DURING REUSE OF A SINGLE USE MEDICAL DEVICE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 13APR2023, STATED THAT CUSTOMER WAS A DISTRIBUTION CENTER. STATED THAT 5 PHYSICAL SAMPLES OF EACH BATCH WOULD BE PROVIDED. THE USER PURCHASED 5350 UNITS AND INSPECTED SOME OF THEM HAD THE ISSUE, THUS THE CUSTOMER WAS WORRIED ABOUT THE QUALITY OF THE WHOLE BATCHES AND REQUEST TO CHANGE AND RETURN 5350 UNITS WITH ANOTHER BATCH. PER SAMPLE RECEIVED BY THE BD REPRESENTATIVE ON 14APR2023, STATED THAT THE MANUFACTURING LOT# OF SAMPLES RECEIVED FOR THE BATCH MYGN5842 WAS 2HJ031 AND FOR THE BATCH MYGR5328 WERE 2JD072, 2HH041, 2HU135 AND 2HL031 (BATCH# MYGN5842 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6),BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6) AND BATCH# MYGR5328 AND SERIAL# (B)(6)). PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 15APR2023, STATED THAT THE CUSTOMER WAS AN OLD CUSTOMER, THEY HAVE USED BD FOLEY FOR ALMOST 18 YEARS SINCE 2006. THIS WAS THE FIRST INCIDENT EXPERIENCED BY THEM. THE MANUFACTURING LOT# OF SAMPLES FOR THE BATCH MYGN5842 WAS 2FD047 AND FOR THE BATCH MYGR5328 WERE 2HU136, 2FR083, 2JC013 AND 2HU134 (BATCH# MYGN5842 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6) AND BATCH# MYGR5328 AND SERIAL# (B)(6)). STATED THAT THIS CUSTOMER PER CHASE FOLEY FOR REPROCESSING. WHEN THEY RECEIVED FOLEY, THEY WOULD JUST CHECK PRODUCT MODEL, SIZE, QUANTITY AND EXPIRATION DATE. WHILE DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. THERE WAS NO SAMPLING INSPECTION PERFORMED, THEY TESTED THE BALLOON OF ALL CATHETERS. STATED THAT DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL FROM IBC ON 21APR2023, STATED THAT THE SAMPLES REPORTED BY CUSTOMER WERE THOSE PICKED UP BY CUSTOMER FOR FUNCTION TESTING DURING THEIR REPRODUCTION PROCESS AND THESE PHYSICAL SAMPLES COULD NOT BE RETURNED FOR INVESTIGATION, BUT THE PHOTOS ATTACHED TO WERE BELONGING TO THEM. WHILE FOR THE SAMPLES RECEIVED, THESE WERE REPRESENTATIVE SAMPLES FROM THE SAME REPORTED BATCHES RETURNED FOR INVESTIGATION, THEY WERE NOT THE REAL AFFECTED SAMPLES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11 SECTION A THROUGH F THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03199 | 2023-05-05 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER TIEMANN MODEL COUDE | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0102L16 | 0102L16 |  | IT WAS REPORTED, THAT THE CUSTOMER HAD ENCOUNTERED ISSUE WITH THE BALLOONS OF SILICONE INDWELLING CATHETERS LOSING FLUID, AND DEFLATING EARLIER THIS WEEK. THE FOLEY CATHETER WAS PLACED ON (B)(6) 2023 IN SURGERY, AND IT WAS DOCUMENTED, THAT THE BALLOON WAS INFLATED. HOWEVER, AT THE TIME OF CATHETER REMOVAL, THE BALLOON WAS FOUND TO BE COMPLETELY DEFLATED (LOT# NGGR3688). THIS WAS THE FIRST TIME THEY HAD HEARD OF THIS ISSUE. BUT THE REGISTERED NURSE ON THIS PATIENT CARE UNIT NOTED, THIS MULTIPLE TIMES RECENTLY (LOT# UNK). ALTHOUGH THEY DID NOT HAVE ANY SPECIFICS OR THE CATHETER. THEY DID RE-INFLATE THE BALLOON YESTERDAY AND PLACED IT IN A BAG TO SAVE IT. TODAY, THERE WAS SOME FLUID IN THE BAG, WHICH WOULD INDICATE THE BALLOON WAS SLOWLY LEAKING. PER ADDITIONAL INFORMATION RECEIVED ON 15 APR 2023, IT WAS REPORTED, THAT THE CUSTOMER WERE SHOWING THAT THEY HAD A TOTAL OF FOUR DIFFERENT FOLEYS IN THEIR OFFICE WAITING THAT THEY CAN RETURN FOR INVESTIGATION. THE 3 CATHETERS THAT SEEM TO HAVE AN ISSUE WITH DEFLATING BALLOONS, ONE FOLEY HAD SLIPPED OUT OF THE PATIENT. AND NURSE STATED, THAT THEY TESTED THE BALLOON AND FELT A LEAK (LOT# NGFX5338), OTHER FOLEY (LOT# NGGX0360), AND AN OTHER FOLEY (LOT# NGG20749). PER FOLLOW UP VIA PHONE 17 APR 2023, CUSTOMER HAD RETURNED FOUR EACH FOR TESTING. NO MEDICAL INTERVENTION WAS REPORTED FOR ANY OF THE PATIENTS INVOLVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03187 | 2023-05-05 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT THE CATHETER SPONTANEOUSLY FELL OUT OF THE PATIENT DUE TO FIXATION WATER LEAKAGE. THE LOCATION OF LEAKAGE WAS UNKNOWN. THE EVENT HAPPENED JUST AFTER PLACEMENT AND THERE WAS NO BALLOON RUPTURE OR TEAR. NO MATERIALS POSSIBLY CAME INTO CONTACT WITH THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01721 | 2023-05-05 | 2023-04-16 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | IT WAS REPORTED THAT THIS ELDERLY FEMALE PATIENT WAS ADMITTED TO HOSPITAL DUE TO TRAUMA IN A CAR ACCIDENT. ON (B)(6) 2023, THE CATHETER WAS INSERTED UNDER ULTRASOUND GUIDANCE. THE INSERTION WAS DONE SMOOTHLY AND THE CATHETER WAS POSITIONED AT T6 BY X-RAY. THE CATHETER WAS USED PROPERLY FOR INFUSION. ON (B)(6) 2023, THE PATIENT REMOVED THE EXTENSION TUBE BY HERSELF AND IT BROKE AT THE CONNECTION WITH THE CATHETER. 43 CM OF THE CATHETER BODY WAS NOT SEEN. NEITHER X-RAY NOR CT SCANNING THE WHOLE BODY DETECTED THE RUPTURED PORTION OF THE CATHETER. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. BASED ON A REVIEW OF AVAILABLE COMPLAINT INFORMATION AND AN EVALUATION OF THE AVAILABLE SAMPLE EVIDENCE, THE FOLLOWING WAS CONCLUDED: THE COMPLAINT OF A DAMAGED GROSHONG CATHETER WAS CONFIRMED; HOWEVER, THE ROOT CAUSE WAS NOT IDENTIFIED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE PHOTOGRAPH WHICH DEPICTED A 4FR S/L GROSHONG NXT CLEARVUE CATHETER. THE DEPICTED PORTION OF THE DEVICE INCLUDED AN ASSEMBLED CONNECTOR. A PORTION OF BLUE CATHETER APPEARED EVIDENT WITHIN THE CLEAR STRAIN-RELIEF SLEEVE PORTION OF THE DISTAL CONNECTOR SEGMENT. NO ADDITIONAL CATHETER WAS EVIDENT IN THE PHOTOGRAPH. CATHETER DAMAGE WAS EVIDENT IN THE SUBMITTED PHOTOGRAPH; HOWEVER, INSPECTION OF THE PHOTOGRAPH WAS INSUFFICIENT TO IDENTIFY THE CAUSE OF THE DAMAGE. CONSEQUENTLY THIS COMPLAINT IS CONFIRMED AS ¿CAUSE UNKNOWN¿ AT THIS TIME. POTENTIAL CONTRIBUTING FACTORS INCLUDE TENSILE (PULLING) STRESS AND SHARP INSTRUMENT CONTACT. H3 OTHER TEXT: EVALUATION FINDINGS ARE IN SECTION H.11. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03180 | 2023-05-05 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® CLEAN-CATH® INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 421714 | 421714 |  | IT WAS REPORTED THAT THE CATHETER WAS BROKEN ON REMOVAL FROM THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
MW5117349 | 2023-05-05 | 2023-04-27 | I, n, j, u, r, y | Voluntary report | BARD POWERPORT IMPLANTABLE PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | BECTON DICKINSON/BARD ACCESS SYSTEMS, INC. |  | 1708000 |  | PORT-A-CATH FRACTURED IN THE MIDPOINT OF THE CATHETER. THE RESERVOIR AND PROXIMAL PORTION OF THE CATHETER WERE REMOVED; THE DISTAL PORTION COULD NOT SAFELY BE REMOVED AND WAS LEFT IN THE SUBCLAVIAN VEIN AND WILL REQUIRE VASCULAR INTERVENTIONAL RADIOLOGY AT AN OUTSIDE FACILITY FOR REMOVAL. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01728 | 2023-05-05 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7617405 |  | IT WAS REPORTED BY THE CUSTOMER THAT, THE FLUID LEAK FROM THE CONNECTION BETWEEN TAPER CONNECTOR AND CATHETER. NO PATIENT HARM REPORTED, NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01708 | 2023-05-05 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT SOMETIMES POST PORT PLACEMENT PROCEDURE, PATIENT ALLEGEDLY EXPERIENCED DISCOMFORT IN THE THROAT. IT WAS FURTHER REPORTED THAT PATIENT ALLEGEDLY EXPERIENCED BLEEDING AT THE INSERTION SITE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03164 | 2023-05-05 | 2023-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123614A |  | IT WAS REPORTED THAT THE FOLEY CATHETER FELL OUT FROM THE PATIENT DUE TO BALLOON LEAKS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER). THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03161 | 2023-05-05 | 2023-04-13 | I, n, j, u, r, y | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT DURING PLACEMENT TO THE PATIENT, THE BALLOON HAD BURST. A CT SCAN SHOWED NO STONES IN THE PATIENT BODY. ACCORDING TO THE INQUIRY SHEET ATTACHED, THE EVENT HAS OCCURRED JUST AFTER PLACEMENT, THERE WAS BALLOON RUPTURE OR TEAR AND THE WATER LEAKED FROM THE BALLOON. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. THE CATHETER WAS INSERTED BY HAND AND SO GLOVES POSSIBLY CAME INTO CONTACT WITH THE CATHETER. THERE WAS NO PATIENT'S URINE STONE AND THE CATHETER WAS NOT PULLED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS]: 1. METHOD FOR USE: (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS: (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] CONTRAINDICATIONS]: 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER. [SHAPE, CONFIGURATION AND PRINCIPLES]. MATERIAL: BALLOON CATHETER: SILICONE; SILVER COATING. THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. SIZES OF CATHETERS: AVAILABLE IN SIZES 12 TO 24 EVERY 2FR. 1. BALLOON CATHETER. [INTENDED USE & EFFECT- EFFICACY]: THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. THE SURFACE OF CATHETER IS COATED FIRST WITH A TRACE AMOUNT OF METALLIC SILVER, AND THEN WITH POLYURETHANE ONTO THAT, TO INHIBIT PROLIFERATION OF BACTERIA THAT MAY ADHERE TO THE CATHETER. [DIRECTIONS FOR USE]: 1. METHOD OF USE: THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. (3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLE-LESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER TO DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE: (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. (4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY.] (5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS]: 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS). 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS: 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. TROUBLESHOOTING: WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD: WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD: 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN'T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD: 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. 1) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. 2) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. 3) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS: 1) MALFUNCTION: CATHETER KINKING, DAMAGE, RUPTURE. DIFFICULTY OR FAILURE TO REMOVE THE DEVICE. OCCLUSION OF CATHETER INNER LUMENS. ENCRUSTATION. ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE. DEVICE DAMAGE DUE TO INAPPROPRIATE USE. 2) ADVERSE EVENTS: URINARY-TRACT INFECTION. HEMORRHAGE, HEMATURIA. ALLERGY REACTION TO THE DEVICE. CALCULUS FORMATION. EDEMA. PAIN. DISCOMFORT. INJURY OF BLADDER OR URETHRAL. URETHRITIS, URINARY INCONTINENCE. RETAINED BALLOON FRAGMENTS. [STORAGE METHOD AND EXPIRATION DATE]: 1. STORAGE: STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE: INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX.." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3: OTHER TEXT: THE DEVICE WAS NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03162 | 2023-05-05 | 2023-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER INSERTION TRAY | FOLEY CATHETER INSERTION TRAY | C.R. BARD INC. (COVINGTON) -1018233 | 802010 | 802010 |  | IT WAS REPORTED THAT THE FOLEY CATHETER INSERTION TRAY SYRINGES WERE NOT PREFILLED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "ERROR OF INSPECTOR". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03167 | 2023-05-05 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE BURST PACKETS IN A BOX OF 30 MAGIC 3 HYDROPHILIC INTERMITTENT CATHETERS WERE ALREADY BURST WHEN THEY HAVE ARRIVED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE, "INCORRECT PROCESS PARAMETERS". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03185 | 2023-05-05 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT WHITE DEPRESSED DAMAGES WITH THE BALLOON PART FOR A TOTAL OF 5350 FOLEY CATHETERS FROM TWO BATCHES BATCH# MYGN5842 AND BATCH# MYGR5328, WERE FOUND DURING USE AND INSPECTION. THE QUALITY DEPARTMENT OF COMPANY HAD DETERMINED THAT THERE WERE RISKS OF BALLOON BURSTING AND THAT THESE PRODUCTS COULD NOT BE USED IN THE PRODUCTION. THE ISSUE WAS DETECTED DURING REUSE OF A SINGLE USE MEDICAL DEVICE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, STATED THAT CUSTOMER WAS A DISTRIBUTION CENTER. STATED THAT 5 PHYSICAL SAMPLES OF EACH BATCH WOULD BE PROVIDED. THE USER PURCHASED 5350 UNITS AND INSPECTED SOME OF THEM HAD THE ISSUE, THUS THE CUSTOMER WAS WORRIED ABOUT THE QUALITY OF THE WHOLE BATCHES AND REQUEST TO CHANGE AND RETURN 5350 UNITS WITH ANOTHER BATCH. PER SAMPLE RECEIVED BY THE BD REPRESENTATIVE ON (B)(6) 2023, STATED THAT THE MANUFACTURING LOT# OF SAMPLES RECEIVED FOR THE BATCH MYGN5842 WAS 2HJ031. THE BATCH MYGR5328 WERE 2JD072, 2HH041, 2HU135 AND 2HL031. BATCH# MYGN5842 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6) ,BATCH# MYGR5328 AND SERIAL# (B)(6), AND BATCH# MYGR5328 AND SERIAL# (B)(6). PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, STATED THAT THE CUSTOMER WAS AN OLD CUSTOMER, THEY HAVE USED BD FOLEY FOR ALMOST 18 YEARS SINCE 2006. THIS WAS THE FIRST INCIDENT EXPERIENCED BY THEM. THE MANUFACTURING LOT# OF SAMPLES FOR THE BATCH MYGN5842 WAS 2FD047 AND FOR THE BATCH MYGR5328 WERE 2HU136, 2FR083, 2JC013 AND 2HU134 (BATCH# MYGN5842 AND SERIAL# (B)(6) ,BATCH# MYGR5328 AND SERIAL# (B)(6) , BATCH# MYGR5328 AND SERIAL# (B)(6) ,BATCH# MYGR5328 AND SERIAL# (B)(6) AND BATCH# MYGR5328 AND SERIAL# (B)(6). STATED THAT THIS CUSTOMER PER CHASE FOLEY FOR REPROCESSING. WHEN THEY RECEIVED FOLEY, THEY WOULD JUST CHECK PRODUCT MODEL, SIZE, QUANTITY AND EXPIRATION DATE. WHILE DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. THERE WAS NO SAMPLING INSPECTION PERFORMED, THEY TESTED THE BALLOON OF ALL CATHETERS. STATED THAT DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL FROM IBC ON (B)(6) 2023, STATED THAT THE SAMPLES REPORTED BY CUSTOMER WERE THOSE PICKED UP BY CUSTOMER FOR FUNCTION TESTING DURING THEIR REPRODUCTION PROCESS AND THESE PHYSICAL SAMPLES COULD NOT BE RETURNED FOR INVESTIGATION, BUT THE PHOTOS ATTACHED TO WERE BELONGING TO THEM. WHILE FOR THE SAMPLES RECEIVED, THESE WERE REPRESENTATIVE SAMPLES FROM THE SAME REPORTED BATCHES RETURNED FOR INVESTIGATION, THEY WERE NOT THE REAL AFFECTED SAMPLES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03165 | 2023-05-05 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETER BENDED AND DID NOT GO THROUGH THE PROSTATE AS NEEDED , ALSO STATE THAT THEY WERE NOT LUBRICATED ENOUGH. PER FOLLOW UP VIA PHONE ON 21APR2023, IT WAS STATED THAT THEIR SUPPLIER WAS LIBERATOR, SINCE THEY HAD SWITCHED FROM THE LIQUID CATHETERS TO THE GEL ALTERNATIVES, THEY HAD SEEN IMPROVEMENTS DURING USAGE. THEIR UROLOGIST RECOMMENDED THE CURE ULTRA CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR, VISCOMETER FAILURE". THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03166 | 2023-05-05 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 63614 |  | IT WAS REPORTED THAT IN THE LAST ORDER, OF THE 2/3 FOLEY CATHETERS IN EACH BOX, THERE WAS NO LUBRICANT IN THE SACHETS OR THEY HAD COME SLIGHTLY OPEN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
MW5117357 | 2023-05-05 | 2023-04-28 | I, n, j, u, r, y | Voluntary report | POWERGLIDE PRO MIDLINE CATHETER | CATHETER, INTRAVASCULAR, THERAPEUTIC, SHORT-TERM LESS THAN 30 DAYS | BARD ACCESS SYSTEMS |  | F120080T |  | THIS RN WAS ASKED TO RE-VISIT A POSSIBLE PICC LINE FOR CONTINUED ELECTROLYTE SUPPLEMENTATION AND FREQUENT LAB DRAWS. PATIENT DECLINED "I HAVE A HEART ISSUE, I DON'T LIKE THE IDEA OF THAT". PATIENT EDUCATED THAT PICC LINE DOES NOT GO INTO THE HEART, BUT DECLINED. DISCUSSED THE OPTION OF A MIDLINE. EXPLAINED THAT IT IS 8CM AND WOULD NOT GO PAST SHOULDER. PATIENT AGREED. ARM PREPPED WITH A MINUTE CHG SCRUB, AND POWER MIDLINE ATTEMPTED. ULTRASOUND USED. BLOOD RETURN NOTED AND ADVANCEMENT GUIDEWIRE WITHOUT DIFFICULTY. RESISTANCE MET WHILE ADVANCING CATHETER. ADVANCEMENT STOPPED IMMEDIATELY AND REMOVAL OF ENTIRE POWERGLIDE SYSTEM. CATHETER NOTED TO NOT BE INTACT. APPROXIMATELY HALF OF THE CATHETER NOTED IN THE POWERGLIDE SYSTEM. TORNIQUET APPLIED AND DIRECT PRESSURE. PHLEBOTOMIST IN ROOM AND ASSISTED WITH PRESSURE. PROVIDERS NOTIFIED. PATIENT WITHOUT SHORTNESS OF BREATH. NO COMPLAINTS. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
MW5117317 | 2023-05-05 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | SURESTEP INTERMITTENT CATHETER TRAY 16 FR STRAIGHT TIP | CATHETER, UROLOGICAL | C. R. BARD, INC. |  |  |  | TWO BARD INTERMITTENT URINARY CATHETER KITS FROM DIFFERENT LOTS, WHEN OPENED, HAD THE CATHETER STICKING OUT OF STERILE WRAP WHEN THE FIRST FLAP WAS UNFOLDED. REFERENCE REPORT: MW5117316. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03181 | 2023-05-05 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® RUBBER UTILITY CATHETER | UTILITY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 277712 | 277712 |  | IT WAS REPORTED THAT THE FOREIGN BODY CONTAMINATION APPEARS TO BE BROKEN TIP OF ANOTHER URETHRAL CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03178 | 2023-05-05 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER WAS LEAKING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03171 | 2023-05-05 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER | MAGIC 3 INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 51612 |  | IT WAS REPORTED THAT THE PATIENT CALLED TO ADVISE THAT THE PATIENT SO FAR HAD FOUND 6 CATHETERS WITH WATER WITHIN THE SACHET AND WITH NO CATHETER INSIDE IT. THEY STATED THAT THEY PUT ONE OF THESE IN THEIR BAG AND TRIED TO USE IT BUT THERE WAS NO CATHETER WITHIN IT. THE PATIENT HADN'T OPENED THE OTHER 5 FAULTY ONES BUT FELT THAT THERE MIGHT BE NO CATHETERS WITHIN THE PACK. THE PATIENT HADN'T GONE THROUGH ALL THE BOXES SENT BUT THE REPRESENTATIVE ASKED THE PATIENT TO CALL BACK IF THEY FOUND ANYMORE FAULTY ONES. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "LACK OF OPERATOR TRAINING". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03175 | 2023-05-05 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 1758SI14 |  | IT WAS REPORTED THAT THE FOLEY CATHETER GOT BROKEN. THE FOLEY CATHETER WAS USED IN THE OPERATING ROOM. SINCE URINE LEAKED DURING AN OPERATION, CHECKING THE CATHETER, THERE WAS A CRACK ALONG THE CATHETER SHAFT VERTICALLY, AND URINE LEAKED FROM THERE. CONSIDERING THE INVASIVENESS TO THE PATIENT, THE CATHETER WAS NOT REPLACED AT THE SITE AND WAS REMOVED LATER IN THE WARD. | THE REPORTED EVENT WAS CONFIRMED CAUSE UNKNOWN. THE DEVICE HAD NOT MET SPECIFICATIONS. THE PRODUCT CAUSED THE REPORTED FAILURE. THE PRODUCT WAS USED FOR PATIENT TREATMENT OR DIAGNOSIS. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 2-WAY CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION NOTED AN OPEN CRACK MEASURING 0.0990" ON THE SHAFT OF THE CATHETER AND LOCATED 0.3025" FROM THE BIFURCATION. THIS IS OUT OF SPECIFICATION WHICH STATES, VISIBLE BREAKAGE OR CRACKS NOT ALLOWED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE INADEQUATE MATERIAL SELECTION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATES THE FOLLOWING: 1. METHOD FOR USE: DO NOT INFLATE THE BALLOON IN THE URETHRA. [HE URETHRA MAY BE INJURED. DO NOT PULL THE CATHETER HARD. THE BLADDER/URETHRA MAY BE INJURED. 2. APPLICABLE PATIENTS: PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER. WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED. 1. METHOD FOR USE: DO NOT REUSE. DO NOT RE-STERILIZE. THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON. DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON. 2. APPLICABLE PATIENTS: PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03176 | 2023-05-05 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123516A |  | IT WAS REPORTED THAT DURING A PRETEST, WATER WAS UNABLE TO REMOVE EASILY. AFTER SEVERAL TIMES OF ATTEMPT, THE WATER WAS ALLEGEDLY ABLE TO REMOVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. NOTE: THE INITIAL REPORTER ZIP CODE IN TAB E IS (B)(6). | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03177 | 2023-05-05 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L14 |  | IT WAS REPORTED THAT 2 BIOCATH CATHETER BALLOONS FAILED TO INFLATE BOTH FROM SAME BATCH AND SAME LOT NUMBER BELOW. PATIENT WAS CATHETERIZED AND THE BALLOON WOULD NOT INFLATE. A SECOND CATHETER WAS USED (FROM THE SAME BATCH) AND THE SAME THING HAPPENED - BALLOON WOULD NOT INFLATE. A THIRD ATTEMPT WAS MADE WITH ANOTHER BIOCATH WHICH WAS SUCCESSFUL THIS WAS A DIFFERENT BATCH NUMBER. PLANNED PATIENT VISIT FOR SUPRAPUBIC CATHETER CHANGE. INTRODUCED MYSELF TO PATIENT AND EXPLAINED THE REASON FOR MY VISIT, PATIENT CONSENT GAINED FOR CATHETER CHANGE. ASEPTIC TECHNIQUE CARRIED OUT AND CATHETER CHANGED PROMPTLY AS PER SUPRAPUBIC CATHETER GUIDANCE ON TIMINGS. ON NEW SUPRAPUBIC CATHETER INSERTION, WITNESSED URINE DRAINING INTO CATHETER LEG BAG AS IT WAS SUCCESSFULLY IN THE BLADDER, HELD CATHETER IN PLACE WHILST ATTEMPTED TO INFLATE CATHETER BALLOON. UNFORTUNATELY THE CATHETER BALLOON WOULD NOT INFLATE, IT WENT INTO THE PORT AND JUST ACCUMULATED AT THE PORT END, MYSELF AND THE PATIENT COULD SEE THE CATHETER LATEX MATERIAL STRETCHING AS I ATTEMPTED TO PUT THE FLUID IN BUT IT WOULD NOT GO IN AS NORMAL. QUICKLY OBTAINED A SECOND CATHETER AND CARRIED OUT THE SAME PROCESS, AWAITED URINE WHICH DRAINED NICELY INTO THE CATHETER LEG BAG THEN TRIED TO INFLATE THE NEW CATHETER BALLOON BUT THE SAME THING HAPPPEND. NOTICED THESE WERE FROM THE BATCH WITH THE SAME LOT NUMBER - BARD BIOCATH HYDROGEL COATED LATEX CATHETERS 14CH WITH 10ML BALLOONS, REF. D226514 LOT. MYGT1420 EXPIRY, 28/04/2027. I CARRIED OUT A CATHETER CHANGE ONCE MORE AS PATIENT VERY KEEN FOR ME TO TRY A THIRD TIME AS DIDN'T WISH TO GO TO A;E AND WITH THE CONCERNS THAT THE SP SITE MAY CLOSE UP IF NOT ACTING QUICKLY. I ASSURED THE PATIENT THAT THIS CATHETER (SAME BRAND ETC) WAS FROM A DIFFERENT BATCH ACCORDING TO THE DIFFERING LOT NUMBER. SUCCESSFUL SP CATHETERISATION, PATENT AND DRAINED OVER 100ML. ADVISED PATIENT OF DATIX AS APPEARED TO BE A FAULTY BATCH. PATIENT GRATEFUL FOR THE VISIT, LEFT SAFE AND WELL. ADVISED TO CONTACT DN'S IF ANY FURTHER ISSUES OCCUR AND DISPOSED OF THE FAULTY BATCH. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. NOTE: THE INITIAL REPORTER ZIP CODE IS (B)(6). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03168 | 2023-05-05 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® TEMPERATURE-SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 117416 | 117416 |  | IT WAS REPORTED THAT THE CUSTOMER HAD ENCOUNTERED ISSUE WITH THE BALLOONS OF SILICONE INDWELLING CATHETERS LOSING FLUID AND DEFLATING EARLIER THIS WEEK. THE FOLEY CATHETER WAS PLACED ON (B)(6) 2023 IN SURGERY, AND IT WAS DOCUMENTED THAT THE BALLOON WAS INFLATED. HOWEVER, AT THE TIME OF CATHETER REMOVAL, THE BALLOON WAS FOUND TO BE COMPLETELY DEFLATED (LOT# NGGR3688). THIS WAS THE FIRST TIME THEY HAD HEARD OF THIS ISSUE, BUT THE REGISTERED NURSE ON THIS PATIENT CARE UNIT NOTED THIS MULTIPLE TIMES RECENTLY (LOT# UNK). ALTHOUGH THEY DID NOT HAVE ANY SPECIFICS OR THE CATHETER. THEY DID RE-INFLATE THE BALLOON YESTERDAY AND PLACED IT IN A BAG TO SAVE IT. TODAY, THERE WAS SOME FLUID IN THE BAG, WHICH WOULD INDICATE THE BALLOON WAS SLOWLY LEAKING. PER ADDITIONAL INFORMATION RECEIVED ON 15APR2023, IT WAS REPORTED THAT THE CUSTOMER WERE SHOWING THAT THEY HAD A TOTAL OF FOUR DIFFERENT FOLEYS IN THEIR OFFICE WAITING THAT THEY CAN RETURN FOR INVESTIGATION. THE 3 CATHETERS THAT SEEM TO HAVE AN ISSUE WITH DEFLATING BALLOONS, ONE FOLEY HAD SLIPPED OUT OF THE PATIENT, AND NURSE STATED THAT THEY TESTED THE BALLOON AND FELT A LEAK (LOT# NGFX5338) OTHER FOLEY (LOT# NGGX0360) AND AN OTHER FOLEY (LOT# NGG20749). PER FOLLOW UP VIA PHONE 17APR2023, CUSTOMER HAD RETURNED FOUR EACH FOR TESTING. NO MEDICAL INTERVENTION WAS REPORTED FOR ANY OF THE PATIENTS INVOLVED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03169 | 2023-05-05 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® TIGERTAIL¿ FLEXIBLE TIP URETERAL CATHETERS | URETERAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 139005 | 139005 |  | IT WAS REPORTED THAT THE GUIDEWIRE WAS TOO SNUG IN THE LUMEN OF A TIGERTAIL CATHETER. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION AND FURTHER INVESTIGATION WAS NOT CONCLUSIVE. THOUGH A SPECIFIC CAUSE COULD NOT BE DETERMINED, A POTENTIAL ROOT CAUSE FOR THIS EVENT COULD BE, "INCORRECT GEOMETRY", OR "OPERATOR DID NOT ATTACH ADAPTER AS PER MANUFACTURING PROCEDURE." THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "AFTER PLACEMENT OF THE CYSTOSCOPE INTO THE BLADDER, THE URETERAL ORIFICE IS LOCATED AND EXAMINED AND A CHOICE OF URETERAL CATHETER IS MADE. UNDER DIRECT VISION, THE GUIDEWIRE IS POSITIONED WITHIN THE URETERAL ORIFICE." "THE URETERAL CATHETER IS INSERTED OVER THE GUIDEWIRE AND ADVANCED INTO THE URETER. THE GUIDEWIRE IS REMOVED AND THE CATHETER ADAPTER IS SECURED TO THE CATHETER. GENTLE ASPIRATION TO EVACUATE AIR FROM THE CATHETER CAN BE PERFORMED AT THIS TIME. CONTRAST MEDIA IS THEN INJECTED THROUGH THE URETERAL CATHETER INTO THE URETER. UNDER FLUOROSCOPY, THE URETER IS VISUALIZED AND REQUIRED INTERVENTION DETERMINED." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
MW5117316 | 2023-05-05 | 2023-05-02 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | SURESTEP INTERMITTENT CATHETER TRAY 16 FR STRAIGHT TIP | CATHETER, UROLOGICAL | C. R. BARD, INC. |  |  |  | TWO BARD INTERMITTENT URINARY CATHETER KITS FROM DIFFERENT LOTS, WHEN OPENED, HAD THE CATHETER STICKING OUT OF STERILE WRAP WHEN THE FIRST FLAP WAS UNFOLDED. REFERENCE REPORT: MW5117317. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03174 | 2023-05-05 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123618A |  | IT WAS REPORTED THAT THERE WAS FLUID LEAKAGE FROM THE FOLEY CATHETER. | THE REPORTED EVENT WAS CONFIRMED ¿ CAUSE UNKNOWN. THE REPORTED FAILURE WAS ABLE TO BE REPRODUCED. THE PRODUCT HAD CAUSED THE REPORTED FAILURE. SAMPLE WAS EVALUATED AND OBSERVED WATER LEAKED FROM THE BALLOON OF THE CATHETER. DISSECTED THE CATHETER AND OBSERVED TEAR ON SAC BODY OF THE CATHETER BALLOON. HOWEVER, THE EXACT CAUSE OF HOW AND WHEN THE PROBLEM OCCURRED COULD NOT BE DETERMINED. POTENTIAL ROOT CAUSE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/MECHANICAL FAILURE/OPERATOR ERROR/THIN RUBBERIZE LAYER. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "STRUCTURE & COMPOSITION: THE PRODUCT INCLUDE THREE SERIES: BARDEX, BARDIC AND BARDIA. THE COMPOSITION OF EACH SERIES IS AS FOLLOWS: BARDEX, BARDIC AND BARDIA BI-LUMEN SERIES ARE COMPOSED OF CATHETER BODY, INFLATION CONE INTERFACE, BALLOON AND VALVE. BARDIA TRI-LUMEN IS COMPOSED TO TUBE BODY, INFLATION CONE INTERFACE, DEFLATION CONE INTERFACE, FLUSH CONE INTERFACE, BALLOON AND VALVE. ¿THE PRODUCT IS WITH SILICONE COATING. THE MATERIAL OF CATHETER BODY, INFLATION CONICAL INTERFACE, DEFLATION CONICAL INTERFACE, FLUSHING CONICAL INTERFACE, BALLOON IS NATURAL LATEX. THE MATERIAL OF VALVE IS POLYPROPYLENE. STERILIZED BY ¿ RADIATION. DISPOSABLE. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. VALVE TYPE: USE LUER SLIP SYRINGE. DO NOT USE NEEDLE. TO DEFLATE CATHETER BALLOON: GENTLE INSERT A SYRINGE IN THE CATHETER VALVE. NEVER USE EXCESSIVE FORCE TO MAKE THE SYRINGE ¿STICK¿ IN THE VALVE. IF YOU NOTICE SLOW OR NO DEFLATION, RE-SEAT THE SYRINGE GENTLY. USE ONLY GENTLE ASPIRATION TO ENCOURAGE DEFLATION IF NEEDED. VIGOROUS ASPIRATION MAY COLLAPSE THE INFLATION LUMEN. IF PERMITTED BY HOSPITAL PRACTICAL, THE VALVE MAY BE CUT OFF. IF THIS FAILS, PLEASE CONTACT AN ADEQUATELY TRAINED PROFESSIONAL FOR ASSISTANCE, AS DIRECTED BY HOSPITAL PRACTICAL. SHOULD BALLOON RUPTURE OCCUR, CARE SHOULD BE TAKEN TO ASSURE THAT ALL BALLOON FRAGMENTS HAVE BEEN REMOVED FROM THE PATIENT. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX. DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01729 | 2023-05-05 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO2 CATHETER WITH SHERLOCK 3CG TIP POSITIONING SYSTEM, STYLET | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 1194108D2 |  | IT WAS REPORTED VIA MEDWATCH, PICC WAS MALFUNCTIONING, ONCE REMOVED, THERE WAS BREAK BETWEEN THE 5-6CM MARK ON THE CATHETER. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01720 | 2023-05-07 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK PICC PLUS SLIDING POST, CRESCENT PAD | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | VPPCSPCE |  | IT WAS REPORTED THAT THE CLAW OF THE RETAINER WAS FOUND TO BE DAMAGED. THERE WAS NO REPORTED PATIENT INJURY. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED STATLOCK RETAINER WAS CONFIRMED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE STATLOCK SECUREMENT DEVICE. THE INVESTIGATION FINDINGS ARE CONSISTENT WITH DAMAGE CAUSED BY PUSHING THE SECUREMENT TAB AT AN ANGLE WHILE ATTEMPTING TO CLOSE THE DEVICE. THIS ISSUE CAN BE AVOIDED BY CENTERING THE THUMB ABOUT THE MIDDLE OF THE TAB AND PRESSING DIRECTLY DOWNWARD. THE RETURNED PRODUCT SAMPLE WAS EVALUATED AND A SECUREMENT TAB WAS OBSERVED TO BE DAMAGED. THE FOLLOWING OBSERVATIONS WERE MADE WHICH WERE CONSISTENT WITH THIS FAILURE TYPE: AN ATTEMPT TO TEST THE FUNCTIONALITY OF THE LOCK WAS UNSUCCESSFUL AND THE LATCH(S) WAS OBSERVED TO BE MISALIGNED WITH THE CATCH BASE THE TAB LATCH(S) WAS OBSERVED TO BE BENT THE TAB CATCH(S) EXHIBITED PLASTIC DEFORMATION AT THE POINT OF CONTACT WITH THE LATCH THIS FAILURE TYPE WAS RECREATED IN THE LABORATORY USING NON-COMPLAINANT SAMPLES BY CLOSING THE TAB AT AN ANGLE, OR BY PRESSING OUTSIDE THE CENTER OF THE TAB, AND THE FEATURES OBSERVED ON THE LABORATORY SAMPLES WERE SIMILAR TO THOSE ON THE RETURNED COMPLAINANT SAMPLE. THE NATURE OF THE DAMAGE OBSERVED ON THE LATCH, AND THEIR POINTS OF CONTACT ON THE CATCH BASE AS WELL AS THE CHARACTERISTICS OF THE HINGE DAMAGE, ARE EVIDENCE THAT THE DAMAGE WAS CAUSED BY MISALIGNMENT OF THE DOORS WHEN CLOSING THE DEVICE. PUSHING THE TAB AT AN ANGLE CAN CAUSE THE LATCH TO MISS THE CATCH BASE AND CAN PERMANENTLY BEND THE LATCH. AN EXAMINATION OF THE SAMPLE REVEALED NO POTENTIAL DAMAGE/DEFECT RELATED TO MANUFACTURE OF THE PRODUCT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H.11. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01743 | 2023-05-08 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT PRIOR TO PORT PLACEMENT PROCEDURE, PORT ALLEGEDLY DIDN'T FLUSH. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRY DATE: 09/2024) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01741 | 2023-05-08 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP MRI 6CF INT W SP, ATT, SL | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 8806060 |  | IT WAS REPORTED THAT PRIOR TO A PORT PLACEMENT PROCEDURE, THE DEVICE WAS ALLEGEDLY DAMAGED. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURES AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 06/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01740 | 2023-05-08 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT PRIOR TO PORT PLACEMENT PROCEDURE, PORT ALLEGEDLY DIDN'T FLUSH. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 09/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01736 | 2023-05-08 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT ONE MONTH AND FOURTEEN DAYS POST PORT PLACEMENT PROCEDURE, WHEN THE PORT SYSTEM WAS ATTEMPTED TO BE USED, LEAKAGE OF THE MEDICINE WAS FOUND. AN X-RAY EXAMINATION WAS PERFORMED AND REVEALED THE WIRE. THE PORT SYSTEM AND THE WIRE WERE REMOVED. IT WAS FOUND THAT THE STYLET COME OUT FROM THE SLIT OF THE CATHETER. THE PHYSICIAN THOUGHT THAT THE STYLET INTO THE CATHETER WAS REMOVED BUT MAY HAVE REMAINED PARTIALLY IN PLACE. IT WAS POSSIBLE THAT THE STYLET HIT THE CATHETER AFTER THE PORT SYSTEM WAS IMPLANTED, THE STYLET COME OUT OF THE CATHETER WALL AND LEAKAGE OCCURRED. THE PHYSICIAN THOUGHT THAT THE STYLET WAS PARTIALLY BROKEN FROM THE BEGINNING AND WAS REMOVED DURING THE PLACEMENT PROCEDURE, BUT THE BROKEN TIP OF THE STYLET REMAINED IN THE CATHETER, WHICH MAY HAVE CAUSED THIS EVENT. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01735 | 2023-05-08 | 2023-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393190 |  | IT WAS REPORTED THAT THREE MONTHS POST DIALYSIS CATHETER PLACEMENT, THE CATHETER WAS ALLEGEDLY BROKEN AT EXTENSION PART. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY HAD LEAKAGE AT EXTENSION PART. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. FOUR PHOTOS WERE PROVIDED FOR REVIEW. THE PHOTO SHOWS THE DIALYSIS CATHETER BEING IMPLANTED IN THE PATIENT IN WHICH THE BLUE LUER EXTENSION LEG WAS NOTED TO BE FRACTURED. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE ISSUE. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FLUID LEAK ISSUE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 12/2023). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
2020394-2023-00317 | 2023-05-08 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CROSSER | RECANALIZATION CATHETER | BARD PERIPHERAL VASCULAR, INC. |  | CRE14S |  | IT WAS REPORTED THAT DURING A RECANALIZATION PROCEDURE, THE TIP OF THE CATHETER ALLEGEDLY SEPARATED. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE CROSSER CTO RECANALIZATION CATHETERS PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED MATERIAL SEPARATION ISSUE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRATION DATE: 01/2025), G3, H6 (METHOD) . H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03255 | 2023-05-08 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED ON (B)(6) 2023. ON (B)(6) AT 07:15 IT WAS FOUND THAT THE FOLEY CATHETER BALLOON WAS BURST AND CATHETER FELL OUT. NO HARM WAS CAUSED TO PATIENT AND IT WAS REPLACED WITH A NEW DEVICE. AS PER DESCRIPTION OF EVENT CAUSE ANALYSIS, STATED THAT IT MIGHT BE CAUSED BY PRODUCT MATERIALS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01745 | 2023-05-08 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT DURING PASSAGE OF THE CATHETER, THE CATHETER WAS TRIMMED AND ATTEMPTED TO ATTACH TO THE CONNECTOR. BUT THE CONNECTOR COULD NOT BE ENGAGED FIRMLY WITH GAP LEFT IN ONE SIDE. THE CONNECTOR WAS THEN REPLACED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03254 | 2023-05-08 | 2023-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 63618 | 63618 |  | IT WAS REPORTED THAT THE CUSTOMER STATED A PATIENT WAS UNABLE TO USE THE INTERMITTENT CATHETER AS THERE WAS NO ENOUGH LUBRICANT. IT WAS CONFIRMED THAT ALL FOUR CATHETERS DID NOT HAVE ENOUGH LUBRICANT SO, COULD NOT BE USED. THE PATIENT HAS USED THIS PRODUCT FOR A LONG TIME SO, WAS VERY FAMILIAR WITH IT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS AVAILABLE FOR EVALUATION AND FURTHER INVESTIGATION WAS NOT CONCLUSIVE. A POTENTIAL ROOT CAUSE COULD BE MECHANICAL FAILURE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "FOR UROLOGICAL USE ONLY. INTENDED FOR USE BY PATIENTS FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION AND MEASUREMENT. THE DEVICES ARE PASSED TO THE URINARY BLADDER VIA THE URETHRA. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/PRECAUTIONS AND ADVERSE REACTIONS. THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. 1. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. 2. PRIOR TO OPENING THE SEALED CATHETER POUCH, APPLY PRESSURE TO THE FOIL PACKET TO RELEASE THE WATER. ENSURE ALL WATER IS RELEASED FROM THE FOIL PACKET. 3. WET THE CATHETER BY HOLDING THE PACKAGE WITH THE PRINTED SIDE UP AND TIP THE PACKAGE END-TO-END THREE TO SIX TIMES TO WET THE CATHETER. THIS MOVEMENT IS REQUIRED SO THAT THE WATER TRANSFERS BACK AND FORTH OVER THE CATHETER TO FULLY WET THE HYDROPHILIC COATING. 4. PEEL OPEN THE PACK AT THE FUNNEL/HANDLE END JUST ENOUGH TO EXPOSE 5 CM (2¿) OF THE CATHETER, OR FOR PRODUCTS WITH AN INSERTION SLEEVE, EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. 5. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. 6. WASH YOUR HANDS AGAIN. 7. USING THE CATHETER WITH INSERTION SLEEVE (GRIPPER): A. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. B. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. C. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. WITHOUT INSERTION SLEEVE (GRIPPER): A. HOLD THE HANDLE/FUNNEL END AND REMOVE THE CATHETER FROM THE PACKAGING. B. GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA AND ADVANCE THE CATHETER UNTIL URINE STARTS TO FLOW. 8. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. 9. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAPAND WATER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01742 | 2023-05-08 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT SOMETIME POST PORT PLACEMENT, THE DEVICE ALLEGEDLY HAD A FLUID LEAK. IT WAS FURTHER REPORTED THAT PATIENT TOOK HIS DRUG OFF AND HELD PRESSURE WITH CLEAN GAUZE FOR TWENTY TO THIRTY MINUTES TO GET IT STOPPED. THE PORT WAS STILL IN PATIENT'S BODY. THERE WAS NO REPORTED PATIENT INJURY. | H10: ADDITIONAL INFORMATION WAS RECEIVED, AND THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE REPORTABLE AS MALFUNCTION. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03249 | 2023-05-08 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 175816 |  | IT WAS REPORTED THAT THE DISTRICT NURSE COULD NOT INSERT THIS FOLEY CATHETER AND HAD TO USE AN ALTERNATIVE. THE PATIENT ADVISED THAT ON TWO OCCASIONS, DIFFERENT DISTRICT NURSES HAD ATTEMPTED TO INSERT THESE BUT FOUND IT WAS IMPOSSIBLE AND INSTEAD HAD TO USE TWO CATHETERS FROM THEIR STOCK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03246 | 2023-05-08 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167L20 |  | IT WAS REPORTED, THAT 3DAYS AFTER PLACEMENT. THE CATHETER WAS FOUND TO HAVE FALLEN OUT OF THE PATIENT. WHEN THE CATHETER WAS INSERTED TO THE PATIENT AND PLACED TO INFLATE THE BALLOON, THERE WAS ALLEGEDLY NO ISSUES FELT. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, THERE WERE NO OBVIOUS VISIBLE DEFECTS TO THE RETURNED SAMPLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03245 | 2023-05-08 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165SI16 |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON GOT BURST. THE MEDICAL INSTITUTIONS AND THE PATIENT WERE HURRYING TO FIND OUT THE CAUSE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03239 | 2023-05-08 | 2023-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. ALL-SILICONE TEMP SENSING FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | A319514AM | 119314 |  | IT WAS REPORTED THAT PATIENT HAD A TEMPERATURE SENSING FOLEY CATHETER IN AND IT WAS NOTED THAT THERE WAS NO URINE OUTPUT FOR THE CURRENT HOUR COMPARED TO THE PREVIOUS HOUR. WHEN LOOKING AT THE 14FR BALLOON END, THE BALLOON WAS RUPTURED AND NOT INTACT, THIS PARTICULAR FOLEY WAS ONLY PLACED SEVERAL HOURS PRIOR IN THE SAME DAY. NEW 14FR BALLOON TEMPERATURE PROBE FOLEY INSERTED WITHOUT DIFFICULTY AND PATIENT DID NOT EXPERIENCE ANY FURTHER ISSUES. PER FOLLOW UP VIA EMAIL ON 04MAY2023, IT WAS REPORTED THAT THE MODEL# A319514AM AND LOT# NGGYR024, AS OF NOW, THIS WAS THE ONLY REPORT THEY HAVE RECEIVED THAT HAS HAD THIS ISSUE. STATED THAT THE PATIENT HAD NO INJURY AS A RESULT AND THE CATHETER WAS DISPOSED OF WHEN IT WAS REMOVED FROM THE PATIENT. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "WALL THICKNESS TOO THIN". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PROPER TECHNIQUES FOR URINARY CATHETER INSERTION ¿ PERFORM HAND HYGIENE IMMEDIATELY BEFORE AND AFTER INSERTION ¿ INSERT URINARY CATHETERS USING ASEPTIC TECHNIQUE AND STERILE EQUIPMENT ¿ USE THE SMALLEST FOLEY CATHETER POSSIBLE, CONSISTENT WITH GOOD DRAINAGE ¿ DOCUMENT THE INDICATIONS FOR CATHETER INSERTION, DATE AND TIME OF CATHETER INSERTION, INDIVIDUAL WHO INSERTED CATHETER, AND DATE AND TIME OF CATHETER REMOVAL IN PATIENT RECORD PROPER TECHNIQUES FOR URINARY CATHETER MAINTENANCE ¿ SECURE THE FOLEY CATHETER, USE THE STATLOCK® FOLEY STABILIZATION DEVICE IF PROVIDED ¿ MAINTAIN A CLOSED DRAINAGE SYSTEM BY UTILIZING PRE-CONNECTED, SEALED CATHETER-TUBING JUNCTIONS ¿ MAINTAIN UNOBSTRUCTED URINE FLOW AND KEEP THE CATHETER AND COLLECTION TUBE FREE FROM KINKING ¿ KEEP THE COLLECTION BAG BELOW THE LEVEL OF THE BLADDER OR HIPS AT ALL TIMES ¿ EMPTY THE COLLECTION BAG REGULARLY (E.G., PRIOR TO TRANSPORT) USING A SEPARATE, CLEAN COLLECTION CONTAINER FOR EACH PATIENT ¿ ROUTINE HYGIENE (E.G., CLEANSING OF THE MEATAL SURFACE DURING DAILY BATHING OR SHOWERING) IS APPROPRIATE ¿ LEAVE FOLEY CATHETER IN PLACE ONLY AS LONG AS NEEDED STERILE UNLESS PACKAGE IS OPENED OR DAMAGED, EXCEPT FOR ANY INDIVIDUALLY PACKAGED COMPONENTS WITHIN THE TRAY WHICH ARE NOT LABELED AS STERILE. THESE COMPONENTS ARE NOT TERMINALLY STERILIZED. CAUTION: FEDERAL (U.S.A.) LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN. SINGLE USE ONLY. DO NOT RESTERILIZE. FOR UROLOGICAL USE ONLY. WARNING: ON CATHETER, DO NOT USE OINTMENTS OR LUBRICANTS HAVING A PETROLATUM BASE. THEY WILL DAMAGE THE CATHETER AND MAY CAUSE BALLOON TO BURST. WARNING: AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. HANDLE AND DISPOSE OF IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICES AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. VISUALLY INSPECT THE PRODUCT FOR ANY IMPERFECTIONS OR SURFACE DETERIORATION PRIOR TO USE. IF PACKAGE IS OPENED OR IF ANY IMPERFECTION OR SURFACE DETERIORATION IS OBSERVED, DO NOT USE. PLEASE CONSULT PRODUCT LABEL AND INSERT FOR ANY INDICATIONS, CONTRAINDICATIONS, HAZARDS, WARNINGS, CAUTIONS AND DIRECTIONS FOR USE." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03237 | 2023-05-08 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 998606 | 998606 |  | IT WAS REPORTED THAT WHEN PREPARING THE URETER STENOSIS DILATATION WITH BALLOON, THE CUSTOMER INFLATED THE BALLOON USING A PRESSURE PUMP AND FOUND THAT THE BALLOON LEAKED WATER. THIS MADE THE PRODUCT UNUSABLE AND THE CUSTOMER IMMEDIATELY DISCONTINUED THE USE OF THE PRODUCT, CAUSING NO IMPACT ON THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | OTHER |  | 2023-05
1018233-2023-03241 | 2023-05-08 | 2023-04-18 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT GOT URINARY TRACT INFECTIONS 4 TIMES NOT SURE IF IT WAS FROM THE PUREWICK FEMALE EXTERNAL CATHETER. IT WAS UNKNOWN IF THE DEVICE CONTRIBUTED TO THE URINARY TRACT INFECTION AT THIS TIME AND MEDICAL INTERVENTION WAS UNKNOWN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INADEQUATE MATERIAL SELECTION - MATERIALS OF CONSTRUCTION ARE NOT BIOCOMPATIBLE ". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE COMPLETED DUE TO NO LOT NUMBER PROVIDED. THE DHR REVIEW COULD NOT BE COMPLETED DUE TO NO LOT NUMBER PROVIDED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "PERI-CARE AND PLACEMENT: 3. PERFORM PERINEAL CARE AND ASSESS SKIN INTEGRITY (DOCUMENT PER HOSPITAL PROTOCOL). SEPARATE LEGS, GLUTEUS MUSCLES, AND LABIA. PALPATE PUBIC BONE AS ANATOMICAL MARKER. 4. WITH SOFT GAUZE SIDE FACING PATIENT, ALIGN DISTAL END OF THE PUREWICKTM FEMALE EXTERNAL CATHETER AT GLUTEAL CLEFT. GENTLY TUCK SOFT GAUZE SIDE BETWEEN SEPARATED GLUTEUS AND LABIA. ENSURE THAT THE TOP OF THE GAUZE IS ALIGNED WITH THE PUBIC BONE. SLOWLY PLACE LEGS BACK TOGETHER ONCE THE PUREWICKTM FEMALE EXTERNAL CATHETER IS POSITIONED. NOTE: PATIENT CAN BE POSITIONED ON BACK, SIDE LYING, FROG LEGGED, OR LYING ON BACK WITH KNEES BENT AND THIGHS APART (LITHOTOMY POSITION) PRIOR TO DEVICE PLACEMENT. WARNINGS: TO AVOID POTENTIAL SKIN INJURY, NEVER PUSH OR PULL THE PUREWICKTM FEMALE EXTERNAL CATHETER AGAINST THE SKIN DURING PLACEMENT OR REMOVAL. NEVER INSERT THE PUREWICKTM FEMALE EXTERNAL CATHETER INTO VAGINA, ANAL CANAL, OR OTHER BODY CAVITIES. DISCONTINUE USE IF AN ALLERGIC REACTION OCCURS. PRECAUTIONS: NOT RECOMMENDED FOR PATIENTS WHO ARE: AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER. HAVING FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. EXPERIENCING SKIN IRRITATION OR BREAKDOWN AT THE SITE. ALWAYS ASSESS SKIN FOR COMPROMISE AND PERFORM PERINEAL CARE PRIOR TO PLACEMENT OF A NEW PUREWICKTM FEMALE EXTERNAL CATHETER. RECOMMENDATIONS: PERFORM EACH STEP WITH CLEAN TECHNIQUE. IN THE HOME SETTING, WASH HANDS THOROUGHLY BEFORE DEVICE PLACEMENT. ENSURE THE PUREWICKTM FEMALE EXTERNAL CATHETER REMAINS IN THE CORRECT POSITION AFTER TURNING THE PATIENT. REMOVE THE PUREWICKTM FEMALE EXTERNAL CATHETER PRIOR TO AMBULATION. PROPERLY PLACING THE PUREWICKTM FEMALE EXTERNAL CATHETER SNUGLY BETWEEN THE LABIA AND GLUTEUS HOLDS THE PUREWICKTM FEMALE EXTERNAL CATHETER IN PLACE FOR MOST PATIENTS. MESH UNDERWEAR MAY BE USEFUL FOR SECURING THE PUREWICKTM FEMALE EXTERNAL CATHETER FOR SOME PATIENTS. ASSESS DEVICE PLACEMENT AND PATIENT¿S SKIN AT LEAST EVERY 2 HOURS. REPLACE THE PUREWICKTM FEMALE EXTERNAL CATHETER EVERY 8-12 HOURS OR WHEN SOILED WITH FECES OR BLOOD. CHANGE SUCTION TUBING PER HOSPITAL PROTOCOL OR AT LEAST EVERY THIRTY (30) DAYS. THE BAW IS ATTACHED TO THE INVESTIGATION OVERVIEW ELEMENT." THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03226 | 2023-05-08 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER DID NOT RETURN FROM THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. NOTE: THE INITIAL REPORTER ZIP CODE IS (B)(6). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03230 | 2023-05-08 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0165L18 |  | IT WAS REPORTED THAT THE CATHETER WAS WITHDRAWN THE DAY AFTER IT WAS CHANGED AND THE BALLOON WAS RUPTURED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03219 | 2023-05-08 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT WHITE DEPRESSED DAMAGES WITH THE BALLOON PART FOR A TOTAL OF 5350 FOLEY CATHETERS FROM TWO BATCHES (BATCH# MYGN5842 AND BATCH# MYGR5328) WERE FOUND DURING USE AND INSPECTION. THE QUALITY DEPARTMENT OF COMPANY HAD DETERMINED THAT THERE WERE RISKS OF BALLOON BURSTING AND THAT THESE PRODUCTS COULD NOT BE USED IN THE PRODUCTION. THE ISSUE WAS DETECTED DURING REUSE OF A SINGLE USE MEDICAL DEVICE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, STATED THAT CUSTOMER WAS A DISTRIBUTION CENTER. STATED THAT 5 PHYSICAL SAMPLES OF EACH BATCH WOULD BE PROVIDED. THE USER PURCHASED 5350 UNITS AND INSPECTED SOME OF THEM HAD THE ISSUE, THUS THE CUSTOMER WAS WORRIED ABOUT THE QUALITY OF THE WHOLE BATCHES AND REQUEST TO CHANGE AND RETURN 5350 UNITS WITH ANOTHER BATCH. PER SAMPLE RECEIVED BY THE BD REPRESENTATIVE ON (B)(6) 2023, STATED THAT THE MANUFACTURING LOT# OF SAMPLES RECEIVED FOR THE BATCH MYGN5842 WAS 2HJ031 AND FOR THE BATCH MYGR5328 WERE 2JD072, 2HH041, 2HU135 AND 2HL031. BATCH# MYGN5842 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL# (B)(6) AND BATCH# MYGR5328 AND SERIAL# (B)(6). PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON (B)(6) 2023, STATED THAT THE CUSTOMER WAS AN OLD CUSTOMER, THEY HAVE USED BD FOLEY FOR ALMOST 18 YEARS SINCE 2006. THIS WAS THE FIRST INCIDENT EXPERIENCED BY THEM. THE MANUFACTURING LOT# OF SAMPLES FOR THE BATCH MYGN5842 WAS 2FD047 AND FOR THE BATCH MYGR5328 WERE 2HU136, 2FR083, 2JC013 AND 2HU134. BATCH# MYGN5842 AND SERIAL# (B)(6) , BATCH# MYGR5328 AND SERIAL# (B)(6), BATCH# MYGR5328 AND SERIAL#(B)(6) , BATCH# MYGR5328 AND SERIAL# (B)(6), AND BATCH# MYGR5328 AND SERIAL# (B)(6). STATED THAT THIS CUSTOMER PER CHASE FOLEY FOR REPROCESSING. WHEN THEY RECEIVED FOLEY, THEY WOULD JUST CHECK PRODUCT MODEL, SIZE, QUANTITY AND EXPIRATION DATE. WHILE DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. THERE WAS NO SAMPLING INSPECTION PERFORMED, THEY TESTED THE BALLOON OF ALL CATHETERS. STATED THAT DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL FROM IBC ON (B)(6) 2023, STATED THAT THE SAMPLES REPORTED BY CUSTOMER WERE THOSE PICKED UP BY CUSTOMER FOR FUNCTION TESTING DURING THEIR REPRODUCTION PROCESS AND THESE PHYSICAL SAMPLES COULD NOT BE RETURNED FOR INVESTIGATION, BUT THE PHOTOS ATTACHED TO THEM. WHILE FOR THE SAMPLES RECEIVED, THESE WERE REPRESENTATIVE SAMPLES FROM THE SAME REPORTED BATCHES RETURNED FOR INVESTIGATION, THEY WERE NOT THE REAL AFFECTED SAMPLES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03227 | 2023-05-08 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO¿ INTERMITTENT URINARY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 51814 | 51814 |  | IT WAS REPORTED THAT PATIENT COULD NOT INSERT THESE MAGIC 3 GO FEMALE INTERMITTENT CATHETERS. PATIENT STATED THAT THEY WERE FLEXIBLE AND SLOPPY. PATIENT THOUGHT THAT THEY WOULD BE TOO SHORT. THE CATHETERS WERE SHIPPED ON 01JAN2023. | UPON FURTHER REVIEW, BD HAS DETERMINED THAT THIS MDR WAS REPORTED IN ERROR AS IT WAS FOUND THAT THERE WAS NO ALLEGATION AGAINST PRODUCT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03244 | 2023-05-08 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN MALE EXTERNAL CATHETER | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT PATIENT STATED IT WAS DIFFICULT TO REMOVE THE MALE EXTERNAL CATHETER THAT THEY HAD BEEN USING OVER 11 YEARS AND WAS REQUESTING INFORMATION ON THE BEST ADHESIVE REMOVER. REPRESENTATIVE ADVISED THERE WAS NO FORMAL RECOMMENDATION FOR AN ADHESIVE REMOVERAS THERE WERE MANY ON THE MARKET IN WHICH THEY CAN CHOOSE FROM. | THE REPORTED ISSUE WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO ¿OPERATOR ERROR". THE LOT NUMBER WAS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE PRODUCT CATALOG NUMBER AND THE LOT NUMBER FOR THIS DEVICE ARE UNKNOWN. THEREFORE, BD IS UNABLE TO DETERMINE THE ASSOCIATED LABELING TO REVIEW. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03223 | 2023-05-08 | 2023-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3 GO MALE COUDE CATHETER WITH SUREGRIP | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50816G | 50816G |  | IT WAS REPORTED THAT THE PATIENT HAD 60 TO 90 MAGIC 3 GO MALE COUDE CATHETERS THAT WOULD NOT DRAIN. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE "INSUFFICIENT TIP STRENGTH DUE TO WALL THICKNESS AND DUROMETER". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INDICATIONS FOR USE: THE MAGIC 3 GO® INTERMITTENT URINARY CATHETER IS INTENDED FOR UROLOGICAL USE ONLY. IT IS INTENDED FOR USE BY ADULT AND PEDIATRIC PATIENTS OF ALL AGES FOR BLADDER MANAGEMENT INCLUDING URINE DRAINAGE, COLLECTION, AND MEASUREMENT. THE DEVICE IS PASSED TO THE URINARY BLADDER VIA THE URETHRA. CONTRAINDICATIONS: NONE KNOWN. WARNINGS THIS IS A SINGLE USE DEVICE. DO NOT RESTERILIZE ANY PORTION OF THIS DEVICE. REUSE AND/OR REPACKAGING MAY CREATE A RISK OF PATIENT OR USER INFECTION, COMPROMISE THE STRUCTURAL INTEGRITY AND/OR ESSENTIAL MATERIAL AND DESIGN CHARACTERISTICS OF THE DEVICE, WHICH MAY LEAD TO DEVICE FAILURE, AND/OR LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: INSPECT THE CATHETER BEFORE USE. DO NOT USE THE PRODUCT IF THE DEVICE OR PACKAGING IS DAMAGED. SELF-CATHETERIZATION SHOULD FOLLOW THE PLAN OF CARE AND ADVICE GIVEN BY YOUR HEALTHCARE PRACTITIONER AND BE CARRIED OUT ONLY IN ACCORDANCE WITH THE INSTRUCTIONS FOR USE PROVIDED. PLEASE CONSULT YOUR HEALTHCARE PRACTITIONER IF ANY CONDITIONS OCCUR WHICH CREATE CONCERN AND/OR DIFFICULTY DURING CATHETERIZATION. ADVERSE REACTIONS: URINARY TRACT INFECTION. BLEEDING FROM THE URETHRA. IRRITATION OF THE URETHRA. REVIEW THE SELF-CATHETERIZATION PROCEDURE WITH A HEALTHCARE PROFESSIONAL. ALWAYS WASH HANDS PRIOR TO USE. OPEN THE CATHETER PACKAGE BY PEELING THE TAB UPWARDS WITH THE AID OF THE FINGER HOLE. IF DESIRED, HANG THE PACKAGE WITH THE ADHESIVE SURFACE ON THE INNER SIDE OF THE TAB TO A NEARBY DRY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. POSITIONED COMFORTABLY WITH THIGHS SPREAD APART, CLEAN THE OPENING TO THE URETHRA ¿ AND THE SURROUNDING AREA. WIPE FROM FRONT TO BACK TO AVOID FECAL CONTAMINATION. WASH YOUR HANDS AGAIN. REMOVE THE CATHETER WITH THE AID OF THE INSERTION HANDLE AND ADVANCE THE CATHETER TIP INTO THE URETHRA. SLOWLY AND GENTLY INSERT THE CATHETER INTO THE URETHRA UNTIL URINE BEGINS TO FLOW APPROXIMATELY 1-1.5" OR 2.5-3.8 CM). WHEN URINE STOPS FLOWING, BEGIN TO WITHDRAW THE CATHETER. IT IS RECOMMENDED TO SLOWLY ROTATE THE CATHETER DURING WITHDRAWAL, STOPPING EACH TIME URINE BEGINS TO FLOW. CHECK THE COLOR, ODOR AND CLARITY OF THE URINE. ANY CHANGES MAY NEED TO BE REPORTED TO A HEALTH CARE PROFESSIONAL. DISPOSAL PLACE THE CATHETER BACK INTO THE PACKAGE AND DISPOSE IN ACCORDANCE WITH ACCEPTED MEDICAL PRACTICE AND APPLICABLE LOCAL, STATE AND FEDERAL LAWS AND REGULATIONS. DO NOT FLUSH DOWN THE TOILET. WASH HANDS." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03202 | 2023-05-08 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | SILICONE FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT STERILE WATER WAS DIFFICULT TO PULL OUT DURING PRE-TEST. | THE REPORTED EVENT WAS UNCONFIRMED BECAUSE THE REPORTED FAILURE COULD NOT BE REPRODUCED. THE REPORTED FAILURE IS WITHIN SPECIFICATION AS THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO ROOT CAUSE COULD BE FOUND BECAUSE THE REPORTED EVENT WAS UNCONFIRMED. VISUAL EVALUATION OF THE RETURNED SAMPLE NOTED ONE OPENED (WITHOUT ORIGINAL PACKAGING), A 3-WAY TEMPERATURE SENSING CATHETER WITH CUT PORTION OF INLET TUBING. VISUAL INSPECTION OF THE SAMPLE NOTED NO OBVIOUS VISIBLE DEFECTS. THE CATHETER BALLOON WAS INFLATED WITH 10 ML METHYLENE BLUE SOLUTION (3 DROPS 1% AQ METHYLENE BLUE PER 100ML DISTILLED WATER) AND THE CATHETER RESTED WITH NO LEAKS NOTED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. AS THE REPORTED EVENT IS UNCONFIRMED A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03209 | 2023-05-08 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWF030 | PWF030 |  | IT WAS REPORTED THAT THE PUREWICK FEMALE EXTERNAL CATHETER WAS NOT PROPERLY LABELED. IT DID NOT STATED THAT THIS PRODUCT DID HAVE LATEX IN IT AND WAS ALMOST USED ON PATIENT THAT WAS ALLERGIC TO LATEX. PER FOLLOW UP VIA EMAIL ON (B)(6) 2023, AT PRESENT THE PWF030 BAG MENTIONS DRY NATURAL RUBBER, WITHOUT SAYING LATEX. IT NEEDS TO INCLUDE THAT WORD, SO THAT NURSES WILL NOT FIT PWF030 TO A PATIENT THEY KNOW TO HAVE ¿LATEX ALLERGY.¿ | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03210 | 2023-05-08 | 2023-04-13 | I, n, j, u, r, y | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 39302 | 39302 |  | IT WAS REPORTED THAT THE PATIENT USED SPIRIT CATHETERS AND BAGS. THE PATIENT WAS HOSPITALIZED WITH AN INFECTION AND HIGH FEVER AND URINARY TRACT INFECTION WITH BLOOD THEM URINE. THE CUSTOMER STATED WAS NOT SURE IF IT IS DUE TO CONDOM CATHETER, BUT THEY HAD ISSUES WITH CATHETERS TWISTING. THE HOSPITAL USING SOMETHING TO HOLD CATHETERS IN PLACE. THE REPRESENTATIVE LET CUSTOMER KNOW ASK WHAT THEY ARE USING. THE REPRESENTATIVE MIGHT BE USING A STAT LOCK AND PROVIDED PRICING. THERE WAS NO MENTION IF LOT INFORMATION WAS AVAILABLE. IT WAS UNKNOWN IF THE DEVICE CAUSED THE INFECTION AND UNKNOWN WHAT MEDICAL INTERVENTION WAS PROVIDED FOR THE INFECTION AT THIS TIME. | THE REPORTED EVENT WAS INCONCLUSIVE. NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "INAPPROPRIATE PACKAGE DESIGN". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. A LABELING REVIEW WAS NOT PERFORMED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED FAILURE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03224 | 2023-05-08 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® SILICONE COATED FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165V14S | 0165V14S |  | IT WAS REPORTED THAT THE SILICON LATEX FOLEY CATHETER BALLOON WAS BURST IN 1 WEEK. BALLOON MIGHT HAD BEEN DEFECTIVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
2020394-2023-00318 | 2023-05-08 | 2023-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80124 | ATG80124 |  | IT WAS REPORTED THAT DURING AN ANGIOPLASTY PROCEDURE, THE BALLOON WAS ALLEGEDLY CAN'T EXIT SMOOTHLY USING VENOVE POST-DILATE. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS AND A VIDEO WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 01/2026). | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03218 | 2023-05-08 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED AFTER THE OPERATION ON (B)(6) 2023 MORNING, WHILE IN THE AFTERNOON OF THE SAME DAY, IT WAS FOUND THAT THE BALLOON WAS BURST AND THE CATHETER FELL OUT. THEN A NEW CATHETER WAS REPLACED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO THIN SAC CAUSING BY SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01739 | 2023-05-08 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN PORT | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | UNK PORT |  | IT WAS REPORTED THAT ONE DAY POST PORT PLACEMENT PROCEDURE, THE DEVICE ALLEGEDLY HAD A LEAK. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. TWO ELECTRONIC PHOTOS WERE PROVIDED FOR REVIEW. THE PHOTO SHOWS THE DRESSING WHICH IS SOAKED IN BLOOD. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED LEAK ISSUE AS NO OBJECTIVE EVIDENCE WAS NOTED TO CONFIRM THE REPORTED EVENT. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: G3, H6 (METHOD) H11: B5, H6 (DEVICE, RESULT, CONCLUSION) H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03220 | 2023-05-08 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123520A |  | IT WAS REPORTED THAT THE CATHETER WAS NORMAL PRIOR TO USE AND WHILE ON USE, IT LEAKED AND FELL OUT. | THE REPORTED EVENT WAS INCONCLUSIVE AS NO SAMPLE WAS RETURNED FOR EVALUATION. IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR UROLOGICAL CARE. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. NO SAMPLE WAS RETURNED FOR EVALUATION. POTENTIAL ROOT CAUSE COULD BE DUE TO BUBBLE VOID ON RUBBERIZE LAYER THAT MAY LEAD TO ISSUE LUMEN TO LUMEN LEAKAGE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS (1) PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (4) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) DO NOT USE IN PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX. [SHAPE, CONFIGURATION AND PRINCIPLES] <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX <SIZES OF CATHETERS> AVAILABLE IN SIZES 12 TO 26 EVERY 2FR 1. BALLOON CATHETER [INTENDED USE & EFFECT-EFFICACY] THIS DEVICE IS AN INDWELLING CATHETER IN THE BLADDER FOR URINARY DRAINAGE. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. 1) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. 2) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT. 3) INSERT CATHETER INTO THE URETHRAL MEATUS AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A NEEDLE-LESS SYRINGE, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. 4) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. 5) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. 2. PRECAUTIONS FOR USE 1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. 2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE BY HANDS. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] 3) WHEN INSERTING CATHETER WITH A STYLET, CONFIRM THAT THE STYLET HAS REACHED THE TIP OF THE CATHETER, AND MAKE SURE TO KEEP THE STYLET IN PLACE INSIDE THE CATHETER DURING INSERTION. 4) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. 5) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. 6) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. 7) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. 8) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) 1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS 1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. 2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING IT USING A CYSTOSCOPE. 3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. 1) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. 2) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. 3) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03216 | 2023-05-08 | 2023-02-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V16S |  | IT WAS REPORTED THAT THE CATHETERIZATION WAS PERFORMED FOR OBSERVING THE URINE VOLUME AND AFTER 4 HOURS, IT WAS FOUND THAT THE CATHETER FELL OUT AND THE BALLOON WAS BURST. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE DUE TO THIN SAC CAUSING BY SHORT LATEX DWELL TIME, LATEX DIP SPEED OUT TOO SLOW. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03217 | 2023-05-08 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123416C |  | IT WAS REPORTED THAT WHITE DEPRESSED DAMAGES WITH THE BALLOON PART FOR A TOTAL OF 5350 FOLEY CATHETERS FROM TWO BATCHES (BATCH# MYGN5842 AND BATCH# MYGR5328) WERE FOUND DURING USE AND INSPECTION. THE QUALITY DEPARTMENT OF COMPANY HAD DETERMINED THAT THERE WERE RISKS OF BALLOON BURSTING AND THAT THESE PRODUCTS COULD NOT BE USED IN THE PRODUCTION. THE ISSUE WAS DETECTED DURING REUSE OF A SINGLE USE MEDICAL DEVICE. PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 13APR2023, STATED THAT CUSTOMER WAS A DISTRIBUTION CENTER. STATED THAT 5 PHYSICAL SAMPLES OF EACH BATCH WOULD BE PROVIDED. THE USER PURCHASED 5350 UNITS AND INSPECTED SOME OF THEM HAD THE ISSUE, THUS THE CUSTOMER WAS WORRIED ABOUT THE QUALITY OF THE WHOLE BATCHES AND REQUEST TO CHANGE AND RETURN 5350 UNITS WITH ANOTHER BATCH. PER SAMPLE RECEIVED BY THE BD REPRESENTATIVE ON 14APR2023, STATED THAT THE MANUFACTURING LOT# OF SAMPLES RECEIVED FOR THE BATCH MYGN5842 WAS 2HJ031 AND FOR THE BATCH MYGR5328 WERE 2JD072, 2HH041, 2HU135 AND 2HL031 (BATCH# MYGN5842 AND SERIAL# (B)(4), BATCH# MYGR5328 AND SERIAL# (B)(4), BATCH# MYGR5328 AND SERIAL# (B)(4), BATCH# MYGR5328 AND SERIAL# (B)(4) AND BATCH# MYGR5328 AND SERIAL# (B)(4)). PER ADDITIONAL INFORMATION VIA EMAIL FROM IBC ON 15APR2023, STATED THAT THE CUSTOMER WAS AN OLD CUSTOMER, THEY HAVE USED BD FOLEY FOR ALMOST 18 YEARS SINCE 2006. THIS WAS THE FIRST INCIDENT EXPERIENCED BY THEM. THE MANUFACTURING LOT# OF SAMPLES FOR THE BATCH MYGN5842 WAS 2FD047 AND FOR THE BATCH MYGR5328 WERE 2HU136, 2FR083, 2JC013 AND 2HU134 (BATCH# MYGN5842 AND SERIAL# (B)(4), BATCH# MYGR5328 AND SERIAL# (B)(4), BATCH# MYGR5328 AND SERIAL# (B)(4), BATCH# MYGR5328 AND SERIAL# (B)(4) AND BATCH# MYGR5328 AND SERIAL# (B)(4)). STATED THAT THIS CUSTOMER PER CHASE FOLEY FOR REPROCESSING. WHEN THEY RECEIVED FOLEY, THEY WOULD JUST CHECK PRODUCT MODEL, SIZE, QUANTITY AND EXPIRATION DATE. WHILE DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. THERE WAS NO SAMPLING INSPECTION PERFORMED, THEY TESTED THE BALLOON OF ALL CATHETERS. STATED THAT DURING THE PROCESS OF REPRODUCTION, THEY WOULD INFLATE BALLOON WITH GAS FOR ALL FOLEY CATHETERS. PER ADDITIONAL INFORMATION RECEIVED VIA EMAIL FROM IBC ON 21APR2023, STATED THAT THE SAMPLES REPORTED BY CUSTOMER WERE THOSE PICKED UP BY CUSTOMER FOR FUNCTION TESTING DURING THEIR REPRODUCTION PROCESS AND THESE PHYSICAL SAMPLES COULD NOT BE RETURNED FOR INVESTIGATION, BUT THE PHOTOS ATTACHED. WHILE FOR THE SAMPLES RECEIVED, THESE WERE REPRESENTATIVE SAMPLES FROM THE SAME REPORTED BATCHES RETURNED FOR INVESTIGATION, THEY WERE NOT THE REAL AFFECTED SAMPLES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03214 | 2023-05-08 | 2023-03-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | UNKNOWN PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT USED PUREWICK FEMALE EXTERNAL CATHETER IN THE HOSPITAL AND EXPRESSED LATEX ALLERGY TO NURSE. PATIENT STATED THAT THEY MUST HAVE USED AN OLD WICK WITH LATEX BECAUSE PATIENT HAD A TERRIBLE ALLERGIC REACTION WITH VAGINAL BLISTERS. THE NURSE STATED THAT THERE WAS NO LONGER MADE WICKS WITH LATEX AND BE SURE BY CHECKED THE PACKAGE ON EACH WICK IT SHOULD SAID THAT NOT MADE WITH NATURAL RUBBER LATEX. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW-UP VIA PHONE ON (B)(6)2023, IT WAS REPORTED THAT THE CUSTOMER HAD LATEX ALLERGY AND WHEN THEY WERE IN THE HOSPITAL A NURSE USED A WICK THAT HAD LATEX AND CAUSED THEM TO HAD AN ALLERGIC REACTION CONSISTING OF BLISTERS ON THEIR VAGINA. THEY WERE TREATED BY THEIR DOCTOR WITH AN ANTIBIOTIC OINTMENT THAT THEY DID NOT KNOW THE NAME OF. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03221 | 2023-05-08 | 2023-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 0167V22S |  | IT WAS REPORTED THAT THE CATHETER WAS CHECKED WITH NO LEAKAGE OR BREAKAGE PRIOR TO USE. WHEN THE CATHETER WAS IMPLANTED IN THE PATIENT AND 25ML SALINE WAS INJECTED FOR FIXATION, IT CAN BE PULLED BACK WITHOUT ANY RESISTANCE. WHEN THE CATHETER WAS REMOVED FROM THE PATIENT, IT WAS FOUND THAT THE BALLOON HAS BURST. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03215 | 2023-05-08 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50616 | 50616 |  | IT WAS REPORTED THAT THE PATIENT STATED THREE OF THE 120 INTERMITTENT CATHETERS RECEIVED WERE SHOVED DOWN INTO SHIPMENT BOX AND THIS BENDED THEM. IT WAS STATED THAT THE PATIENT STILL ABLE TO USE THEM BUT WANTED TO MAKE US AWARE OF THIS. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿COMPONENT KINKED OR DAMAGED FROM SUPPLIER¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. LABELING REVIEW IS NOT REQUIRED BECAUSE LABELING COULD NOT HAVE PREVENTED THE REPORTED ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01768 | 2023-05-09 | 2023-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7274118 |  | IT WAS REPORTED, "PATIENT WITH A 4 FR BILUMEN PICC CATHETER INSERTED ON (B)(6) 2023 IN THE RIGHT ARM, COVERED WITH AN INTEGRAL TRANSPARENT DRESSING, WITH THE LAST CURE DATE OF (B)(6) 2023. PER PROTOCOL, CURING IS PERFORMED WITH A STERILE TECHNIQUE AFTER CHANGING THE NEEDLE-FREE VALVE PROCEEDS TO PERFORM IRRIGATION THROUGH A RED LUMEN, FINDING LIQUID LEAKAGE BETWEEN THE ZERO LEVEL AND THE BODY OF THE DEVICE, WHICH IS WHY IT IS DECIDED TO WITHDRAW IT.". ADDITIONAL INFORMATION MAY 2, 2023: THE TREATMENT HAS BEEN EXCHANGED TO AN ORAL ANTIBIOTIC. NO NEW CATHETER PLACED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01770 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERPICC 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001028 |  | IT WAS REPORTED THE INTRODUCER WILL NOT INSERT INTO VASCULATURE. PATIENT HAD TO BE ACCESSED TWICE. IT WAS STATED THERE WAS NO HARM TO THE PATIENT. NO OTHER INFORMATION WAS REPORTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01769 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 5F DUAL-LUMEN POLYURETHANE CATHETER RADIOLOGY BASIC TRAY | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3275135 |  | IT WAS REPORTED, "WENT TO A CLIENT'S HOME FOR ROUTINE PICC MAINTENANCE. NOTED AIR BUBBLES IN BOTH LUMENS OF CLIENTS PICC LINE. DENIED ANY SHORTNESS OF BREATH, CHEST PAIN OR LIGHT HEADEDNESS. NOTIFIED HOSPITAL TRIAGE OF AIR BUBBLES. HOSPITAL STATED WILL FOLLOW UP WITH CLIENT." NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01767 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER (20G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202252 |  | IT WAS REPORTED BY THE CUSTOMER THAT "DEVICE WOULD NOT SAFETY AFTER PLACING THE CATHETER IN THE PATIENT. THE WIRE AND NEEDLE STAYED OUT TOGETHER WHEN THE NURSE TRIED TO SAFETY. THERE WAS NO NEGATIVE IMPACT TO THE PATIENT. THE CATHETER WAS ABLE TO THREAD OFF BUT WHEN THE NURSE PULLED BACK THE NEEDLE JUST DID NOT SAFETY WHEN THEY HIT THE ACTIVE SAFETY BUTTON." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01772 | 2023-05-09 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PPICC 5F DL RN | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK000960 |  | IT WAS REPORTED BY THE CUSTOMER THAT "NEEDLE IS VERY DULL. HAVING TO PUSH SO HARD TO ACCESS VEINS AND THEY ARE BACK WALLING THE VESSEL. HAVE NOTICED SIGNIFICANT CHANGE IN NEEDLE SHARPNESS." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01773 | 2023-05-09 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERMIDLINE CATHETER BASIC KIT (4F) (SINGLE-LUMEN) (20CM) | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | P6154118 |  | IT WAS REPORTED, THERE IS A FRACTURE AT THE BASE OF THE CANNULA/POWERMIDLINE CATHETER. TREATMENT HAS BEEN INTERRUPTED AFTER COUPLE OF DAYS. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01766 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SV CATHETER 4F RADIOLOGY BASIC TRAY (DUAL-LUMEN) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 3274355 |  | IT WAS REPORTED, "CALLER STATES HER DAUGHTER HAD THIS PICC FOR ABOUT FOUR DAYS THEN IT STARTED LEAKING AND HAD TO BE REMOVED. CALLER STATES AFTER IT WAS REMOVED HER DAUGHTER HAD AN ALLERGIC REACTION TO THE CATHETER. CALLER STATES HER DAUGHTER EXPERIENCED PUFFINESS AND SWELLING. CALLER STATES HER DAUGHTER IS AT HOME, DOING WELL AND TAKING BENADRYL." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | LAY USER/PATIENT |  | 2023-05
2020394-2023-00320 | 2023-05-09 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ATLAS GOLD | PTA BALLOON DILATATION CATHETER | BARD PERIPHERAL VASCULAR, INC. | ATG80126 | ATG80126 |  | IT WAS REPORTED THAT PRIOR TO AN ANGIOPLASTY PROCEDURE, THE OUTSIDE AND INSIDE OF THE PACKAGE ALLEGEDLY GOT FOLDED DURING DELIVERY. IT WAS FURTHER REPORTED THAT THE STERILITY OF THE DEVICE WAS AFFECTED. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, A PHOTO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 03/2026). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01763 | 2023-05-09 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1295108FD4 |  | IT WAS REPORTED BY THE CUSTOMER THAT "ISSUE WITH TWO PICCS THAT WERE INSERTED BY TWO DIFFERENT NURSES. BOTH HAD TROUBLE WITH THE INTRODUCER NOT HAVING A CLEAN SPLIT OF THE RED PART INSIDE THE INTRODUCER. THERE WERE PIECES OF THE RED INNER CORE FRAYED AND THE PICC WAS THEN STUCK IN THE INTRODUCER. THE NURSES HAD TO CUT AWAY THE RED PLASTIC INNER PART." THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01764 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY W/MICROINTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | IT WAS REPORTED THAT THE OPERATOR PRE-FLUSHED THE CATHETER TO INSPECT IT FOR INTEGRITY AND FOUND THAT THE CATHETER HAD SAND HOLES. THE ASSISTANT HELPED THE OPERATOR TO REPLACE THE CATHETER AND A NEW CATHETER WAS INSERTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01771 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER (20G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1202252 |  | IT WAS REPORTED, "THE DEVICE WOULD NOT SAFETY AFTER PLACING THE CATHETER IN PATIENT. THE WIRE AND NEEDLE STAYED OUT TOGETHER WHEN THE NURSE TRIED TO SAFETY IT." NO OTHER INFORMATION WAS PROVIDED. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01775 | 2023-05-09 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD POWERGLIDE MIDLINE | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | UNKNOWN |  | IT WAS REPORTED BY THE CUSTOMER THAT "SEVERAL INSTANCES OF POWERGLIDES NOT WORKING RIGHT AFTER PLACEMENT. EACH TIME THE NURSE STATED THEY PULLED THE CATHETER BACK SLIGHTLY AND THE LINE THEN STARTED WORKING." THE SPECIFIC NUMBER OF PATIENTS OR DEVICES EFFECTED WERE NOT PROVIDED BY THE CUSTOMER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01780 | 2023-05-09 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIG JAPAN HANG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405J |  | IT WAS REPORTED THAT AFTER THE CATHETER WAS INSERTED INTO THE SHEATH AND THE CATHETER WAS IMPLANTED, BLOOD RETURN COULD BE CONFIRMED. AFTER THE STYLET WAS REMOVED, THE CATHETER WAS CUT AND ATTACHED TO THE CATHETER CONNECTOR, AND WHEN BLOOD RETURN WAS CONFIRMED AGAIN, ASPIRATION AND INFUSION COULD NOT BE PERFORMED. THEREFORE, THE CATHETER CONNECTOR WAS REMOVED AND ANOTHER NEW CATHETER CONNECTOR WAS ATTACHED TO THE CATHETER. THEN, ASPIRATION AND INFUSION COULD BE PERFORMED. THE FACILITY IS REQUESTING AN INVESTIGATION TO DETERMINE WHETHER THE PROBLEM IS WITH THE CATHETER CONNECTOR, OR WITH THE PLACEMENT PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01776 | 2023-05-09 | 2023-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THE CATHETER WAS BLOCKED JUST A COUPLE OF HOURS AFTER PLACEMENT BY THE HEAD NURSE. IT WAS SUSPECTED THAT THE VALVE MALFUNCTIONED. CATHETER WAS REMOVED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01777 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT DURING PUNCTURE, THE CUSTOMER OF THE RADIOTHERAPY DEPARTMENT FOUND THAT THE SELDINGER COULD NOT BE INTRODUCED INTO THE PICC CATHETER SINCE THE PROXIMAL END OF THE FORMER WAS CURLED UP. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01778 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL-APART SHEATH INTRODUCER 4.5F X 5CM | INTRODUCER, CATHETER | C.R. BARD, INC. (BASD) -3006260740 | N/A | 0668945 |  | IT WAS REPORTED THAT DURING PUNCTURE, THE CUSTOMER OF THE (B)(6) HOSPITAL FOUND THAT THE SELDINGER COULD NOT BE INTRODUCED INTO THE PICC CATHETER SINCE THE PROXIMAL END OF THE FORMER WAS CURLED UP. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01779 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT DURING PLACEMENT AT HEPATOLOGICAL SURGERY DEPARTMENT THE CATHETER WAS KINKED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01781 | 2023-05-09 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1274108D4 |  | CUSTOMER REPORTED THAT WHEN THE PICC WAS FLUSHED, SALINE CAME OUT AROUND THE GUIDEWIRE VIA THE RUBBER STOPPER ON THE T-PORT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01783 | 2023-05-09 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH, | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED A SMALL LEAK WAS NOTED DURING FLUSHING. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01748 | 2023-05-09 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT PRIOR TO PORT PLACEMENT PROCEDURE, PORT ALLEGEDLY DIDN'T FLUSH. THERE WAS NO PATIENT CONTACT. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. EXPIRATION DATE: 09/2024. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01762 | 2023-05-09 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA SOLO 5F DL FT MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1295108FD4 |  | IT WAS REPORTED BY THE CUSTOMER THAT "ISSUE WITH TWO PICCS THAT WERE INSERTED BY TWO DIFFERENT NURSES. BOTH HAD TROUBLE WITH THE INTRODUCER NOT HAVING A CLEAN SPLIT OF THE RED PART INSIDE THE INTRODUCER. THERE WERE PIECES OF THE RED INNER CORE FRAYED AND THE PICC WAS THEN STUCK IN THE INTRODUCER. THE NURSES HAD TO CUT AWAY THE RED PLASTIC INNER PART." THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01784 | 2023-05-09 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER & NEEDLE GUIDES | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D5 |  | IT WAS REPORTED VIA MEDWATCH, RUBBER STOPPER ON SINGLE-LUMEN PICC LEAKS SALINE WHEN FLUSHED DURING PICC INSERTION. NO PATIENT HARM REPORTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01785 | 2023-05-09 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH, | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED VIA MEDWATCH, RUBBER LUMEN LEAKING WHERE THE WIRE FOR THE PICC IS DURING PICC INSERTION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01786 | 2023-05-09 | 2022-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | IT WAS REPORTED THAT THE HEAD NURSE OF THE NEUROSURGERY DEPARTMENT II REVEALED THAT THE EXTENSION TUBING OF THE GROSHONG PICC CATHETER INSERTED IN ONE PATIENT WAS RUPTURED. THE CATHETER WAS INSERTED IN THIS PATIENT ON (B)(6) 2023. DURING ROUTINE MAINTENANCE ONE MONTH AFTER PLACEMENT, THE NURSE FOUND THAT THE EXTENSION TUBING WAS RUPTURED. THE HEAD NURSE HAD REPLACED THE EXTENSION TUBING WITH A NEW ONE. THE NURSE SAID THAT SHE HAD BEEN CAREFUL IN MAINTAINING THE CATHETER. THE EXTENSION TUBING WAS NOT KINKED, AND BLOOD DRAW AND TUBING PRE-FLUSHING WERE DONE SMOOTHLY. THE FAMILY OF THE PATIENT ALWAYS REMINDED THE PATIENT OF EXPOSING THE UPPER ARM WHERE THE CATHETER WAS PLACED TO PREVENT CATHETER KINKING OR COMPRESSION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01758 | 2023-05-09 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 5608062 | 5608062 |  | IT WAS REPORTED THAT ONE WEEKS TWO DAYS PORT PLACEMENT PROCEDURE, THE PATIENT EXPERIENCED BURNING SENSATION AT THE SITE. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. DEVICE PENDING RETURN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01774 | 2023-05-09 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ACCUCATH ACE INTRAVASCULAR CATHETER (18G) | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | AC1182252 |  | IT WAS REPORTED BY THE CUSTOMER "RN PLACED ACCUCATH BUT DID SO WITHOUT SAFETY ENGAGING AND WITH GUIDEWIRE NOT BEING ABLE TO DEPLOY. GUIDEWIRE SEEMED TO HAVE GOTTEN STUCK IN NEEDLE AND SPREAD TO THE LEFT FROM ONE OF THE NEEDLE HOLES. PATIENT EXPERIENCED PAIN DURING WITHDRAW OF DEVICE. CATHETER REMAINED IN PLACE." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03285 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP INTERMITTENT CATHETER TRAY | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | INTS14 | INTS14 |  | IT WAS REPORTED THAT THERE WAS NO URINE OUTPUT FOLLOWING INTERMITTENT CATHETER INSERTION AND THE CATHETER WAS REMOVED AND REVEALED NO PERFORATIONS IN THE END OF THE TUBE AND IT WAS DEFECTIVE PRODUCT NOTED BY THE NURSE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
16897174 | 2023-05-09 | 2023-04-25 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARD FOLEY CATHETER | FOLEY CATHETER | BARD | Q174 L24 |  |  | PT HAD A FOLEY INSERTED OVERNIGHT; 3WAY FOLEY. PT COMPLAINED OF FEELING WET. FOLEY HAD COME OUT. BALLOON BROWN. NEW FOLEY INSERTED. EXAM OF FOLEY: SLIT IN FOLEY BALLOON - CATHETER INNER PORTION INTACT. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03282 | 2023-05-09 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 0165SI12 | 0165SI12 |  | IT WAS REPORTED THAT URINE DID NOT FLOW INTO THE URINE COLLECTION BAG AFTER CATHETER INSERTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3012307300-2023-05442 | 2023-05-09 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DELTEC POWER PORT-A-CATH II PORTS | PORT ANDCATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVA | ST PAUL |  | 21-4474-24 |  | IT WAS REPORTED THAT THERE WAS A LEAK OF CHEMO SOLUTION FROM THE PUMP THAT RUNS THROUGH THE PORT CATHETER IN A FEW PATIENTS. NO PATIENT INJURY WAS REPORTED. TEXT IN COMPLAINT FORM: LEAKAGE (DROP BY DROP) OF CHEMO PRODUCT AT THE INSERTION SITE IN 2 PATIENTS WITH A PORT. LEAKS OCCURRED IN 3 OTHER PATIENTS, BUT PORTS FROM A DIFFERENT SUPPLIER HAD BEEN PLACED THERE, WHICH LEADS TO SUSPECT THAT IT WAS NOT THE PORTS, BUT RATHER THE HUBER POINT NEEDLES. | D4 UDI AND G5 ARE UNKNOWN. NO PRODUCT INFORMATION HAS BEEN PROVIDED TO DATE. THIS REMEDIATION MDR WAS GENERATED UNDER PROTOCOL (B)(4), AS A RESULT OF WARNING LETTER CMS# (B)(4). NO PRODUCT SAMPLE WAS RECEIVED; THEREFORE, VISUAL AND FUNCTIONAL TESTING COULD NOT BE PERFORMED. THE REPORTED ISSUE COULD NOT BE CONFIRMED AS NO PRODUCT SAMPLE WAS RECEIVED FOR EVALUATION. IF THE PRODUCT IS RETURNED, THE MANUFACTURER WILL REOPEN THIS COMPLAINT FOR FURTHER INVESTIGATION. NO PROBLEMS OR ISSUES WERE IDENTIFIED DURING THE DEVICE HISTORY RECORD REVIEW. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01747 | 2023-05-09 | 2023-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PRISTINE LONG-TERM HEMODIALYSIS CATHETER | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5403190 | 5403190 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE DEVICE SHEATH ALLEGEDLY CRACKED. IT WAS FURTHER REPORTED THAT THE DEVICE PEELED AWAY, AND THE HANDLE OF THE SHEATH BROKE OFF. REPORTEDLY, THE DEVICE WAS REMOVED. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) ARE ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01750 | 2023-05-09 | 2022-11-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NAVARRE OPTI DRAIN | DRAINAGE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NOD8LPT | NOD8LPT |  | IT WAS REPORTED THAT DURING A DRAINAGE CATHETER PLACEMENT PROCEDURE, THE HEAD END OF THE DRAINAGE TUBE WAS ALLEGEDLY WRINKLED. IT WAS FURTHER REPORTED THAT CATHETER ALLEGEDLY UNABLE TO ADVANCE INTO TISSUE. THE PROCEDURE WAS COMPLETED BY USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 8F NAVARRE DRAINAGE CATHETER LOCKING PIGTAIL SEGMENTS WERE RETURNED FOR EVALUATION. GROSS VISUAL, FUNCTIONAL AND DIMENSIONAL EVALUATION WERE PERFORMED. MANUFACTURING SITE EVALUATION OF THE SAMPLES FOUND PRESENCE OF A WRINKLE IN THE SHAFT SEVERAL CENTIMETERS DISTAL OF THE CATHETER MIDPOINT. THE INVESTIGATION IS CONFIRMED FOR THE REPORTED HEAD END OF THE DRAINAGE TUBE WAS WRINKLED AND REPORTED DEFORMATION ISSUE AS SMALL WRINKLE ON THE DISTAL END OF THE CATHETER. ALSO, THE INVESTIGATION IS CONFIRMED FOR THE IDENTIFIED DEFECTIVE COMPONENT ISSUE AS THE STYLET PROTRUSION (DISTAL TIP OF THE TROCAR NEEDLE) IS GREATER THAN THE ACCEPTED TOLERANCE LIMIT. HOWEVER, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED ADVANCEMENT ISSUE AS EXACT CIRCUMSTANCE OF THE TIME OF THE EVENT WAS UNKNOWN. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. IT IS UNKNOWN WHETHER PATIENT AND/OR PROCEDURAL ISSUES CONTRIBUTED TO THE REPORTED EVENT. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: D4 (EXPIRY DATE: 08/2024). | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01751 | 2023-05-09 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. HARD BASE IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0604550 | 0604550 |  | IT WAS REPORTED THAT ONE WEEK AND THREE DAYS POST PORT PLACEMENT, THE TIP OF THE CATHETER WAS ALLEGEDLY FOUND SPLIT. IT WAS FURTHER REPORTED THAT THE CATHETER ALLEGEDLY LEAKED SUBCUTANEOUSLY. THERE WAS NO REPORTED PATIENT INJURY. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 07/2026). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01753 | 2023-05-09 | 2023-02-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE, CATHETER WAS ALLEGEDLY TWISTED. IT WAS FURTHER REPORTED THAT CATHETER WAS ALLEGEDLY DIFFICULT TO PASS OVER THE GUIDEWIRE. FURTHERMORE, THE OUTER PACKAGING ALLEGEDLY HAD SEVERE PLASTIC INDENTATIONS AND CANNOT BE OPENED. REPORTEDLY, THE CATHETER WAS REMOVED AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 19CM HEMOSTAR D/L CATHETER KIT. VARIOUS COMPONENTS WERE RETURNED FOR EVALUATION, INCLUDING THE CATHETER. GROSS VISUAL, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. THE 19CM HEMOSTAR D/L CATHETER WAS NOTED TO BE BENT. IN ADDITION TO THE RETURNED PHYSICAL SAMPLE, FOUR ELECTRONIC PHOTOS WERE PROVIDED FOR REVIEW. THESE FOUR ELECTRONIC PHOTOS WERE REVIEWED. THE PHOTOS SHOW A 19CM HEMOSTAR D/L CATHETER. PHOTOS DEPICTING THE CATHETER IN THE TRAY SHOW IT IS CURVED AND ORIENTED IN A MANNER SIMILAR TO THE SPECIFIED PACKAGING ORIENTATION OF PRE-CURVED CATHETERS, ALTHOUGH THIS IS A STRAIGHT CATHETER. THAT IS, THE CATHETER IS DOUBLED BACK ON ITSELF IN A "U" SHAPE, WITH THE TIP IN A PROTECTIVE TUBE TAPED TO THE RIGID PACKAGING. NO PACKAGING DAMAGE WAS NOTED. IT IS UNKNOWN WHAT "INDENTATIONS" IN THE PACKAGING MAY BE REFERRED TO, AND ALSO THE PHOTOGRAPHS DO NOT CONSTITUTE A COMPLETE UNOBSTRUCTED VIEW OF THE TRAY AND RETAINER COMPONENTS, AND SO THE REPORTED PACKAGING/PACKAGING DEFORMATION ISSUES WILL REMAIN INCONCLUSIVE. THE CATHETER BENDING CAN BE CONFIRMED BASED ON THE SAMPLE EVALUATION. MISASSEMBLY OF THE COMPONENTS IS ALSO CONFIRMED. THE DIFFICULTY OF ADVANCEMENT CANNOT BE CONFIRMED FROM THE AVAILABLE SAMPLES. THE CAUSE OF THIS CONDITION IS RELATED TO INCORRECT PACKAGING DURING MANUFACTURING. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRATION DATE: 03/2024). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01754 | 2023-05-09 | 2023-02-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE VIA SUBCLAVIAN VEIN, THE CATHETER WAS ALLEGEDLY TWISTED. IT WAS FURTHER REPORTED THAT CATHETER WAS ALLEGEDLY DIFFICULT TO PASS OVER THE GUIDEWIRE. IT WAS ALSO REPORTED THAT THE PRODUCT PACKAGING ALLEGEDLY HAD DEFORMATION. REPORTEDLY, THE CATHETER WAS REMOVED AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURES AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: D4 (EXPIRY DATE: 03/2024). H11: H6 (DEVICE). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01755 | 2023-05-09 | 2023-02-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE VIA SUBCLAVIAN VEIN, THE CATHETER WAS ALLEGEDLY TWISTED. IT WAS FURTHER REPORTED THAT CATHETER WAS ALLEGEDLY DIFFICULT TO PASS OVER THE GUIDEWIRE. REPORTEDLY, THE CATHETER WAS REMOVED AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 19CM HEMOSTAR D/L CATHETER KIT. FOLLOWING COMPONENTS WERE RECEIVED: ONE 19CM HEMOSTAR D/L CATHETER, ONE 8FR DILATOR, ONE 12FR DILATOR, ONE 15FR PEEL-APART SHEATH AND VESSEL DILATOR, ONE J-TIP GUIDEWIRE IN A GUIDEWIRE HOOP, ONE TUNNELER, ONE INTRODUCER NEEDLE AND TWO END CAPS. GROSS VISUAL, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. THE 19CM HEMOSTAR D/L CATHETER WAS NOTED TO BE BENT. AN ATTEMPT TO LOAD THE CATHETER TO THE RECEIVED GUIDEWIRE WAS PERFORMED AND WAS SUCCESSFUL. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED CATHETER DEFORMATION ISSUES. HOWEVER THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED FAILURE TO ADVANCE ISSUE AS THE CATHETER WAS OBSERVED SLIDING ALONG THE WIRE WITH NO ISSUE. ALSO, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED PACKAGING PROBLEM ISSUE AS THERE WAS NO OBJECTIVE EVIDENCE PROVIDED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03256 | 2023-05-09 | 2023-02-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE FOLEY CATHETER BALLOON DEFLATED WITHIN ABOUT 30 SECONDS. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03258 | 2023-05-09 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123616A |  | IT WAS REPORTED THAT THE CATHETER WAS DIFFICULT TO REMOVE. THE SYRINGE WAS USED TO PULL AND PUSH WITH AIR REPEATEDLY THEN THE LIQUID CAN BE WITHDRAWN, THEN THE CATHETER CAN BE REMOVED SUCCESSFULLY. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: OVER ASPIRATED, INCORRECT SYRINGE/COLLAPSE LUMEN/SAC CLOSE EYE/VALVE DAMAGE). A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITION THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE NO ADDITIONAL ACTION REQUIRED AT THIS TIME. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03261 | 2023-05-09 | 2023-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 123618C |  | IT WAS REPORTED THAT THE PATIENT HAD AN INDWELLING FOLEY CATHETER FOR URETHRAL INJURY AND THEN THE CATHETER FELL OUT. UPON CHECKING, IT WAS FOUND THAT THE BALLOON FOUND TO BE RUPTURED AND THEN A NEW CATHETER WAS PLACED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. HOWEVER, THE POTENTIAL ROOT CAUSE FOR THIS FAILURE MODE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/ OPERATOR ERROR. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "DO NOT USE THE PRODUCT OF HAVE LATEX ALLERGY. SCOPE OF APPLICATION: FOLEY CATHETER IS INTENDED FOR USE IN THE DRAINAGE OF URINE FROM THE BLADDER OF CHILDREN AND ADULTS. PRECAUTION: THIS PRODUCT CONTAINS NATURAL RUBBER LATEX WHICH MAY CAUSE ALLERGIC REACTION. STERILE UNLESS PACKAGE HAS BEEN OPENED OR DAMAGE. WARNING: DO NOT USE PETROLEUM SUBSTRATE LUBRICANTS WITH THE LATEX- BASED URETHRAL CATHETERS SUCH AS PETROLEUM JELLY AND LABEL LIQUID PARAFFIN, WHICH WILL DAMAGE LATEX AND MAY BURST BALLOON. WATER-SOLUBLE LUBRICANT CAN BE USED. STORAGE: CATHETERS NEED TO BE STORED AT ROOM TEMPERATURE AND KEEP AWAY FROM THE DIRECT LIGHT AND PREFERABLY STORED IN THE ORIGINAL PACKING BOX." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01756 | 2023-05-09 | 2023-02-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833690 | 5833690 |  | IT WAS REPORTED THAT DURING A DIALYSIS CATHETER PLACEMENT PROCEDURE VIA SUBCLAVIAN VEIN, THE CATHETER WAS ALLEGEDLY TWISTED. IT WAS FURTHER REPORTED THAT CATHETER WAS ALLEGEDLY DIFFICULT TO PASS OVER THE GUIDEWIRE. IT WAS ALSO REPORTED THAT THE PRODUCT PACKAGING ALLEGEDLY HAD DEFORMATION. REPORTEDLY, THE CATHETER WAS REMOVED AND THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: ONE 19CM HEMOSTAR D/L CATHETER KIT. FOLLOWING COMPONENTS WERE RECEIVED: ONE 19CM HEMOSTAR D/L CATHETER, ONE 8FR DILATOR, ONE 12FR DILATOR, ONE 15FR PEEL-APART SHEATH AND VESSEL DILATOR, ONE J-TIP GUIDEWIRE IN A GUIDEWIRE HOOP, ONE TUNNELER, ONE INTRODUCER NEEDLE AND TWO END CAPS. GROSS VISUAL, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. THE 19CM HEMOSTAR D/L CATHETER WAS NOTED TO BE BENT. AN ATTEMPT TO LOAD THE CATHETER TO THE RECEIVED GUIDEWIRE WAS PERFORMED AND WAS SUCCESSFUL. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED CATHETER DEFORMATION ISSUES. HOWEVER THE INVESTIGATION IS UNCONFIRMED FOR THE REPORTED FAILURE TO ADVANCE ISSUE AS THE CATHETER WAS OBSERVED SLIDING ALONG THE WIRE WITH NO ISSUE. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED PACKAGING PROBLEM ISSUE AS THERE WAS NO OBJECTIVE EVIDENCE PROVIDED. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. IT IS UNKNOWN WHETHER PATIENT AND/OR PROCEDURAL FACTORS CONTRIBUTED TO THE EVENT. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03265 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDIA® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 |  | 123620A |  | IT WAS REPORTED THAT THE BALLOON LEAKED. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 21APR2023, THE BALLOON WAS UNDAMAGED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03279 | 2023-05-09 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SURESTEP¿ FOLEY TRAY SYSTEM BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | A902916 | A902916 |  | IT WAS REPORTED THAT ONE OF THE FOLEY TRAY WAS MISSING THE SWAB STICK AND THEY WERE UNABLE TO PERFORM PROPER INSERTION TECHNIQUES. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE ¿OPERATOR/MACHINE ERROR ¿. IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. BASED ON PATIENT CODE 2645 THE PRODUCT DOES NOT APPEAR TO HAVE BEEN USED. HOWEVER, THE PRODUCT IS INTENDED TO BE USED FOR TREATMENT PURPOSE BUT IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A LABELING REVIEW IS NOT REQUIRED BECAUSE A REVIEW OF THE LABEL COULD NOT HAVE PREVENTED THE REPORTED ISSUE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03264 | 2023-05-09 | 2022-11-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LATEX FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED FROM A LOCAL INCIDENT THAT DURING A ROUTINE CATHETER CHANGE, THE CATHETER DID NOT DRAIN AND SO WAS REMOVED. THE CATHETER DID NOT HAVE A URINE OUTLET HOLE. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED . IT WAS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO INSUFFICIENT PROCEDURE KNOWLEDGE. A DHR REVIEW IS NOT REQUIRED AS THE LOT NUMBER IS UNKNOWN. THEREFORE, NO ADDITIONAL ACTION REQUIRED AT THIS TIME. UNABLE TO REVIEW THE LABELLING DUE TO UNKNOWN PRODUCT CODE. ALTHOUGH THE PRODUCT FAMILY IS UNKNOWN, THE FOLEY CATHETER IFUS ARE FOUND TO BE ADEQUATE BASED ON PAST REVIEWS. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03275 | 2023-05-09 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER COUDE TIP | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0168L18 | 0168L18 |  | IT WAS REPORTED THAT THE INDWELLING FOLEY CATHETERS WERE DEFECTIVE, PIN-HOLE IN PORT MAKING INFLATION BALLOON NOT POSSIBLE. ALSO STATED THAT THE CATHETERS WERE UNABLE TO BE USED AND NOT ALL IN THE BOX WERE DEFECTIVE BUT THEY HAD 4-5 IN THE SAME BRAND AND SIZE (18F AND 20F) DO THE SAME THING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03274 | 2023-05-09 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | UNKNOWN LATEX FOLEY CATHETER | UNKNOWN LATEX FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE INDWELLING FOLEY CATHETERS WERE DEFECTIVE, PIN-HOLE IN PORT MAKING INFLATION BALLOON NOT POSSIBLE. ALSO STATED THAT THE CATHETERS WERE UNABLE TO BE USED AND NOT ALL IN THE BOX WERE DEFECTIVE BUT THEY HAD 4-5 IN THE SAME BRAND AND SIZE (18F AND 20F) DO THE SAME THING. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03271 | 2023-05-09 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT STERILE WATER LEAKED FROM THE SYRINGE AND FOLEY CATHETER CONNECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03267 | 2023-05-09 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI20 | 0165SI20 |  | IT WAS REPORTED THAT UPON ATTEMPTING TO INFLATE THE BALLOON AFTER INSERTION OF THE CATHETER INTO THE PATIENT, WATER LEAKED FROM THE VALVE AND THE BALLOON DID NOT INFLATE. | THE REPORTED EVENT WAS CONFIRMED ¿ SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. THIS INVESTIGATION DOES NOT REQUIRE A DHR REVIEW DUE TO A CLOSURE LETTER NOT REQUIRED FOR THIS COMPLAINT. A LABELING REVIEW IS NOT REQUIRED AS THE INVESTIGATION WAS CONFIRMED RELATED TO SUPPLIER ISSUE. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03269 | 2023-05-09 | 2023-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® X-FORCE® BALLOON DILATION CATHETER | BALLOON DILATION CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 997706 | 997706 |  | IT WAS REPORTED THAT THE MEDICAL STAFF PERFORMED SURGERY ON THE PATIENT ON (B)(6) 2023. DURING THE OPERATION, IT WAS FOUND THAT THE FRONT END OF THE BALLOON DILATATION CATHETER WAS BLOCKED AND THE OPERATION COULD NOT BE PERFORMED. THEN, A NEW DEVICE WAS REPLACED IMMEDIATELY. THE OPERATION WENT SMOOTHLY AND WAS REPORTED TO THE MEDICAL EQUIPMENT DEPARTMENT. IT DID NOT CAUSE SIGNIFICANT HARM TO THE PATIENT. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03268 | 2023-05-09 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 53614G | 53614G |  | IT WAS REPORTED THAT THESE MAGIC 3 HYDROPHILIC INTERMITTENT CATHETERS WERE NOT ALWAYS LUBRICATED ENOUGH. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". IT IS UNKNOWN WHETHER THE DEVICE HAD MET RELEVANT SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE LOT NUMBER IS UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOME" H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03266 | 2023-05-09 | 2023-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PERSONAL CATHETER¿ INTERMITTENT CATHETER | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 73614 |  | IT WAS REPORTED THAT IN THE LAST ORDER, OF THE 2/3 FOLEY CATHETERS IN EACH BOX, THERE WAS NO LUBRICANT IN THE SACHETS, OR THEY HAD COME SLIGHTLY OPEN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03284 | 2023-05-09 | 2023-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® I.C. FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 0165SI14 | 0165SI14 |  | IT WAS REPORTED THAT ON THE NEXT DAY AFTER PLACEMENT, THE CATHETER SPONTANEOUSLY FELL OFF DUE TO BALLOON BURST. | THE REPORTED EVENT WAS CONFIRMED AND THE CAUSE WAS UNKNOWN. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE USER RELATED (EXAMPLE: CONTACT WITH SHARP OBJECT)/ EXPOSURE TO PETROLATUM BASED PRODUCTS/ MECHANICAL FAILURE/OPERATOR ERROR) AND ALSO MIGHT BE CONTRIBUTED WITH "HIGH MECHANICAL STABILITY TIME (MST) OF INCOMING LATEX" OR "COAGULANT DIP 2- DIPPING SPEED TOO SLOW". THE DEVICE HISTORY RECORD WAS REVIEWED AND FOUND NOTHING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "[WARNINGS] 1. METHOD FOR USE (1) DO NOT INFLATE THE BALLOON IN THE URETHRA. [THE URETHRA MAY BE INJURED.] (2) DO NOT PULL THE CATHETER HARD. [THE BLADDER/URETHRA MAY BE INJURED.] 2. APPLICABLE PATIENTS PATIENTS WITH DELIRIUM WHO MIGHT PULL OUT CATHETER [WHEN PATIENT TUGS AT CATHETER UNCONSCIOUSLY, THE BLADDER AND URETHRA MAY BE DAMAGED.] [CONTRAINDICATIONS] 1. METHOD FOR USE: (1) DO NOT REUSE. (2) DO NOT RESTERILIZE. (3) THIS DEVICE CONTAINS 10% POVIDONE-IODINE. FOR PATIENTS WITH PAST HISTORY OF ALLERGIC HYPERSENSITIVITY TO POVIDONE-IODINE OR IODINE, CONSIDER USING ALTERNATIVE DISINFECTANTS. (4) BE CAREFUL THAT THE CATHETER IS NOT EXPOSED TO OINTMENTS, CONTRAST MEDIUM OR OIL-BASED LUBRICANTS (INCLUDING VEGETABLE OILS SUCH AS OLIVE OIL, MINERAL OILS SUCH AS WHITE PETROLATUM AND ANIMAL OILS). [THEY MAY DAMAGE THE DEVICE AND MAY BURST BALLOON.] (5) DO NOT HOLD THE DEVICE WITH FORCEPS, ETC. AVOID CONTACT WITH ANY BLADES OR SHARP-EDGED INSTRUMENTS. [CATHETER DAMAGE MAY CAUSE BALLOON RUPTURE AND ACCIDENTAL BALLOON REMOVAL OR FAILURE TO DEFLATE OR REMOVE THE BALLOON.] 2. APPLICABLE PATIENTS (1) PATIENTS WHO ARE OR HAVE BEEN ALLERGIC TO NATURAL RUBBER LATEX (2) PATIENTS WITH KNOWN ALLERGY TO SILVER COATED CATHETER [SHAPE, CONFIGURATION AND PRINCIPLES] BARD® I.C. SILVER FOLEY TRAY B CONSISTS OF A BALLOON CATHETER, URINE-COLLECTING BAG FOR CLOSED DRAINAGE SYSTEM, SYRINGE PREFILLED WITH STERILE WATER, WATER-SOLUBLE LUBRICANT, ANTISEPTIC SOLUTION, TWEEZERS, WATERPROOF SHEET, GAUZE PADS, COTTON BALLS, VINYL GLOVES AND STATLOCK® FOLEY. THERE ARE SEVERAL TYPES OF CLOSED DRAINAGE BAGS. THE BAG AND STATLOCK® FOLEY INCLUDED IN THE TRAY WILL DEPEND ON THE PRODUCT. <MATERIAL> BALLOON CATHETER: NATURAL RUBBER LATEX; SILVER COATING THIS PRODUCT IS MADE WITH BACTI-GUARD® SILVER ALLOY COATING. <SIZES OF CATHETERS> AVAILABLE IN SIZES 8 TO 24 EVERY 2FR 1. FOLEY CATHETER 2. ACCESSORIES CLOSED DRAINAGE BAG (THE ILLUSTRATION SHOWS ONE EXAMPLE OF TYPICAL CONFIGURATIONS.) SYRINGE PREFILLED WITH STERILE WATER WATER SOLUBLE LUBRICANT ANTISEPTICS: BARD® 10% POVIDONE-IODINE SOLUTION TWEEZERS GAUZE PADS WATERPROOF SHEET COTTON BALLS GLOVES STATLOCK® FOLEY [INTENDED USE & EFFECT- EFFICACY] THE DEVICE IS A TRAY KIT PRODUCT THAT IS USED FOR THE PURPOSE OF URINARY DRAINAGE AND COMBINES A DISPOSABLE CATHETER DESIGNED TO BE PLACED IN THE BLADDER, AND A URINE DRAINAGE BAG. [DIRECTIONS FOR USE] 1. METHOD OF USE THE DEVICE IS INTENDED FOR SINGLE USE ONLY AND IS NOT REUSABLE. (1) TO SECURE A STERILE FIELD FOR THE PROCEDURE, SPREAD A CLEAN WRAPPING PAPER. (2) PLACE WATERPROOF SHEET BENEATH PATIENT¿S BUTTOCKS. (FIG. 1) (3) PUT ON STERILE GLOVES. OPEN TRAY AND PLACE IT ON THE WRAPPING PAPER. (FIG. 2) FIG.1 FIG.2 (4) CLEANSE THE AREA AROUND THE EXTERNAL URETHRAL MEATUS WITH THE COTTON BALLS IMMERSED IN THE ANTISEPTICS. (FIG. 3) (5) LUBRICATE THE DISTAL END OF THE CATHETER WITH WATER-SOLUBLE LUBRICANT PACKAGED IN THE TRAY. (FIG. 4) FIG.3 FIG.4 (6) INSERT CATHETER INTO THE URETHRAL MEATUS, AND ADVANCE IT UNTIL THE BALLOON ENTERS THE BLADDER AND URINE FLOWS OUT THROUGH THE CATHETER. USING A SYRINGE PACKAGED IN THE TRAY, INFUSE THE SPECIFIED VOLUME OF STERILE WATER INTO THE INFLATION LUMEN TO INFLATE THE BALLOON. (FIG. 5) FIG.5 (7) PULL CATHETER TO SEAT THE BALLOON AT THE LEVEL OF THE BLADDER NECK AND SECURE PLACEMENT. (8) KEEP THE DRAINAGE BAG BELOW THE BLADDER LEVEL WITHOUT TOUCHING THE FLOOR. (FIG. 6) (9) SECURE DRAINAGE TUBE TO BED SHEET WITH CLIP TO ENSURE THAT THERE IS NEITHER TWIST NOR KINK IN THE TUBE. (FIG. 7) FIG.6 FIG.7 (10) TO DEFLATE BALLOON AND REMOVE CATHETER, INSERT A LUER TIP (NEEDLELESS) SYRINGE IN THE INFLATION VALVE TO ALLOW THE WATER DRAIN SPONTANEOUSLY. AFTER BALLOON DEFLATION, WITHDRAW THE CATHETER WHILE CONFIRMING THAT NO RESISTANCE IS ENCOUNTERED. <DIRECTION FOR USE BARD® EZ-LOK® SAMPLING PORT> 1) OCCLUDE DRAINAGE TUBING A MINIMUM OF 10 CM BELOW THE SAMPLING PORT BY KINKING THE TUBING UNTIL URINE FILLS THE TUBING UP TO NEAR (SLIGHTLY ABOVE) THE SAMPLING PORT. 2) SWAB SURFACE OF SITE WITH ANTISEPTIC WIPE. 3) USING ASEPTIC TECHNIQUE, POSITION THE NEEDLE-LESS SYRINGE (SLIP-TIP TYPE OR LUER-LOCK TYPE) IN THE CENTER OF THE SAMPLING PORT. THE SYRINGE SHOULD BE HELD PERPENDICULAR TO THE SURFACE OF THE SAMPLING PORT. PRESS THE SYRINGE AND TWIST TO LOCK THE SYRINGE ONTO THE SAMPLING PORT. (FIG. 8) FIG. 8 4) ASPIRATE DESIRED VOLUME OF URINE. AFTER SAMPLING, DETACH THE SYRINGE. ENSURE THAT THE RUBBER STEM OF THE SAMPLING PORT HAS RETURNED TO ITS ORIGINAL POSITION. 5) UNKINK TUBING. <HOW TO DISCONNECT CATHETER FROM TUBING> CATHETER IS PRE-CONNECTED TO EZ-LOK, AND THE CONNECTING PART IS COVERED WITH RED SEAL (TAMPER-EVIDENT SEAL). REMOVE THE SEAL BY GRASPING THE TAB AT THE END OF THE SEAL AND PULLING ALONG PERFORATIONS, AND THEN DISCONNECT THE CATHETER AND THE TUBING USING ASEPTIC TECHNIQUE. (FIG. 9) FIG. 9 2. PRECAUTIONS FOR USE (1) WHEN RESISTANCE IS ENCOUNTERED IN INSERTING CATHETER, STOP THE PROCEDURE AND REMOVE THE CATHETER. (2) WHEN DEFLATING BALLOON, DO NOT ASPIRATE WITH A SYRINGE. [THE INFLATION LUMEN FOR BALLOON DEFLATION MAY BE OCCLUDED BY NEGATIVE PRESSURE, AND AS A RESULT THE CATHETER CANNOT BE REMOVED.] (3) NO SUBSTANCE EXCEPT STERILE WATER SHOULD BE USED TO INFLATE THE BALLOON. [IF CONTRAST MEDIUM IS USED, BALLOON MAY BURST. IF NORMAL SALINE IS USED, CRYSTALLIZED SALT MAY OCCLUDE THE INFLATION LUMEN TO PREVENT DEFLATION OF THE BALLOON. IF AIR IS USED, AIR MAY ESCAPE TO CAUSE INADVERTENT DEFLATION OF THE BALLOON SO THAT THE CATHETER MAY COME OUT PREMATURELY. ] (4) DO NOT WIPE CATHETER SURFACE WITH ORGANIC SOLVENTS SUCH AS ALCOHOL. (5) DO NOT ASPIRATE URINE THROUGH DRAINAGE FUNNEL WALL. (6) SINCE MOVEMENT OF THE BODY, ETC. MAY TWIST OR BEND CATHETER TO CAUSE OCCLUSION, CARE SHOULD BE TAKEN TO FIX THE CATHETER SECURELY. (7) WHEN URINARY FLOW CANNOT BE NOTED, CONFIRM THAT THE CATHETER IS NEITHER OCCLUDED NOR BROKEN. (8) AVOID FORCE ON THE CONNECTING PARTS AS THEY MAY BE ACCIDENTALLY DISCONNECTED DUE TO THE WEIGHT OF THE DRAINAGE BAG ETC. AND MAY CAUSE URINE SPILL. (9) DO NOT PULL OR TWIST THE OUTLET TUBE. ALSO, DO NOT SQUEEZE THE DRAINAGE BAG. [THE JOINT OF THE DRAINAGE BAG AND THE OUTLET TUBE MAY BE DAMAGED AND URINE LEAKAGE MAY OCCUR.] (10) WHEN DISPOSING OF URINE, OBSERVE THE FOLLOWING; 1) REMOVE THE OUTLET TUBE FROM THE HOUSING OF THE URINE DRAINAGE BAG. 2) LIFT THE GREEN LEVER TO OPEN WITH HOLDING THE OUTLET TUBE. BE CAREFUL NOT TO PULL THE OUTLET TUBE WHEN LIFTING THE GREEN LEVER. 3) WHEN DISPOSAL OF URINE IS COMPLETED, CLOSE THE GREEN LEVER AND PUT THE OUTLET TUBE INTO THE HOUSING. (11) WHEN USING STATLOCK® FOLEY, OBSERVE THE FOLLOWING; 1) DO NOT USE THE STATLOCK® DEVICE WHERE LOSS OF ADHERENCE COULD OCCUR, SUCH AS WITH A CONFUSED PATIENT, UNATTENDED ACCESS DEVICE, DIAPHORETIC OR NONADHERENT SKIN. 2) MINIMIZE CATHETER MANIPULATION DURING STATLOCK® STABILIZATION DEVICE APPLICATION AND REMOVAL. 3) DO NOT USE THE STATLOCK® DEVICE FOR PATIENTS SHOWING ALLERGIC REACTION TO TAPE OR ADHESIVE. 4) CONDUCT SKIN ASSESSMENT PRIOR TO APPLICATION AND REPEAT DAILY PER FACILITY PROTOCOL. 5) THE STATLOCK® DEVICE SHOULD BE ASSESSED DAILY AND CHANGED WHEN CLINICALLY, INDICATED, AT LEAST EVERY SEVEN DAYS. 6) AFTER PLACING THE STATLOCK® DEVICE, ALLOW TO DRY COMPLETELY (10-15 SECONDS) DUE TO ALCOHOL INCLUDED IN SKIN PROTECTANT. 7) USE ALCOHOL PADS WHEN REMOVAL. DO NOT PULL OR FORCE PAD TO REMOVE. [PRECAUTIONS] 1. PRECAUTIONS FOR USE (EXERCISE CAUTION WHEN USING THE DEVICE IN THE FOLLOWING PATIENTS) (1) EXERCISE CAUTION WHEN USING THE DEVICE IN PATIENTS WITH HIGH URINARY CALCIUM LEVELS AS ENCRUSTATION ON THE BALLOON SURFACE, CATHETER OCCLUSION OR DAMAGE MAY OCCUR. 2. IMPORTANT PRECAUTIONS (1) WHEN CATHETER IS INADVERTENTLY REMOVED, INSPECT THE BALLOON AND SHAFT OF CATHETER FOR RUPTURE, DEFECT, ETC. BEFORE INSERTING A NEW CATHETER. (2) WHEN ANY PART OF THE BALLOON AND/OR THE CATHETER IS MISSING, CONSIDER REMOVING THEM USING A CYSTOSCOPE. (3) WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON, TAKE APPROPRIATE MEASURES ACCORDING TO THE SECTION ¿TROUBLESHOOTING¿. <TROUBLESHOOTING> WHEN IT IS DIFFICULT TO REMOVE CATHETER BY DEFLATING THE BALLOON (EXPRESSED AS ¿REMOVAL-DIFFICULT CASE¿ HEREINAFTER), TAKE APPROPRIATE MEASURES ACCORDING TO THE FOLLOWING PROCEDURES. THE FOLLOWING TWO METHODS ARE AVAILABLE FOR REMOVAL-DIFFICULT CASES. A. NON-RUPTURE METHOD (STERILE WATER IS WITHDRAWN WITHOUT BURSTING THE BALLOON.) B. BALLOON-RUPTURE METHOD WITH BALLOON-RUPTURE METHOD, FRAGMENTS OF THE RUPTURED BALLOON MAY REMAIN IN THE BLADDER. THEREFORE, TRY NON-RUPTURE METHOD FIRST. A. NON-RUPTURE METHOD 1) ATTACH LUER TIP SYRINGE TO THE INFLATION VALVE. INJECT AN ADDITIONAL AMOUNT OF STERILE WATER INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), SEVER THE INFLATION FUNNEL OF VALVE. (FIG. 10) FIG. 10 3) IF SITUATION WOULDN¿T BE IMPROVED WITH 2), SEVER THE CATHETER SHAFT WHILE HOLDING IT WITH FORCEPS SO THAT THE DISTAL SEGMENT MAY NOT BE DRAWN INTO THE URETHRA. (FIG. 11) FIG. 11 4) IF SITUATION WOULDN'T BE IMPROVED WITH 3), INSERT A NEEDLE INTO THE INFLATION LUMEN AND PUMP THE PLUNGER. (FIG.12) FIG. 12 5) IF SITUATION WOULDN'T BE IMPROVED WITH 4), INSERT A FINE STEEL WIRE THROUGH THE INFLATION LUMEN OF CATHETER. (FIG.13) FIG. 13 B. BALLOON-RUPTURE METHOD 1) INJECT 100-200ML/CC OF SALINE SOLUTION WARMED TO BODY TEMPERATURE INTO THE BLADDER THROUGH THE DRAINAGE LUMEN, AND THEN INJECT A LARGE AMOUNT OF WATER OR 10-15 ML/CC OF MINERAL OIL INTO THE BALLOON THROUGH THE INFLATION LUMEN WITH A NEEDLE TO INDUCE RUPTURE. AFTER RUPTURE OF THE BALLOON, IRRIGATE THE BLADDER. (FIG. 14) FIG. 14 2) IF SITUATION WOULDN'T BE IMPROVED WITH 1), ATTEMPT FOLLOWING PROCEDURES. A) UNDER THE RADIOSCOPIC OBSERVATION, INFUSE A CONTRAST MEDIUM INTO THE BLADDER, AND BURST THE BALLOON BY SUPRAPUBIC PUNCTURE OF THE BLADDER. (FIG. 15) FIG. 15 B) IN MALE PATIENTS, BURST THE BALLOON BY PUNCTURE WITH A NEEDLE FROM THE PERINEAL (OR SUPRAPUBIC) REGION OR THROUGH THE RECTUM UNDER ULTRASONOGRAPHIC GUIDANCE. (FIG. 16) FIG. 16 C) IN FEMALE PATIENTS, BURST THE BALLOON BY INSERTION OF A NEEDLE ALONG THE URETHRA. (FIG. 17) FIG. 17 3. MALFUNCTION AND ADVERSE EVENTS 1) MALFUNCTION - CATHETER KINKING, DAMAGE, RUPTURE - DIFFICULTY OR FAILURE TO REMOVE THE DEVICE - OCCLUSION OF CATHETER INNER LUMENS - ENCRUSTATION - ACCIDENTAL REMOVAL OF THE DEVICE DUE TO LEAKAGE OF STERILE WATER OR BALLOON RUPTURE - DEVICE DAMAGE DUE TO INAPPROPRIATE USE 2) ADVERSE EVENTS - URINARY-TRACT INFECTION - HEMORRHAGE, HEMATURIA - ALLERGY REACTION TO THE DEVICE - CALCULUS FORMATION - EDEMA - PAIN - DISCOMFORT - INJURY OF BLADDER OR URETHRAL - URETHRITIS, URINARY INCONTINENCE - RETAINED BALLOON FRAGMENTS [STORAGE METHOD AND EXPIRATION DATE] 1. STORAGE STORE IN A DRY, COOL PLACE AWAY FROM HEAT, MOISTURE, AND DIRECT SUNLIGHT. 2. EXPIRATION DATE INDICATED ON THE DIRECT PACKAGE AND THE OUTER BOX." CORRECTION: H H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01805 | 2023-05-10 | 2023-04-25 | I, n, j, u, r, y | Manufacturer report | POWERPICC CATHETER, 4F SL, 3CG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | CK000775A |  | CUSTOMER REPORTED: OUR TEAM INSERTED A BARD 4FR SINGLE LUMEN PICC IN A PATIENT. UNFORTUNATELY, THE PATIENT DID SUFFER FROM AN AIR EMBOLISM SECONDARY TO THE PICC LINE. THE PICC LINE WAS REMOVED AS THE INTEGRITY OF THE LINE WAS IN QUESTION AS A RESULT OF THE CLINICAL OUTCOME OF THIS PATIENT. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01797 | 2023-05-10 | 2023-03-15 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | 7655405 |  | IT WAS REPORTED BY THE CUSTOMER THAT THE PATIENT UNDERWENT CHEST X-RAY ON (B)(6) 2023, TO CHECK THE CORRECT POSITION OF THE CATHETER TIP. THEY FOUND SWELLING OF THE ARM ON THE PUNCTURE SIDE DURING INFUSION WITH THE PICC CATHETER AFTER UNOBSTRUCTED MAINTENANCE. THE INFUSION WAS STOPPED IMMEDIATELY, AND THERE WAS NO THROMBUS IN THE BLOOD VESSEL COLOR ULTRASOUND EXAMINATION. THE PICC CATHETER WAS PULLED OUT ACCORDING TO THE DOCTOR'S ADVICE, AND THE CATHETER WAS FOUND TO BE BROKEN IN THE BODY. CATHETER REMOVED VIA INTERVENTIONAL RADIOLOGY. THE PATIENT RECEIVED CHEMOTHERAPY ON MARCH 16 AND WAS DISCHARGED AFTER TREATMENT ON MARCH 17. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01804 | 2023-05-10 | 2023-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1385108D5 |  | CUSTOMER REPORTED A 5FR TL POWERPICC PROVENA INSERTED IN THE LEFT UPPER ARM BRACHIAL VEIN ON (B)(6) 2023. TWO LUMENS WERE TPA'D TWICE ON (B)(6) WITHOUT SUCCESS. CHEST X-RAY ON (B)(6) SHOWED A KINK NEAR THE INSERTION SITE. CATHETER WAS DISCOVERED TO BE KINKED NEAR THE INSERTION SITE ON (B)(6) 2023. PICC WAS REMOVED AND A NEW PICC PLACED. PICC DWELL TIME WAS 11 DAYS. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01803 | 2023-05-10 | 2023-03-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | CUSTOMER REPORTED A 4FR DL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2023. RED LUMEN WAS TPA'D ON (B)(6) BUT DIDN'T WORK. CATHETER WAS DISCOVERED TO BE KINKED NEAR THE INSERTION SITE ON (B)(6) 2023. PICC WAS PULLED BACK ON (B)(6) BUT KINK REMAINED. PICC EXCHANGED BY IR ON (B)(6) 2023. PICC DWELL TIME WAS 9 DAYS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01801 | 2023-05-10 | 2023-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | CUSTOMER REPORTED A 4FR DL POWERPICC PROVENA INSERTED IN THE LEFT UPPER ARM BASILIC VEIN ON (B)(6) 2023. THE CATHETER WAS TPA'D IN BOTH LUMENS. CATHETER WAS DISCOVERED TO BE KINKED NEAR THE INSERTION SITE ON (B)(6) 2023. THE PICC WAS NOTED TO BE TOO SHORT. PICC DWELL TIME WAS 23 DAYS. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01799 | 2023-05-10 | 2023-04-18 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN BASIC TRAY WITH SAFETY INTRODUCER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405 |  | IT WAS REPORTED THAT THIS FEMALE PATIENT WAS IN THE ADVANCED STAGE OF BREAST CANCER AND HAD A GROSHONG CATHETER PLACED FROM THE RIGHT BRACHIAL VEIN IN NOVEMBER 2022 AT A FACILITY. THE BLOOD WAS SEEN UPON ONE PUNCTURE. 36 CM OF THE CATHETER WAS INSERTED. THE PATIENT HAD NO HISTORY OF CATHETER PLACEMENT. THE CATHETER WAS USED TO INFUSE MEDICATIONS TO TREAT THE CANCER AND ALBUMIN, ETC. AND WAS MAINTAINED ONCE OUT OF THIS HOSPITAL. THE CATHETER APPEARED INTACT. THE CIRCUMFERENCE OF THE ARM HAVING THE CATHETER WAS 26.5 BEFORE INFUSION OF CHEMOTHERAPY DRUGS ON APRIL 17. THE CATHETER MAINTENANCE AND BLOOD DRAW WERE DONE SMOOTHLY. NO EFFUSION OR HEMATOCELE WAS SEEN IN THE ADDRESSING AND VITAMIN B WAS INFUSED. ON APRIL 18, THE ARM CIRCUMFERENCE INCREASED TO 31 AND NO BLOOD COULD BE DRAWN. X-RAY SHOWED THAT THE CATHETER TIP WAS IN A POSITION SLIGHTLY DEEPER THAN CAJ. A DAMAGE OF THE INTERNAL PORTION OF THE CATHETER WAS SUSPECTED. THE SPECIALIST NURSE RETRACTED THE CATHETER ONCE EVERY ABOUT 2 CM. AT THE FIRST 2 CM, THE CATHETER WAS RUPTURED WITHOUT FORCEFUL PULL. THE NURSE IMMEDIATELY BANDAGED THE PATIENT¿S ARM PROXIMAL TO THE HEART DUE TO THE CONCERN ABOUT LOOSENING FROM SWOLLEN ARM. ONE HOUR LATER, A VASCULAR SURGEON WAS INVITED TO PERFORM OPERATION, BUT NO CATHETER WAS FOUND FROM THE OPERATIVE WOUND TO THE AXILLA. TO AVOID A LARGER TRAUMA, THE WOUND WAS SUTURED. THREE HOURS LATER, THE PATIENT WAS SENT TO THE DSA DEPARTMENT OF THE PROVINCIAL PEOPLE¿S HOSPITAL TO RETRIEVE THE CATHETER IN AN INTERVENTIONAL MANNER. THE COMPLETE 35 CM CATHETER WAS REMOVED A HALF-HOUR LATER. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01787 | 2023-05-10 | 2023-04-13 | I, n, j, u, r, y | Manufacturer report | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 9808560 | 9808560 |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS AND ELEVEN MONTHS POST A PORT PLACEMENT, COMPUTED TOMOGRAPHIC EXAMINATION REVEALED THAT THE CATHETER WAS ALLEGEDLY BROKEN. IT WAS FURTHER REPORTED THAT APPROXIMATELY TEN CENTIMETERS OF THE DISTAL CATHETER SEGMENT REMAINED IN THE RIGHT VENTRICLE. REPORTEDLY, THE REMOVAL OF THE PORT BODY AND THE DISTAL CATHETER SEGMENT WILL BE CONSIDERED IN THE FUTURE. THERE CURRENT STATUS OF THE PATIENT IS UNKNOWN. | MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE POWERPORT ISP MRI IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, TACTILE AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED AT THE DISTAL END OF THE ATTACHED CATHETER THAT WAS ELLIPTICAL IN SHAPE. THE EDGES OF THE BREAK WAS NOTED TO BE UNEVEN AND THE SURFACE WERE NOTED TO BE ROUND AND SMOOTH IN ONE REGION AND GRANULAR IN THE OTHER REGION. THE DISTAL CATHETER SEGMENT WAS NOT RETURNED. THE CATH-LOCK WAS NOTED TO BE LOADED BACKWARDS. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE, MATERIAL SEPARATION AND THE IDENTIFIED DISASSEMBLED BY USER ISSUES. HOWEVER THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED MIGRATION ISSUES AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G. PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01800 | 2023-05-10 | 2023-03-13 | I, n, j, u, r, y | Manufacturer report | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (W/MICROINTRODUCER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7655405 |  | ON (B)(6) 2023, THE PATIENT UNDERWENT CHEST RADIOGRAM TO CHECK THAT THE PICC CATHETER WAS BROKEN IN VIVO, AND THE SPECIFIC LOCATION WAS UNKNOWN, SO HE WAS ASKED FOR URGENT CONSULTATION IN THE DEPARTMENT OF VASCULAR SURGERY AND TRANSFERRED TO THE DEPARTMENT OF VASCULAR SURGERY FOR TREATMENT. THE BROKEN CATHETER WAS REMOVED UNDER INTERVENTION ON (B)(6) AND THE PATIENT WAS TRANSFERRED BACK TO THE DEPARTMENT OF ONCOLOGY FOR FURTHER CHEMOTHERAPY ON (B)(6) AND WAS DISCHARGED AFTER TREATMENT ON (B)(6). | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01798 | 2023-05-10 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GROSHONG NXT WITH STRIPED CATH 4F SINGLE LUMEN BASIG JAPAN HANG | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 7617405J |  | IT WAS REPORTED THAT ON (B)(6)2023, THE LEAKAGE OCCURRED DUE TO THE CATHETER WAS DISLODGED IN THE CATHETER CONNECTOR. IT WAS FURTHER REPORTED THAT IMPLANTED DATE IS BEING CONFIRMING. THERE WAS NO REPORTED PATIENT INJURY. ADDITIONAL INFORMATION 5/2/2023: CATHETER WAS INSERTED ON (B)(6)2023 AND THE INCIDENT OCCURRED ON (B)(6)2023. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01795 | 2023-05-10 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC SOLO | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | CK001027 |  | IT WAS REPORTED BY THE CUSTOMER THERE IS A BUMP ON THE ACTUAL PICC. ADDITIONAL INFORMATION (B)(6) 2023: DEVICE WAS NOT USED ON A PATIENT, AND THERE WAS NO HARM OR NEGATIVE IMPACT REPORTED. | THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01793 | 2023-05-10 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED BY MEDWATCH THAT THREE DIFFERENT PICC'S WERE REQUIRED TO BE ATTEMPTED TO BE PLACED ON A PATIENT SECONDARY TO THE LINES LEAKING AT THE RUBBER STOPPER WHILE BEING FLUSHED. THIS REPORT ADDRESSES THE THIRD DEVICE. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01792 | 2023-05-10 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED BY MEDWATCH THAT THREE DIFFERENT PICC'S WERE REQUIRED TO BE ATTEMPTED TO BE PLACED ON A PATIENT SECONDARY TO THE LINES LEAKING AT THE RUBBER STOPPER WHILE BEING FLUSHED. THIS REPORT ADDRESSES THE SECOND DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01794 | 2023-05-10 | 2023-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NIAGARA DIALYSIS CATHETER KIT 13.5F X 15CM (PRE-CURVED) (SHORT-TERM) (DL) | CATHETER, HEMODIALYSIS, NON-IMPLANTED | C.R. BARD, INC. (BASD) -3006260740 | N/A | 5594150 |  | IT WAS REPORTED THAT THE PRODUCT WAS NOT OPENED OR USED. HAIRS WERE SEEN INSIDE THE PACKAGE THROUGH THE TRANSPARENT PACKAGING. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01790 | 2023-05-10 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC 4F SL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 1174108D5 |  | IT WAS REPORTED VIA MEDWATCH, SALINE LEAKS FROM RUBBER STOPPER IN PICC LINE DURING INSERTION PROCEDURE. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01791 | 2023-05-10 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1275108FD5 |  | IT WAS REPORTED BY MEDWATCH THAT THREE DIFFERENT PICC'S WERE REQUIRED TO BE ATTEMPTED TO BE PLACED ON A PATIENT SECONDARY TO THE LINES LEAKING AT THE RUBBER STOPPER WHILE BEING FLUSHED. THIS REPORT ADDRESSES THE FIRST DEVICE. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01789 | 2023-05-10 | 2023-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HEMOSTAR | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 | 5833730 | 5833730 |  | IT WAS REPORTED THAT PRIOR TO A DIALYSIS CATHETER PLACEMENT PROCEDURE, THE INTRODUCER NEEDLE WAS ALLEGEDLY FRACTURED. IT WAS FURTHER REPORTED THAT NEEDLE ALLEGEDLY HAD A LEAK. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT CONTACT. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. HOWEVER, A VIDEO WAS PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRATION DATE: 03/2024). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01788 | 2023-05-10 | 2022-10-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | GLIDEPATH | DIALYSIS CATHETER | C.R. BARD, INC. (BASD) -3006260740 |  | 6393230 |  | IT WAS REPORTED THAT PRIOR TO THE DIALYSIS CATHETER PLACEMENT PROCEDURE, A LINE OF THE CATHETER ALLEGEDLY PERFORATED AT THE SITE OF THE PUSH. IT WAS FURTHER REPORTED THAT METAL LEAD WAS ALLEGEDLY IMPOSSIBLE TO MOVE EVEN WHEN USING THE STRAIGHT END OF THE CATHETER, SO THE CATHETER OUTLET LINE WAS USED. REPORTEDLY MINOR BLEEDING WAS NOTED FROM THE LUMEN. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO REPORTED PATIENT INJURY. | H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUIRED AS THIS IS THE ONLY COMPLAINT REPORTED TO DATE FOR THIS PRODUCT AND LOT. INVESTIGATION SUMMARY: ONE GLIDEPATH D/L 23CM CATHETER WAS RETURNED FOR EVALUATION. GROSS VISUAL AND FUNCTIONAL EVALUATIONS WERE PERFORMED. A SPLIT WAS NOTED ON THE RED LUER EXTENSION LEG. UPON INFUSION OF THE RED LUER, A SMALL LEAK WAS OBSERVED FROM THE SPLIT ON THE EXTENSION LEG. THEREFORE THE INVESTIGATION IS CONFIRMED FOR THE REPORTED CATHETER PUNCTURE AND LEAK ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H10: B5, D4 (EXPIRY DATE: 03/2023), G3, H6 (METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3012307300-2023-05457 | 2023-05-10 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DELTEC POWER PORT-A-CATH II PORTS | SET, ADMINISTRATION, INTRAVASCULAR | ST PAUL |  |  |  | IT WAS REPORTED THAT THE PUMP WAS LEAKING. NO PATIENT INJURY WAS REPORTED. | THIS MDR WAS GENERATED UNDER PROTOCOL (B)(4), AS A RESULT OF WARNING LETTER CMS# (B)(4). NO PRODUCT SAMPLE WAS RECEIVED; THEREFORE, VISUAL AND FUNCTIONAL TESTING COULD NOT BE PERFORMED. THE REPORTED ISSUE COULD NOT BE CONFIRMED AS NO PRODUCT SAMPLE WAS RECEIVED FOR EVALUATION. IF THE PRODUCT IS RETURNED, THE MANUFACTURER WILL REOPEN THIS COMPLAINT FOR FURTHER INVESTIGATION. NO PROBLEMS OR ISSUES WERE IDENTIFIED DURING THIS DEVICE HISTORY RECORD REVIEW. D4: CATALOG NUMBER IS UNKNOWN. D4: UDI INFORMATION IS UNKNOWN. G5: PREMARKET (510K) NUMBER IS UNKNOWN. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03295 | 2023-05-10 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE STERILE WATER LEAKED FROM THE SHAFT IMMEDIATELY AFTER THE FOLEY CATHETER INSERTION. THE USER DISCONTINUED TO USE THE FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01818 | 2023-05-11 | 2023-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | STATLOCK CV PLUS SLIDING POST, TRICOT HYDROCOLLOID | DEVICE, INTRAVASCULAR CATHETER SECUREMENT | C.R. BARD, INC. (BASD) -3006260740 | N/A | CV0220CE |  | IT WAS REPORTED BY THE CUSTOMER "ON (B)(6) 2023, THE FIRST DAY OF THE SECOND CYCLE OF NEOADJUVANT CHEMOTHERAPY, THE CATHETER FIXATOR WAS OPENED DURING THE PICC CATHETER MAINTENANCE FOR THE PATIENT. HOWEVER, THE FASTENER WAS FOUND TO BE LOOSELY CONNECTED AND NOT IN USE, SO THE CATHETER FIXATOR WAS REPLACED." NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03324 | 2023-05-11 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE DOCTOR HAD ISSUES INSERTING THE PREFILLED FOLEY CATHETER, AND WAS UNABLE TO FILL THE BALLOON WITH THE PREFILLED WATER (BATCH# MYGT5336). A SECOND CATHETER WAS USED WITH A DIFFERENT BATCH# AND HAD THE SAME ISSUES BUT WAS EVENTUALLY ABLE TO GET BALLOON INFLATED (BATCH# MYGT5333). | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01808 | 2023-05-11 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 |  | 7707540J |  | IT WAS REPORTED THAT SOME TIME POST A PORT PLACEMENT, THE CATHETER WAS ALLEGEDLY BROKEN. THERE WAS NO REPORTED PATIENT INJURY. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE X-PORT ISP IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: ONE X-PORT IMPLANTABLE PORT ATTACHED TO A GROSHONG CATHETER WAS RECEIVED FOR EVALUATION. VISUAL, MICROSCOPIC, FUNCTIONAL AND TACTILE EVALUATIONS WERE PERFORMED. A COMPLETE CIRCUMFERENTIAL BREAK WAS NOTED ON THE PROXIMAL END OF THE DISTAL CATHETER SEGMENT AND THE DISTAL END OF THE ATTACHED CATHETER WHICH APPEARED ELLIPTICAL IN SHAPE. THE EDGES OF THE BREAKS WERE NOTED TO BE UNEVEN. SMALL SPLITS WERE OBSERVED ON THE BORDER OF BOTH BREAK SURFACES AND WERE NOTED TO BE GRANULAR. THEREFORE, THE INVESTIGATION IS CONFIRMED FOR THE REPORTED FRACTURE AND THE IDENTIFIED MATERIAL SEPARATION AND DENT ISSUES. THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BASED UPON AVAILABLE INFORMATION. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : SEE H10. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03309 | 2023-05-11 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE INTERMITTENT CATHETERS WERE TOO MESSY. NO ADDITIONAL INFORMATION PROVIDED. | THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS RETURNED FOR EVALUATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR, VISCOMETER FAILURE". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: "INSTRUCTIONS FOR USE INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03315 | 2023-05-11 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED, THAT WHEN THE PATIENT RETURNED TO WARD AFTER SURGERY. THE CATHETER SPONTANEOUSLY FELL OUT. ACCORDING TO THE INQUIRY SHEET ATTACHED, THE EVENT OCCURRED AFTER 5 HOURS OF PLACEMENT AND THERE WAS NO BALLOON RUPTURE / TEAR. NO MATERIALS POSSIBLY CAME IN CONTACT WITH THE CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03312 | 2023-05-11 | 2023-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE USER FACED DIFFICULTY WHILE DRAINING WATER FROM THE FOLEY CATHETER DURING PRETEST. THE WATER COULD BE INJECTED DURING THE PRETEST BUT COULD NOT BE WITHDRAWN. WHEN THE REPRESENTATIVE LOOKED AT THE SHAPE OF THE BALLOON, THE BALLOON WAS FOUND TO BE ALREADY DRAINED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03316 | 2023-05-11 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE BALLOON WAS NOT INFLATED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03321 | 2023-05-11 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT THE WATER LEAKED FROM THE INFLATION VALVE OF THE FOLEY CATHETER DURING A PRETEST BEFORE USE. | THE REPORTED EVENT WAS CONFIRMED AS SUPPLIER RELATED DUE TO WATER LEAKAGE AT THE INFLATION VALVE OF THE CATHETERS DURING INFLATING THE BALLOON WITH A SYRINGE. THE POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO A DEFECTIVE TIP ON THE SYRINGE- SMALLER TIP AT THE TAPER END CAUSED THE WATER TO LEAK DURING INFLATION FROM SUPPLIER. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS A CLOSURE LETTER WAS NOT REQUIRED FOR THIS COMPLAINT. THE REPORTED EVENT WAS UNDERSTOOD, CHARACTERIZED, AND CONFIRMED AS UNRELATED TO THE LABELING ISSUE. IT WAS CONFIRMED RELATED TO A SUPPLIER, THEREFORE LABELING REVIEW WAS NOT REQUIRED FOR THIS REPORTED EVENT. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT: THE ACTUAL/SUSPECTED DEVICE WAS INSPECTED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01807 | 2023-05-11 | 2023-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | M.R.I. LOW-PROFILE IMPLANTABLE PORT,HICKMAN SINGLE-LUMEN, 6.6F | PORT & CATHETER, IMPLANTED, SUBCUTANEOUS, INTRAVASCULAR | C.R. BARD, INC. (BASD) -3006260740 | 0603880 | 0603880 |  | IT WAS REPORTED THAT ONE YEAR EIGHT MONTHS AND TWENTY SEVEN DAYS POST PORT PLACEMENT, A LEAK WAS OBSERVED AT THE LEVEL OF THE PROXIMAL THIRD OF THE CATHETER. REPORTEDLY THE PORT SYSTEM WAS REMOVED AND REPLACED. THERE WAS NO REPORTED PATIENT INJURY. | AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. (EXPIRY DATE: 06/2023). SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03323 | 2023-05-11 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MAGIC3¿ COUDE INTERMITTENT CATHETER WITH SURE-GRIP¿ SLEEVE | INTERMITTENT CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 50614 | 50614 |  | IT WAS REPORTED THAT THE SAMPLES SENT WERE EXACTLY SAME ACCORDING TO THE PACKAGE, HOWEVER THE MAGIC 3 HYDROPHILIC INTERMITTENT COUDE CATHETER SAMPLES SENT BY THE LIBERATOR WERE NOTICEABLY THICKER IN DIAMETER THAN THE SAMPLES RECEIVED FROM THE DOCTOR¿S OFFICE AND THE PATIENT WAS NOT USING THE SAMPLES DUE TO DIFFERENCES AND THE REPRESENTATIVE OFFERED TO SEND NEW DIFFERENT SAMPLES BUT THE PATIENT DECLINED AND THIS SEEMS LIKE A VERY SERIOUS ISSUE. | THE REPORTED EVENT WAS INCONCLUSIVE AND NO SAMPLE INVESTIGATION. A POTENTIAL ROOT CAUSE FOR THIS FAILURE COULD BE DUE TO "OPERATOR ERROR". IT WAS UNKNOWN WHETHER THE DEVICE HAD MET SPECIFICATIONS. THE PRODUCT WAS USED FOR TREATMENT PURPOSES. IT WAS UNKNOWN WHETHER THE PRODUCT HAD CAUSED THE REPORTED FAILURE. THE DHR REVIEW COULD NOT BE PERFORMED WITHOUT A LOT NUMBER. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: INTENDED USE: THE CATHETER IS INTENDED FOR URINARY BLADDER DRAINAGE IN PATIENTS REQUIRING CATHETERIZATION FOR MANAGEMENT OF INCONTINENCE, VOIDING DYSFUNCTION, AND SURGICAL PROCEDURES. PLEASE CONTACT YOUR PHYSICIAN TO DETERMINE WHICH PRODUCT OPTIONS ARE BEST FOR YOU, PAYING CLOSE ATTENTION TO PRODUCT WARNINGS/ PRECAUTIONS AND ADVERSE REACTIONS. WASH YOUR HANDS THOROUGHLY WITH SOAP AND WATER. RELEASE THE STERILE WATER FROM THE FOIL PACKET. TIP THE CATHETER POUCH END-TO-END THREE TO SIX TIMES SO THE WATER MOVES BACK AND FORTH TO THOROUGHLY WET THE CATHETER SURFACE. PEEL OPEN THE PACK AT THE FUNNEL END JUST ENOUGH TO EXPOSE THE INSERTION SLEEVE. DON¿T REMOVE THE CATHETER YET. USE THE ADHESIVE TAB AT THE FUNNEL END OF THE PACK TO STICK THE PACK TO A NEARBY VERTICAL SURFACE WHILE PREPARING TO CATHETERIZE. WASH THE AREA AROUND THE MEATUS BEFORE CATHETERIZING. WASH YOUR HANDS AGAIN. HOLD THE INSERTION SLEEVE WITH YOUR DOMINANT HAND AND SQUEEZE IT TO GRIP THE CATHETER SHAFT AS YOU REMOVE THE CATHETER FROM THE PACK. NEXT, HOLD THE CATHETER FUNNEL ABOVE THE INSERTION SLEEVE WITH YOUR OTHER HAND AND SLIDE THE INSERTION SLEEVE DOWN THE SHAFT, STOPPING AT ABOUT 6¿ FROM THE TIP. RELEASE THE FUNNEL. USING THE INSERTION SLEEVE TO HOLD THE CATHETER FIRMLY, GENTLY PASS THE TIP OF THE CATHETER INTO YOUR URETHRA UNTIL THE INSERTION SLEEVE NEARS THE MEATUS. REPEAT UNTIL URINE STARTS TO FLOW. TRY TO KEEP THE CATHETER STEADY UNTIL URINE STOPS FLOWING. WHEN URINE STOPS FLOWING, SLOWLY WITHDRAW THE CATHETER, STOPPING IF FLOW STARTS AGAIN, UNTIL THE LAST FEW DROPS HAVE DRAINED. FINISH BY DISPOSING OF THE CATHETER AND ITS PACKAGING. WASH YOUR HANDS WITH SOAP AND WATER. WARNING: THIS IS A SINGLE USE DEVICE. DO NOT REUSE. REUSE OF A SINGLE USE DEVICE INCREASES THE RISK OF CATHETER ACQUIRED URINARY TRACT INFECTIONS. URETHRAL CATHETER FOR UROLOGICAL USE ONLY. DISCARD AFTER USE. MADE OF SILICONE ELASTOMER." H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01831 | 2023-05-11 | 2023-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REFURB, SHERLOCK 3CG TCS | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | 9770131R |  | IT WAS REPORTED WHEN PLACING A PICC LINE, THE WAVE LINE LOOKS FINE, BUT THE LOLLIPOP IS NOT TURNING GREEN AND OFTEN GIVES INACCURATE LOCATION. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01828 | 2023-05-11 | 2023-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1385108D5 |  | CUSTOMER REPORTED A 5FR TL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2023. ON (B)(6) 2023, THE CATHETER WAS TPA'D IN BOTH LUMENS BUT DIDN'T WORK. X-RAY SHOWED A KINK IN THE SUBCLAVIAN REGION ON (B)(6) 2023. PICC DWELL TIME WAS 17 DAYS. PICC WAS EXCHANGED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01827 | 2023-05-11 | 2023-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1385108D5 |  | CUSTOMER REPORTED A 5FR TL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2023 WAS FOUND TO BE KINKED ON (B)(6) 2023 IN THE AXILLARY REGION. PICC DWELL TIME WAS 13 DAY. PICC WAS EXCHANGED BY (B)(6). | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A KINK IN THE CATHETER WAS UNCONFIRMED AS THE REPORTED KINK COULD NOT BE CONCLUSIVELY SEEN IN THE RETURNED IMAGE. A SINGLE AP RADIOGRAPH SHOWING THE PATIENT¿S UPPER ARM AND CHEST WAS RETURNED FOR EVALUATION. A CATHETER, LIKELY THE IMPLICATED 5FR T/L POWERPICC PROVENA, WAS SEEN IN THE IMAGE AS A RIGHT-SIDED PICC. THE CATHETER TIP WAS LOCATED AT THE SVC CONFLUENCE. THE CATHETER WAS SEEN OVERLAPPING ITSELF IN THE RIGHT AXILLA. A DEFINITIVE KINK OF COIL IN THE CATHETER SHAFT COULD NOT BE CONFIRMED FROM THIS IMAGE. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01825 | 2023-05-11 | 2022-12-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1385108D5 |  | CUSTOMER REPORTED A 5FR TL POWERPICC PROVENA INSERTED IN THE LEFT UPPER ARM BRACHIAL VEIN ON (B)(6) 2022 WAS FOUND TO BE KINKED ON (B)(6) 2022 NEAR THE INSERTION SITE. PICC DWELL TIME WAS 10 DAYS. TPA WAS INSTILLED IN 2 LUMENS ON 12/20 WITHOUT GOOD RESULT SO AN X-RAY WAS REQUESTED TO ASSESS FOR A KINK. THE KINK RESULTED IN A DELAY IN INFUSING ALBUMIN. PICC WAS EXCHANGED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01824 | 2023-05-11 | 2023-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1385108D5 |  | CUSTOMER REPORTED A 5FR TL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2023. CATHETER WAS TPA'D IN ONE LUMEN ON 4/8 AND ONE ON 4/13. CATHETER CONTINUED TO HAVE PERSISTENT ISSUES WITH FUNCTIONALITY. CHEST X-RAY WAS REQUESTED TO ASSESS FOR KINK. KINK NEAR THE INSERTION SITE CONFIRMED ON 4/16. PICC WAS PULLED BACK 2CM ON 4/16 AND KINK RESOLVED. HOWEVER, PICC TIP IS NOW IN THE UPPER SVC AND "MAY NEED REPLACEMENT". PICC DWELL TIME WAS 17 DAYS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01822 | 2023-05-11 | 2023-02-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | CUSTOMER REPORTED A 4FR DL POWERPICC PROVENA INSERTED IN THE LEFT UPPER ARM BRACHIAL VEIN ON (B)(6) 2022, WAS FOUND TO BE KINKED ON (B)(6) 2023, IN THE UPPER 1/3 OF THE UPPER EXTREMITY PER X-RAY. PICC WAS SLUGGISH TO FLUSH AND HAD NO BLOOD RETURN, RESULTING IN TPA ADMINISTRATION ON 1/28 AND 1/31. CHEST X-RAY AND LEFT UPPER EXTREMITY X-RAY TAKEN ON 2/4 DUE TO RECURRENCE OF BOTH LUMENS BEING SLUGGISH TO FLUSH WITH NO BLOOD RETURN. X-RAY SHOWED KINKING. PICC DWELL TIME WAS 68 DAYS. A NEW PICC WAS INSERTED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01826 | 2023-05-11 | 2023-02-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1385108D5 |  | CUSTOMER REPORTED A 5FR TL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2023 WAS FOUND TO BE KINKED ON (B)(6) 2023 NEAR THE INSERTION SITE. ON (B)(6) , TWO LUMENS WERE SLUGGISH TO FLUSH WITH NO BLOOD RETURN SO TPA WAS INSTILLED. ON (B)(6), LUMENS WERE AGAIN SLUGGISH TO FLUSH WITHOUT BLOOD RETURN SO AN X-RAY OF THE RIGHT UPPER EXTREMITY WAS TAKEN AND SHOWED A CATHETER KINK NEAR THE INSERTION SITE. PICC DWELL TIME WAS 41 DAYS. CUSTOMER STATES THIS WAS USER ERROR. DRESSING WAS CHANGED AND HUB REPOSITIONED TO HAVE MORE GRADUAL TURN. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01823 | 2023-05-11 | 2023-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | S1385108D5 | S1385108D5 |  | CUSTOMER REPORTED A 5FR TL POWERPICC PROVENA INSERTED IN THE LEFT UPPER ARM BASILIC VEIN ON (B)(6) 2023. CATHETER WAS TPA'D 3 TIMES DURING IT'S USE. CATHETER WAS DISCOVERED TO BE KINKED NEAR THE INSERTION SITE ON (B)(6) 2023. PICC WAS PULLED BACK ON (B)(6) BUT KINK REMAINED. A NEW PICC WAS INSERTED. PICC DWELL TIME WAS 10 DAYS. | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01821 | 2023-05-11 | 2023-04-18 | I, n, j, u, r, y | Manufacturer report | BARD POWERGLIDE MIDLINE CATHETER | CATHETER,INTRAVASCULAR,THERAPEUTIC,SHORT-TERM LESS THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | UNKNOWN |  | IT WAS REPORTED VIA MS&S: CALLER STATES SHE HAD A POWERGLIDE PRO INSERTED ON (B)(6) 2023, FOR CONTINUOUS INFUSION OF BENADRYL. CALLER STATES SHE IS NOW HAVING PAIN DURING INFUSION AND SLIGHT SWELLING IN THE ARM THAT THE MIDLINE WAS PLACED. CALLER STATES THE SITE IS NOT RED OR FEVERISH. CALLER STATES THEY HAVE SCHEDULED HOME HEALTH TO EVALUATE TODAY BUT NOT SURE WHAT TIME THEY WILL ARRIVE. MS&S RESPONSE: INFORMED SHE NEED TO REACH OUT TO THE HOME HEALTH NURSE IMMEDIATELY FOR EVALUATION. IF HOME HEALTH IS NOT AVAILABLE TO ACCESS THE MIDLINE, WE RECOMMEND GOING TO THE EMERGENCY ROOM. SUGGESTED SHE ASK HOME HEALTH ABOUT DISCONTINUING THE USE OF THE MIDLINE UNTIL SHE WAS ABLE TO BE EVALUATED AND CALLER STATED SHE IS UNABLE TO STOP THE INFUSION. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03336 | 2023-05-11 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0165L16 |  | IT WAS REPORTED THAT DURING A PRETEST, WATER LEAKED FROM BETWEEN THE SYRINGE AND THE INFLATION VALVE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03334 | 2023-05-11 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® ALL-SILICONE TEMPERATURE-SENSING PEDIATRIC FOLEY CATHETER | FOLEY CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 119108 | 119108 |  | IT WAS REPORTED THAT WHEN THE CATHETER WAS TESTED PRIOR TO INSERTION, THERE WAS NO BALLOON DEFLATION, BUT DURING THE PROCEDURE, THE CATHETER FELL OUT AND WHEN CHECKED, THERE WAS A HOLE. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03329 | 2023-05-11 | 2023-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® I.C. TEMPERATURE-SENSING ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 119314 | 119314 |  | IT WAS REPORTED THAT THE CATHETER FELL OFF DURING THE SURGERY. ACCORDING TO THE INQUIRY SHEET, THE EVENT OCCURRED AFTER 1 DAY OF PLACEMENT AND THERE WAS NO BALLOON RUPTURE/TEAR. NO MATERIALS CAME IN CONTACT WITH THE CATHETER. PER FOLLOW-UP INFORMATION RECEIVED FROM IBC ON 26APR2023, STATED THAT THE DETAILS REGARDING ANY WATER LEAK WERE NOT PROVIDED FROM CUSTOMER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03330 | 2023-05-11 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® FLIP-FLO¿ CATHETER VALVE | CATHETER VALVE | C.R. BARD, INC. (COVINGTON) -1018233 | BFF5 | BFF5 |  | IT WAS REPORTED THAT THE PATIENT RECEIVED 5 BFF5 BUT 2 OF THE BOXES WERE EMPTY. ALL 5 BOXES WERE SEALED, BUT PATIENT ONLY RECEIVED 3 VALVES. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01819 | 2023-05-11 | 2023-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | S1385108D5 | S1385108D5 |  | CUSTOMER REPORTED A 5FR TL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BRACHIAL VEIN ON (B)(6) 2023. RED LUMEN TPA'D ON (B)(6) AND REPEATED X1 (2 DOSES) BUT DIDN'T WORK. KINK NEAR THE INSERTION SITE CONFIRMED ON (B)(6). PICC WAS PULLED BACK 2CM ON 4/14 AND KINK RESOLVED. HOWEVER, PICC TIP IS NOW IN THE UPPER SVC (TOO SHORT, 2CM ABOVE CARINA) AND "MAY NEED REPLACEMENT". PICC DWELL TIME WAS 23 DAYS. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IS PENDING EVALUATION. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
16910699 | 2023-05-11 | 2023-02-22 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | BARD® TOUCHLESS® PLUS UNISEX PRE-LUBRICATED VINYL URETHRAL CATHETER KIT | TRAY, CATHETERIZATION, STERILE URETHRAL, WITH OR WITHOUT CATHETER (KIT) | C. R. BARD, INC. | 4A5144 | 4A5144 |  | THE PACK OF THREE POVIDONE-IODINE SWAB STICKS THAT ARE ALWAYS INCLUDED AND USUALLY SATURATED WITH POVIDONE IODINE, WERE COMPLETELY DRY IN THIS STRAIGHT CATHETER KIT. THE STRAIGHT CATHETER PACKAGE WAS NOT OPENED AND THE PACKAGE FOR THE IODINE WAS ALSO NOT OPENED AT ALL UNTIL USE. THE PATIENT THAT THE CATHETER WAS TO BE USED FOR WAS NOT AFFECTED, BUT A NEW KIT HAD TO BE USED AS I HAD TO BREAK STERILITY TO GET A NEW KIT WITH IODINE TO CLEAN THE PATIENT. I COULD ALSO HAVE JUST GOTTEN IODINE OR BETADINE TO CLEAN BUT WE DO NOT HAVE THOSE ROUTINELY STOCKED IN ROOMS SO IN ORDER TO MINIMIZE THE TIME GATHERING SUPPLIES, I USED ANOTHER KIT THAT WAS ALREADY PRESENT. OTHERWISE THE REST OF THE KIT HAD ALL EXPECTED COMPONENTS. |  | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03339 | 2023-05-11 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARD® RUBBER UTILITY CATHETER | INTERMITTENT CATHETER | C.R. BARD INC. (COVINGTON) -1018233 | 277716 | 277716 |  | IT WAS REPORTED THAT THE CATHETER WAS BIGGER THAN THE 16FR, THEY USED TO USE AND IT WAS UNCOMFORTABLE INSERTING. NO MEDICAL INTERVENTION WAS REPORTED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | LAY USER/PATIENT |  | 2023-05
1018233-2023-03350 | 2023-05-12 | 2022-11-22 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | PWFX30 | PWFX30 |  | IT WAS REPORTED THAT THE PATIENT FELT THE PUREWICK FEMALE EXTERNAL CATHETER WAS UNCOMFORTABLE IN THE HOSPITAL. NO MEDICAL INTERVENTION WAS REPORTED. PER FOLLOW UP VIA PHONE ON 06DEC2022, IT WAS REPORTED THAT THE CUSTOMER ONLY USED THE PRODUCT WHILE IN HOSPITAL. IT WAS STATED THAT THE CUSTOMER EXPERIENCED AN URINARY TRACT INFECTION WITH THE USE OF PURE WICK AND WAS GIVEN ANTIBIOTICS TO TREAT THE INFECTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01868 | 2023-05-12 | 2023-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F DL FT MAX BARRIER PLUS KIT W/BIOPATCH | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1275108FD5 |  | IT WAS REPORTED VIA MEDWATCH "BD POWERPICC PROVENA FT CATHETERS WITH SHERLOCK 3CG TIP POSITIONING SYSTEM (TPS) STYLET REF S1275108FD5, LOT # REGY0842 NOTED LEAKAGE AT RUBBER-LIKE STOPPER AND STYLET ONCE PLACED. STYLET ALSO NOTABLY DIFFICULT TO REMOVE FROM CATHETER. STYLET RETAINED. PICC SUCCESSFULLY PLACED." THIS REPORT ADDRESSES THE "LEAKAGE." NO OTHER INFORMATION WAS PROVIDED. | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03342 | 2023-05-12 | 2023-04-21 | I, n, j, u, r, y | Manufacturer report | SPIRIT® HYDROCOLLOID ADHESIVE SHEATH | MALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 35301 | 35301 |  | IT WAS REPORTED THAT THE PATIENT STATED THAT THIS ISSUE HAS BEEN GOING ON FOR ABOUT THREE MONTHS AND THE PATIENT WAS STILL HAVING WHEN REMOVING. THE PATIENT¿S SKIN BECAME SENSITIVE AND IRRITATED. THE PATIENT STATED THAT THEIR DOCTOR GAVE A MATERIAL TO PUT ON THEIR PENIS FOR A COUPLE OF WEEKS AND NOW IT IS REALLY RED. THE PATIENT HAD BEEN USING A BLOW DRYER FOR THE PAST FIVE WEEKS TO KEEP THE CONDOM CATHETER ON. IT WAS STATED THAT THE PATIENT WAS THINKING THAT WITH WASHING, THEY WOULD BE OK. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | Y | N | LAY USER/PATIENT |  | 2023-05
1018233-2023-03351 | 2023-05-12 | 2023-03-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRI-SIL® ALL-SILICONE FOLEY CATHETER | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 175816 | 175816 |  | IT WAS REPORTED THAT UPON AN ATTEMPT TO REMOVE A FOLEY CATHETER, THERE WAS NOTED TO BE A DEFECT IN THE DISTAL MOST PART OF THE BALLOON. WHEN THE 10CC OF WATER WERE REMOVED FROM THE BALLOON, THE SILICONE APPEARS NOT TO HAVE RETRACTED ALL THE WAY BACK INTO ORIGINAL POSITION AND CREATED A SILICONE FOLD. THE PRESENCE OF THE FOLD CAUSED THE CATHETER TO BECOME LODGED JUST INSIDE THE PATIENT'S URINARY MEATUS, PREVENTING NURSING STAFF FROM BEING ABLE TO REMOVE THE LAST BIT OF THE FOLEY CATHETER ALL THE WAY. PER FOLLOW UP VIA EMAIL ON 03MAY2023, IT WAS REPORTED THAT THE PROVIDER DID NOT HAVE TO DO ADDITIONAL SURGERY BUT DID ADDITIONAL MANEUVERS TO ENSURE NO MATERIAL REMAINED IN THE PATIENT. IT WAS ALSO NEVER LODGED IN THE PATIENT¿S URETHRAL MEATUS ¿ THIS WAS A BALLOON ANGIOPLASTY PROCEDURE, NOT A FOLEY CATHETER. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03348 | 2023-05-12 | 2023-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® FOLEY CATHETER SILICONE COATED | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | 0167V20S |  | IT WAS REPORTED THAT WHEN THE NURSING WAS PERFORMED FOR THE PATIENT'S PERINEUM, THE CATHETER WAS FOUND COMPLETELY FALLING OUT FROM THE URETHRAL OPENING. UPON CHECKING, THERE WAS NO SALINE IN THE BALLOON OF THE URINARY CATHETER AND THE BALLOON WAS COMPLETELY DEFLATED AND THE TIME OF CATHETER FALLING OUT WAS UNKNOWN. AS PER THE FOLLOW UP INFORMATION RECEIVED ON 06MAY2023, STATED THAT THERE WAS NO DAMAGE ON THE BALLOON AND IF THE BALLOON FELL OUT DUE TO WATER LEAKAGE WAS UNKNOWN. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03347 | 2023-05-12 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BARDEX® LUBRICATH® FOLEY CATHETER CREEVY MODEL | FOLEY CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 | 01267520 | 1267520 |  | IT WAS REPORTED THAT THE BALLOON HAD BURST AND DROPPED. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3012307300-2023-05637 | 2023-05-12 | 2022-01-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DELTEC POWER PORT-A-CATH II PORTS | PORT AND CATHETER, IMPLANTED, SUBCUTANEOUS | ST PAUL | 21-8466-24 | 21-8466-24 |  | IT WAS REPORTED THAT THE DEVICE SEAL WAS DEFECTIVE AND MEDICATION WAS LEAKING. NO PATIENT INJURY WAS REPORTED. | OTHER TEXT: E4 IS UNKNOWN. THIS MDR WAS GENERATED UNDER PROTOCOL (B)(4), AS A RESULT OF WARNING LETTER (B)(4). NO PRODUCT SAMPLE WAS RECEIVED; THEREFORE, VISUAL AND FUNCTIONAL TESTING COULD NOT BE PERFORMED. THE REPORTED ISSUE COULD NOT BE CONFIRMED AS NO PRODUCT SAMPLE WAS RECEIVED FOR EVALUATION. IF THE PRODUCT IS RETURNED, THE MANUFACTURER WILL REOPEN THIS COMPLAINT FOR FURTHER INVESTIGATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01849 | 2023-05-12 | 2022-10-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1274108D5 |  | IT WAS REPORTED BY THE CUSTOMER, "A 4FR DL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BRACHIAL VEIN ON (B)(6) 2022 WAS FOUND TO BE KINKED ON (B)(6) 2022 IN THE UPPER 1/3 OF THE UPPER EXTREMITY. PICC DWELL TIME WAS 5 DAYS. PICC WAS PULLED BACK BUT IT DID NOT RESOLVE THE KINK SO THE PICC WAS REMOVED." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A CATHETER KINK WAS CONFIRMED BUT THE CAUSE IS UNKNOWN. A SINGLE AP RADIOGRAPH OF THE PATIENT¿S RIGHT ARM WAS RETURNED FOR EVALUATION. A CATHETER, CONSISTENT WITH THE IMPLICATED 4FR D/L POWERPICC PROVENA, WAS SEEN IN THE RETURNED IMAGE. THREE POSSIBLE KINKS WERE NOTED IN THE CATHETER SHAFT AT THE PATIENT¿S SKIN, THE MUSCLE ENTRANCE, AND THE VENOUS INSERTION SITE. POSSIBLE CONTRIBUTING FACTORS FOR A KINK IN THE CATHETER COULD INCLUDE THE ANATOMIC VARIABLES, CATHETER PLACEMENT, OR CATHETER SECUREMENT. THE IFU STATES, ¿AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT.¿ AND ¿AVOID SHARP OR ACUTE ANGLES DURING IMPLANTATION WHICH COULD COMPROMISE THE PATENCY OF THE CATHETER LUMEN¿. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. THE REPORT OF A KINKED CATHETER HAS BEEN DOCUMENTED AND WILL BE MONITORED AS PART OF COMPLAINT TRENDING. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01850 | 2023-05-12 | 2022-10-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1385108D5 |  | IT WAS REPORTED BY THE CUSTOMER, "A 5FR TL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BRACHIAL VEIN ON (B)(6)2022 WAS FOUND TO BE KINKED ON (B)(6)2022 IN THE UPPER 1/3 OF THE UPPER EXTREMITY. PICC DWELL TIME WAS 13 DAYS. PICC WAS PULLED BACK." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01851 | 2023-05-12 | 2022-10-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 5F TL MAX BARRIER PLUS KIT W/BIOPATCH, PROBE COVER | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S1385108D5 |  | IT WAS REPORTED BY THE CUSTOMER, "A 5FR TL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2022 WAS FOUND TO BE KINKED ON (B)(6) 2022 IN THE UPPER 1/3 OF THE UPPER EXTREMITY. PICC DWELL TIME WAS 15 DAYS. PICC WAS PULLED BACK, WHICH RESOLVED THE KINK BUT THE TIP WAS TOO SHORT." | SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01855 | 2023-05-12 | 2023-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA CATHETER RADIOLOGY BASIC KIT (5F) (TRIPLE-LUMEN) (N135CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | N/A | S3385335 |  | CUSTOMER REPORTED A 5FR TL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2023 WAS SUSPECTED TO BE KINKED DUE TO PERSISTENT SLUGGISHNESS AND INTERMITTENT BLOOD RETURN. PICC DWELL TIME WAS 0 DAYS. PICC REMOVED AND A TUNNELED CVC PLACED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01853 | 2023-05-12 | 2023-01-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER, BELDAR NG) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1274108D5 |  | CUSTOMER REPORTED A 4FR DL POWERPICC PROVENA INSERTED IN THE LEFT UPPER ARM BRACHIAL VEIN ON(B)(6)2023 WAS FOUND TO BE KINKED ON (B)(6)2023 IN THE UPPER 1/3 OF THE UPPER EXTREMITY PER X-RAY. BOTH LUMENS WERE SLUGGISH TO FLUSH AND TPA WAS INSTILLED TWICE. PICC DWELL TIME WAS 2 DAYS. A NEW PICC WAS PLACED IN THE RIGHT UPPER EXTREMITY BY IR. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A KINK IN THE CATHETER WAS UNCONFIRMED AS THE REPORTED KINK COULD NOT BE CONCLUSIVELY SEEN IN THE RETURNED IMAGE. A SINGLE AP RADIOGRAPH SHOWING THE PATIENT¿S LEFT ARM/HUMERUS WAS RETURNED FOR EVALUATION. A CATHETER, CONSISTENT WITH THE IMPLICATED POWERPICC PROVENA, WAS SEEN IN THE IMAGE. THE CATHETER SHAFT WAS OVERLAPPING ITSELF WITHIN THIS VIEW. A DEFINITIVE KINK, BREAK, OR COIL COULD NOT BE CONFIRMED FROM THE IMAGE. HOWEVER, THE RADIOGRAPH DID NOT EXCLUDE THE POSSIBILITY OF KINKING AT THE VENOUS INSERTION SITE. IT IS POSSIBLE THAT THE CATHETER WAS KINKED, BUT IT COULD NOT BE INDEPENDENTLY CONFIRMED FROM THE PROVIDED IMAGE. POSSIBLE CONTRIBUTING FACTORS FOR A KINK IN THE CATHETER COULD INCLUDE CATHETER PLACEMENT, PATIENT ANATOMY, OR CATHETER SECUREMENT. THE IFU STATES, ¿AVOID PLACEMENT OR SECUREMENT OF THE CATHETER WHERE KINKING MAY OCCUR, TO MINIMIZE STRESS ON THE CATHETER, PATENCY PROBLEMS OR PATIENT DISCOMFORT.¿ AND ¿AVOID SHARP OR ACUTE ANGLES DURING IMPLANTATION WHICH COULD COMPROMISE THE PATENCY OF THE CATHETER LUMEN.¿ A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE SPECIFIC LOT INVOLVED FOUND NO POTENTIALLY RELATED ISSUES ENCOUNTERED DURING THE MANUFACTURE AND ASSEMBLY OF THAT PRODUCT LOT. H3 OTHER TEXT : EVALUATION FINDINGS ARE IN SECTION H11 | N | Y | HEALTH PROFESSIONAL |  | 2023-05
1018233-2023-03341 | 2023-05-12 | 2023-04-21 | I, n, j, u, r, y | Manufacturer report | PUREWICK FEMALE EXTERNAL CATHETER | PUREWICK FEMALE EXTERNAL CATHETER | C.R. BARD, INC. (COVINGTON) -1018233 |  | UNK |  | IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN ALLERGIC REACTION TO THEIR OWN URINE FROM THE PUREWICK FEMALE EXTERNAL CATHETER. THE CATHETER WAS USED ON THE PATIENT FOR 10 HOURS CAUSING THE REACTION. PATIENT WAS PRESCRIBED POWDER AND ANTIBIOTIC CREAM FOR THE REACTION. | THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : THE DEVICE WAS NOT RETURNED. | Y | N | LAY USER/PATIENT |  | 2023-05
3006260740-2023-01854 | 2023-05-12 | 2023-01-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA CATHETER RADIOLOGY BASIC KIT (5F) (TRIPLE-LUMEN) (N135CM) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S3385335 |  | CUSTOMER REPORTED A 5FR TL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BASILIC VEIN ON (B)(6) 2023 WAS FOUND TO BE KINKED ON (B)(6) 2023 IN THE UPPER 1/3 OF THE UPPER EXTREMITY. PICC DWELL TIME WAS 3 DAYS. PICC WAS REMOVED AND A CVC WAS PLACED BY IR. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL EVALUATE. RESULTS ARE EXPECTED SOON. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01856 | 2023-05-12 | 2023-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1274108D5 |  | CUSTOMER REPORTED A 4FR DL POWERPICC PROVENA INSERTED IN THE RIGHT UPPER ARM BRACHIAL VEIN ON (B)(6)2023. ON (B)(6), THE CATHETER WAS TPA'D IN BOTH LUMENS. A CHEST X-RAY WAS REQUESTED AND SHOWED A KINK IN THE UPPER 1/3 OF THE UPPER ARM ON (B)(6)2023. PICC DWELL TIME WAS 41 DAYS. PICC WAS PULLED BACK. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01857 | 2023-05-12 | 2023-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | POWERPICC PROVENA 4F DL MBPLUS (3CG, BIOPATCH, PROBE COVER) | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | S1274108D4 |  | IT WAS REPORTED BY THE CUSTOMER LEAKING AT THE T-PORT WHERE THE STYLE GOES THROUGH A RUBBER STOPPER. NO HARM REPORTED, JUST A SAFETY CONCERN FOR THE CLINICIAN. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01858 | 2023-05-12 | 2023-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT 4FR SL PWERML | MIDLINE CATHETER | C.R. BARD, INC. (BASD) -3006260740 | NA | PP4154108DBPX |  | IT WAS REPORTED BY THE CUSTOMER "NOTICED LEAK IN EXTENSION SET NEAR T-LOCK EXTENSION SET OF MIDLINE CATHETER WHEN FLUSHING SALINE PRIOR TO STARTING MIDLINE INSERTION PROCEDURE. STOPPED PROCEDURE AND WENT TO GET A NEW MIDLINE. DID NOT USE KIT WITH LEAK NOTED IN EXTENSION. NO HARM REPORTED; SAFETY CONCERN FOR CLINICIAN." | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. H3 OTHER TEXT : DEVICE NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2023-05
3006260740-2023-01861 | 2023-05-12 | 2023-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DOT PWRMIDLINE CATHETER 4F SL | CATHETER,INTRAVASCULAR,THERAPEUTIC,LONG-TERM GREATER THAN 30 DAYS | C.R. BARD, INC. (BASD) -3006260740 | NA | CK001029 |  | IT WAS REPORTED BY THE CUSTOMER THAT THERE WAS A HOLE IN THE WRAPPING ON THE OUTSIDE OF THE TRAY. THERE WAS NO PATIENT HARM ONLY HAVING TO OPEN ANOTHER KIT. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A DAMAGED DRAPE OR CSR WRAP IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE PHOTOGRAPH OF WHAT APPEARS TO BE A DRAPE OR CSR WRAP WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED WHAT APPEARS TO BE A DRAPE OR CSR WRAP THAT APPEARS TO BE DAMAGED WITH A PIECE OF MATERIAL MISSING. WHILE THE EXACT CAUSE OF THE OBSERVED DAMAGED DRAPE OR CSR WRAP COULD NOT BE DETERMINED FROM THE RETURNED PHOTOGRAPH, POSSIBLE CAUSES OF THE DAMAGED DRAPE OR CSR WRAP INCLUDE DAMAGED DURING STORAGE, PACKAGING, HANDLING, OR USE, AND SHARP INSTRUMENT DAMAGE. | N | Y | HEALTH PROFESSIONAL |  | 2023-05